## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Browns Crossing
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2001-07-16 14:27 | 19.8 | 8.58 | 7.1 | 90 | 48 | 0
Sonde | 2001-07-16 14:31 | 19.9 | 8.6 | 8.8 | 112 | 40
Grade |  | A | A | Fail |  | Fail
Audit 2 | 2001-07-18 15:00 | 18.9 | 9 | 7.6 | 94 | 50 | 0
Sonde | 2001-07-18 15:01 | 18.8 | 8.7 | 8.3 | 103 | 41
Grade |  | A | A | B |  | Fail
Audit 3 | 2001-07-20 09:35 | 18.3 | 8.6 | 7.9 | 97 | 48 | 0
Sonde | 2001-07-20 09:31 | 18.2 | 8.5 | 8.2 | 101 | 31
Grade |  | A | A | A |  | Fail
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Formatting Note: Audits 1-3 represnet the duration of the deployment and are macro (ed) into the main audit sheet. Audits 4-5 represent the middle of the deployment and are independent of the main audit sheet. Conductivity readings were off and average of 52.5 uS/cm per reading based on pre and post deployment data. Conductivity probes was difficult to calibrate and readings for this datasonde's conductivity probe have been recording elevated readings during previous deployments. A corrected reading graph has been included in this DS Report. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-16 14:27 | 19.8 | 8.58 | 7.1 | 90 | 48
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-18 15:00 | 18.9 | 9 | 7.6 | 94 | 50
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-20 09:35 | 18.3 | 8.6 | 7.9 | 97 | 48
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Old Audit
[5 empty] | Department of Environmental Quality Laboratory | [6 empty] | SITE ELEVATION |  | 4430
[5 empty] | Water Quality Monitoring Section Datalogger Monitoring Report | [6 empty] | SITE CODE |  | 17 |  |  | d2= temp g2=DO
[16 empty] | example | =g2/vlookup(d2,'DS QA BLANK'!table,'DS QA BLANK'!$s$2+1)
| LASAR # | SITENAME | [6 empty] | FILENAME | CASE #
| 25837 | Deschutes River @ Rd 42 Bridge d/s of Crane Prairie | [10 empty] | DO Sat Table
Elev
[12 empty] | 0 | 500 | 750 | 1000 | 1250 | 1500 | 1750 | 2000 | 2250 | 2500 | 2750 | 3000 | 3250 | 3500 | 3750 | 4000 | 4250 | 4500 | 4750 | 5000 | Temp
| NOTE: Logged data is reported in database-ready format. Data available via electronic transfer (floppy) upon request | [9 empty] | Temp | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
| or can be retrieved from agency database/STORET. | [5 empty] | ELEVATION | 4430
[11 empty] | 0 | 14.6 | 14.2 | 14.2 | 14 | 14 | 13.8 | 13.65 | 13.5 | 13.3 | 13.3 | 13.1 | 12.95 | 12.8 | 12.8 | 12.65 | 12.5 | 12.35 | 12.2 | 12.2 | 12.05 | 0
| TEMPERATURE AUDIT RESULTS | [5 empty] | pH AUDIT RESULTS |  |  |  | 0.5 | 14.4 | 14 | 14 | 13.8 | 13.8 | 13.65 | 13.5 | 13.3 | 13.1 | 13.1 | 12.95 | 12.8 | 12.65 | 12.65 | 12.5 | 12.35 | 12.2 | 12.05 | 12.05 | 11.9 | 0.5
[11 empty] | 1 | 14.2 | 13.8 | 13.8 | 13.65 | 13.65 | 13.5 | 13.3 | 13.1 | 12.95 | 12.95 | 12.8 | 12.65 | 12.5 | 12.5 | 12.35 | 12.2 | 12.05 | 11.9 | 11.9 | 11.75 | 1
# | Audit | DS Value | Abs. Difference | Status |  | # | Audit | DS Value | Abs. Difference | Status | 1.5 | 14 | 13.65 | 13.65 | 13.5 | 13.5 | 13.3 | 13.1 | 12.95 | 12.8 | 12.8 | 12.65 | 12.5 | 12.35 | 12.35 | 12.2 | 12.05 | 11.9 | 11.75 | 11.75 | 11.6 | 1.5
0 | 19.8 | 19.9 | 0.1 | GRADE A |  | 0 | 8.6 | 8.2 | 0.4 | GRADE B | 2 | 13.8 | 13.5 | 13.5 | 13.3 | 13.3 | 13.1 | 12.95 | 12.8 | 12.65 | 12.65 | 12.5 | 12.35 | 12.2 | 12.2 | 12.05 | 11.9 | 11.75 | 11.6 | 11.52 | 11.37 | 2
0 | 18 | 18.9 | 0.9 | GRADE A |  | 0 | 8.6 | 8.6 | 0 | GRADE A | 2.5 | 13.65 | 13.3 | 13.3 | 13.1 | 13.1 | 12.95 | 12.8 | 12.65 | 12.5 | 12.5 | 12.35 | 12.2 | 12.05 | 11.9 | 11.9 | 11.75 | 11.6 | 11.45 | 11.3 | 11.2 | 2.5
0 | 18.5 | 18.8 | 0.3 | GRADE A |  | 0 | 9 | 8.7 | 0.3 | GRADE A | 3 | 13.5 | 13.1 | 13.1 | 12.95 | 12.95 | 12.8 | 12.65 | 12.5 | 12.35 | 12.35 | 12.2 | 12.05 | 11.9 | 11.75 | 11.75 | 11.6 | 11.45 | 11.3 | 11.2 | 11.1 | 3
0 | 17.8 | 18.2 | 0.4 | GRADE A |  | 0 | 8.8 | 8.6 | 0.2 | GRADE A | 3.5 | 13.3 | 12.95 | 12.95 | 12.8 | 12.75 | 12.65 | 12.5 | 12.35 | 12.2 | 12.2 | 12.05 | 11.9 | 11.75 | 11.6 | 11.52 | 11.45 | 11.3 | 11.2 | 11.1 | 10.95 | 3.5
0 | 18.3 | 18.2 | 0.1 | GRADE A |  | 0 | 8.6 | 8.2 | 0.4 | GRADE B | 4 | 13.1 | 12.8 | 12.8 | 12.65 | 12.5 | 12.5 | 12.35 | 12.2 | 12.05 | 12.05 | 11.9 | 11.75 | 11.6 | 11.45 | 11.3 | 11.2 | 11.2 | 11.1 | 10.95 | 10.8 | 4
0 | 0 |  | 0 | GRADE A |  | 0 | 0 |  | 0 | GRADE A | 4.5 | 12.95 | 12.65 | 12.65 | 12.5 | 12.35 | 12.27 | 12.2 | 12.05 | 11.9 | 11.83 | 11.68 | 11.53 | 11.45 | 11.3 | 11.2 | 11.1 | 11.02 | 10.87 | 10.8 | 10.7 | 4.5
0 | 0 |  | 0 | GRADE A |  | 0 | 0 |  | 0 | GRADE A | 5 | 12.8 | 12.5 | 12.5 | 12.35 | 12.2 | 12.05 | 12.05 | 11.9 | 11.75 | 11.6 | 11.45 | 11.3 | 11.3 | 11.2 | 11.1 | 10.95 | 10.8 | 10.7 | 10.7 | 10.6 | 5
| Criteria: | [5 empty] | Criteria: |  |  |  | 5.5 | 12.65 | 12.35 | 12.35 | 12.2 | 12.05 | 11.9 | 11.82 | 11.68 | 11.6 | 11.45 | 11.3 | 11.2 | 11.15 | 11.1 | 10.95 | 10.8 | 10.7 | 10.6 | 10.55 | 10.45 | 5.5
| GRADE A | GRADE B | FAIL |  |  |  | GRADE A | GRADE B | FAIL |  | 6 | 12.5 | 12.2 | 12.2 | 12.05 | 11.9 | 11.75 | 11.6 | 11.45 | 11.45 | 11.3 | 11.2 | 11.1 | 10.95 | 10.95 | 10.8 | 10.7 | 10.6 | 10.5 | 10.4 | 10.3 | 6
| =<1.5 | 1.51 - 2.00 | >2.01 |  |  |  | =<0.3 | 0.31 - 0.50 | >0.50 |  | 6.5 | 12.35 | 12.05 | 12.05 | 11.9 | 11.75 | 11.6 | 11.45 | 11.3 | 11.25 | 11.2 | 11.1 | 10.95 | 10.8 | 10.75 | 10.7 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 6.5
[11 empty] | 7 | 12.2 | 11.9 | 11.9 | 11.75 | 11.6 | 11.45 | 11.3 | 11.2 | 11.2 | 11.1 | 10.95 | 10.8 | 10.7 | 10.6 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 10.1 | 7
| CONDUCTIVITY AUDIT RESULTS | [5 empty] | DO AUDIT RESULTS |  |  |  | 7.5 | 12.05 | 11.75 | 11.75 | 11.6 | 11.45 | 11.3 | 11.2 | 11.1 | 11.1 | 10.95 | 10.8 | 10.7 | 10.6 | 10.5 | 10.45 | 10.35 | 10.25 | 10.15 | 10.05 | 9.92 | 7.5
[11 empty] | 8 | 11.9 | 11.6 | 11.6 | 11.45 | 11.3 | 11.2 | 11.1 | 10.95 | 10.95 | 10.8 | 10.7 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 8
# | Audit | DS Value | Abs % Difference | Status |  | # | Audit | DS Value | Abs. Difference | Status | 8.5 | 11.75 | 11.45 | 11.45 | 11.3 | 11.2 | 11.1 | 10.95 | 10.8 | 10.75 | 10.7 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 8.5 | 0:00 | 0
0 | 48 | 39.5 | 17.7083 | FAIL |  | 0 | 7.1 | 8.8 | 1.7 | GRADE C | 9 | 11.6 | 11.3 | 11.3 | 11.2 | 11.1 | 10.95 | 10.8 | 10.7 | 10.6 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9 | 0:15 | 0.010417
0 | 29 | 41.5 | 43.1034 | FAIL |  | 0 | 7.7 | 8.3 | 0.6 | GRADE B | 9.5 | 11.45 | 11.2 | 11.2 | 11.1 | 10.95 | 10.8 | 10.7 | 10.6 | 10.5 | 10.45 | 10.35 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.5 | 0:30 | 0.020833
0 | 50 | 40.5 | 19 | FAIL |  | 0 | 7.6 | 8.3 | 0.7 | GRADE B | 10 | 11.3 | 11.1 | 11.1 | 10.95 | 10.8 | 10.7 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.4 | 10 | 0:45 | 0.03125
0 | 53 | 40.5 | 23.5849 | FAIL |  | 0 | 8.8 | 8.5 | 0.3 | GRADE A | 10.5 | 11.2 | 10.95 | 10.95 | 10.8 | 10.7 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 10.1 | 10.05 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.35 | 9.25 | 10.5 | 1:00 | 0.041667
0 | 48 | 30.5 | 36.4583 | FAIL |  | 0 | 7.9 | 8.5 | 0.6 | GRADE B | 11 | 11.1 | 10.8 | 10.8 | 10.7 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.4 | 9.2 | 9.1 | 11 | 1:15 | 0.052083
0 | 0 |  | #DIV/0! | #DIV/0! |  | 0 | 0 |  | 0 | GRADE A | 11.5 | 10.95 | 10.7 | 10.7 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.65 | 9.55 | 9.47 | 9.45 | 9.35 | 9.25 | 9.1 | 9 | 11.5 | 1:30 | 0.0625
0 | 0 |  | #DIV/0! | #DIV/0! |  | 0 | 0 |  | 0 | GRADE A | 12 | 10.8 | 10.6 | 10.6 | 10.5 | 10.4 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.4 | 9.4 | 9.2 | 9.1 | 9 | 8.9 | 12 | 1:45 | 0.072917
| Criteria: | [5 empty] | Criteria: |  |  |  | 12.5 | 10.7 | 10.5 | 10.5 | 10.4 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.4 | 9.3 | 9.25 | 9.1 | 9 | 8.9 | 8.8 | 12.5 | 2:00 | 0.083333
| GRADE A | GRADE B | FAIL |  |  |  | GRADE A | GRADE B | GRADE C | FAIL | 13 | 10.6 | 10.4 | 10.4 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.4 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.75 | 13 | 2:15 | 0.09375
| =<10% | 10.1% - 15% | >15% |  |  |  | =<0.3 | .31-1.0 | >1.1-2.0 | >2.01 | 13.5 | 10.5 | 10.3 | 10.3 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.4 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.77 | 8.72 | 8.65 | 13.5 | 2:30 | 0.10417
| RUN COMMENTS: | Conductivity readings were off and average of 52.5 uS/cm per reading based on pre and post deployment data. Conductivity probes was difficult to calibrate and readings for this datasonde's conductivity probe have been recording elevated readings during previous deployments. A corrected reading graph has been included in this DS Report. | [8 empty] | 14 | 10.4 | 10.2 | 10.2 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.4 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.7 | 8.6 | 8.5 | 14 | 2:45 | 0.11458
[11 empty] | 14.5 | 10.3 | 10.1 | 10.1 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.4 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.75 | 8.6 | 8.5 | 8.45 | 14.5 | 3:00 | 0.125
A | AUDIT RESULTS: | [9 empty] | 15 | 10.2 | 10 | 10 | 9.85 | 9.7 | 9.6 | 9.5 | 9.45 | 9.4 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.75 | 8.7 | 8.5 | 8.45 | 8.4 | 15 | 3:15 | 0.13542
U | [10 empty] | 15.5 | 10.1 | 9.85 | 9.8 | 9.65 | 9.6 | 9.5 | 9.45 | 9.4 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.77 | 8.72 | 8.65 | 8.55 | 8.45 | 8.4 | 8.3 | 15.5 | 3:30 | 0.14583
D | [10 empty] | 16 | 10 | 9.7 | 9.6 | 9.5 | 9.5 | 9.45 | 9.4 | 9.25 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 16 | 3:45 | 0.15625
I | [10 empty] | 16.5 | 9.85 | 9.6 | 9.5 | 9.45 | 9.43 | 9.35 | 9.25 | 9.1 | 9.05 | 8.95 | 8.9 | 8.8 | 8.73 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.12 | 16.5 | 4:00 | 0.16667
T | STORET | DATE | TIME | TEMP | pH |  | CONDUCTIVITY | DO | DO SAT | COMMENTS | 17 | 9.7 | 9.5 | 9.45 | 9.4 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.8 | 8.75 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8 | 17 | 4:15 | 0.17708
# | # | (mm/dd/yyyy) | (hh:mm) | (deg C) | (SU) |  | (umhos/cm) | (mg/L) | % |  | 17.5 | 9.6 | 9.45 | 9.4 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.75 | 8.75 | 8.65 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.12 | 8.05 | 7.9 | 17.5 | 4:30 | 0.1875
| 25837 | 2001-07-16 | 14:27 | 19.8 | 8.58 |  | 48 | 7.1 | 90.45% |  | 18 | 9.5 | 9.4 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.75 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8.1 | 8 | 7.85 | 18 | 4:45 | 0.19792
| 25837 | 2001-07-17 | 11:40 | 18 | 8.6 |  | 29 | 7.7 | 94.48% |  | 18.5 | 9.45 | 9.3 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.75 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8.05 | 7.9 | 7.85 | 7.8 | 18.5 | 5:00 | 0.20833
| 25837 | 2001-07-18 | 15:00 | 18.5 | 9 |  | 50 | 7.6 | 94.41% |  | 19 | 9.4 | 9.2 | 9.1 | 9 | 8.9 | 8.8 | 8.75 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8.1 | 7.9 | 7.85 | 7.85 | 7.7 | 19 | 5:15 | 0.21875
| 25837 | 2001-07-19 | 10:26 | 17.8 | 8.8 |  | 53 | 8.8 | 107.3% |  | 19.5 | 9.3 | 9.1 | 9 | 8.9 | 8.8 | 8.75 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.17 | 8.12 | 8.05 | 7.95 | 7.85 | 7.8 | 7.7 | 7.6 | 19.5 | 5:30 | 0.22917
| 25837 | 2001-07-20 | 9:35 | 18.3 | 8.6 |  | 48 | 7.9 | 96.93% |  | 20 | 9.2 | 9 | 8.9 | 8.8 | 8.75 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.1 | 8.05 | 7.95 | 7.85 | 7.8 | 7.7 | 7.6 | 7.55 | 20 | 5:45 | 0.23958
| 25837 | [7 empty] | #VALUE! |  | 20.5 | 9.1 | 8.9 | 8.8 | 8.75 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8 | 8 | 7.8 | 7.8 | 7.7 | 7.6 | 7.55 | 7.5 | 20.5 | 6:00 | 0.25
| 25837 | [7 empty] | #VALUE! |  | 21 | 9 | 8.8 | 8.75 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8.1 | 7.9 | 7.95 | 7.7 | 7.7 | 7.6 | 7.55 | 7.5 | 7.45 | 21 | 6:15 | 0.26042
| 25837 | [7 empty] | #VALUE! |  | 21.5 | 8.9 | 8.75 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8.1 | 8 | 7.85 | 7.8 | 7.65 | 7.6 | 7.55 | 7.5 | 7.45 | 7.4 | 21.5 | 6:30 | 0.27083
[11 empty] | 22 | 8.8 | 8.7 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8.1 | 8 | 7.9 | 7.8 | 7.7 | 7.6 | 7.55 | 7.5 | 7.45 | 7.4 | 7.35 | 22 | 6:45 | 0.28125
[11 empty] | 22.5 | 8.75 | 8.6 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8.1 | 8 | 7.9 | 7.85 | 7.7 | 7.6 | 7.55 | 7.5 | 7.42 | 7.4 | 7.35 | 7.3 | 22.5 | 7:00 | 0.29167
[11 empty] | 23 | 8.7 | 8.5 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8.1 | 8 | 7.9 | 7.85 | 7.8 | 7.57 | 7.55 | 7.5 | 7.42 | 7.35 | 7.33 | 7.3 | 7.25 | 23 | 7:15 | 0.30208
[11 empty] | 23.5 | 8.6 | 8.45 | 8.4 | 8.3 | 8.2 | 8.15 | 8.1 | 8 | 7.9 | 7.85 | 7.8 | 7.7 | 7.5 | 7.5 | 7.42 | 7.35 | 7.3 | 7.25 | 7.22 | 7.2 | 23.5 | 7:30 | 0.3125
[11 empty] | 24 | 8.5 | 8.4 | 8.3 | 8.2 | 8.15 | 8.1 | 8 | 7.9 | 7.85 | 7.8 | 7.65 | 7.57 | 7.45 | 7.45 | 7.35 | 7.3 | 7.25 | 7.2 | 7.15 | 7.12 | 24 | 7:45 | 0.32292
[11 empty] | 24.5 | 8.45 | 8.3 | 8.2 | 8.15 | 8.1 | 8 | 7.9 | 7.85 | 7.8 | 7.7 | 7.55 | 7.5 | 7.4 | 7.4 | 7.3 | 7.25 | 7.2 | 7.15 | 7.1 | 7.05 | 24.5 | 8:00 | 0.33333
[11 empty] | 25 | 8.4 | 8.2 | 8.15 | 8.1 | 8 | 7.9 | 7.85 | 7.8 | 7.7 | 7.6 | 7.5 | 7.45 | 7.35 | 7.35 | 7.25 | 7.2 | 7.15 | 7.1 | 7.03 | 6.97 | 25 | 8:15 | 0.34375
[11 empty] | 25.5 | 8.3 | 8.15 | 8.1 | 8 | 7.9 | 7.85 | 7.8 | 7.7 | 7.57 | 7.52 | 7.45 | 7.4 | 7.3 | 7.3 | 7.2 | 7.15 | 7.1 | 7.03 | 6.95 | 6.9 | 25.5 | 8:30 | 0.35417
[11 empty] | 26 | 8.2 | 8.1 | 8 | 7.9 | 7.85 | 7.8 | 7.7 | 7.6 | 7.5 | 7.45 | 7.4 | 7.35 | 7.25 | 7.22 | 7.15 | 7.1 | 7.03 | 6.95 | 6.9 | 6.85 | 26 | 8:45 | 0.36458
[11 empty] | 26.5 | 8.15 | 8 | 7.9 | 7.85 | 7.8 | 7.7 | 7.6 | 7.55 | 7.45 | 7.4 | 7.35 | 7.3 | 7.2 | 7.15 | 7.1 | 7.03 | 6.95 | 6.9 | 6.85 | 6.75 | 26.5 | 9:00 | 0.375
[11 empty] | 27 | 8.1 | 7.9 | 7.85 | 7.8 | 7.7 | 7.6 | 7.55 | 7.5 | 7.4 | 7.35 | 7.3 | 7.25 | 7.15 | 7.1 | 7.03 | 6.95 | 6.9 | 6.85 | 6.75 | 6.7 | 27 | 9:15 | 0.38542
[11 empty] | 27.5 | 8 | 7.85 | 7.8 | 7.7 | 7.6 | 7.55 | 7.5 | 7.47 | 7.35 | 7.3 | 7.25 | 7.2 | 7.1 | 7.05 | 6.95 | 6.9 | 6.85 | 6.77 | 6.7 | 6.65 | 27.5 | 9:30 | 0.39583
[11 empty] | 28 | 7.9 | 7.8 | 7.7 | 7.6 | 7.55 | 7.5 | 7.45 | 7.4 | 7.3 | 7.25 | 7.2 | 7.15 | 7.05 | 7 | 6.9 | 6.85 | 6.77 | 6.7 | 6.65 | 6.6 | 28 | 9:45 | 0.40625
[11 empty] | 28.5 | 7.85 | 7.7 | 7.6 | 7.55 | 7.5 | 7.43 | 7.4 | 7.35 | 7.25 | 7.2 | 7.15 | 7.1 | 7 | 6.95 | 6.85 | 6.8 | 6.72 | 6.65 | 6.6 | 6.55 | 28.5 | 10:00 | 0.41667
[11 empty] | 29 | 1: 7.8; 2: 7.6; 3: 7.55; 4: 7.5; 5: 7.45 | 29 | 10:15 | 0.42708
[11 empty] | 29.5 | 1: 7.7; 2: 7.55; 3: 7.5; 4: 7.45; 5: 7.4 | 29.5 | 10:30 | 0.4375
[11 empty] | 30 | 1: 7.6; 2: 7.5; 3: 7.45; 4: 7.4; 5: 7.35 | 30 | 10:45 | 0.44792
[11 empty] | 30.5 | 1: 7.55; 2: 7.45; 3: 7.4; 4: 7.35; 5: 7.3 | 30.5 | 11:00 | 0.45833
[11 empty] | 31 | 1: 7.5; 2: 7.4; 3: 7.35; 4: 7.3; 5: 7.25 | 31 | 11:15 | 0.46875
[11 empty] | 31.5 | 1: 7.45; 2: 7.35; 3: 7.3; 4: 7.25; 5: 7.2 | 31.5 | 11:30 | 0.47917
[11 empty] | 32 | 1: 7.4; 2: 7.3; 3: 7.25; 4: 7.2; 5: 7.15 | 32 | 11:45 | 0.48958
[11 empty] | 32.5 | 1: 7.35; 2: 7.25; 3: 7.2; 4: 7.15; 5: 7.1 | 32.5 | 12:00 | 0.5
[11 empty] | 33 | 1: 7.3; 2: 7.2; 3: 7.15; 4: 7.1; 5: 7.05 | 33 | 12:15 | 0.51042
[11 empty] | 33.5 | 1: 7.25; 2: 7.15; 3: 7.1; 4: 7.05 | 33.5 | 12:30 | 0.52083
[11 empty] | 34 | 1: 7.2; 2: 7.1; 3: 7.05 | 34 | 12:45 | 0.53125
[11 empty] | 34.5 | 1: 7.15; 2: 7.05 | 34.5 | 13:00 | 0.54167
[11 empty] | 35 | 1: 7.1 | 35 | 13:15 | 0.55208
[11 empty] | 35.5 | 1: 7.05 | 35.5 | 13:30 | 0.5625
[11 empty] | 36 | [20 empty] | 36 | 13:45 | 0.57292
[33 empty] | 14:00 | 0.58333
[32 empty] | 36.5 | 14:15 | 0.59375
[32 empty] | 37 | 14:30 | 0.60417
[33 empty] | 14:45 | 0.61458
[33 empty] | 15:00 | 0.625
[33 empty] | 15:15 | 0.63542
[33 empty] | 15:30 | 0.64583
[33 empty] | 15:45 | 0.65625
[33 empty] | 16:00 | 0.66667
[33 empty] | 16:15 | 0.67708
[33 empty] | 16:30 | 0.6875
[33 empty] | 16:45 | 0.69792
[33 empty] | 17:00 | 0.70833
[33 empty] | 17:15 | 0.71875
[33 empty] | 17:30 | 0.72917
[33 empty] | 17:45 | 0.73958
[33 empty] | 18:00 | 0.75
[33 empty] | 18:15 | 0.76042
[33 empty] | 18:30 | 0.77083
[33 empty] | 18:45 | 0.78125
[33 empty] | 19:00 | 0.79167
[33 empty] | 19:15 | 0.80208
[33 empty] | 19:30 | 0.8125
[33 empty] | 19:45 | 0.82292
[33 empty] | 20:00 | 0.83333
[33 empty] | 20:15 | 0.84375
[33 empty] | 20:30 | 0.85417
[33 empty] | 20:45 | 0.86458
[33 empty] | 21:00 | 0.875
[33 empty] | 21:15 | 0.88542
[33 empty] | 21:30 | 0.89583
[33 empty] | 21:45 | 0.90625
[33 empty] | 22:00 | 0.91667
[33 empty] | 22:15 | 0.92708
[33 empty] | 22:30 | 0.9375
[33 empty] | 22:45 | 0.94792
[33 empty] | 23:00 | 0.95833
[33 empty] | 23:15 | 0.96875
[33 empty] | 23:30 | 0.97917
[33 empty] | 23:45 | 0.98958

## audit 4
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-17 11:40 | 18 | 8.6 | 7.7 | 94 | 29
Sonde Value | 2001-07-17 11:31 | 18.9 | 8.6 | 8.3 | 103 | 42
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-19 10:26 | 17.8 | 8.8 | 8.8 | 107 | 53
Sonde Value | 2001-07-19 10:31 | 18.2 | 8.6 | 8.5 | 104 | 41
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | Do Sat
|  | °C |  | uS/cm | mg/L | %
2001-07-16 | 14:31 | 19.87 | 8.58 | 39.5 | 8.81 | 112.2%
2001-07-16 | 14:46 | 19.93 | 8.67 | 39.5 | 8.59 | 109.4%
2001-07-16 | 15:01 | 19.94 | 8.67 | 40.5 | 8.43 | 107.4%
2001-07-16 | 15:16 | 19.94 | 8.67 | 40.5 | 8.35 | 106.4%
2001-07-16 | 15:31 | 19.91 | 8.68 | 40.5 | 8.29 | 105.6%
2001-07-16 | 15:46 | 19.82 | 8.68 | 40.5 | 8.23 | 104.8%
2001-07-16 | 16:01 | 19.67 | 8.65 | 40.5 | 8.18 | 104.2%
2001-07-16 | 16:16 | 19.63 | 8.65 | 40.5 | 8.15 | 103.8%
2001-07-16 | 16:31 | 19.69 | 8.67 | 40.5 | 8.11 | 103.3%
2001-07-16 | 16:46 | 19.68 | 8.69 | 40.5 | 8.06 | 102.7%
2001-07-16 | 17:01 | 19.65 | 8.69 | 40.5 | 8.05 | 102.5%
2001-07-16 | 17:16 | 19.63 | 8.69 | 40.5 | 8.04 | 102.4%
2001-07-16 | 17:31 | 19.59 | 8.67 | 40.5 | 8.01 | 102%
2001-07-16 | 17:46 | 19.56 | 8.69 | 40.5 | 8 | 101.9%
2001-07-16 | 18:01 | 19.55 | 8.69 | 40.5 | 7.98 | 101.7%
2001-07-16 | 18:16 | 19.54 | 8.7 | 40.5 | 7.96 | 101.4%
2001-07-16 | 18:31 | 19.53 | 8.7 | 40.5 | 7.96 | 101.4%
2001-07-16 | 18:46 | 19.52 | 8.71 | 40.5 | 7.96 | 101.4%
2001-07-16 | 19:01 | 19.52 | 8.72 | 40.5 | 7.95 | 101.3%
2001-07-16 | 19:16 | 19.52 | 8.72 | 40.5 | 7.96 | 101.4%
2001-07-16 | 19:31 | 19.49 | 8.73 | 40.5 | 7.94 | 100.5%
2001-07-16 | 19:46 | 19.48 | 8.73 | 40.5 | 7.94 | 100.5%
2001-07-16 | 20:01 | 19.49 | 8.74 | 40.5 | 7.93 | 100.4%
2001-07-16 | 20:16 | 19.5 | 8.74 | 40.5 | 7.92 | 100.9%
2001-07-16 | 20:31 | 19.51 | 8.74 | 40.5 | 7.91 | 100.8%
2001-07-16 | 20:46 | 19.49 | 8.73 | 40.5 | 7.91 | 100.1%
2001-07-16 | 21:01 | 19.47 | 8.73 | 40.5 | 7.91 | 100.1%
2001-07-16 | 21:16 | 19.46 | 8.73 | 40.5 | 7.9 | 100%
2001-07-16 | 21:31 | 19.45 | 8.73 | 40.5 | 7.91 | 100.1%
2001-07-16 | 21:46 | 19.43 | 8.72 | 40.5 | 7.91 | 100.1%
2001-07-16 | 22:01 | 19.41 | 8.72 | 40.5 | 7.91 | 100.1%
2001-07-16 | 22:16 | 19.4 | 8.72 | 40.5 | 7.91 | 100.1%
2001-07-16 | 22:31 | 19.4 | 8.72 | 40.5 | 7.91 | 100.1%
2001-07-16 | 22:46 | 19.38 | 8.72 | 40.5 | 7.91 | 100.1%
2001-07-16 | 23:01 | 19.36 | 8.72 | 40.5 | 7.92 | 100.3%
2001-07-16 | 23:16 | 19.34 | 8.71 | 40.5 | 7.93 | 100.4%
2001-07-16 | 23:31 | 19.29 | 8.71 | 40.5 | 7.93 | 100.4%
2001-07-16 | 23:46 | 19.26 | 8.71 | 40.5 | 7.93 | 100.4%
2001-07-17 | 0:01 | 19.23 | 8.71 | 40.5 | 7.94 | 100.5%
2001-07-17 | 0:16 | 19.21 | 8.7 | 40.5 | 7.94 | 100.5%
2001-07-17 | 0:31 | 19.19 | 8.7 | 40.5 | 7.95 | 100.6%
2001-07-17 | 0:46 | 19.18 | 8.69 | 40.5 | 7.95 | 100.6%
2001-07-17 | 1:01 | 19.15 | 8.68 | 40.5 | 7.95 | 100.6%
2001-07-17 | 1:16 | 19.12 | 8.67 | 40.5 | 7.96 | 100.8%
2001-07-17 | 1:31 | 19.09 | 8.67 | 40.5 | 7.96 | 100.8%
2001-07-17 | 1:46 | 19.06 | 8.66 | 40.5 | 7.96 | 100.8%
2001-07-17 | 2:01 | 19.04 | 8.65 | 40.5 | 7.96 | 100.8%
2001-07-17 | 2:16 | 19 | 8.65 | 40.5 | 7.97 | 100.9%
2001-07-17 | 2:31 | 18.97 | 8.64 | 40.5 | 7.98 | 99.13%
2001-07-17 | 2:46 | 18.93 | 8.63 | 41.5 | 7.98 | 99.13%
2001-07-17 | 3:01 | 18.91 | 8.62 | 40.5 | 7.99 | 99.25%
2001-07-17 | 3:16 | 18.89 | 8.61 | 41.5 | 7.99 | 99.25%
2001-07-17 | 3:31 | 18.87 | 8.59 | 41.5 | 7.99 | 99.25%
2001-07-17 | 3:46 | 18.85 | 8.58 | 41.5 | 7.99 | 99.25%
2001-07-17 | 4:01 | 18.83 | 8.57 | 40.5 | 8 | 99.38%
2001-07-17 | 4:16 | 18.8 | 8.56 | 41.5 | 8 | 99.38%
2001-07-17 | 4:31 | 18.78 | 8.54 | 41.5 | 8.01 | 99.5%
2001-07-17 | 4:46 | 18.75 | 8.54 | 41.5 | 8.01 | 99.5%
2001-07-17 | 5:01 | 18.72 | 8.54 | 42.5 | 8.02 | 99.63%
2001-07-17 | 5:16 | 18.68 | 8.54 | 41.5 | 8.02 | 99.63%
2001-07-17 | 5:31 | 18.66 | 8.53 | 41.5 | 8.03 | 99.75%
2001-07-17 | 5:46 | 18.65 | 8.53 | 41.5 | 8.03 | 99.75%
2001-07-17 | 6:01 | 18.63 | 8.52 | 41.5 | 8.04 | 99.88%
2001-07-17 | 6:16 | 18.61 | 8.52 | 41.5 | 8.04 | 99.88%
2001-07-17 | 6:31 | 18.6 | 8.52 | 41.5 | 8.05 | 100%
2001-07-17 | 6:46 | 18.58 | 8.52 | 41.5 | 8.06 | 100.1%
2001-07-17 | 7:01 | 18.56 | 8.52 | 41.5 | 8.07 | 100.2%
2001-07-17 | 7:16 | 18.53 | 8.52 | 41.5 | 8.08 | 100.4%
2001-07-17 | 7:31 | 18.5 | 8.52 | 41.5 | 8.09 | 100.5%
2001-07-17 | 7:46 | 18.48 | 8.51 | 42.5 | 8.1 | 99.39%
2001-07-17 | 8:01 | 18.48 | 8.53 | 41.5 | 8.12 | 99.63%
2001-07-17 | 8:16 | 18.48 | 8.55 | 42.5 | 8.12 | 99.63%
2001-07-17 | 8:31 | 18.46 | 8.54 | 41.5 | 8.13 | 99.75%
2001-07-17 | 8:46 | 18.47 | 8.54 | 42.5 | 8.16 | 100.1%
2001-07-17 | 9:01 | 18.5 | 8.55 | 41.5 | 8.17 | 101.5%
2001-07-17 | 9:16 | 18.51 | 8.54 | 42.5 | 8.15 | 101.2%
2001-07-17 | 9:31 | 18.53 | 8.54 | 42.5 | 8.16 | 101.4%
2001-07-17 | 9:46 | 18.52 | 8.52 | 42.5 | 8.17 | 101.5%
2001-07-17 | 10:01 | 18.63 | 8.55 | 42.5 | 8.18 | 101.6%
2001-07-17 | 10:16 | 18.63 | 8.55 | 42.5 | 8.16 | 101.4%
2001-07-17 | 10:31 | 18.63 | 8.55 | 42.5 | 8.19 | 101.7%
2001-07-17 | 10:46 | 18.73 | 8.56 | 41.5 | 8.28 | 102.9%
2001-07-17 | 11:01 | 18.85 | 8.59 | 42.5 | 8.33 | 103.5%
2001-07-17 | 11:16 | 18.88 | 8.58 | 42.5 | 8.31 | 103.2%
2001-07-17 | 11:31 | 18.89 | 8.57 | 41.5 | 8.28 | 102.9%
2001-07-17 | 11:46 | 18.92 | 8.58 | 41.5 | 8.24 | 102.4%
2001-07-17 | 12:01 | 18.91 | 8.59 | 42.5 | 8.25 | 102.5%
2001-07-17 | 12:16 | 19.08 | 8.62 | 41.5 | 8.24 | 104.3%
2001-07-17 | 12:31 | 19.2 | 8.64 | 41.5 | 8.26 | 104.6%
2001-07-17 | 12:46 | 19.15 | 8.62 | 42.5 | 8.24 | 104.3%
2001-07-17 | 13:01 | 19.2 | 8.63 | 42.5 | 8.24 | 104.3%
2001-07-17 | 13:16 | 19.25 | 8.64 | 41.5 | 8.2 | 103.8%
2001-07-17 | 13:31 | 19.26 | 8.64 | 41.5 | 8.23 | 104.2%
2001-07-17 | 13:46 | 19.27 | 8.64 | 41.5 | 8.21 | 103.9%
2001-07-17 | 14:01 | 19.29 | 8.65 | 41.5 | 8.2 | 103.8%
2001-07-17 | 14:16 | 19.3 | 8.66 | 41.5 | 8.21 | 103.9%
2001-07-17 | 14:31 | 19.27 | 8.66 | 41.5 | 8.21 | 103.9%
2001-07-17 | 14:46 | 19.3 | 8.66 | 41.5 | 8.2 | 103.8%
2001-07-17 | 15:01 | 19.29 | 8.67 | 41.5 | 8.16 | 103.3%
2001-07-17 | 15:16 | 19.31 | 8.68 | 41.5 | 8.18 | 103.5%
2001-07-17 | 15:31 | 19.32 | 8.69 | 41.5 | 8.15 | 103.2%
2001-07-17 | 15:46 | 19.41 | 8.71 | 41.5 | 8.19 | 103.7%
2001-07-17 | 16:01 | 19.4 | 8.71 | 41.5 | 8.15 | 103.2%
2001-07-17 | 16:16 | 19.4 | 8.71 | 42.5 | 8.15 | 103.2%
2001-07-17 | 16:31 | 19.38 | 8.71 | 41.5 | 8.09 | 102.4%
2001-07-17 | 16:46 | 19.34 | 8.69 | 41.5 | 8.09 | 102.4%
2001-07-17 | 17:01 | 19.34 | 8.69 | 41.5 | 8.06 | 102%
2001-07-17 | 17:16 | 19.32 | 8.67 | 41.5 | 8.04 | 101.8%
2001-07-17 | 17:31 | 19.32 | 8.67 | 41.5 | 8.03 | 101.6%
2001-07-17 | 17:46 | 19.31 | 8.67 | 41.5 | 8.03 | 101.6%
2001-07-17 | 18:01 | 19.3 | 8.67 | 41.5 | 8.02 | 101.5%
2001-07-17 | 18:16 | 19.3 | 8.67 | 41.5 | 8.02 | 101.5%
2001-07-17 | 18:31 | 19.28 | 8.67 | 41.5 | 8.02 | 101.5%
2001-07-17 | 18:46 | 19.28 | 8.67 | 41.5 | 8.02 | 101.5%
2001-07-17 | 19:01 | 19.27 | 8.66 | 41.5 | 8.04 | 101.8%
2001-07-17 | 19:16 | 19.27 | 8.68 | 41.5 | 8.03 | 101.6%
2001-07-17 | 19:31 | 19.25 | 8.68 | 41.5 | 8.02 | 101.5%
2001-07-17 | 19:46 | 19.23 | 8.67 | 41.5 | 8.01 | 101.4%
2001-07-17 | 20:01 | 19.21 | 8.66 | 41.5 | 8 | 101.3%
2001-07-17 | 20:16 | 19.18 | 8.65 | 42.5 | 8 | 101.3%
2001-07-17 | 20:31 | 19.17 | 8.64 | 42.5 | 8 | 101.3%
2001-07-17 | 20:46 | 19.15 | 8.63 | 41.5 | 8 | 101.3%
2001-07-17 | 21:01 | 19.12 | 8.62 | 42.5 | 8 | 101.3%
2001-07-17 | 21:16 | 19.11 | 8.62 | 42.5 | 8 | 101.3%
2001-07-17 | 21:31 | 19.1 | 8.61 | 41.5 | 8 | 101.3%
2001-07-17 | 21:46 | 19.08 | 8.61 | 41.5 | 8.01 | 101.4%
2001-07-17 | 22:01 | 19.05 | 8.61 | 41.5 | 8 | 101.3%
2001-07-17 | 22:16 | 19.04 | 8.6 | 42.5 | 8.01 | 101.4%
2001-07-17 | 22:31 | 19.02 | 8.59 | 42.5 | 8.01 | 101.4%
2001-07-17 | 22:46 | 19 | 8.59 | 42.5 | 8.02 | 101.5%
2001-07-17 | 23:01 | 18.98 | 8.58 | 42.5 | 8.03 | 99.75%
2001-07-17 | 23:16 | 18.96 | 8.56 | 40.5 | 8.03 | 99.75%
2001-07-17 | 23:31 | 18.94 | 8.55 | 40.5 | 8.03 | 99.75%
2001-07-17 | 23:46 | 18.93 | 8.55 | 40.5 | 8.03 | 99.75%
2001-07-18 | 0:01 | 18.91 | 8.54 | 40.5 | 8.03 | 99.75%
2001-07-18 | 0:16 | 18.89 | 8.54 | 40.5 | 8.03 | 99.75%
2001-07-18 | 0:31 | 18.87 | 8.53 | 40.5 | 8.04 | 99.88%
2001-07-18 | 0:46 | 18.85 | 8.53 | 40.5 | 8.04 | 99.88%
2001-07-18 | 1:01 | 18.84 | 8.53 | 40.5 | 8.04 | 99.88%
2001-07-18 | 1:16 | 18.83 | 8.53 | 40.5 | 8.04 | 99.88%
2001-07-18 | 1:31 | 18.81 | 8.52 | 40.5 | 8.04 | 99.88%
2001-07-18 | 1:46 | 18.79 | 8.53 | 40.5 | 8.04 | 99.88%
2001-07-18 | 2:01 | 18.78 | 8.53 | 40.5 | 8.05 | 100%
2001-07-18 | 2:16 | 18.76 | 8.55 | 40.5 | 8.05 | 100%
2001-07-18 | 2:31 | 18.75 | 8.54 | 40.5 | 8.05 | 100%
2001-07-18 | 2:46 | 18.74 | 8.55 | 40.5 | 8.05 | 100%
2001-07-18 | 3:01 | 18.73 | 8.55 | 40.5 | 8.05 | 100%
2001-07-18 | 3:16 | 18.72 | 8.57 | 40.5 | 8.05 | 100%
2001-07-18 | 3:31 | 18.7 | 8.57 | 40.5 | 8.06 | 100.1%
2001-07-18 | 3:46 | 18.69 | 8.57 | 40.5 | 8.06 | 100.1%
2001-07-18 | 4:01 | 18.67 | 8.57 | 40.5 | 8.06 | 100.1%
2001-07-18 | 4:16 | 18.65 | 8.57 | 40.5 | 8.06 | 100.1%
2001-07-18 | 4:31 | 18.63 | 8.57 | 40.5 | 8.07 | 100.2%
2001-07-18 | 4:46 | 18.61 | 8.56 | 40.5 | 8.07 | 100.2%
2001-07-18 | 5:01 | 18.58 | 8.55 | 40.5 | 8.08 | 100.4%
2001-07-18 | 5:16 | 18.56 | 8.54 | 40.5 | 8.08 | 100.4%
2001-07-18 | 5:31 | 18.53 | 8.53 | 40.5 | 8.09 | 100.5%
2001-07-18 | 5:46 | 18.5 | 8.51 | 40.5 | 8.09 | 100.5%
2001-07-18 | 6:01 | 18.47 | 8.48 | 40.5 | 8.1 | 99.39%
2001-07-18 | 6:16 | 18.45 | 8.47 | 40.5 | 8.11 | 99.51%
2001-07-18 | 6:31 | 18.42 | 8.47 | 40.5 | 8.12 | 99.63%
2001-07-18 | 6:46 | 18.4 | 8.45 | 40.5 | 8.13 | 99.75%
2001-07-18 | 7:01 | 18.38 | 8.45 | 40.5 | 8.15 | 100%
2001-07-18 | 7:16 | 18.37 | 8.46 | 40.5 | 8.17 | 100.2%
2001-07-18 | 7:31 | 18.36 | 8.48 | 40.5 | 8.19 | 100.5%
2001-07-18 | 7:46 | 18.33 | 8.48 | 40.5 | 8.18 | 100.4%
2001-07-18 | 8:01 | 18.31 | 8.49 | 40.5 | 8.22 | 100.9%
2001-07-18 | 8:16 | 18.33 | 8.51 | 40.5 | 8.24 | 101.1%
2001-07-18 | 8:31 | 18.31 | 8.52 | 40.5 | 8.24 | 101.1%
2001-07-18 | 8:46 | 18.3 | 8.54 | 40.5 | 8.26 | 101.3%
2001-07-18 | 9:01 | 18.31 | 8.56 | 40.5 | 8.29 | 101.7%
2001-07-18 | 9:16 | 18.33 | 8.59 | 40.5 | 8.31 | 102%
2001-07-18 | 9:31 | 18.34 | 8.6 | 40.5 | 8.32 | 102.1%
2001-07-18 | 9:46 | 18.34 | 8.61 | 40.5 | 8.32 | 102.1%
2001-07-18 | 10:01 | 18.31 | 8.61 | 40.5 | 8.34 | 102.3%
2001-07-18 | 10:16 | 18.29 | 8.6 | 40.5 | 8.34 | 102.3%
2001-07-18 | 10:31 | 18.29 | 8.6 | 40.5 | 8.32 | 102.1%
2001-07-18 | 10:46 | 18.3 | 8.61 | 40.5 | 8.32 | 102.1%
2001-07-18 | 11:01 | 18.3 | 8.61 | 40.5 | 8.33 | 102.2%
2001-07-18 | 11:16 | 18.3 | 8.62 | 40.5 | 8.35 | 102.5%
2001-07-18 | 11:31 | 18.37 | 8.63 | 40.5 | 8.4 | 103.1%
2001-07-18 | 11:46 | 18.43 | 8.65 | 40.5 | 8.42 | 103.3%
2001-07-18 | 12:01 | 18.4 | 8.64 | 40.5 | 8.37 | 102.7%
2001-07-18 | 12:16 | 18.4 | 8.65 | 40.5 | 8.37 | 102.7%
2001-07-18 | 12:31 | 18.43 | 8.65 | 40.5 | 8.41 | 103.2%
2001-07-18 | 12:46 | 18.51 | 8.66 | 40.5 | 8.37 | 104%
2001-07-18 | 13:01 | 18.5 | 8.66 | 40.5 | 8.38 | 104.1%
2001-07-18 | 13:16 | 18.43 | 8.64 | 40.5 | 8.36 | 102.6%
2001-07-18 | 13:31 | 18.61 | 8.66 | 40.5 | 8.35 | 103.7%
2001-07-18 | 13:46 | 18.78 | 8.68 | 40.5 | 8.37 | 104%
2001-07-18 | 14:01 | 18.83 | 8.69 | 40.5 | 8.35 | 103.7%
2001-07-18 | 14:16 | 18.83 | 8.7 | 40.5 | 8.32 | 103.4%
2001-07-18 | 14:31 | 18.84 | 8.7 | 40.5 | 8.34 | 103.6%
2001-07-18 | 14:46 | 18.85 | 8.69 | 40.5 | 8.33 | 103.5%
2001-07-18 | 15:01 | 18.83 | 8.7 | 40.5 | 8.33 | 103.5%
2001-07-18 | 15:16 | 18.84 | 8.71 | 40.5 | 8.31 | 103.2%
2001-07-18 | 15:31 | 18.8 | 8.7 | 40.5 | 8.29 | 103%
2001-07-18 | 15:46 | 18.78 | 8.7 | 40.5 | 8.27 | 102.7%
2001-07-18 | 16:01 | 18.77 | 8.7 | 40.5 | 8.25 | 102.5%
2001-07-18 | 16:16 | 18.78 | 8.71 | 40.5 | 8.25 | 102.5%
2001-07-18 | 16:31 | 18.78 | 8.72 | 40.5 | 8.24 | 102.4%
2001-07-18 | 16:46 | 18.79 | 8.71 | 40.5 | 8.22 | 102.1%
2001-07-18 | 17:01 | 18.79 | 8.72 | 40.5 | 8.2 | 101.9%
2001-07-18 | 17:16 | 18.77 | 8.71 | 40.5 | 8.18 | 101.6%
2001-07-18 | 17:31 | 18.78 | 8.71 | 40.5 | 8.16 | 101.4%
2001-07-18 | 17:46 | 18.78 | 8.71 | 40.5 | 8.15 | 101.2%
2001-07-18 | 18:01 | 18.76 | 8.71 | 40.5 | 8.14 | 101.1%
2001-07-18 | 18:16 | 18.78 | 8.72 | 40.5 | 8.13 | 101%
2001-07-18 | 18:31 | 18.79 | 8.72 | 40.5 | 8.13 | 101%
2001-07-18 | 18:46 | 18.8 | 8.72 | 40.5 | 8.11 | 100.7%
2001-07-18 | 19:01 | 18.8 | 8.73 | 40.5 | 8.1 | 100.6%
2001-07-18 | 19:16 | 18.8 | 8.71 | 40.5 | 8.09 | 100.5%
2001-07-18 | 19:31 | 18.8 | 8.72 | 40.5 | 8.09 | 100.5%
2001-07-18 | 19:46 | 18.8 | 8.74 | 40.5 | 8.09 | 100.5%
2001-07-18 | 20:01 | 18.81 | 8.73 | 40.5 | 8.08 | 100.4%
2001-07-18 | 20:16 | 18.81 | 8.72 | 40.5 | 8.08 | 100.4%
2001-07-18 | 20:31 | 18.81 | 8.72 | 40.5 | 8.07 | 100.2%
2001-07-18 | 20:46 | 18.8 | 8.72 | 40.5 | 8.06 | 100.1%
2001-07-18 | 21:01 | 18.78 | 8.72 | 40.5 | 8.06 | 100.1%
2001-07-18 | 21:16 | 18.76 | 8.71 | 40.5 | 8.06 | 100.1%
2001-07-18 | 21:31 | 18.76 | 8.72 | 40.5 | 8.07 | 100.2%
2001-07-18 | 21:46 | 18.74 | 8.72 | 40.5 | 8.07 | 100.2%
2001-07-18 | 22:01 | 18.73 | 8.72 | 40.5 | 8.07 | 100.2%
2001-07-18 | 22:16 | 18.72 | 8.72 | 40.5 | 8.07 | 100.2%
2001-07-18 | 22:31 | 18.7 | 8.72 | 40.5 | 8.09 | 100.5%
2001-07-18 | 22:46 | 18.67 | 8.72 | 40.5 | 8.08 | 100.4%
2001-07-18 | 23:01 | 18.65 | 8.7 | 40.5 | 8.08 | 100.4%
2001-07-18 | 23:16 | 18.62 | 8.71 | 40.5 | 8.09 | 100.5%
2001-07-18 | 23:31 | 18.59 | 8.7 | 40.5 | 8.09 | 100.5%
2001-07-18 | 23:46 | 18.56 | 8.68 | 40.5 | 8.09 | 100.5%
2001-07-19 | 0:01 | 18.53 | 8.68 | 40.5 | 8.1 | 100.6%
2001-07-19 | 0:16 | 18.51 | 8.66 | 40.5 | 8.1 | 100.6%
2001-07-19 | 0:31 | 18.49 | 8.66 | 40.5 | 8.11 | 99.51%
2001-07-19 | 0:46 | 18.47 | 8.66 | 40.5 | 8.11 | 99.51%
2001-07-19 | 1:01 | 18.46 | 8.66 | 40.5 | 8.11 | 99.51%
2001-07-19 | 1:16 | 18.44 | 8.65 | 40.5 | 8.12 | 99.63%
2001-07-19 | 1:31 | 18.42 | 8.64 | 40.5 | 8.11 | 99.51%
2001-07-19 | 1:46 | 18.4 | 8.64 | 40.5 | 8.11 | 99.51%
2001-07-19 | 2:01 | 18.38 | 8.64 | 40.5 | 8.1 | 99.39%
2001-07-19 | 2:16 | 18.36 | 8.62 | 40.5 | 8.11 | 99.51%
2001-07-19 | 2:31 | 18.34 | 8.61 | 40.5 | 8.1 | 99.39%
2001-07-19 | 2:46 | 18.32 | 8.62 | 40.5 | 8.12 | 99.63%
2001-07-19 | 3:01 | 18.3 | 8.62 | 40.5 | 8.13 | 99.75%
2001-07-19 | 3:16 | 18.27 | 8.6 | 40.5 | 8.13 | 99.75%
2001-07-19 | 3:31 | 18.24 | 8.6 | 40.5 | 8.14 | 99.88%
2001-07-19 | 3:46 | 18.2 | 8.59 | 40.5 | 8.14 | 99.88%
2001-07-19 | 4:01 | 18.16 | 8.58 | 40.5 | 8.15 | 100%
2001-07-19 | 4:16 | 18.13 | 8.57 | 40.5 | 8.16 | 100.1%
2001-07-19 | 4:31 | 18.1 | 8.57 | 40.5 | 8.15 | 100%
2001-07-19 | 4:46 | 18.08 | 8.55 | 40.5 | 8.16 | 100.1%
2001-07-19 | 5:01 | 18.06 | 8.54 | 40.5 | 8.17 | 100.2%
2001-07-19 | 5:16 | 18.05 | 8.54 | 40.5 | 8.16 | 100.1%
2001-07-19 | 5:31 | 18.03 | 8.53 | 40.5 | 8.17 | 100.2%
2001-07-19 | 5:46 | 18 | 8.53 | 40.5 | 8.16 | 100.1%
2001-07-19 | 6:01 | 17.99 | 8.52 | 40.5 | 8.16 | 99.51%
2001-07-19 | 6:16 | 17.96 | 8.51 | 40.5 | 8.17 | 99.63%
2001-07-19 | 6:31 | 17.95 | 8.52 | 40.5 | 8.2 | 100%
2001-07-19 | 6:46 | 17.94 | 8.54 | 40.5 | 8.17 | 99.63%
2001-07-19 | 7:01 | 17.93 | 8.56 | 40.5 | 8.18 | 99.76%
2001-07-19 | 7:16 | 17.92 | 8.58 | 40.5 | 8.19 | 99.88%
2001-07-19 | 7:31 | 17.93 | 8.6 | 40.5 | 8.2 | 100%
2001-07-19 | 7:46 | 17.93 | 8.61 | 40.5 | 8.2 | 100%
2001-07-19 | 8:01 | 17.91 | 8.6 | 40.5 | 8.21 | 100.1%
2001-07-19 | 8:16 | 17.89 | 8.58 | 40.5 | 8.21 | 100.1%
2001-07-19 | 8:31 | 17.87 | 8.6 | 40.5 | 8.18 | 99.76%
2001-07-19 | 8:46 | 17.88 | 8.59 | 40.5 | 7.8 | 95.12%
2001-07-19 | 9:01 | 17.91 | 8.6 | 40.5 | 7.84 | 95.61%
2001-07-19 | 9:16 | 17.93 | 8.59 | 40.5 | 7.89 | 96.22%
2001-07-19 | 9:31 | 17.97 | 8.58 | 40.5 | 7.84 | 95.61%
2001-07-19 | 9:46 | 18 | 8.59 | 40.5 | 7.84 | 96.2%
2001-07-19 | 10:01 | 18.06 | 8.6 | 40.5 | 7.84 | 96.2%
2001-07-19 | 10:16 | 18.11 | 8.61 | 40.5 | 7.95 | 97.55%
2001-07-19 | 10:31 | 18.18 | 8.59 | 40.5 | 8.48 | 104%
2001-07-19 | 10:46 | 18.3 | 8.59 | 40.5 | 8.45 | 103.7%
2001-07-19 | 11:01 | 18.36 | 8.6 | 40.5 | 8.43 | 103.4%
2001-07-19 | 11:16 | 18.44 | 8.62 | 40.5 | 8.46 | 103.8%
2001-07-19 | 11:31 | 18.56 | 8.63 | 40.5 | 8.45 | 105%
2001-07-19 | 11:46 | 18.64 | 8.65 | 40.5 | 8.42 | 104.6%
2001-07-19 | 12:01 | 18.66 | 8.65 | 40.5 | 8.41 | 104.5%
2001-07-19 | 12:16 | 18.71 | 8.68 | 40.5 | 8.36 | 103.9%
2001-07-19 | 12:31 | 18.6 | 8.67 | 40.5 | 8.32 | 103.4%
2001-07-19 | 12:46 | 18.66 | 8.71 | 40.5 | 8.36 | 103.9%
2001-07-19 | 13:01 | 18.58 | 8.68 | 40.5 | 8.29 | 103%
2001-07-19 | 13:16 | 18.73 | 8.72 | 40.5 | 8.35 | 103.7%
2001-07-19 | 13:31 | 18.9 | 8.72 | 40.5 | 8.26 | 102.6%
2001-07-19 | 13:46 | 18.74 | 8.72 | 40.5 | 8.29 | 103%
2001-07-19 | 14:01 | 18.7 | 8.73 | 40.5 | 8.26 | 102.6%
2001-07-19 | 14:16 | 18.69 | 8.7 | 40.5 | 8.16 | 101.4%
2001-07-19 | 14:31 | 18.63 | 8.67 | 40.5 | 8.2 | 101.9%
2001-07-19 | 14:46 | 18.67 | 8.69 | 40.5 | 8.14 | 101.1%
2001-07-19 | 15:01 | 18.65 | 8.67 | 40.5 | 8.12 | 100.9%
2001-07-19 | 15:16 | 18.62 | 8.65 | 39.5 | 8.11 | 100.7%
2001-07-19 | 15:31 | 18.61 | 8.65 | 39.5 | 8.14 | 101.1%
2001-07-19 | 15:46 | 18.65 | 8.69 | 40.5 | 8.15 | 101.2%
2001-07-19 | 16:01 | 18.68 | 8.72 | 40.5 | 8.16 | 101.4%
2001-07-19 | 16:16 | 18.71 | 8.76 | 40.5 | 8.18 | 101.6%
2001-07-19 | 16:31 | 18.76 | 8.76 | 40.5 | 8.17 | 101.5%
2001-07-19 | 16:46 | 18.79 | 8.77 | 40.5 | 8.14 | 101.1%
2001-07-19 | 17:01 | 18.81 | 8.77 | 40.5 | 8.12 | 100.9%
2001-07-19 | 17:16 | 18.83 | 8.76 | 40.5 | 8.13 | 101%
2001-07-19 | 17:31 | 18.86 | 8.78 | 40.5 | 8.12 | 100.9%
2001-07-19 | 17:46 | 18.89 | 8.78 | 40.5 | 8.1 | 100.6%
2001-07-19 | 18:01 | 18.9 | 8.79 | 40.5 | 8.1 | 100.6%
2001-07-19 | 18:16 | 18.92 | 8.8 | 40.5 | 8.09 | 100.5%
2001-07-19 | 18:31 | 18.89 | 8.79 | 40.5 | 8.08 | 100.4%
2001-07-19 | 18:46 | 18.87 | 8.79 | 40.5 | 8.08 | 100.4%
2001-07-19 | 19:01 | 18.85 | 8.77 | 40.5 | 8.06 | 100.1%
2001-07-19 | 19:16 | 18.83 | 8.75 | 40.5 | 8.06 | 100.1%
2001-07-19 | 19:31 | 18.83 | 8.74 | 40.5 | 8.07 | 100.2%
2001-07-19 | 19:46 | 18.84 | 8.74 | 40.5 | 8.06 | 100.1%
2001-07-19 | 20:01 | 18.84 | 8.75 | 39.5 | 8.05 | 100%
2001-07-19 | 20:16 | 18.85 | 8.75 | 39.5 | 8.04 | 99.88%
2001-07-19 | 20:31 | 18.85 | 8.75 | 39.5 | 8.04 | 99.88%
2001-07-19 | 20:46 | 18.87 | 8.74 | 39.5 | 8.02 | 99.63%
2001-07-19 | 21:01 | 18.87 | 8.73 | 39.5 | 8.02 | 99.63%
2001-07-19 | 21:16 | 18.86 | 8.75 | 39.5 | 8.02 | 99.63%
2001-07-19 | 21:31 | 18.85 | 8.74 | 39.5 | 8.02 | 99.63%
2001-07-19 | 21:46 | 18.83 | 8.73 | 39.5 | 8.03 | 99.75%
2001-07-19 | 22:01 | 18.82 | 8.75 | 39.5 | 8.02 | 99.63%
2001-07-19 | 22:16 | 18.81 | 8.73 | 39.5 | 8.03 | 99.75%
2001-07-19 | 22:31 | 18.79 | 8.73 | 39.5 | 8.03 | 99.75%
2001-07-19 | 22:46 | 18.77 | 8.73 | 39.5 | 8.03 | 99.75%
2001-07-19 | 23:01 | 18.76 | 8.73 | 39.5 | 8.03 | 99.75%
2001-07-19 | 23:16 | 18.74 | 8.73 | 39.5 | 8.03 | 99.75%
2001-07-19 | 23:31 | 18.72 | 8.71 | 39.5 | 8.03 | 99.75%
2001-07-19 | 23:46 | 18.71 | 8.71 | 39.5 | 8.04 | 99.88%
2001-07-20 | 0:01 | 18.69 | 8.7 | 39.5 | 8.04 | 99.88%
2001-07-20 | 0:16 | 18.67 | 8.71 | 39.5 | 8.04 | 99.88%
2001-07-20 | 0:31 | 18.66 | 8.71 | 39.5 | 8.04 | 99.88%
2001-07-20 | 0:46 | 18.64 | 8.7 | 39.5 | 8.04 | 99.88%
2001-07-20 | 1:01 | 18.62 | 8.69 | 39.5 | 8.04 | 99.88%
2001-07-20 | 1:16 | 18.6 | 8.69 | 39.5 | 8.05 | 100%
2001-07-20 | 1:31 | 18.58 | 8.69 | 39.5 | 8.05 | 100%
2001-07-20 | 1:46 | 18.55 | 8.69 | 39.5 | 8.06 | 100.1%
2001-07-20 | 2:01 | 18.52 | 8.69 | 39.5 | 8.06 | 100.1%
2001-07-20 | 2:16 | 18.5 | 8.67 | 39.5 | 8.06 | 100.1%
2001-07-20 | 2:31 | 18.49 | 8.67 | 39.5 | 8.06 | 98.9%
2001-07-20 | 2:46 | 18.46 | 8.66 | 39.5 | 8.06 | 98.9%
2001-07-20 | 3:01 | 18.45 | 8.65 | 39.5 | 8.06 | 98.9%
2001-07-20 | 3:16 | 18.45 | 8.66 | 39.5 | 8.06 | 98.9%
2001-07-20 | 3:31 | 18.43 | 8.65 | 39.5 | 8.07 | 99.02%
2001-07-20 | 3:46 | 18.41 | 8.64 | 39.5 | 8.07 | 99.02%
2001-07-20 | 4:01 | 18.38 | 8.62 | 39.5 | 8.08 | 99.14%
2001-07-20 | 4:16 | 18.36 | 8.63 | 39.5 | 8.08 | 99.14%
2001-07-20 | 4:31 | 18.34 | 8.63 | 39.5 | 8.08 | 99.14%
2001-07-20 | 4:46 | 18.33 | 8.63 | 39.5 | 8.09 | 99.26%
2001-07-20 | 5:01 | 18.31 | 8.62 | 39.5 | 8.09 | 99.26%
2001-07-20 | 5:16 | 18.29 | 8.6 | 39.5 | 8.09 | 99.26%
2001-07-20 | 5:31 | 18.27 | 8.58 | 39.5 | 8.09 | 99.26%
2001-07-20 | 5:46 | 18.25 | 8.56 | 39.5 | 8.1 | 99.39%
2001-07-20 | 6:01 | 18.23 | 8.56 | 39.5 | 8.1 | 99.39%
2001-07-20 | 6:16 | 18.22 | 8.56 | 39.5 | 8.11 | 99.51%
2001-07-20 | 6:31 | 18.21 | 8.56 | 39.5 | 8.12 | 99.63%
2001-07-20 | 6:46 | 18.21 | 8.54 | 39.5 | 8.12 | 99.63%
2001-07-20 | 7:01 | 18.21 | 8.53 | 39.5 | 8.12 | 99.63%
2001-07-20 | 7:16 | 18.21 | 8.54 | 39.5 | 8.13 | 99.75%
2001-07-20 | 7:31 | 18.21 | 8.55 | 39.5 | 8.15 | 100%
2001-07-20 | 7:46 | 18.2 | 8.55 | 39.5 | 8.16 | 100.1%
2001-07-20 | 8:01 | 18.19 | 8.54 | 39.5 | 8.17 | 100.2%
2001-07-20 | 8:16 | 18.18 | 8.55 | 39.5 | 8.18 | 100.4%
2001-07-20 | 8:31 | 18.19 | 8.56 | 39.5 | 8.2 | 100.6%
2001-07-20 | 8:46 | 18.22 | 8.56 | 39.5 | 8.21 | 100.7%
2001-07-20 | 9:01 | 18.22 | 8.56 | 39.5 | 8.21 | 100.7%
2001-07-20 | 9:16 | 18.23 | 8.55 | 39.5 | 8.2 | 100.6%
2001-07-20 | 9:31 | 18.24 | 8.54 | 30.5 | 8.21 | 100.7%