## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2001-07-16 19:08 | 16.3 | 7.9 | 8.8 | 104 | 59.3 | 0
Sonde | 2001-07-16 09:15 | 16.23 | 7.56 | 9.66 | 114 | 53.1
Grade |  | A | B | B |  | B
Audit 2 | 2001-07-18 09:40 | 14.3 | 7.8 | 8.1 | 91 | 62 | 0
Sonde | 2001-07-18 09:45 | 14.51 | 7.69 | 8.46 | 96 | 52.9
Grade |  | A | A | B |  | B
Audit 3 | 2001-07-20 11:30 | 16.5 | 7.8 | 8.1 | 96 | 59 | 0
Sonde | 2001-07-20 11:30 | 15.04 | 7.68 | 8.52 | 97 | 52.4
Grade |  | A | A | B |  | B
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Formatting Note: Audits 1-3 represnet the duration of the deployment and are macro (ed) into the main audit sheet. Audits 4-5 represent the middle of the deployment and are independent of the main audit sheet. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-16 19:08 | 16.3 | 7.9 | 8.8 | 104 | 59.3
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-18 09:40 | 14.3 | 7.8 | 8.1 | 91 | 62
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-20 11:30 | 16.5 | 7.8 | 8.1 | 96 | 59
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-17 15:34 | 15 | 8 | 8.8 | 101 | 61
| 2001-07-17 15:30 | 15.01 | 7.81 | 9.44 | 108 | 53.3
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-19 15:30 | 15.3 | 8 | 8.8 | 101 | 61
Sonde Values | 2001-07-19 15:30 | 15.46 | 7.8 | 9.07 | 104 | 53
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | DO Sat
|  | °C |  | uS/cm | mg/L | %
2001-07-16 | 19:15 | 16.23 | 7.56 | 53.1 | 9.66 | 113.6%
2001-07-16 | 19:30 | 16.23 | 7.75 | 53 | 9.73 | 114.5%
2001-07-16 | 19:45 | 16.23 | 7.78 | 53 | 9.65 | 113.5%
2001-07-16 | 20:00 | 16.22 | 7.81 | 53 | 9.68 | 113.9%
2001-07-16 | 20:15 | 16.22 | 7.8 | 52.9 | 9.66 | 113.6%
2001-07-16 | 20:30 | 16.19 | 7.8 | 52.9 | 9.61 | 113.1%
2001-07-16 | 20:45 | 16.16 | 7.8 | 52.9 | 9.57 | 112.6%
2001-07-16 | 21:00 | 16.16 | 7.8 | 52.8 | 9.59 | 112.8%
2001-07-16 | 21:15 | 16.12 | 7.79 | 52.8 | 9.54 | 112.2%
2001-07-16 | 21:30 | 16.14 | 7.8 | 52.8 | 9.47 | 111.4%
2001-07-16 | 21:45 | 16.14 | 7.82 | 52.7 | 9.44 | 111.1%
2001-07-16 | 22:00 | 16.11 | 7.81 | 52.7 | 9.53 | 112.1%
2001-07-16 | 22:15 | 16.06 | 7.81 | 52.7 | 9.46 | 111.3%
2001-07-16 | 22:30 | 16.06 | 7.81 | 52.6 | 9.48 | 111.5%
2001-07-16 | 22:45 | 16.06 | 7.81 | 52.6 | 9.44 | 111.1%
2001-07-16 | 23:00 | 16.02 | 7.8 | 52.6 | 9.41 | 110.7%
2001-07-16 | 23:15 | 16 | 7.79 | 52.5 | 9.41 | 110.7%
2001-07-16 | 23:30 | 15.95 | 7.79 | 52.6 | 9.39 | 108.6%
2001-07-16 | 23:45 | 15.91 | 7.79 | 52.6 | 9.33 | 107.9%
2001-07-17 | 0:00 | 15.89 | 7.78 | 52.5 | 9.24 | 106.8%
2001-07-17 | 0:15 | 15.87 | 7.76 | 52.4 | 9.26 | 107.1%
2001-07-17 | 0:30 | 15.85 | 7.74 | 52.4 | 9.19 | 106.2%
2001-07-17 | 0:45 | 15.79 | 7.72 | 52.4 | 9.21 | 106.5%
2001-07-17 | 1:00 | 15.73 | 7.71 | 52.4 | 9.17 | 106%
2001-07-17 | 1:15 | 15.71 | 7.69 | 52.4 | 9.05 | 104.6%
2001-07-17 | 1:30 | 15.71 | 7.68 | 52.3 | 9.08 | 105%
2001-07-17 | 1:45 | 15.65 | 7.67 | 52.4 | 9.11 | 105.3%
2001-07-17 | 2:00 | 15.58 | 7.65 | 52.3 | 8.98 | 103.8%
2001-07-17 | 2:15 | 15.55 | 7.64 | 52.3 | 8.95 | 103.5%
2001-07-17 | 2:30 | 15.51 | 7.62 | 52.3 | 8.94 | 103.4%
2001-07-17 | 2:45 | 15.46 | 7.62 | 52.3 | 8.89 | 101.6%
2001-07-17 | 3:00 | 15.43 | 7.58 | 52.3 | 8.86 | 101.3%
2001-07-17 | 3:15 | 15.36 | 7.58 | 52.3 | 8.86 | 101.3%
2001-07-17 | 3:30 | 15.31 | 7.55 | 52.3 | 8.75 | 100%
2001-07-17 | 3:45 | 15.23 | 7.55 | 52.3 | 8.81 | 100.7%
2001-07-17 | 4:00 | 15.2 | 7.54 | 52.3 | 8.72 | 99.66%
2001-07-17 | 4:15 | 15.15 | 7.52 | 52.3 | 8.66 | 98.97%
2001-07-17 | 4:30 | 15.09 | 7.51 | 52.3 | 8.67 | 99.09%
2001-07-17 | 4:45 | 15.05 | 7.49 | 52.3 | 8.61 | 98.4%
2001-07-17 | 5:00 | 14.99 | 7.49 | 52.3 | 8.57 | 97.39%
2001-07-17 | 5:15 | 14.91 | 7.46 | 52.3 | 8.55 | 97.16%
2001-07-17 | 5:30 | 14.89 | 7.45 | 52.3 | 8.52 | 96.82%
2001-07-17 | 5:45 | 14.83 | 7.44 | 52.4 | 8.49 | 96.48%
2001-07-17 | 6:00 | 14.78 | 7.43 | 52.3 | 8.42 | 95.68%
2001-07-17 | 6:15 | 14.75 | 7.43 | 52.3 | 8.45 | 96.02%
2001-07-17 | 6:30 | 14.68 | 7.42 | 52.4 | 8.42 | 95.68%
2001-07-17 | 6:45 | 14.63 | 7.41 | 52.4 | 8.43 | 95.8%
2001-07-17 | 7:00 | 14.59 | 7.42 | 52.4 | 8.4 | 95.45%
2001-07-17 | 7:15 | 14.56 | 7.42 | 52.4 | 8.42 | 95.68%
2001-07-17 | 7:30 | 14.49 | 7.42 | 52.5 | 8.48 | 95.28%
2001-07-17 | 7:45 | 14.45 | 7.43 | 52.6 | 8.45 | 94.94%
2001-07-17 | 8:00 | 14.43 | 7.43 | 52.6 | 8.49 | 95.39%
2001-07-17 | 8:15 | 14.42 | 7.43 | 52.7 | 8.49 | 95.39%
2001-07-17 | 8:30 | 14.44 | 7.45 | 52.6 | 8.53 | 95.84%
2001-07-17 | 8:45 | 14.44 | 7.46 | 52.7 | 8.53 | 95.84%
2001-07-17 | 9:00 | 14.46 | 7.46 | 52.7 | 8.6 | 96.63%
2001-07-17 | 9:15 | 14.47 | 7.46 | 52.7 | 8.55 | 96.07%
2001-07-17 | 9:30 | 14.48 | 7.48 | 52.7 | 8.62 | 96.85%
2001-07-17 | 9:45 | 14.54 | 7.49 | 52.7 | 8.63 | 98.07%
2001-07-17 | 10:00 | 14.59 | 7.51 | 52.8 | 8.67 | 98.52%
2001-07-17 | 10:15 | 14.56 | 7.51 | 52.9 | 8.66 | 98.41%
2001-07-17 | 10:30 | 14.62 | 7.53 | 52.9 | 8.68 | 98.64%
2001-07-17 | 10:45 | 14.64 | 7.54 | 53 | 8.73 | 99.2%
2001-07-17 | 11:00 | 14.72 | 7.55 | 53 | 8.76 | 99.55%
2001-07-17 | 11:15 | 14.78 | 7.56 | 53.1 | 8.85 | 100.6%
2001-07-17 | 11:30 | 14.88 | 7.59 | 53.1 | 8.81 | 100.1%
2001-07-17 | 11:45 | 14.87 | 7.6 | 53.1 | 8.84 | 100.5%
2001-07-17 | 12:00 | 14.93 | 7.61 | 53.2 | 8.92 | 101.4%
2001-07-17 | 12:15 | 14.99 | 7.62 | 53.2 | 8.92 | 101.4%
2001-07-17 | 12:30 | 15.04 | 7.64 | 53.2 | 9 | 102.9%
2001-07-17 | 12:45 | 15.05 | 7.64 | 53.3 | 9.06 | 103.5%
2001-07-17 | 13:00 | 15.13 | 7.67 | 53.3 | 9.05 | 103.4%
2001-07-17 | 13:15 | 15.09 | 7.69 | 53.3 | 9.15 | 104.6%
2001-07-17 | 13:30 | 15.18 | 7.72 | 53.3 | 9.12 | 104.2%
2001-07-17 | 13:45 | 15.21 | 7.71 | 53.3 | 9.26 | 105.8%
2001-07-17 | 14:00 | 15.14 | 7.73 | 53.3 | 9.24 | 105.6%
2001-07-17 | 14:15 | 15.22 | 7.76 | 53.3 | 9.25 | 105.7%
2001-07-17 | 14:30 | 15.17 | 7.8 | 53.4 | 9.35 | 106.9%
2001-07-17 | 14:45 | 15.1 | 7.79 | 53.3 | 9.42 | 107.7%
2001-07-17 | 15:00 | 15.12 | 7.82 | 53.3 | 9.43 | 107.8%
2001-07-17 | 15:15 | 15.03 | 7.81 | 53.3 | 9.47 | 108.2%
2001-07-17 | 15:30 | 15.01 | 7.81 | 53.3 | 9.44 | 107.9%
2001-07-17 | 15:45 | 14.95 | 7.82 | 53.3 | 9.55 | 108.5%
2001-07-17 | 16:00 | 14.96 | 7.82 | 53.3 | 9.6 | 109.1%
2001-07-17 | 16:15 | 14.97 | 7.84 | 53.3 | 9.6 | 109.1%
2001-07-17 | 16:30 | 15.01 | 7.86 | 53.3 | 9.71 | 111%
2001-07-17 | 16:45 | 15.06 | 7.9 | 53.3 | 9.68 | 110.6%
2001-07-17 | 17:00 | 15.08 | 7.94 | 53.3 | 9.8 | 112%
2001-07-17 | 17:15 | 15.07 | 7.95 | 53.3 | 9.84 | 112.5%
2001-07-17 | 17:30 | 15.1 | 7.97 | 53.3 | 9.86 | 112.7%
2001-07-17 | 17:45 | 15.08 | 7.98 | 53.3 | 9.88 | 112.9%
2001-07-17 | 18:00 | 15.1 | 7.98 | 53.2 | 9.88 | 112.9%
2001-07-17 | 18:15 | 15.11 | 8 | 53.2 | 9.89 | 113%
2001-07-17 | 18:30 | 15.1 | 7.98 | 53.1 | 9.83 | 112.3%
2001-07-17 | 18:45 | 15.09 | 7.99 | 53.1 | 9.86 | 112.7%
2001-07-17 | 19:00 | 15.06 | 7.98 | 53.1 | 9.84 | 112.5%
2001-07-17 | 19:15 | 15.06 | 7.96 | 53.1 | 9.8 | 112%
2001-07-17 | 19:30 | 15.05 | 7.96 | 53 | 9.79 | 111.9%
2001-07-17 | 19:45 | 15.06 | 7.95 | 52.9 | 9.85 | 112.6%
2001-07-17 | 20:00 | 15.05 | 7.94 | 52.9 | 9.78 | 111.8%
2001-07-17 | 20:15 | 15.04 | 7.93 | 52.9 | 9.7 | 110.9%
2001-07-17 | 20:30 | 15.04 | 7.93 | 52.9 | 9.69 | 110.7%
2001-07-17 | 20:45 | 15.06 | 7.91 | 52.8 | 9.69 | 110.7%
2001-07-17 | 21:00 | 15.07 | 7.92 | 52.8 | 9.68 | 110.6%
2001-07-17 | 21:15 | 15.08 | 7.92 | 52.7 | 9.68 | 110.6%
2001-07-17 | 21:30 | 15.08 | 7.9 | 52.7 | 9.62 | 109.9%
2001-07-17 | 21:45 | 15.07 | 7.9 | 52.7 | 9.65 | 110.3%
2001-07-17 | 22:00 | 15.06 | 7.89 | 52.7 | 9.6 | 109.7%
2001-07-17 | 22:15 | 15.05 | 7.89 | 52.6 | 9.55 | 109.1%
2001-07-17 | 22:30 | 15.07 | 7.89 | 52.5 | 9.56 | 109.3%
2001-07-17 | 22:45 | 15.04 | 7.89 | 52.6 | 9.54 | 109%
2001-07-17 | 23:00 | 15.04 | 7.88 | 52.6 | 9.49 | 108.5%
2001-07-17 | 23:15 | 15.03 | 7.87 | 52.5 | 9.51 | 108.7%
2001-07-17 | 23:30 | 15.03 | 7.86 | 52.5 | 9.47 | 108.2%
2001-07-17 | 23:45 | 15.02 | 7.85 | 52.5 | 9.4 | 107.4%
2001-07-18 | 0:00 | 15.01 | 7.83 | 52.5 | 9.3 | 106.3%
2001-07-18 | 0:15 | 15.02 | 7.82 | 52.4 | 9.38 | 107.2%
2001-07-18 | 0:30 | 15 | 7.83 | 52.4 | 9.34 | 106.7%
2001-07-18 | 0:45 | 14.97 | 7.8 | 52.3 | 9.28 | 105.5%
2001-07-18 | 1:00 | 14.94 | 7.81 | 52.4 | 9.22 | 104.8%
2001-07-18 | 1:15 | 14.92 | 7.78 | 52.4 | 9.24 | 105%
2001-07-18 | 1:30 | 14.87 | 7.78 | 52.4 | 9.25 | 105.1%
2001-07-18 | 1:45 | 14.84 | 7.77 | 52.4 | 9.12 | 103.6%
2001-07-18 | 2:00 | 14.8 | 7.74 | 52.6 | 9.14 | 103.9%
2001-07-18 | 2:15 | 14.8 | 7.72 | 52.3 | 9.02 | 102.5%
2001-07-18 | 2:30 | 14.74 | 7.72 | 52.4 | 9.05 | 102.8%
2001-07-18 | 2:45 | 14.71 | 7.68 | 52.4 | 9.01 | 102.4%
2001-07-18 | 3:00 | 14.7 | 7.66 | 52.4 | 9.03 | 102.6%
2001-07-18 | 3:15 | 14.68 | 7.65 | 52.4 | 8.91 | 101.2%
2001-07-18 | 3:30 | 14.67 | 7.62 | 52.4 | 8.93 | 101.5%
2001-07-18 | 3:45 | 14.65 | 7.62 | 52.4 | 8.9 | 101.1%
2001-07-18 | 4:00 | 14.63 | 7.63 | 52.4 | 8.79 | 99.89%
2001-07-18 | 4:15 | 14.63 | 7.6 | 52.4 | 8.73 | 99.2%
2001-07-18 | 4:30 | 14.59 | 7.59 | 52.4 | 8.78 | 99.77%
2001-07-18 | 4:45 | 14.56 | 7.56 | 52.5 | 8.68 | 98.64%
2001-07-18 | 5:00 | 14.55 | 7.56 | 52.5 | 8.69 | 98.75%
2001-07-18 | 5:15 | 14.52 | 7.53 | 52.5 | 8.63 | 98.07%
2001-07-18 | 5:30 | 14.5 | 7.53 | 52.5 | 8.55 | 97.16%
2001-07-18 | 5:45 | 14.46 | 7.52 | 52.6 | 8.61 | 96.74%
2001-07-18 | 6:00 | 14.44 | 7.55 | 52.6 | 8.54 | 95.96%
2001-07-18 | 6:15 | 14.42 | 7.51 | 52.7 | 8.52 | 95.73%
2001-07-18 | 6:30 | 14.42 | 7.52 | 52.6 | 8.47 | 95.17%
2001-07-18 | 6:45 | 14.41 | 7.47 | 52.6 | 8.49 | 95.39%
2001-07-18 | 7:00 | 14.4 | 7.48 | 52.6 | 8.45 | 94.94%
2001-07-18 | 7:15 | 14.41 | 7.47 | 52.6 | 8.46 | 95.06%
2001-07-18 | 7:30 | 14.43 | 7.47 | 52.6 | 8.49 | 95.39%
2001-07-18 | 7:45 | 14.42 | 7.48 | 52.7 | 8.41 | 94.49%
2001-07-18 | 8:00 | 14.43 | 7.48 | 52.7 | 8.49 | 95.39%
2001-07-18 | 8:15 | 14.42 | 7.48 | 52.7 | 8.46 | 95.06%
2001-07-18 | 8:30 | 14.42 | 7.48 | 52.8 | 8.45 | 94.94%
2001-07-18 | 8:45 | 14.45 | 7.5 | 52.8 | 8.52 | 95.73%
2001-07-18 | 9:00 | 14.5 | 7.52 | 52.8 | 8.57 | 97.39%
2001-07-18 | 9:15 | 14.5 | 7.53 | 52.8 | 8.61 | 97.84%
2001-07-18 | 9:30 | 14.51 | 7.53 | 52.9 | 8.64 | 98.18%
2001-07-18 | 9:45 | 14.51 | 7.69 | 52.9 | 8.46 | 96.14%
2001-07-18 | 10:00 | 14.51 | 7.68 | 52.9 | 8.51 | 96.7%
2001-07-18 | 10:15 | 14.52 | 7.68 | 52.9 | 8.53 | 96.93%
2001-07-18 | 10:30 | 14.53 | 7.69 | 53 | 8.6 | 97.73%
2001-07-18 | 10:45 | 14.56 | 7.7 | 53 | 8.67 | 98.52%
2001-07-18 | 11:00 | 14.51 | 7.71 | 53.1 | 8.66 | 98.41%
2001-07-18 | 11:15 | 14.56 | 7.73 | 53.1 | 8.7 | 98.86%
2001-07-18 | 11:30 | 14.54 | 7.71 | 53.1 | 8.76 | 99.55%
2001-07-18 | 11:45 | 14.58 | 7.74 | 53.1 | 8.78 | 99.77%
2001-07-18 | 12:00 | 14.58 | 7.75 | 53.2 | 8.87 | 100.8%
2001-07-18 | 12:15 | 14.71 | 7.77 | 53.1 | 8.83 | 100.3%
2001-07-18 | 12:30 | 14.75 | 7.79 | 53.2 | 8.91 | 101.2%
2001-07-18 | 12:45 | 14.81 | 7.81 | 53.3 | 8.9 | 101.1%
2001-07-18 | 13:00 | 14.97 | 7.86 | 53.3 | 9 | 102.3%
2001-07-18 | 13:15 | 15.05 | 7.89 | 53.3 | 9.05 | 103.4%
2001-07-18 | 13:30 | 15.07 | 7.89 | 53.3 | 9.13 | 104.3%
2001-07-18 | 13:45 | 15.11 | 7.9 | 53.4 | 9.11 | 104.1%
2001-07-18 | 14:00 | 15.23 | 7.94 | 53.4 | 9.14 | 104.5%
2001-07-18 | 14:15 | 15.28 | 7.96 | 53.4 | 9.23 | 105.5%
2001-07-18 | 14:30 | 15.33 | 7.95 | 53.4 | 9.22 | 105.4%
2001-07-18 | 14:45 | 15.29 | 7.95 | 53.4 | 9.26 | 105.8%
2001-07-18 | 15:00 | 15.29 | 7.95 | 53.4 | 9.19 | 105%
2001-07-18 | 15:15 | 15.31 | 7.96 | 53.4 | 9.32 | 106.5%
2001-07-18 | 15:30 | 15.36 | 8 | 53.4 | 9.31 | 106.4%
2001-07-18 | 15:45 | 15.38 | 8.03 | 53.4 | 9.3 | 106.3%
2001-07-18 | 16:00 | 15.42 | 8.05 | 53.5 | 9.33 | 106.6%
2001-07-18 | 16:15 | 15.42 | 8.06 | 53.5 | 9.38 | 107.2%
2001-07-18 | 16:30 | 15.4 | 8.03 | 53.5 | 9.42 | 107.7%
2001-07-18 | 16:45 | 15.42 | 8.05 | 53.5 | 9.55 | 109.1%
2001-07-18 | 17:00 | 15.42 | 8.05 | 53.5 | 9.51 | 108.7%
2001-07-18 | 17:15 | 15.47 | 8.12 | 53.5 | 9.52 | 108.8%
2001-07-18 | 17:30 | 15.53 | 8.13 | 53.4 | 9.57 | 110.6%
2001-07-18 | 17:45 | 15.54 | 8.16 | 53.4 | 9.63 | 111.3%
2001-07-18 | 18:00 | 15.59 | 8.17 | 53.4 | 9.6 | 111%
2001-07-18 | 18:15 | 15.63 | 8.18 | 53.4 | 9.68 | 111.9%
2001-07-18 | 18:30 | 15.61 | 8.11 | 53.3 | 9.68 | 111.9%
2001-07-18 | 18:45 | 15.63 | 8.13 | 53.3 | 9.63 | 111.3%
2001-07-18 | 19:00 | 15.65 | 8.18 | 53.3 | 9.67 | 111.8%
2001-07-18 | 19:15 | 15.65 | 8.18 | 53.2 | 9.63 | 111.3%
2001-07-18 | 19:30 | 15.67 | 8.16 | 53.2 | 9.61 | 111.1%
2001-07-18 | 19:45 | 15.68 | 8.12 | 53.2 | 9.6 | 111%
2001-07-18 | 20:00 | 15.68 | 8.09 | 53.2 | 9.64 | 111.4%
2001-07-18 | 20:15 | 15.71 | 8.08 | 53.1 | 9.55 | 110.4%
2001-07-18 | 20:30 | 15.72 | 8.06 | 53.1 | 9.46 | 109.4%
2001-07-18 | 20:45 | 15.74 | 8.04 | 53.1 | 9.5 | 109.8%
2001-07-18 | 21:00 | 15.74 | 8.03 | 53.1 | 9.51 | 109.9%
2001-07-18 | 21:15 | 15.79 | 8.03 | 53 | 9.48 | 109.6%
2001-07-18 | 21:30 | 15.8 | 8.01 | 53 | 9.35 | 108.1%
2001-07-18 | 21:45 | 15.8 | 8.02 | 52.9 | 9.38 | 108.4%
2001-07-18 | 22:00 | 15.84 | 8 | 52.9 | 9.47 | 109.5%
2001-07-18 | 22:15 | 15.82 | 7.99 | 52.8 | 9.46 | 109.4%
2001-07-18 | 22:30 | 15.8 | 8.01 | 52.9 | 9.27 | 107.2%
2001-07-18 | 22:45 | 15.79 | 8 | 52.8 | 9.4 | 108.7%
2001-07-18 | 23:00 | 15.74 | 7.96 | 52.8 | 9.33 | 107.9%
2001-07-18 | 23:15 | 15.72 | 7.95 | 52.8 | 9.32 | 107.7%
2001-07-18 | 23:30 | 15.65 | 7.95 | 52.8 | 9.3 | 107.5%
2001-07-18 | 23:45 | 15.6 | 7.94 | 52.7 | 9.23 | 106.7%
2001-07-19 | 0:00 | 15.57 | 7.91 | 52.8 | 9.27 | 107.2%
2001-07-19 | 0:15 | 15.54 | 7.9 | 52.6 | 9.19 | 106.2%
2001-07-19 | 0:30 | 15.49 | 7.88 | 52.6 | 9.18 | 104.9%
2001-07-19 | 0:45 | 15.46 | 7.88 | 52.6 | 9.16 | 104.7%
2001-07-19 | 1:00 | 15.4 | 7.84 | 52.6 | 9.05 | 103.4%
2001-07-19 | 1:15 | 15.35 | 7.85 | 52.6 | 9.07 | 103.7%
2001-07-19 | 1:30 | 15.33 | 7.81 | 52.5 | 9.04 | 103.3%
2001-07-19 | 1:45 | 15.31 | 7.77 | 52.4 | 8.98 | 102.6%
2001-07-19 | 2:00 | 15.26 | 7.77 | 52.5 | 9.02 | 103.1%
2001-07-19 | 2:15 | 15.23 | 7.76 | 52.4 | 8.89 | 101.6%
2001-07-19 | 2:30 | 15.2 | 7.74 | 52.4 | 8.85 | 101.1%
2001-07-19 | 2:45 | 15.16 | 7.72 | 52.4 | 8.82 | 100.8%
2001-07-19 | 3:00 | 15.11 | 7.7 | 52.4 | 8.8 | 100.6%
2001-07-19 | 3:15 | 15.1 | 7.7 | 52.4 | 8.8 | 100.6%
2001-07-19 | 3:30 | 15.05 | 7.65 | 52.4 | 8.76 | 100.1%
2001-07-19 | 3:45 | 15.01 | 7.65 | 52.4 | 8.71 | 99.54%
2001-07-19 | 4:00 | 14.96 | 7.64 | 52.4 | 8.68 | 98.64%
2001-07-19 | 4:15 | 14.88 | 7.62 | 52.4 | 8.6 | 97.73%
2001-07-19 | 4:30 | 14.83 | 7.61 | 52.3 | 8.56 | 97.27%
2001-07-19 | 4:45 | 14.81 | 7.59 | 52.3 | 8.6 | 97.73%
2001-07-19 | 5:00 | 14.75 | 7.6 | 52.3 | 8.51 | 96.7%
2001-07-19 | 5:15 | 14.7 | 7.55 | 52.4 | 8.56 | 97.27%
2001-07-19 | 5:30 | 14.66 | 7.55 | 52.3 | 8.48 | 96.36%
2001-07-19 | 5:45 | 14.6 | 7.54 | 52.3 | 8.45 | 96.02%
2001-07-19 | 6:00 | 14.58 | 7.53 | 52.3 | 8.4 | 95.45%
2001-07-19 | 6:15 | 14.53 | 7.52 | 52.3 | 8.35 | 94.89%
2001-07-19 | 6:30 | 14.48 | 7.51 | 52.3 | 8.4 | 94.38%
2001-07-19 | 6:45 | 14.44 | 7.52 | 52.3 | 8.37 | 94.04%
2001-07-19 | 7:00 | 14.39 | 7.54 | 52.4 | 8.4 | 94.38%
2001-07-19 | 7:15 | 14.36 | 7.55 | 52.4 | 8.36 | 93.93%
2001-07-19 | 7:30 | 14.34 | 7.58 | 52.4 | 8.37 | 94.04%
2001-07-19 | 7:45 | 14.32 | 7.58 | 52.4 | 8.47 | 95.17%
2001-07-19 | 8:00 | 14.37 | 7.62 | 52.4 | 8.47 | 95.17%
2001-07-19 | 8:15 | 14.37 | 7.64 | 52.4 | 8.44 | 94.83%
2001-07-19 | 8:30 | 14.39 | 7.68 | 52.5 | 8.54 | 95.96%
2001-07-19 | 8:45 | 14.44 | 7.69 | 52.5 | 8.55 | 96.07%
2001-07-19 | 9:00 | 14.44 | 7.68 | 52.5 | 8.52 | 95.73%
2001-07-19 | 9:15 | 14.46 | 7.7 | 52.5 | 8.56 | 96.18%
2001-07-19 | 9:30 | 14.49 | 7.72 | 52.5 | 8.52 | 95.73%
2001-07-19 | 9:45 | 14.51 | 7.72 | 52.5 | 8.57 | 97.39%
2001-07-19 | 10:00 | 14.55 | 7.77 | 52.6 | 8.62 | 97.95%
2001-07-19 | 10:15 | 14.64 | 7.8 | 52.6 | 8.65 | 98.3%
2001-07-19 | 10:30 | 14.69 | 7.81 | 52.6 | 8.67 | 98.52%
2001-07-19 | 10:45 | 14.76 | 7.85 | 52.6 | 8.64 | 98.18%
2001-07-19 | 11:00 | 14.88 | 7.88 | 52.7 | 8.73 | 99.2%
2001-07-19 | 11:15 | 14.92 | 7.92 | 52.7 | 8.7 | 98.86%
2001-07-19 | 11:30 | 14.97 | 7.87 | 52.8 | 8.78 | 99.77%
2001-07-19 | 11:45 | 15.03 | 7.89 | 52.8 | 8.76 | 100.1%
2001-07-19 | 12:00 | 15.19 | 7.92 | 52.9 | 8.81 | 100.7%
2001-07-19 | 12:15 | 15.14 | 7.89 | 52.9 | 8.82 | 100.8%
2001-07-19 | 12:30 | 15.24 | 7.96 | 53 | 8.88 | 101.5%
2001-07-19 | 12:45 | 15.31 | 7.95 | 52.9 | 8.92 | 101.9%
2001-07-19 | 13:00 | 15.27 | 7.9 | 53.1 | 8.93 | 102.1%
2001-07-19 | 13:15 | 15.37 | 7.98 | 53.1 | 9.01 | 103%
2001-07-19 | 13:30 | 15.51 | 8.03 | 53.1 | 9 | 104%
2001-07-19 | 13:45 | 15.59 | 8.05 | 53.1 | 9.1 | 105.2%
2001-07-19 | 14:00 | 15.54 | 8.01 | 53.1 | 9.1 | 105.2%
2001-07-19 | 14:15 | 15.56 | 8.02 | 53.1 | 9.22 | 106.6%
2001-07-19 | 14:30 | 15.66 | 8.04 | 53.1 | 9.2 | 106.4%
2001-07-19 | 14:45 | 15.64 | 8 | 53.1 | 9.21 | 106.5%
2001-07-19 | 15:00 | 15.57 | 7.87 | 53.1 | 9.15 | 105.8%
2001-07-19 | 15:15 | 15.49 | 7.84 | 53 | 9.03 | 103.2%
2001-07-19 | 15:30 | 15.46 | 7.8 | 53 | 9.07 | 103.7%
2001-07-19 | 15:45 | 15.41 | 7.82 | 53 | 9.16 | 104.7%
2001-07-19 | 16:00 | 15.42 | 7.8 | 53 | 9.31 | 106.4%
2001-07-19 | 16:15 | 15.4 | 7.8 | 52.9 | 9.17 | 104.8%
2001-07-19 | 16:30 | 15.36 | 7.78 | 53 | 9.29 | 106.2%
2001-07-19 | 16:45 | 15.37 | 7.87 | 53 | 9.34 | 106.7%
2001-07-19 | 17:00 | 15.4 | 7.94 | 53 | 9.44 | 107.9%
2001-07-19 | 17:15 | 15.46 | 8 | 53 | 9.52 | 108.8%
2001-07-19 | 17:30 | 15.53 | 8.01 | 52.9 | 9.41 | 108.8%
2001-07-19 | 17:45 | 15.56 | 8.05 | 52.9 | 9.59 | 110.9%
2001-07-19 | 18:00 | 15.51 | 8.03 | 53 | 9.59 | 110.9%
2001-07-19 | 18:15 | 15.5 | 8.03 | 52.9 | 9.61 | 111.1%
2001-07-19 | 18:30 | 15.5 | 7.99 | 52.8 | 9.54 | 110.3%
2001-07-19 | 18:45 | 15.49 | 8.03 | 52.8 | 9.52 | 108.8%
2001-07-19 | 19:00 | 15.45 | 7.97 | 52.8 | 9.41 | 107.5%
2001-07-19 | 19:15 | 15.42 | 7.93 | 52.7 | 9.5 | 108.6%
2001-07-19 | 19:30 | 15.4 | 7.95 | 52.8 | 9.42 | 107.7%
2001-07-19 | 19:45 | 15.38 | 7.92 | 52.8 | 9.5 | 108.6%
2001-07-19 | 20:00 | 15.36 | 7.88 | 52.8 | 9.43 | 107.8%
2001-07-19 | 20:15 | 15.39 | 7.87 | 52.7 | 9.39 | 107.3%
2001-07-19 | 20:30 | 15.39 | 7.85 | 52.7 | 9.36 | 107%
2001-07-19 | 20:45 | 15.4 | 7.82 | 52.6 | 9.26 | 105.8%
2001-07-19 | 21:00 | 15.38 | 7.81 | 52.5 | 9.16 | 104.7%
2001-07-19 | 21:15 | 15.37 | 7.77 | 52.6 | 9.26 | 105.8%
2001-07-19 | 21:30 | 15.33 | 7.79 | 52.7 | 9.27 | 105.9%
2001-07-19 | 21:45 | 15.34 | 7.76 | 52.5 | 9.23 | 105.5%
2001-07-19 | 22:00 | 15.34 | 7.74 | 52.4 | 9.13 | 104.3%
2001-07-19 | 22:15 | 15.34 | 7.73 | 52.2 | 9.16 | 104.7%
2001-07-19 | 22:30 | 15.33 | 7.72 | 52.2 | 9.08 | 103.8%
2001-07-19 | 22:45 | 15.29 | 7.72 | 52.2 | 9.02 | 103.1%
2001-07-19 | 23:00 | 15.26 | 7.72 | 52.1 | 9.06 | 103.5%
2001-07-19 | 23:15 | 15.23 | 7.71 | 52.1 | 9.07 | 103.7%
2001-07-19 | 23:30 | 15.22 | 7.71 | 52.1 | 9.09 | 103.9%
2001-07-19 | 23:45 | 15.2 | 7.7 | 52.1 | 9 | 102.9%
2001-07-20 | 0:00 | 15.2 | 7.67 | 51.9 | 8.96 | 102.4%
2001-07-20 | 0:15 | 15.18 | 7.66 | 52 | 9 | 102.9%
2001-07-20 | 0:30 | 15.15 | 7.65 | 51.9 | 8.96 | 102.4%
2001-07-20 | 0:45 | 15.14 | 7.65 | 51.9 | 8.82 | 100.8%
2001-07-20 | 1:00 | 15.12 | 7.62 | 51.9 | 8.84 | 101%
2001-07-20 | 1:15 | 15.12 | 7.62 | 51.8 | 8.81 | 100.7%
2001-07-20 | 1:30 | 15.1 | 7.62 | 51.8 | 8.8 | 100.6%
2001-07-20 | 1:45 | 15.1 | 7.59 | 51.8 | 8.76 | 100.1%
2001-07-20 | 2:00 | 15.08 | 7.57 | 51.8 | 8.64 | 98.74%
2001-07-20 | 2:15 | 15.05 | 7.58 | 51.8 | 8.68 | 99.2%
2001-07-20 | 2:30 | 15.07 | 7.56 | 51.7 | 8.66 | 98.97%
2001-07-20 | 2:45 | 15.05 | 7.57 | 51.8 | 8.59 | 98.17%
2001-07-20 | 3:00 | 15.01 | 7.54 | 51.7 | 8.56 | 97.83%
2001-07-20 | 3:15 | 14.97 | 7.53 | 51.8 | 8.5 | 96.59%
2001-07-20 | 3:30 | 14.95 | 7.54 | 51.8 | 8.47 | 96.25%
2001-07-20 | 3:45 | 14.9 | 7.51 | 51.8 | 8.47 | 96.25%
2001-07-20 | 4:00 | 14.88 | 7.5 | 51.7 | 8.43 | 95.8%
2001-07-20 | 4:15 | 14.87 | 7.49 | 51.7 | 8.42 | 95.68%
2001-07-20 | 4:30 | 14.83 | 7.48 | 51.8 | 8.35 | 94.89%
2001-07-20 | 4:45 | 14.81 | 7.47 | 51.8 | 8.38 | 95.23%
2001-07-20 | 5:00 | 14.8 | 7.45 | 51.8 | 8.32 | 94.55%
2001-07-20 | 5:15 | 14.77 | 7.45 | 51.8 | 8.36 | 95%
2001-07-20 | 5:30 | 14.75 | 7.44 | 51.8 | 8.15 | 92.61%
2001-07-20 | 5:45 | 14.72 | 7.42 | 51.8 | 8.24 | 93.64%
2001-07-20 | 6:00 | 14.69 | 7.42 | 51.8 | 8.22 | 93.41%
2001-07-20 | 6:15 | 14.67 | 7.41 | 51.8 | 8.13 | 92.39%
2001-07-20 | 6:30 | 14.65 | 7.41 | 51.9 | 8.18 | 92.95%
2001-07-20 | 6:45 | 14.64 | 7.4 | 51.8 | 8.12 | 92.27%
2001-07-20 | 7:00 | 14.62 | 7.4 | 51.8 | 8.08 | 91.82%
2001-07-20 | 7:15 | 14.6 | 7.4 | 51.8 | 8.07 | 91.7%
2001-07-20 | 7:30 | 14.58 | 7.4 | 51.9 | 8.03 | 91.25%
2001-07-20 | 7:45 | 14.55 | 7.38 | 51.9 | 8.12 | 92.27%
2001-07-20 | 8:00 | 14.54 | 7.39 | 52 | 8.02 | 91.14%
2001-07-20 | 8:15 | 14.53 | 7.42 | 52 | 8.12 | 92.27%
2001-07-20 | 8:30 | 14.55 | 7.46 | 52 | 8.18 | 92.95%
2001-07-20 | 8:45 | 14.55 | 7.49 | 52 | 8.34 | 94.77%
2001-07-20 | 9:00 | 14.56 | 7.5 | 52 | 8.28 | 94.09%
2001-07-20 | 9:15 | 14.55 | 7.5 | 52.1 | 8.23 | 93.52%
2001-07-20 | 9:30 | 14.56 | 7.5 | 52.1 | 8.31 | 94.43%
2001-07-20 | 9:45 | 14.59 | 7.52 | 52.1 | 8.28 | 94.09%
2001-07-20 | 10:00 | 14.66 | 7.56 | 52.2 | 8.36 | 95%
2001-07-20 | 10:15 | 14.7 | 7.59 | 52.2 | 8.42 | 95.68%
2001-07-20 | 10:30 | 14.78 | 7.63 | 52.2 | 8.42 | 95.68%
2001-07-20 | 10:45 | 14.82 | 7.63 | 52.2 | 8.48 | 96.36%
2001-07-20 | 11:00 | 14.85 | 7.63 | 52.2 | 8.5 | 96.59%
2001-07-20 | 11:15 | 14.89 | 7.64 | 52.3 | 8.41 | 95.57%
2001-07-20 | 11:30 | 15.04 | 7.68 | 52.4 | 8.52 | 97.37%