## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2001-07-16 15:10 | 21.5 | 8.5 | 8.4 | 111 | 54 | 0
Sonde | 2001-07-16 15:16 | 21.43 | 8.49 | 8.96 | 118 | 52
Grade |  | A | A | B |  | A
Audit 2 | 2001-07-18 13:31 | 19.7 | 8.3 | 8.7 | 111 | 54 | 0
Sonde | 2001-07-18 13:31 | 19.5 | 8.37 | 9.82 | 118 | 52
Grade |  | A | A | Fail |  | A
Audit 3 | 2001-07-20 10:02 | 15.1 | 7.7 | 8.6 | 99 | 55 | 0
Sonde | 2001-07-20 10:01 | 14.88 | 7.73 | 9.19 | 105 | 53
Grade |  | A | A | B |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note: Audits 1-3 are macro (ed) into this main audit sheet and have audits that represent the duration of the deployment. Audits 4-5 represent the middle of the deployment and have the sonde values on the individual sheets. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-16 15:10 | 21.5 | 8.5 | 8.4 | 111 | 54
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-18 13:31 | 19.7 | 8.3 | 8.7 | 111 | 54
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-20 10:02 | 15.1 | 7.7 | 8.6 | 99 | 55
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-17 09:20 | 14.3 | 7.9 | 8.7 | 99 | 56
Sonde Values | 2001-07-17 09:16 | 14.06 | 7.7 | 9.14 | 104 | 54
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-19 10:15 | 15.4 | 7.8 | 8.9 | 102 | 56
Sonde Values | 2001-07-19 10:16 | 15.29 | 7.89 | 9.41 | 108 | 53
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | DO Sat
|  | °C |  | uS/cm | mg/L | %
2001-07-16 | 15:16 | 21.43 | 8.49 | 52 | 8.96 | 117.9%
2001-07-16 | 15:31 | 21.64 | 8.5 | 52 | 8.84 | 117.1%
2001-07-16 | 15:46 | 21.79 | 8.52 | 52 | 8.81 | 116.7%
2001-07-16 | 16:01 | 21.86 | 8.53 | 52 | 8.76 | 116%
2001-07-16 | 16:16 | 21.9 | 8.54 | 52 | 8.72 | 115.5%
2001-07-16 | 16:31 | 21.93 | 8.55 | 52 | 8.69 | 115.1%
2001-07-16 | 16:46 | 21.94 | 8.55 | 52 | 8.65 | 114.6%
2001-07-16 | 17:01 | 21.88 | 8.55 | 52 | 8.62 | 114.2%
2001-07-16 | 17:16 | 21.8 | 8.56 | 52 | 8.59 | 113.8%
2001-07-16 | 17:31 | 21.72 | 8.56 | 52 | 8.57 | 113.5%
2001-07-16 | 17:46 | 21.58 | 8.55 | 51 | 8.54 | 113.1%
2001-07-16 | 18:01 | 21.46 | 8.55 | 51 | 8.51 | 112%
2001-07-16 | 18:16 | 21.28 | 8.54 | 51 | 8.48 | 111.6%
2001-07-16 | 18:31 | 21.03 | 8.53 | 51 | 8.45 | 111.2%
2001-07-16 | 18:46 | 20.82 | 8.52 | 51 | 8.43 | 109.5%
2001-07-16 | 19:01 | 20.56 | 8.5 | 51 | 8.4 | 109.1%
2001-07-16 | 19:16 | 20.27 | 8.47 | 51 | 8.39 | 107.6%
2001-07-16 | 19:31 | 19.99 | 8.45 | 51 | 8.38 | 106.8%
2001-07-16 | 19:46 | 19.7 | 8.41 | 51 | 8.34 | 106.2%
2001-07-16 | 20:01 | 19.43 | 8.36 | 51 | 8.29 | 104.9%
2001-07-16 | 20:16 | 19.13 | 8.3 | 51 | 8.22 | 104.1%
2001-07-16 | 20:31 | 18.84 | 8.23 | 51 | 8.15 | 101.2%
2001-07-16 | 20:46 | 18.58 | 8.16 | 51 | 8.09 | 100.5%
2001-07-16 | 21:01 | 18.34 | 8.08 | 51 | 8.02 | 98.4%
2001-07-16 | 21:16 | 18.09 | 8 | 51 | 7.98 | 97.91%
2001-07-16 | 21:31 | 17.88 | 7.93 | 51 | 7.94 | 96.83%
2001-07-16 | 21:46 | 17.7 | 7.86 | 51 | 7.91 | 96.46%
2001-07-16 | 22:01 | 17.52 | 7.8 | 51 | 7.88 | 96.1%
2001-07-16 | 22:16 | 17.35 | 7.75 | 51 | 7.84 | 94.46%
2001-07-16 | 22:31 | 17.18 | 7.71 | 52 | 7.82 | 94.22%
2001-07-16 | 22:46 | 17.04 | 7.67 | 52 | 7.81 | 94.1%
2001-07-16 | 23:01 | 16.88 | 7.64 | 52 | 7.8 | 92.86%
2001-07-16 | 23:16 | 16.73 | 7.61 | 52 | 7.79 | 92.74%
2001-07-16 | 23:31 | 16.6 | 7.59 | 52 | 7.79 | 92.74%
2001-07-16 | 23:46 | 16.46 | 7.57 | 52 | 7.79 | 92.19%
2001-07-17 | 0:01 | 16.32 | 7.56 | 52 | 7.8 | 92.31%
2001-07-17 | 0:16 | 16.19 | 7.54 | 52 | 7.8 | 92.31%
2001-07-17 | 0:31 | 16.07 | 7.53 | 52 | 7.8 | 92.31%
2001-07-17 | 0:46 | 15.94 | 7.52 | 52 | 7.82 | 91.46%
2001-07-17 | 1:01 | 15.81 | 7.51 | 52 | 7.83 | 91.58%
2001-07-17 | 1:16 | 15.69 | 7.49 | 52 | 7.84 | 91.7%
2001-07-17 | 1:31 | 15.56 | 7.49 | 52 | 7.85 | 91.81%
2001-07-17 | 1:46 | 15.43 | 7.48 | 52 | 7.87 | 90.46%
2001-07-17 | 2:01 | 15.31 | 7.47 | 52 | 7.88 | 90.57%
2001-07-17 | 2:16 | 15.19 | 7.46 | 52 | 7.89 | 90.69%
2001-07-17 | 2:31 | 15.08 | 7.46 | 52 | 7.91 | 90.92%
2001-07-17 | 2:46 | 14.97 | 7.45 | 52 | 7.93 | 90.63%
2001-07-17 | 3:01 | 14.84 | 7.44 | 52 | 7.95 | 90.86%
2001-07-17 | 3:16 | 14.73 | 7.44 | 53 | 7.97 | 91.09%
2001-07-17 | 3:31 | 14.61 | 7.43 | 53 | 7.99 | 91.31%
2001-07-17 | 3:46 | 14.5 | 7.43 | 53 | 8.01 | 91.54%
2001-07-17 | 4:01 | 14.39 | 7.43 | 53 | 8.03 | 91.25%
2001-07-17 | 4:16 | 14.28 | 7.42 | 53 | 8.05 | 91.48%
2001-07-17 | 4:31 | 14.17 | 7.42 | 53 | 8.07 | 91.7%
2001-07-17 | 4:46 | 14.05 | 7.42 | 53 | 8.1 | 92.05%
2001-07-17 | 5:01 | 13.94 | 7.42 | 53 | 8.12 | 91.24%
2001-07-17 | 5:16 | 13.83 | 7.41 | 53 | 8.14 | 91.46%
2001-07-17 | 5:31 | 13.72 | 7.41 | 53 | 8.16 | 91.69%
2001-07-17 | 5:46 | 13.61 | 7.41 | 53 | 8.19 | 92.02%
2001-07-17 | 6:01 | 13.51 | 7.41 | 53 | 8.21 | 92.25%
2001-07-17 | 6:16 | 13.4 | 7.41 | 53 | 8.25 | 91.67%
2001-07-17 | 6:31 | 13.3 | 7.41 | 54 | 8.29 | 92.11%
2001-07-17 | 6:46 | 13.22 | 7.42 | 54 | 8.35 | 92.78%
2001-07-17 | 7:01 | 13.17 | 7.43 | 54 | 8.42 | 93.56%
2001-07-17 | 7:16 | 13.14 | 7.45 | 54 | 8.5 | 94.44%
2001-07-17 | 7:31 | 13.15 | 7.47 | 54 | 8.59 | 95.44%
2001-07-17 | 7:46 | 13.19 | 7.5 | 54 | 8.68 | 96.44%
2001-07-17 | 8:01 | 13.27 | 7.53 | 54 | 8.77 | 97.44%
2001-07-17 | 8:16 | 13.37 | 7.56 | 54 | 8.86 | 98.44%
2001-07-17 | 8:31 | 13.51 | 7.6 | 54 | 8.93 | 100.3%
2001-07-17 | 8:46 | 13.67 | 7.63 | 54 | 9 | 101.1%
2001-07-17 | 9:01 | 13.86 | 7.67 | 54 | 9.07 | 101.9%
2001-07-17 | 9:16 | 14.06 | 7.7 | 54 | 9.14 | 103.9%
2001-07-17 | 9:31 | 14.3 | 7.75 | 54 | 9.21 | 104.7%
2001-07-17 | 9:46 | 14.54 | 7.79 | 54 | 9.26 | 105.8%
2001-07-17 | 10:01 | 14.8 | 7.83 | 54 | 9.3 | 106.3%
2001-07-17 | 10:16 | 14.88 | 7.86 | 54 | 9.33 | 106.6%
2001-07-17 | 10:31 | 15.05 | 7.9 | 53 | 9.36 | 107.6%
2001-07-17 | 10:46 | 15.26 | 7.94 | 53 | 9.38 | 107.8%
2001-07-17 | 11:01 | 15.59 | 7.98 | 53 | 9.42 | 110.2%
2001-07-17 | 11:16 | 15.94 | 8.03 | 53 | 9.44 | 110.4%
2001-07-17 | 11:31 | 16.28 | 8.06 | 53 | 9.45 | 111.8%
2001-07-17 | 11:46 | 16.74 | 8.11 | 53 | 9.45 | 112.5%
2001-07-17 | 12:01 | 17.17 | 8.16 | 53 | 9.46 | 114%
2001-07-17 | 12:16 | 17.41 | 8.19 | 53 | 9.44 | 113.7%
2001-07-17 | 12:31 | 17.6 | 8.22 | 53 | 9.43 | 115%
2001-07-17 | 12:46 | 17.58 | 8.24 | 53 | 9.4 | 114.6%
2001-07-17 | 13:01 | 17.74 | 8.27 | 53 | 9.39 | 114.5%
2001-07-17 | 13:16 | 17.84 | 8.3 | 52 | 9.39 | 114.5%
2001-07-17 | 13:31 | 17.91 | 8.33 | 52 | 9.39 | 114.5%
2001-07-17 | 13:46 | 18.16 | 8.36 | 52 | 9.38 | 115.1%
2001-07-17 | 14:01 | 18.54 | 8.39 | 52 | 9.37 | 116.4%
2001-07-17 | 14:16 | 18.64 | 8.41 | 52 | 9.34 | 116%
2001-07-17 | 14:31 | 19.02 | 8.43 | 52 | 9.31 | 117.8%
2001-07-17 | 14:46 | 19.17 | 8.44 | 52 | 9.3 | 117.7%
2001-07-17 | 15:01 | 19.31 | 8.46 | 52 | 9.25 | 117.1%
2001-07-17 | 15:16 | 19.56 | 8.47 | 52 | 9.23 | 117.6%
2001-07-17 | 15:31 | 19.71 | 8.49 | 52 | 9.2 | 117.2%
2001-07-17 | 15:46 | 19.91 | 8.51 | 52 | 9.17 | 116.8%
2001-07-17 | 16:01 | 20.04 | 8.52 | 52 | 9.13 | 117.1%
2001-07-17 | 16:16 | 20.09 | 8.52 | 52 | 9.09 | 116.5%
2001-07-17 | 16:31 | 20.22 | 8.53 | 52 | 9.05 | 116%
2001-07-17 | 16:46 | 20.31 | 8.55 | 51 | 9.01 | 115.5%
2001-07-17 | 17:01 | 20.42 | 8.55 | 51 | 8.98 | 115.1%
2001-07-17 | 17:16 | 20.49 | 8.56 | 51 | 8.95 | 114.7%
2001-07-17 | 17:31 | 20.51 | 8.56 | 51 | 8.91 | 115.7%
2001-07-17 | 17:46 | 20.32 | 8.55 | 51 | 8.83 | 113.2%
2001-07-17 | 18:01 | 20.11 | 8.51 | 51 | 8.73 | 111.9%
2001-07-17 | 18:16 | 19.82 | 8.48 | 51 | 8.64 | 110.1%
2001-07-17 | 18:31 | 19.57 | 8.45 | 51 | 8.59 | 109.4%
2001-07-17 | 18:46 | 19.32 | 8.42 | 51 | 8.54 | 108.1%
2001-07-17 | 19:01 | 19.08 | 8.39 | 51 | 8.51 | 107.7%
2001-07-17 | 19:16 | 18.93 | 8.36 | 50 | 8.48 | 105.3%
2001-07-17 | 19:31 | 18.76 | 8.32 | 51 | 8.44 | 104.8%
2001-07-17 | 19:46 | 18.59 | 8.27 | 51 | 8.39 | 104.2%
2001-07-17 | 20:01 | 18.4 | 8.21 | 51 | 8.33 | 102.2%
2001-07-17 | 20:16 | 18.18 | 8.14 | 51 | 8.29 | 101.7%
2001-07-17 | 20:31 | 17.97 | 8.07 | 51 | 8.23 | 100.4%
2001-07-17 | 20:46 | 17.76 | 8 | 51 | 8.19 | 99.88%
2001-07-17 | 21:01 | 17.56 | 7.94 | 51 | 8.15 | 99.39%
2001-07-17 | 21:16 | 17.38 | 7.88 | 51 | 8.11 | 97.71%
2001-07-17 | 21:31 | 17.24 | 7.82 | 51 | 8.08 | 97.35%
2001-07-17 | 21:46 | 17.1 | 7.78 | 51 | 8.04 | 96.87%
2001-07-17 | 22:01 | 16.99 | 7.74 | 51 | 8.01 | 95.36%
2001-07-17 | 22:16 | 16.88 | 7.71 | 51 | 7.98 | 95%
2001-07-17 | 22:31 | 16.79 | 7.67 | 51 | 7.95 | 94.64%
2001-07-17 | 22:46 | 16.7 | 7.65 | 51 | 7.93 | 94.4%
2001-07-17 | 23:01 | 16.62 | 7.63 | 51 | 7.92 | 94.29%
2001-07-17 | 23:16 | 16.55 | 7.6 | 52 | 7.9 | 94.05%
2001-07-17 | 23:31 | 16.49 | 7.59 | 52 | 7.88 | 93.25%
2001-07-17 | 23:46 | 16.42 | 7.57 | 52 | 7.87 | 93.14%
2001-07-18 | 0:01 | 16.35 | 7.56 | 52 | 7.87 | 93.14%
2001-07-18 | 0:16 | 16.31 | 7.54 | 52 | 7.86 | 93.02%
2001-07-18 | 0:31 | 16.25 | 7.53 | 52 | 7.86 | 93.02%
2001-07-18 | 0:46 | 16.2 | 7.52 | 52 | 7.85 | 92.9%
2001-07-18 | 1:01 | 16.14 | 7.52 | 52 | 7.85 | 92.9%
2001-07-18 | 1:16 | 16.07 | 7.51 | 52 | 7.86 | 93.02%
2001-07-18 | 1:31 | 16.01 | 7.5 | 52 | 7.86 | 93.02%
2001-07-18 | 1:46 | 15.96 | 7.49 | 52 | 7.86 | 91.93%
2001-07-18 | 2:01 | 15.9 | 7.48 | 52 | 7.86 | 91.93%
2001-07-18 | 2:16 | 15.85 | 7.48 | 52 | 7.87 | 92.05%
2001-07-18 | 2:31 | 15.81 | 7.47 | 52 | 7.88 | 92.16%
2001-07-18 | 2:46 | 15.76 | 7.47 | 52 | 7.88 | 92.16%
2001-07-18 | 3:01 | 15.7 | 7.47 | 52 | 7.89 | 92.28%
2001-07-18 | 3:16 | 15.66 | 7.46 | 52 | 7.89 | 92.28%
2001-07-18 | 3:31 | 15.61 | 7.46 | 53 | 7.9 | 92.4%
2001-07-18 | 3:46 | 15.57 | 7.46 | 53 | 7.91 | 92.51%
2001-07-18 | 4:01 | 15.52 | 7.45 | 53 | 7.91 | 92.51%
2001-07-18 | 4:16 | 15.48 | 7.45 | 53 | 7.92 | 91.03%
2001-07-18 | 4:31 | 15.44 | 7.45 | 53 | 7.93 | 91.15%
2001-07-18 | 4:46 | 15.39 | 7.44 | 53 | 7.94 | 91.26%
2001-07-18 | 5:01 | 15.35 | 7.44 | 53 | 7.95 | 91.38%
2001-07-18 | 5:16 | 15.31 | 7.44 | 53 | 7.96 | 91.49%
2001-07-18 | 5:31 | 15.26 | 7.44 | 53 | 7.97 | 91.61%
2001-07-18 | 5:46 | 15.22 | 7.44 | 53 | 7.98 | 91.72%
2001-07-18 | 6:01 | 15.17 | 7.43 | 53 | 7.99 | 91.84%
2001-07-18 | 6:16 | 15.13 | 7.43 | 53 | 8.01 | 92.07%
2001-07-18 | 6:31 | 15.1 | 7.44 | 53 | 8.03 | 92.3%
2001-07-18 | 6:46 | 15.08 | 7.44 | 54 | 8.07 | 92.76%
2001-07-18 | 7:01 | 15.06 | 7.45 | 54 | 8.11 | 93.22%
2001-07-18 | 7:16 | 15.04 | 7.46 | 54 | 8.15 | 93.68%
2001-07-18 | 7:31 | 15.03 | 7.47 | 54 | 8.21 | 94.37%
2001-07-18 | 7:46 | 15.03 | 7.49 | 54 | 8.28 | 95.17%
2001-07-18 | 8:01 | 15.05 | 7.52 | 54 | 8.37 | 96.21%
2001-07-18 | 8:16 | 15.09 | 7.55 | 54 | 8.47 | 97.36%
2001-07-18 | 8:31 | 15.16 | 7.59 | 54 | 8.58 | 98.62%
2001-07-18 | 8:46 | 15.27 | 7.63 | 54 | 8.68 | 99.77%
2001-07-18 | 9:01 | 15.36 | 7.67 | 54 | 8.77 | 100.8%
2001-07-18 | 9:16 | 15.45 | 7.71 | 54 | 8.86 | 101.8%
2001-07-18 | 9:31 | 15.57 | 7.76 | 54 | 8.94 | 104.6%
2001-07-18 | 9:46 | 15.7 | 7.8 | 54 | 9.02 | 105.5%
2001-07-18 | 10:01 | 15.85 | 7.85 | 54 | 9.09 | 106.3%
2001-07-18 | 10:16 | 15.97 | 7.9 | 53 | 9.14 | 106.9%
2001-07-18 | 10:31 | 16.17 | 7.95 | 53 | 9.2 | 108.9%
2001-07-18 | 10:46 | 16.37 | 7.99 | 53 | 9.25 | 109.5%
2001-07-18 | 11:01 | 16.53 | 8.02 | 53 | 9.27 | 110.4%
2001-07-18 | 11:16 | 16.9 | 8.07 | 53 | 9.3 | 110.7%
2001-07-18 | 11:31 | 17.29 | 8.12 | 53 | 9.32 | 112.3%
2001-07-18 | 11:46 | 17.57 | 8.16 | 53 | 9.32 | 113.7%
2001-07-18 | 12:01 | 17.86 | 8.2 | 53 | 9.34 | 113.9%
2001-07-18 | 12:16 | 18.14 | 8.23 | 53 | 9.34 | 114.6%
2001-07-18 | 12:31 | 18.48 | 8.26 | 53 | 9.34 | 114.6%
2001-07-18 | 12:46 | 18.71 | 8.29 | 53 | 9.33 | 115.9%
2001-07-18 | 13:01 | 19.01 | 8.32 | 53 | 9.33 | 118.1%
2001-07-18 | 13:16 | 19.21 | 8.35 | 52 | 9.31 | 117.8%
2001-07-18 | 13:31 | 19.5 | 8.37 | 52 | 9.28 | 118.2%
2001-07-18 | 13:46 | 19.75 | 8.39 | 52 | 9.27 | 118.1%
2001-07-18 | 14:01 | 19.92 | 8.41 | 52 | 9.23 | 117.6%
2001-07-18 | 14:16 | 20.19 | 8.44 | 52 | 9.23 | 118.3%
2001-07-18 | 14:31 | 20.34 | 8.46 | 52 | 9.21 | 118.1%
2001-07-18 | 14:46 | 20.53 | 8.43 | 52 | 9.14 | 118.7%
2001-07-18 | 15:01 | 20.52 | 8.45 | 52 | 9.11 | 118.3%
2001-07-18 | 15:16 | 20.53 | 8.49 | 52 | 9.09 | 118.1%
2001-07-18 | 15:31 | 20.52 | 8.51 | 52 | 9.06 | 117.7%
2001-07-18 | 15:46 | 20.48 | 8.52 | 52 | 9.04 | 115.9%
2001-07-18 | 16:01 | 20.49 | 8.54 | 52 | 9.05 | 116%
2001-07-18 | 16:16 | 20.55 | 8.55 | 51 | 9.03 | 117.3%
2001-07-18 | 16:31 | 20.6 | 8.57 | 51 | 9.01 | 117%
2001-07-18 | 16:46 | 20.69 | 8.58 | 51 | 9 | 116.9%
2001-07-18 | 17:01 | 20.8 | 8.59 | 51 | 8.96 | 116.4%
2001-07-18 | 17:16 | 20.88 | 8.6 | 51 | 8.91 | 115.7%
2001-07-18 | 17:31 | 20.85 | 8.59 | 51 | 8.88 | 115.3%
2001-07-18 | 17:46 | 20.82 | 8.58 | 51 | 8.83 | 114.7%
2001-07-18 | 18:01 | 20.79 | 8.58 | 51 | 8.8 | 114.3%
2001-07-18 | 18:16 | 20.73 | 8.59 | 51 | 8.78 | 114%
2001-07-18 | 18:31 | 20.66 | 8.58 | 50 | 8.76 | 113.8%
2001-07-18 | 18:46 | 20.48 | 8.57 | 50 | 8.71 | 111.7%
2001-07-18 | 19:01 | 20.33 | 8.55 | 50 | 8.65 | 110.9%
2001-07-18 | 19:16 | 20.16 | 8.53 | 50 | 8.61 | 110.4%
2001-07-18 | 19:31 | 19.99 | 8.51 | 50 | 8.56 | 109%
2001-07-18 | 19:46 | 19.8 | 8.48 | 50 | 8.51 | 108.4%
2001-07-18 | 20:01 | 19.63 | 8.45 | 50 | 8.45 | 107.6%
2001-07-18 | 20:16 | 19.44 | 8.4 | 50 | 8.38 | 106.1%
2001-07-18 | 20:31 | 19.24 | 8.34 | 50 | 8.3 | 105.1%
2001-07-18 | 20:46 | 19.07 | 8.27 | 50 | 8.21 | 103.9%
2001-07-18 | 21:01 | 18.87 | 8.2 | 50 | 8.14 | 101.1%
2001-07-18 | 21:16 | 18.7 | 8.12 | 51 | 8.06 | 100.1%
2001-07-18 | 21:31 | 18.56 | 8.03 | 51 | 8.01 | 99.5%
2001-07-18 | 21:46 | 18.4 | 7.95 | 51 | 7.94 | 97.42%
2001-07-18 | 22:01 | 18.26 | 7.88 | 51 | 7.89 | 96.81%
2001-07-18 | 22:16 | 18.09 | 7.81 | 51 | 7.86 | 96.44%
2001-07-18 | 22:31 | 17.92 | 7.74 | 51 | 7.82 | 95.37%
2001-07-18 | 22:46 | 17.75 | 7.7 | 51 | 7.79 | 95%
2001-07-18 | 23:01 | 17.59 | 7.66 | 51 | 7.78 | 94.88%
2001-07-18 | 23:16 | 17.43 | 7.63 | 51 | 7.76 | 93.49%
2001-07-18 | 23:31 | 17.28 | 7.59 | 51 | 7.75 | 93.37%
2001-07-18 | 23:46 | 17.13 | 7.56 | 51 | 7.75 | 93.37%
2001-07-19 | 0:01 | 16.99 | 7.54 | 52 | 7.74 | 92.14%
2001-07-19 | 0:16 | 16.85 | 7.52 | 52 | 7.75 | 92.26%
2001-07-19 | 0:31 | 16.7 | 7.51 | 52 | 7.75 | 92.26%
2001-07-19 | 0:46 | 16.56 | 7.49 | 52 | 7.76 | 92.38%
2001-07-19 | 1:01 | 16.42 | 7.48 | 52 | 7.77 | 91.95%
2001-07-19 | 1:16 | 16.29 | 7.47 | 52 | 7.77 | 91.95%
2001-07-19 | 1:31 | 16.16 | 7.46 | 52 | 7.79 | 92.19%
2001-07-19 | 1:46 | 16.03 | 7.46 | 52 | 7.8 | 92.31%
2001-07-19 | 2:01 | 15.91 | 7.45 | 52 | 7.82 | 91.46%
2001-07-19 | 2:16 | 15.78 | 7.44 | 52 | 7.83 | 91.58%
2001-07-19 | 2:31 | 15.66 | 7.44 | 52 | 7.84 | 91.7%
2001-07-19 | 2:46 | 15.53 | 7.43 | 52 | 7.86 | 91.93%
2001-07-19 | 3:01 | 15.42 | 7.43 | 52 | 7.88 | 90.57%
2001-07-19 | 3:16 | 15.3 | 7.42 | 52 | 7.89 | 90.69%
2001-07-19 | 3:31 | 15.19 | 7.42 | 52 | 7.92 | 91.03%
2001-07-19 | 3:46 | 15.09 | 7.41 | 52 | 7.93 | 91.15%
2001-07-19 | 4:01 | 15 | 7.41 | 52 | 7.95 | 91.38%
2001-07-19 | 4:16 | 14.91 | 7.41 | 52 | 7.96 | 90.97%
2001-07-19 | 4:31 | 14.81 | 7.41 | 52 | 8 | 91.43%
2001-07-19 | 4:46 | 14.72 | 7.41 | 52 | 8.01 | 91.54%
2001-07-19 | 5:01 | 14.62 | 7.41 | 53 | 8.03 | 91.77%
2001-07-19 | 5:16 | 14.51 | 7.4 | 53 | 8.05 | 92%
2001-07-19 | 5:31 | 14.4 | 7.4 | 53 | 8.08 | 91.82%
2001-07-19 | 5:46 | 14.3 | 7.4 | 53 | 8.1 | 92.05%
2001-07-19 | 6:01 | 14.21 | 7.4 | 53 | 8.14 | 92.5%
2001-07-19 | 6:16 | 14.11 | 7.4 | 53 | 8.17 | 92.84%
2001-07-19 | 6:31 | 14.03 | 7.41 | 53 | 8.22 | 93.41%
2001-07-19 | 6:46 | 13.94 | 7.42 | 53 | 8.28 | 93.03%
2001-07-19 | 7:01 | 13.86 | 7.43 | 53 | 8.35 | 93.82%
2001-07-19 | 7:16 | 13.81 | 7.44 | 53 | 8.43 | 94.72%
2001-07-19 | 7:31 | 13.77 | 7.46 | 54 | 8.51 | 95.62%
2001-07-19 | 7:46 | 13.77 | 7.49 | 54 | 8.61 | 96.74%
2001-07-19 | 8:01 | 13.83 | 7.52 | 54 | 8.71 | 97.87%
2001-07-19 | 8:16 | 13.93 | 7.56 | 54 | 8.81 | 98.99%
2001-07-19 | 8:31 | 14.08 | 7.59 | 54 | 8.91 | 101.2%
2001-07-19 | 8:46 | 14.16 | 7.63 | 54 | 8.99 | 102.2%
2001-07-19 | 9:01 | 14.27 | 7.67 | 54 | 9.07 | 103.1%
2001-07-19 | 9:16 | 14.43 | 7.71 | 54 | 9.14 | 103.9%
2001-07-19 | 9:31 | 14.62 | 7.75 | 54 | 9.22 | 105.4%
2001-07-19 | 9:46 | 14.81 | 7.79 | 54 | 9.29 | 106.2%
2001-07-19 | 10:01 | 15.04 | 7.84 | 54 | 9.36 | 107.6%
2001-07-19 | 10:16 | 15.29 | 7.89 | 53 | 9.41 | 108.2%
2001-07-19 | 10:31 | 15.66 | 7.94 | 53 | 9.45 | 110.5%
2001-07-19 | 10:46 | 15.98 | 7.99 | 53 | 9.48 | 110.9%
2001-07-19 | 11:01 | 16.41 | 8.03 | 53 | 9.51 | 112.5%
2001-07-19 | 11:16 | 16.64 | 8.07 | 53 | 9.52 | 113.3%
2001-07-19 | 11:31 | 16.74 | 8.1 | 53 | 9.5 | 113.1%
2001-07-19 | 11:46 | 16.9 | 8.14 | 53 | 9.5 | 113.1%
2001-07-19 | 12:01 | 17.08 | 8.18 | 53 | 9.5 | 114.5%
2001-07-19 | 12:16 | 17.37 | 8.22 | 53 | 9.5 | 114.5%
2001-07-19 | 12:31 | 17.74 | 8.25 | 53 | 9.5 | 115.9%
2001-07-19 | 12:46 | 18.13 | 8.29 | 53 | 9.49 | 116.4%
2001-07-19 | 13:01 | 18.39 | 8.32 | 53 | 9.47 | 116.2%
2001-07-19 | 13:16 | 18.56 | 8.33 | 52 | 9.43 | 117.1%
2001-07-19 | 13:31 | 18.56 | 8.33 | 52 | 9.36 | 116.3%
2001-07-19 | 13:46 | 18.53 | 8.33 | 52 | 9.3 | 115.5%
2001-07-19 | 14:01 | 18.34 | 8.32 | 52 | 9.22 | 113.1%
2001-07-19 | 14:16 | 18.16 | 8.29 | 52 | 9.15 | 112.3%
2001-07-19 | 14:31 | 18.17 | 8.3 | 52 | 9.13 | 112%
2001-07-19 | 14:46 | 18.29 | 8.32 | 52 | 9.15 | 112.3%
2001-07-19 | 15:01 | 18.23 | 8.32 | 52 | 9.12 | 111.9%
2001-07-19 | 15:16 | 18.11 | 8.3 | 52 | 9.07 | 111.3%
2001-07-19 | 15:31 | 17.99 | 8.28 | 52 | 9.04 | 110.2%
2001-07-19 | 15:46 | 17.88 | 8.26 | 52 | 8.98 | 109.5%
2001-07-19 | 16:01 | 17.78 | 8.23 | 52 | 8.93 | 108.9%
2001-07-19 | 16:16 | 17.91 | 8.24 | 52 | 8.95 | 109.1%
2001-07-19 | 16:31 | 17.99 | 8.26 | 52 | 9 | 109.8%
2001-07-19 | 16:46 | 17.95 | 8.26 | 52 | 9 | 109.8%
2001-07-19 | 17:01 | 17.88 | 8.24 | 52 | 8.97 | 109.4%
2001-07-19 | 17:16 | 17.75 | 8.22 | 51 | 8.93 | 108.9%
2001-07-19 | 17:31 | 17.65 | 8.21 | 51 | 8.91 | 108.7%
2001-07-19 | 17:46 | 17.74 | 8.21 | 51 | 8.93 | 108.9%
2001-07-19 | 18:01 | 17.81 | 8.23 | 51 | 8.97 | 109.4%
2001-07-19 | 18:16 | 17.8 | 8.24 | 51 | 8.98 | 109.5%
2001-07-19 | 18:31 | 17.79 | 8.24 | 51 | 8.96 | 109.3%
2001-07-19 | 18:46 | 17.75 | 8.22 | 51 | 8.93 | 108.9%
2001-07-19 | 19:01 | 17.68 | 8.2 | 51 | 8.88 | 108.3%
2001-07-19 | 19:16 | 17.66 | 8.19 | 51 | 8.86 | 108%
2001-07-19 | 19:31 | 17.62 | 8.17 | 51 | 8.86 | 108%
2001-07-19 | 19:46 | 17.54 | 8.16 | 51 | 8.84 | 107.8%
2001-07-19 | 20:01 | 17.51 | 8.14 | 51 | 8.78 | 107.1%
2001-07-19 | 20:16 | 17.45 | 8.1 | 51 | 8.71 | 104.9%
2001-07-19 | 20:31 | 17.37 | 8.06 | 51 | 8.64 | 104.1%
2001-07-19 | 20:46 | 17.25 | 8 | 51 | 8.56 | 103.1%
2001-07-19 | 21:01 | 17.14 | 7.94 | 51 | 8.48 | 102.2%
2001-07-19 | 21:16 | 17.04 | 7.88 | 51 | 8.41 | 101.3%
2001-07-19 | 21:31 | 16.97 | 7.83 | 51 | 8.34 | 99.29%
2001-07-19 | 21:46 | 16.91 | 7.78 | 51 | 8.29 | 98.69%
2001-07-19 | 22:01 | 16.82 | 7.75 | 51 | 8.23 | 97.98%
2001-07-19 | 22:16 | 16.72 | 7.71 | 51 | 8.18 | 97.38%
2001-07-19 | 22:31 | 16.63 | 7.67 | 51 | 8.14 | 96.9%
2001-07-19 | 22:46 | 16.52 | 7.64 | 51 | 8.11 | 96.55%
2001-07-19 | 23:01 | 16.42 | 7.62 | 51 | 8.08 | 95.62%
2001-07-19 | 23:16 | 16.32 | 7.59 | 52 | 8.06 | 95.38%
2001-07-19 | 23:31 | 16.22 | 7.57 | 51 | 8.04 | 95.15%
2001-07-19 | 23:46 | 16.11 | 7.55 | 51 | 8.03 | 95.03%
2001-07-20 | 0:01 | 16 | 7.53 | 52 | 8.02 | 94.91%
2001-07-20 | 0:16 | 15.88 | 7.52 | 52 | 8.02 | 93.8%
2001-07-20 | 0:31 | 15.76 | 7.5 | 52 | 8.02 | 93.8%
2001-07-20 | 0:46 | 15.63 | 7.49 | 52 | 8.03 | 93.92%
2001-07-20 | 1:01 | 15.53 | 7.48 | 52 | 8.03 | 93.92%
2001-07-20 | 1:16 | 15.42 | 7.47 | 52 | 8.03 | 92.3%
2001-07-20 | 1:31 | 15.31 | 7.45 | 52 | 8.03 | 92.3%
2001-07-20 | 1:46 | 15.22 | 7.45 | 52 | 8.04 | 92.41%
2001-07-20 | 2:01 | 15.15 | 7.44 | 52 | 8.05 | 92.53%
2001-07-20 | 2:16 | 15.08 | 7.43 | 52 | 8.05 | 92.53%
2001-07-20 | 2:31 | 15.03 | 7.43 | 52 | 8.06 | 92.64%
2001-07-20 | 2:46 | 14.96 | 7.43 | 52 | 8.08 | 92.34%
2001-07-20 | 3:01 | 14.87 | 7.42 | 52 | 8.09 | 92.46%
2001-07-20 | 3:16 | 14.77 | 7.42 | 52 | 8.09 | 92.46%
2001-07-20 | 3:31 | 14.66 | 7.42 | 52 | 8.12 | 92.8%
2001-07-20 | 3:46 | 14.55 | 7.41 | 52 | 8.13 | 92.91%
2001-07-20 | 4:01 | 14.45 | 7.41 | 52 | 8.15 | 92.61%
2001-07-20 | 4:16 | 14.36 | 7.41 | 52 | 8.17 | 92.84%
2001-07-20 | 4:31 | 14.28 | 7.4 | 52 | 8.19 | 93.07%
2001-07-20 | 4:46 | 14.22 | 7.41 | 52 | 8.21 | 93.3%
2001-07-20 | 5:01 | 14.16 | 7.4 | 52 | 8.22 | 93.41%
2001-07-20 | 5:16 | 14.1 | 7.4 | 52 | 8.23 | 93.52%
2001-07-20 | 5:31 | 14.04 | 7.4 | 52 | 8.25 | 93.75%
2001-07-20 | 5:46 | 13.98 | 7.4 | 52 | 8.25 | 92.7%
2001-07-20 | 6:01 | 13.93 | 7.4 | 53 | 8.28 | 93.03%
2001-07-20 | 6:16 | 13.88 | 7.4 | 53 | 8.29 | 93.15%
2001-07-20 | 6:31 | 13.84 | 7.4 | 53 | 8.32 | 93.48%
2001-07-20 | 6:46 | 13.8 | 7.41 | 53 | 8.34 | 93.71%
2001-07-20 | 7:01 | 13.77 | 7.41 | 53 | 8.38 | 94.16%
2001-07-20 | 7:16 | 13.74 | 7.42 | 53 | 8.4 | 94.38%
2001-07-20 | 7:31 | 13.73 | 7.42 | 53 | 8.44 | 94.83%
2001-07-20 | 7:46 | 13.71 | 7.43 | 53 | 8.46 | 95.06%
2001-07-20 | 8:01 | 13.69 | 7.44 | 53 | 8.5 | 95.51%
2001-07-20 | 8:16 | 13.73 | 7.46 | 53 | 8.58 | 96.4%
2001-07-20 | 8:31 | 13.87 | 7.49 | 53 | 8.7 | 97.75%
2001-07-20 | 8:46 | 14.07 | 7.54 | 53 | 8.82 | 100.2%
2001-07-20 | 9:01 | 14.32 | 7.59 | 53 | 8.94 | 101.6%
2001-07-20 | 9:16 | 14.46 | 7.63 | 53 | 9.01 | 102.4%
2001-07-20 | 9:31 | 14.65 | 7.67 | 53 | 9.08 | 103.8%
2001-07-20 | 9:46 | 14.77 | 7.7 | 53 | 9.14 | 104.5%
2001-07-20 | 10:01 | 14.88 | 7.73 | 53 | 9.19 | 105%