## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2001-08-13 15:02 | 17.3 | 8.41 | 8.6 | 101 | 64 | 0
Sonde | 2001-08-13 15:00 | 17.32 | 7.74 | 9.35 | 110 | 64
Grade |  | A | Fail | B |  | A
Audit 2 | 2001-08-15 12:54 | 17.2 | 8.31 | 8.7 | 102 | 64 | 0
Sonde | 2001-08-15 13:00 | 17.28 | 8.06 | 9.54 | 112 | 63
Grade |  | A | A | B |  | A
Audit 3 | 2001-08-17 13:51 | 17.5 | 8.3 | 8.5 | 101 | 64 | 0
Sonde | 2001-08-17 13:35 | 17.43 | 8.11 | 9.5 | 112 | 63
Grade |  | A | A | B |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2001-08-17 |  |  |  | F
Time | Computer set to atomic time |  | 15:00 |  |  |  | F
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution |  | 7 |  | 7.07 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Data Note: Continous Data collected week of 8/14/2001 due to failure of equipment at same site week of 7/23/2001. Sonde calibration, auditing and data collection performed by Bonnie Lamb of ER-Bend office. Formatting Note: Auits 1-3 represent the duration of the deployment and are marco (ed) into the main audit sheet. Audits 4-5 are during the middle of the deployment and are represented only by audit sheets 4 and 5. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-08-13 15:02 | 17.3 | 8.41 | 8.6 | 101 | 64
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
NOTE: May not have added enough sulfuric acid - pipette not graduated but didn't look like enough |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-08-15 12:54 | 17.2 | 8.31 | 8.7 | 102 | 64
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
NOTE: May not have added enough sulfuric acid - pipette not graduated but didn't look like enough |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-08-17 13:51 | 17.5 | 8.3 | 8.5 | 101 | 64
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
NOTE: May not have added enough sulfuric acid - pipette not graduated but didn't look like enough |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-08-14 11:14 | 16.1 | 8.12 | 9 | 103 | 64
Sonde Values | 2001-08-14 11:15 | 16.14 | 7.87 | 9.67 | 111 | 63
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
NOTE: May not have added enough sulfuric acid - pipette not graduated but didn't look like enough |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-08-16 13:35 | 17.6 | 8.35 | 8.9 | 105 | 64
Sonde Values | 2001-08-16 13:30 | 17.57 | 8.12 | 9.48 | 112 | 63
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
NOTE: May not have added enough sulfuric acid - pipette not graduated but didn't look like enough |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | SpCond | Cond | DOsat | DO Conc | DOchrg | Depth | pH | pH | Battery
|  | °C | uS/cm | uS/cm | % | mg/L |  | feet |  | mV | volts
2001-08-13 | 9:31 | 23.63 | 3 | 3 | 91.7 | 7.77 | 60.4 | 0.106 | 8.82 | -151.9 | 5.6
2001-08-13 | 9:45 | 23.93 | 3 | 3 | 90.1 | 7.6 | 60.4 | 0.102 | 8.72 | -146.4 | 5.6
2001-08-13 | 10:00 | 23.96 | 4 | 4 | 90 | 7.58 | 60.4 | 0.13 | 8.65 | -142 | 5.6
2001-08-13 | 10:15 | 22.21 | 4 | 3 | 90.1 | 7.85 | 60.4 | 0.643 | 8.45 | -130 | 5.5
2001-08-13 | 10:30 | 22.44 | 3 | 3 | 89.7 | 7.78 | 60.4 | 1.039 | 8.49 | -132.4 | 5.5
2001-08-13 | 10:45 | 23.78 | 3 | 3 | 89.6 | 7.58 | 60.9 | 1.252 | 8.5 | -133.7 | 5.5
2001-08-13 | 11:00 | 25.78 | 3 | 3 | 89.6 | 7.3 | 60.9 | 1.25 | 8.46 | -131.5 | 5.5
2001-08-13 | 11:15 | 27.66 | 3 | 3 | 89.8 | 7.07 | 62.1 | 1.249 | 8.39 | -128 | 5.5
2001-08-13 | 11:30 | 29.33 | 3 | 3 | 90.1 | 6.89 | 64.5 | 1.247 | 8.32 | -124.3 | 5.5
2001-08-13 | 11:45 | 30.79 | 3 | 3 | 90.3 | 6.73 | 66.2 | 1.245 | 8.27 | -122 | 5.5
2001-08-13 | 12:00 | 32.08 | 4 | 4 | 90.5 | 6.6 | 68.6 | 1.243 | 8.24 | -120.4 | 5.5
2001-08-13 | 12:15 | 31.93 | 4 | 5 | 90.5 | 6.62 | 69.1 | 0.884 | 8.13 | -114 | 5.5
2001-08-13 | 12:30 | 29.55 | 4 | 5 | 90.2 | 6.87 | 68.6 | 0.694 | 7.99 | -105.2 | 5.5
2001-08-13 | 12:45 | 28.48 | 5 | 5 | 90 | 6.99 | 68.6 | 0.164 | 7.97 | -103.8 | 5.5
2001-08-13 | 13:00 | 30.58 | 5 | 5 | 90.2 | 6.75 | 68.6 | 0.158 | 8.11 | -112.5 | 5.4
2001-08-13 | 13:15 | 34.04 | 6 | 7 | 90.5 | 6.39 | 70.3 | 0.154 | 8.16 | -116 | 5.4
2001-08-13 | 13:30 | 37.55 | 8 | 10 | 91.1 | 6.08 | 74.4 | 0.155 | 8.2 | -119.1 | 5.4
2001-08-13 | 13:45 | 39.83 | 18 | 23 | 91.6 | 5.89 | 76.2 | 0.149 | 8.16 | -117.5 | 5.4
2001-08-13 | 14:00 | 41.39 | 28 | 36 | 91.7 | 5.76 | 77.3 | 0.138 | 8.14 | -116.3 | 5.4
2001-08-13 | 14:15 | 42.37 | 31 | 41 | 91.9 | 5.68 | 77.3 | 0.133 | 8.11 | -115.1 | 5.4
2001-08-13 | 14:30 | 42.96 | 30 | 41 | 92.3 | 5.65 | 77.3 | 0.133 | 8.11 | -114.9 | 5.3
2001-08-13 | 14:45 | 42.58 | 29 | 38 | 92.2 | 5.68 | 78.5 | 0.135 | 8.07 | -112.4 | 5.2
2001-08-13 | 15:01 | 17.32 | 64 | 54 | 97.4 | 9.35 | 82.6 | 0.159 | 7.74 | -88.6 | 5.5
2001-08-13 | 15:16 | 17.38 | 64 | 54 | 97.4 | 9.34 | 79.7 | 0.13 | 8.01 | -103.9 | 5.5
2001-08-13 | 15:31 | 17.44 | 63 | 54 | 97.1 | 9.29 | 78.5 | 0.122 | 8.08 | -107.8 | 5.5
2001-08-13 | 15:46 | 17.5 | 63 | 54 | 97.1 | 9.29 | 77.3 | 0.139 | 8.16 | -112.2 | 5.5
2001-08-13 | 16:01 | 17.53 | 63 | 54 | 96.7 | 9.25 | 76.2 | 0.138 | 8.21 | -115.4 | 5.5
2001-08-13 | 16:16 | 17.57 | 63 | 54 | 96.3 | 9.19 | 75.6 | 0.151 | 8.24 | -116.9 | 5.5
2001-08-13 | 16:31 | 17.61 | 63 | 54 | 95.7 | 9.13 | 75.6 | 0.136 | 8.26 | -118 | 5.5
2001-08-13 | 16:46 | 17.65 | 63 | 54 | 95.2 | 9.08 | 74.4 | 0.113 | 8.28 | -119 | 5.5
2001-08-13 | 17:01 | 17.65 | 63 | 54 | 94.8 | 9.04 | 74.4 | 0.126 | 8.28 | -119.1 | 5.5
2001-08-13 | 17:16 | 17.66 | 63 | 54 | 94.5 | 9.01 | 73.2 | 0.128 | 8.28 | -119 | 5.5
2001-08-13 | 17:31 | 17.67 | 63 | 55 | 94.2 | 8.98 | 73.2 | 0.122 | 8.27 | -118.8 | 5.5
2001-08-13 | 17:46 | 17.66 | 64 | 55 | 93.8 | 8.94 | 73.2 | 0.118 | 8.26 | -117.8 | 5.5
2001-08-13 | 18:01 | 17.66 | 64 | 55 | 93.2 | 8.88 | 72.7 | 0.131 | 8.23 | -116.3 | 5.5
2001-08-13 | 18:16 | 17.66 | 64 | 55 | 92.9 | 8.86 | 72.7 | 0.144 | 8.22 | -116 | 5.5
2001-08-13 | 18:31 | 17.65 | 64 | 55 | 92.3 | 8.8 | 72.7 | 0.153 | 8.2 | -114.4 | 5.5
2001-08-13 | 18:46 | 17.64 | 64 | 55 | 91.8 | 8.75 | 72.7 | 0.165 | 8.16 | -112.4 | 5.5
2001-08-13 | 19:01 | 17.62 | 64 | 55 | 91.5 | 8.73 | 71.5 | 0.168 | 8.14 | -111.1 | 5.5
2001-08-13 | 19:16 | 17.58 | 64 | 55 | 91.2 | 8.71 | 71.5 | 0.166 | 8.1 | -109 | 5.5
2001-08-13 | 19:31 | 17.54 | 64 | 55 | 90.9 | 8.69 | 71.5 | 0.173 | 8.06 | -106.9 | 5.5
2001-08-13 | 19:46 | 17.5 | 64 | 55 | 90.8 | 8.68 | 71.5 | 0.173 | 8.02 | -104.4 | 5.5
2001-08-13 | 20:01 | 17.44 | 64 | 55 | 90.6 | 8.67 | 70.3 | 0.177 | 7.98 | -102.1 | 5.5
2001-08-13 | 20:16 | 17.39 | 64 | 55 | 90.5 | 8.67 | 70.3 | 0.187 | 7.93 | -99.4 | 5.5
2001-08-13 | 20:31 | 17.33 | 64 | 55 | 90.2 | 8.66 | 70.3 | 0.192 | 7.88 | -96.6 | 5.5
2001-08-13 | 20:46 | 17.27 | 64 | 54 | 90.1 | 8.66 | 70.3 | 0.198 | 7.84 | -94.3 | 5.5
2001-08-13 | 21:01 | 17.19 | 64 | 54 | 90.1 | 8.67 | 70.3 | 0.2 | 7.8 | -92 | 5.5
2001-08-13 | 21:16 | 17.11 | 64 | 54 | 90 | 8.68 | 70.3 | 0.204 | 7.77 | -90.1 | 5.5
2001-08-13 | 21:31 | 17.03 | 64 | 54 | 90.1 | 8.7 | 70.3 | 0.212 | 7.72 | -87.6 | 5.5
2001-08-13 | 21:46 | 16.94 | 64 | 54 | 90.1 | 8.71 | 70.3 | 0.215 | 7.7 | -86 | 5.5
2001-08-13 | 22:01 | 16.85 | 64 | 54 | 90 | 8.72 | 69.1 | 0.225 | 7.67 | -84.5 | 5.5
2001-08-13 | 22:16 | 16.77 | 64 | 54 | 90.1 | 8.75 | 69.1 | 0.23 | 7.65 | -83.3 | 5.5
2001-08-13 | 22:31 | 16.68 | 64 | 54 | 90.2 | 8.78 | 69.1 | 0.236 | 7.63 | -82 | 5.5
2001-08-13 | 22:46 | 16.59 | 64 | 54 | 90.3 | 8.8 | 69.1 | 0.24 | 7.61 | -81 | 5.5
2001-08-13 | 23:01 | 16.5 | 64 | 54 | 90.3 | 8.82 | 69.1 | 0.244 | 7.59 | -79.9 | 5.5
2001-08-13 | 23:16 | 16.4 | 64 | 54 | 90.4 | 8.84 | 69.1 | 0.242 | 7.57 | -78.9 | 5.4
2001-08-13 | 23:31 | 16.31 | 64 | 53 | 90.5 | 8.87 | 69.1 | 0.25 | 7.56 | -78.2 | 5.5
2001-08-13 | 23:46 | 16.22 | 64 | 53 | 90.6 | 8.9 | 69.1 | 0.251 | 7.54 | -77.4 | 5.5
2001-08-14 | 0:01 | 16.14 | 64 | 53 | 90.6 | 8.92 | 69.1 | 0.249 | 7.53 | -76.4 | 5.5
2001-08-14 | 0:16 | 16.05 | 64 | 53 | 90.7 | 8.94 | 69.1 | 0.249 | 7.52 | -76.1 | 5.4
2001-08-14 | 0:31 | 15.98 | 64 | 53 | 90.7 | 8.96 | 69.1 | 0.25 | 7.51 | -75.6 | 5.4
2001-08-14 | 0:46 | 15.91 | 64 | 53 | 90.8 | 8.97 | 68.6 | 0.25 | 7.5 | -74.9 | 5.4
2001-08-14 | 1:01 | 15.84 | 64 | 53 | 90.7 | 8.99 | 68.6 | 0.248 | 7.5 | -74.6 | 5.5
2001-08-14 | 1:16 | 15.78 | 64 | 53 | 90.8 | 9 | 68.6 | 0.248 | 7.49 | -74.2 | 5.4
2001-08-14 | 1:31 | 15.73 | 64 | 52 | 90.9 | 9.02 | 68.6 | 0.246 | 7.49 | -74 | 5.4
2001-08-14 | 1:46 | 15.68 | 64 | 52 | 90.8 | 9.03 | 68.6 | 0.243 | 7.47 | -73.4 | 5.4
2001-08-14 | 2:01 | 15.63 | 64 | 52 | 90.8 | 9.03 | 68.6 | 0.237 | 7.47 | -73.2 | 5.4
2001-08-14 | 2:16 | 15.59 | 64 | 52 | 90.9 | 9.05 | 68.6 | 0.236 | 7.46 | -72.8 | 5.4
2001-08-14 | 2:31 | 15.56 | 63 | 52 | 90.9 | 9.06 | 68.6 | 0.239 | 7.46 | -72.6 | 5.4
2001-08-14 | 2:46 | 15.52 | 63 | 52 | 90.9 | 9.06 | 68.6 | 0.248 | 7.45 | -72.3 | 5.4
2001-08-14 | 3:01 | 15.49 | 63 | 52 | 90.9 | 9.07 | 68.6 | 0.243 | 7.46 | -72.3 | 5.4
2001-08-14 | 3:16 | 15.46 | 63 | 52 | 90.9 | 9.08 | 68.6 | 0.244 | 7.45 | -71.9 | 5.4
2001-08-14 | 3:31 | 15.44 | 63 | 52 | 90.8 | 9.07 | 68.6 | 0.247 | 7.45 | -71.9 | 5.4
2001-08-14 | 3:46 | 15.41 | 63 | 52 | 90.9 | 9.08 | 67.4 | 0.248 | 7.44 | -71.6 | 5.4
2001-08-14 | 4:01 | 15.39 | 63 | 52 | 90.9 | 9.09 | 68.6 | 0.251 | 7.44 | -71.5 | 5.4
2001-08-14 | 4:16 | 15.38 | 63 | 52 | 90.9 | 9.1 | 67.4 | 0.246 | 7.44 | -71.5 | 5.4
2001-08-14 | 4:31 | 15.36 | 63 | 52 | 90.8 | 9.09 | 67.4 | 0.248 | 7.44 | -71.3 | 5.4
2001-08-14 | 4:46 | 15.35 | 63 | 52 | 91 | 9.1 | 67.4 | 0.25 | 7.44 | -71.3 | 5.4
2001-08-14 | 5:01 | 15.33 | 63 | 51 | 90.9 | 9.1 | 67.4 | 0.25 | 7.43 | -70.9 | 5.4
2001-08-14 | 5:16 | 15.32 | 63 | 51 | 91 | 9.11 | 67.4 | 0.254 | 7.43 | -71.1 | 5.4
2001-08-14 | 5:31 | 15.3 | 63 | 51 | 91 | 9.11 | 67.4 | 0.256 | 7.43 | -70.9 | 5.4
2001-08-14 | 5:46 | 15.29 | 63 | 51 | 91 | 9.12 | 67.4 | 0.255 | 7.43 | -71 | 5.4
2001-08-14 | 6:01 | 15.28 | 63 | 51 | 91 | 9.13 | 67.4 | 0.255 | 7.43 | -70.9 | 5.4
2001-08-14 | 6:16 | 15.28 | 63 | 51 | 91.1 | 9.13 | 67.4 | 0.256 | 7.43 | -71.1 | 5.4
2001-08-14 | 6:31 | 15.28 | 63 | 51 | 91.2 | 9.14 | 67.4 | 0.257 | 7.43 | -71 | 5.4
2001-08-14 | 6:46 | 15.28 | 63 | 51 | 91.5 | 9.17 | 67.4 | 0.261 | 7.43 | -71.1 | 5.4
2001-08-14 | 7:01 | 15.28 | 63 | 51 | 91.7 | 9.19 | 67.4 | 0.27 | 7.44 | -71.7 | 5.4
2001-08-14 | 7:16 | 15.27 | 63 | 51 | 92 | 9.22 | 67.4 | 0.274 | 7.46 | -72.3 | 5.4
2001-08-14 | 7:31 | 15.28 | 63 | 51 | 92.4 | 9.26 | 67.4 | 0.278 | 7.46 | -72.5 | 5.4
2001-08-14 | 7:46 | 15.28 | 63 | 51 | 92.7 | 9.29 | 67.4 | 0.281 | 7.47 | -73 | 5.4
2001-08-14 | 8:01 | 15.3 | 63 | 51 | 93.1 | 9.33 | 67.4 | 0.273 | 7.49 | -74 | 5.4
2001-08-14 | 8:16 | 15.32 | 63 | 51 | 93.5 | 9.37 | 67.4 | 0.274 | 7.5 | -75 | 5.4
2001-08-14 | 8:31 | 15.36 | 63 | 51 | 94 | 9.41 | 67.4 | 0.25 | 7.52 | -76.2 | 5.4
2001-08-14 | 8:46 | 15.42 | 63 | 51 | 94.7 | 9.47 | 67.4 | 0.221 | 7.55 | -77.8 | 5.4
2001-08-14 | 9:01 | 15.48 | 63 | 51 | 95.3 | 9.51 | 67.4 | 0.199 | 7.58 | -79.5 | 5.4
2001-08-14 | 9:16 | 15.54 | 63 | 51 | 95.9 | 9.56 | 67.4 | 0.183 | 7.61 | -81 | 5.4
2001-08-14 | 9:31 | 15.6 | 63 | 51 | 96.3 | 9.59 | 67.4 | 0.184 | 7.63 | -82.3 | 5.5
2001-08-14 | 9:46 | 15.66 | 63 | 51 | 96.7 | 9.61 | 67.4 | 0.188 | 7.67 | -84.3 | 5.5
2001-08-14 | 10:01 | 15.72 | 63 | 51 | 97.2 | 9.65 | 67.4 | 0.187 | 7.7 | -86 | 5.5
2001-08-14 | 10:16 | 15.8 | 63 | 52 | 97.5 | 9.67 | 67.4 | 0.185 | 7.73 | -87.8 | 5.5
2001-08-14 | 10:31 | 15.88 | 63 | 52 | 97.9 | 9.69 | 67.4 | 0.183 | 7.76 | -89.4 | 5.5
2001-08-14 | 10:46 | 15.96 | 63 | 52 | 98.1 | 9.69 | 67.4 | 0.18 | 7.8 | -91.5 | 5.5
2001-08-14 | 11:01 | 16.05 | 63 | 52 | 98.2 | 9.69 | 67.4 | 0.187 | 7.83 | -93.4 | 5.5
2001-08-14 | 11:16 | 16.14 | 63 | 52 | 98.3 | 9.67 | 67.4 | 0.163 | 7.87 | -95.6 | 5.5
2001-08-14 | 11:31 | 16.25 | 62 | 52 | 98.5 | 9.67 | 67.4 | 0.16 | 7.91 | -97.8 | 5.5
2001-08-14 | 11:46 | 16.35 | 62 | 52 | 98.7 | 9.67 | 67.4 | 0.159 | 7.94 | -99.8 | 5.5
2001-08-14 | 12:01 | 16.43 | 63 | 52 | 98.7 | 9.65 | 67.4 | 0.177 | 7.97 | -101.1 | 5.5
2001-08-14 | 12:16 | 16.54 | 63 | 52 | 98.7 | 9.63 | 67.4 | 0.154 | 8 | -102.9 | 5.5
2001-08-14 | 12:31 | 16.66 | 63 | 53 | 98.7 | 9.61 | 67.4 | 0.134 | 8.04 | -105.5 | 5.5
2001-08-14 | 12:46 | 16.76 | 63 | 53 | 98.9 | 9.6 | 67.4 | 0.133 | 8.07 | -107.1 | 5.5
2001-08-14 | 13:01 | 16.84 | 63 | 53 | 98.7 | 9.57 | 67.4 | 0.157 | 8.09 | -108.2 | 5.5
2001-08-14 | 13:16 | 16.93 | 63 | 53 | 98.7 | 9.55 | 67.4 | 0.129 | 8.11 | -109.4 | 5.5
2001-08-14 | 13:31 | 17.03 | 63 | 53 | 98.7 | 9.53 | 67.4 | 0.082 | 8.14 | -111.1 | 5.5
2001-08-14 | 13:46 | 17.12 | 63 | 53 | 98.6 | 9.51 | 67.4 | 0.092 | 8.16 | -112.5 | 5.5
2001-08-14 | 14:01 | 17.2 | 63 | 53 | 98.7 | 9.5 | 67.4 | 0.105 | 8.18 | -113.3 | 5.5
2001-08-14 | 14:16 | 17.28 | 63 | 53 | 98.5 | 9.47 | 67.4 | 0.103 | 8.19 | -114.3 | 5.5
2001-08-14 | 14:31 | 17.36 | 63 | 54 | 98.3 | 9.43 | 66.2 | 0.109 | 8.22 | -115.8 | 5.5
2001-08-14 | 14:46 | 17.43 | 63 | 54 | 97.9 | 9.38 | 66.2 | 0.132 | 8.23 | -116.5 | 5.5
2001-08-14 | 15:01 | 17.51 | 63 | 54 | 97.7 | 9.34 | 66.2 | 0.042 | 8.25 | -117.6 | 5.5
2001-08-14 | 15:16 | 17.58 | 63 | 54 | 97.5 | 9.31 | 66.2 | 0.075 | 8.28 | -119 | 5.5
2001-08-14 | 15:31 | 17.64 | 63 | 54 | 97.2 | 9.27 | 66.2 | 0.095 | 8.28 | -119.4 | 5.5
2001-08-14 | 15:46 | 17.71 | 63 | 54 | 97 | 9.23 | 66.2 | 0.027 | 8.29 | -119.9 | 5.5
2001-08-14 | 16:01 | 17.74 | 63 | 54 | 96.6 | 9.19 | 66.2 | 0.051 | 8.3 | -120.2 | 5.5
2001-08-14 | 16:16 | 17.75 | 63 | 54 | 96 | 9.13 | 66.2 | 0.083 | 8.29 | -119.6 | 5.5
2001-08-14 | 16:31 | 17.76 | 63 | 54 | 95.5 | 9.08 | 66.2 | 0.064 | 8.28 | -119.4 | 5.5
2001-08-14 | 16:46 | 17.78 | 63 | 54 | 95.1 | 9.04 | 66.2 | 0.063 | 8.28 | -119.4 | 5.5
2001-08-14 | 17:01 | 17.78 | 63 | 54 | 94.3 | 8.97 | 66.2 | 0.066 | 8.26 | -118.2 | 5.5
2001-08-14 | 17:16 | 17.79 | 63 | 54 | 94 | 8.94 | 66.2 | 0.069 | 8.25 | -117.8 | 5.5
2001-08-14 | 17:31 | 17.79 | 63 | 55 | 93.6 | 8.9 | 65 | 0.071 | 8.24 | -117.2 | 5.5
2001-08-14 | 17:46 | 17.8 | 63 | 55 | 93.4 | 8.88 | 65 | 0.09 | 8.22 | -116.1 | 5.5
2001-08-14 | 18:01 | 17.8 | 63 | 55 | 93.3 | 8.87 | 65 | 0.085 | 8.21 | -115.2 | 5.5
2001-08-14 | 18:16 | 17.8 | 63 | 55 | 93.1 | 8.85 | 65 | 0.032 | 8.2 | -114.6 | 5.5
2001-08-14 | 18:31 | 17.8 | 63 | 55 | 93.1 | 8.85 | 65 | 0.017 | 8.18 | -113.8 | 5.5
2001-08-14 | 18:46 | 17.8 | 63 | 55 | 92.9 | 8.83 | 65 | 0.029 | 8.17 | -113 | 5.5
2001-08-14 | 19:01 | 17.79 | 63 | 55 | 92.6 | 8.8 | 65 | 0.057 | 8.14 | -111.5 | 5.5
2001-08-14 | 19:16 | 17.79 | 63 | 55 | 92 | 8.75 | 65 | 0.085 | 8.1 | -109.3 | 5.5
2001-08-14 | 19:31 | 17.77 | 63 | 55 | 91.6 | 8.71 | 65 | 0.095 | 8.07 | -107.2 | 5.5
2001-08-14 | 19:46 | 17.74 | 63 | 55 | 91.2 | 8.68 | 65 | 0.101 | 8.03 | -104.8 | 5.5
2001-08-14 | 20:01 | 17.71 | 64 | 55 | 90.9 | 8.65 | 65 | 0.111 | 7.98 | -102.2 | 5.5
2001-08-14 | 20:16 | 17.67 | 64 | 55 | 90.5 | 8.62 | 65 | 0.119 | 7.93 | -99.6 | 5.5
2001-08-14 | 20:31 | 17.63 | 64 | 55 | 90.3 | 8.61 | 65 | 0.129 | 7.89 | -96.9 | 5.5
2001-08-14 | 20:46 | 17.59 | 64 | 55 | 90.2 | 8.61 | 65 | 0.139 | 7.85 | -94.6 | 5.4
2001-08-14 | 21:01 | 17.54 | 64 | 55 | 90.1 | 8.61 | 65 | 0.145 | 7.8 | -92.1 | 5.5
2001-08-14 | 21:16 | 17.49 | 64 | 54 | 90.1 | 8.62 | 65 | 0.147 | 7.77 | -90 | 5.4
2001-08-14 | 21:31 | 17.43 | 64 | 54 | 90.1 | 8.63 | 64.5 | 0.154 | 7.73 | -88 | 5.4
2001-08-14 | 21:46 | 17.37 | 64 | 54 | 90 | 8.63 | 64.5 | 0.151 | 7.71 | -86.6 | 5.4
2001-08-14 | 22:01 | 17.3 | 64 | 54 | 90.1 | 8.65 | 64.5 | 0.153 | 7.68 | -85 | 5.4
2001-08-14 | 22:16 | 17.23 | 64 | 54 | 90.2 | 8.67 | 64.5 | 0.156 | 7.65 | -83.6 | 5.4
2001-08-14 | 22:31 | 17.16 | 64 | 54 | 90.2 | 8.69 | 64.5 | 0.169 | 7.64 | -82.8 | 5.4
2001-08-14 | 22:46 | 17.09 | 64 | 54 | 90.3 | 8.71 | 64.5 | 0.152 | 7.61 | -81.4 | 5.4
2001-08-14 | 23:01 | 17.03 | 64 | 54 | 90.2 | 8.71 | 64.5 | 0.12 | 7.6 | -80.4 | 5.4
2001-08-14 | 23:16 | 16.97 | 64 | 54 | 90.2 | 8.72 | 64.5 | 0.092 | 7.58 | -79.7 | 5.4
2001-08-14 | 23:31 | 16.91 | 64 | 54 | 90.1 | 8.72 | 64.5 | 0.084 | 7.57 | -78.7 | 5.4
2001-08-14 | 23:46 | 16.85 | 64 | 54 | 90 | 8.73 | 64.5 | 0.104 | 7.55 | -77.6 | 5.4
2001-08-15 | 0:01 | 16.78 | 64 | 54 | 90.1 | 8.74 | 64.5 | 0.159 | 7.53 | -76.8 | 5.4
2001-08-15 | 0:16 | 16.72 | 64 | 54 | 90.3 | 8.78 | 64.5 | 0.176 | 7.52 | -76.2 | 5.4
2001-08-15 | 0:31 | 16.66 | 64 | 53 | 90.3 | 8.79 | 64.5 | 0.146 | 7.51 | -75.7 | 5.4
2001-08-15 | 0:46 | 16.61 | 64 | 53 | 90.3 | 8.8 | 64.5 | 0.116 | 7.51 | -75.7 | 5.4
2001-08-15 | 1:01 | 16.56 | 63 | 53 | 90.3 | 8.8 | 64.5 | 0.108 | 7.5 | -75.1 | 5.4
2001-08-15 | 1:16 | 16.51 | 63 | 53 | 90.3 | 8.82 | 64.5 | 0.142 | 7.49 | -74.3 | 5.4
2001-08-15 | 1:31 | 16.46 | 63 | 53 | 90.4 | 8.83 | 64.5 | 0.179 | 7.48 | -73.7 | 5.4
2001-08-15 | 1:46 | 16.41 | 63 | 53 | 90.5 | 8.85 | 64.5 | 0.193 | 7.47 | -73.4 | 5.4
2001-08-15 | 2:01 | 16.37 | 63 | 53 | 90.7 | 8.88 | 64.5 | 0.143 | 7.47 | -73.2 | 5.4
2001-08-15 | 2:16 | 16.34 | 63 | 53 | 90.5 | 8.87 | 64.5 | 0.105 | 7.47 | -73.1 | 5.4
2001-08-15 | 2:31 | 16.31 | 63 | 53 | 90.5 | 8.87 | 64.5 | 0.134 | 7.47 | -73.1 | 5.4
2001-08-15 | 2:46 | 16.28 | 63 | 53 | 90.6 | 8.89 | 64.5 | 0.178 | 7.46 | -72.6 | 5.4
2001-08-15 | 3:01 | 16.24 | 63 | 53 | 90.6 | 8.9 | 64.5 | 0.159 | 7.46 | -72.6 | 5.4
2001-08-15 | 3:16 | 16.21 | 63 | 53 | 90.6 | 8.91 | 64.5 | 0.113 | 7.46 | -72.6 | 5.4
2001-08-15 | 3:31 | 16.18 | 63 | 52 | 90.6 | 8.9 | 64.5 | 0.11 | 7.45 | -72.3 | 5.4
2001-08-15 | 3:46 | 16.16 | 63 | 52 | 90.6 | 8.91 | 63.3 | 0.163 | 7.45 | -72 | 5.4
2001-08-15 | 4:01 | 16.14 | 63 | 52 | 90.7 | 8.92 | 63.3 | 0.179 | 7.44 | -71.7 | 5.4
2001-08-15 | 4:16 | 16.12 | 63 | 52 | 90.8 | 8.94 | 63.3 | 0.13 | 7.44 | -71.7 | 5.3
2001-08-15 | 4:31 | 16.1 | 63 | 52 | 90.8 | 8.94 | 64.5 | 0.096 | 7.44 | -71.5 | 5.3
2001-08-15 | 4:46 | 16.08 | 63 | 52 | 90.7 | 8.94 | 64.5 | 0.101 | 7.44 | -71.5 | 5.3
2001-08-15 | 5:01 | 16.06 | 63 | 52 | 90.7 | 8.94 | 63.3 | 0.139 | 7.44 | -71.3 | 5.4
2001-08-15 | 5:16 | 16.04 | 63 | 52 | 90.7 | 8.94 | 63.3 | 0.175 | 7.44 | -71.2 | 5.4
2001-08-15 | 5:31 | 16.03 | 63 | 52 | 90.8 | 8.96 | 63.3 | 0.179 | 7.43 | -71 | 5.3
2001-08-15 | 5:46 | 16.02 | 63 | 52 | 90.9 | 8.97 | 63.3 | 0.13 | 7.44 | -71.3 | 5.4
2001-08-15 | 6:01 | 15.99 | 63 | 52 | 90.8 | 8.96 | 63.3 | 0.1 | 7.44 | -71.3 | 5.4
2001-08-15 | 6:16 | 15.98 | 63 | 52 | 90.8 | 8.96 | 63.3 | 0.111 | 7.43 | -71 | 5.4
2001-08-15 | 6:31 | 15.97 | 63 | 52 | 90.8 | 8.97 | 63.3 | 0.15 | 7.43 | -71.2 | 5.4
2001-08-15 | 6:46 | 15.95 | 63 | 52 | 91.2 | 9.01 | 63.3 | 0.18 | 7.43 | -71 | 5.3
2001-08-15 | 7:01 | 15.94 | 63 | 52 | 91.5 | 9.05 | 63.3 | 0.197 | 7.44 | -71.4 | 5.3
2001-08-15 | 7:16 | 15.93 | 63 | 52 | 92 | 9.09 | 63.3 | 0.168 | 7.45 | -72 | 5.3
2001-08-15 | 7:31 | 15.92 | 63 | 52 | 92.4 | 9.13 | 63.3 | 0.128 | 7.47 | -73.1 | 5.4
2001-08-15 | 7:46 | 15.91 | 63 | 52 | 92.7 | 9.17 | 63.3 | 0.118 | 7.48 | -74 | 5.4
2001-08-15 | 8:01 | 15.93 | 63 | 52 | 93.2 | 9.21 | 63.3 | 0.15 | 7.5 | -74.6 | 5.3
2001-08-15 | 8:16 | 15.94 | 63 | 52 | 93.7 | 9.26 | 63.3 | 0.179 | 7.51 | -75.4 | 5.3
2001-08-15 | 8:31 | 15.96 | 63 | 52 | 94.2 | 9.31 | 63.3 | 0.196 | 7.53 | -76.5 | 5.3
2001-08-15 | 8:46 | 16.01 | 63 | 52 | 94.9 | 9.37 | 63.3 | 0.202 | 7.56 | -78.1 | 5.4
2001-08-15 | 9:01 | 16.05 | 63 | 52 | 95.5 | 9.41 | 63.3 | 0.189 | 7.58 | -79.5 | 5.4
2001-08-15 | 9:16 | 16.11 | 63 | 52 | 96 | 9.45 | 63.3 | 0.153 | 7.62 | -81.4 | 5.4
2001-08-15 | 9:31 | 16.17 | 63 | 52 | 96.5 | 9.49 | 63.3 | 0.127 | 7.64 | -82.9 | 5.4
2001-08-15 | 9:46 | 16.23 | 63 | 52 | 97 | 9.53 | 64.5 | 0.105 | 7.67 | -84.5 | 5.4
2001-08-15 | 10:01 | 16.31 | 63 | 52 | 97.4 | 9.55 | 64.5 | 0.092 | 7.7 | -86.2 | 5.4
2001-08-15 | 10:16 | 16.39 | 63 | 52 | 97.8 | 9.58 | 64.5 | 0.133 | 7.73 | -88.1 | 5.4
2001-08-15 | 10:31 | 16.47 | 63 | 52 | 98.2 | 9.59 | 64.5 | 0.133 | 7.76 | -89.6 | 5.4
2001-08-15 | 10:46 | 16.54 | 63 | 52 | 98.5 | 9.61 | 64.5 | 0.135 | 7.79 | -91.2 | 5.4
2001-08-15 | 11:01 | 16.61 | 63 | 53 | 98.5 | 9.6 | 64.5 | 0.153 | 7.82 | -93 | 5.4
2001-08-15 | 11:16 | 16.7 | 63 | 53 | 98.6 | 9.59 | 64.5 | 0.132 | 7.85 | -94.7 | 5.4
2001-08-15 | 11:31 | 16.79 | 63 | 53 | 98.9 | 9.6 | 64.5 | 0.105 | 7.89 | -97 | 5.4
2001-08-15 | 11:46 | 16.87 | 63 | 53 | 99 | 9.59 | 64.5 | 0.082 | 7.92 | -98.9 | 5.4
2001-08-15 | 12:01 | 16.96 | 63 | 53 | 99.3 | 9.6 | 64.5 | 0.077 | 7.95 | -100.6 | 5.4
2001-08-15 | 12:16 | 17.05 | 63 | 53 | 99.3 | 9.59 | 64.5 | 0.067 | 7.98 | -102.2 | 5.4
2001-08-15 | 12:31 | 17.14 | 63 | 53 | 99.6 | 9.59 | 64.5 | 0.059 | 8.02 | -104.2 | 5.4
2001-08-15 | 12:46 | 17.22 | 63 | 53 | 99.5 | 9.57 | 64.5 | 0.102 | 8.04 | -105.5 | 5.4
2001-08-15 | 13:01 | 17.28 | 63 | 54 | 99.3 | 9.54 | 64.5 | 0.11 | 8.06 | -106.9 | 5.4
2001-08-15 | 13:16 | 17.36 | 63 | 54 | 99.1 | 9.5 | 64.5 | 0.13 | 8.09 | -108.5 | 5.4
2001-08-15 | 13:31 | 17.44 | 63 | 54 | 98.7 | 9.45 | 64.5 | 0.099 | 8.12 | -110.2 | 5.4
2001-08-15 | 13:46 | 17.52 | 63 | 54 | 98.6 | 9.43 | 64.5 | 0.072 | 8.15 | -112.1 | 5.4
2001-08-15 | 14:01 | 17.6 | 63 | 54 | 98.6 | 9.41 | 64.5 | 0.051 | 8.18 | -113.4 | 5.4
2001-08-15 | 14:16 | 17.67 | 63 | 54 | 98.5 | 9.39 | 64.5 | 0.043 | 8.19 | -114.2 | 5.4
2001-08-15 | 14:31 | 17.75 | 63 | 54 | 98.5 | 9.37 | 64.5 | 0.038 | 8.22 | -115.7 | 5.4
2001-08-15 | 14:46 | 17.81 | 63 | 54 | 98.2 | 9.33 | 64.5 | 0.032 | 8.23 | -116.5 | 5.4
2001-08-15 | 15:01 | 17.88 | 63 | 54 | 98.1 | 9.3 | 64.5 | 0.065 | 8.25 | -117.4 | 5.4
2001-08-15 | 15:16 | 17.95 | 63 | 54 | 97.8 | 9.27 | 64.5 | 0.061 | 8.25 | -117.6 | 5.4
2001-08-15 | 15:31 | 18.02 | 63 | 55 | 97.6 | 9.23 | 64.5 | 0.055 | 8.27 | -118.7 | 5.4
2001-08-15 | 15:46 | 18.06 | 63 | 55 | 97.2 | 9.19 | 64.5 | 0.051 | 8.28 | -119.3 | 5.4
2001-08-15 | 16:01 | 18.1 | 63 | 55 | 96.9 | 9.15 | 64.5 | 0.049 | 8.28 | -119.4 | 5.4
2001-08-15 | 16:16 | 18.13 | 63 | 55 | 96.5 | 9.11 | 64.5 | 0.047 | 8.28 | -119.5 | 5.4
2001-08-15 | 16:31 | 18.16 | 63 | 55 | 96 | 9.06 | 63.3 | 0.045 | 8.28 | -119.5 | 5.4
2001-08-15 | 16:46 | 18.18 | 63 | 55 | 95.6 | 9.02 | 63.3 | 0.046 | 8.28 | -119.4 | 5.4
2001-08-15 | 17:01 | 18.21 | 63 | 55 | 95.1 | 8.97 | 63.3 | 0.051 | 8.28 | -119.3 | 5.4
2001-08-15 | 17:16 | 18.23 | 63 | 55 | 94.7 | 8.93 | 63.3 | 0.058 | 8.27 | -119.1 | 5.4
2001-08-15 | 17:31 | 18.25 | 63 | 55 | 94.3 | 8.88 | 63.3 | 0.01 | 8.27 | -118.7 | 5.4
2001-08-15 | 17:46 | 18.26 | 63 | 55 | 93.9 | 8.84 | 63.3 | 0.015 | 8.25 | -117.9 | 5.4
2001-08-15 | 18:01 | 18.25 | 63 | 55 | 93.4 | 8.79 | 63.3 | 0.032 | 8.24 | -117.2 | 5.4
2001-08-15 | 18:16 | 18.25 | 63 | 55 | 93.2 | 8.77 | 63.3 | 0.032 | 8.22 | -116.1 | 5.4
2001-08-15 | 18:31 | 18.24 | 63 | 55 | 92.7 | 8.73 | 63.3 | 0.028 | 8.2 | -114.9 | 5.4
2001-08-15 | 18:46 | 18.22 | 63 | 55 | 92.2 | 8.69 | 63.3 | 0.048 | 8.18 | -113.5 | 5.4
2001-08-15 | 19:01 | 18.21 | 63 | 55 | 91.9 | 8.66 | 63.3 | 0.052 | 8.15 | -112.1 | 5.4
2001-08-15 | 19:16 | 18.19 | 63 | 55 | 91.5 | 8.63 | 63.3 | 0.055 | 8.13 | -110.5 | 5.4
2001-08-15 | 19:31 | 18.16 | 63 | 55 | 91.4 | 8.63 | 63.3 | 0.1 | 8.09 | -108.4 | 5.4
2001-08-15 | 19:46 | 18.12 | 64 | 55 | 91.1 | 8.6 | 63.3 | 0.099 | 8.05 | -106.2 | 5.4
2001-08-15 | 20:01 | 18.08 | 64 | 55 | 90.8 | 8.58 | 63.3 | 0.08 | 8.01 | -104 | 5.4
2001-08-15 | 20:16 | 18.04 | 64 | 55 | 90.4 | 8.56 | 63.3 | 0.076 | 7.97 | -101.9 | 5.4
2001-08-15 | 20:31 | 17.99 | 64 | 55 | 90.2 | 8.54 | 62.1 | 0.07 | 7.92 | -98.9 | 5.4
2001-08-15 | 20:46 | 17.93 | 64 | 55 | 89.9 | 8.52 | 62.1 | 0.072 | 7.87 | -96.3 | 5.4
2001-08-15 | 21:01 | 17.87 | 64 | 55 | 89.8 | 8.52 | 62.1 | 0.077 | 7.83 | -93.6 | 5.4
2001-08-15 | 21:16 | 17.81 | 64 | 55 | 89.7 | 8.53 | 62.1 | 0.083 | 7.79 | -91.2 | 5.4
2001-08-15 | 21:31 | 17.74 | 64 | 55 | 89.7 | 8.54 | 62.1 | 0.084 | 7.75 | -89 | 5.4
2001-08-15 | 21:46 | 17.67 | 64 | 55 | 89.7 | 8.55 | 62.1 | 0.084 | 7.72 | -87.3 | 5.4
2001-08-15 | 22:01 | 17.6 | 64 | 55 | 89.7 | 8.57 | 62.1 | 0.087 | 7.69 | -85.6 | 5.4
2001-08-15 | 22:16 | 17.52 | 64 | 55 | 89.8 | 8.59 | 62.1 | 0.092 | 7.66 | -84.2 | 5.4
2001-08-15 | 22:31 | 17.44 | 64 | 55 | 89.8 | 8.6 | 62.1 | 0.095 | 7.64 | -82.8 | 5.4
2001-08-15 | 22:46 | 17.36 | 64 | 55 | 89.9 | 8.62 | 62.1 | 0.13 | 7.61 | -81.1 | 5.3
2001-08-15 | 23:01 | 17.27 | 64 | 54 | 90.1 | 8.66 | 62.1 | 0.166 | 7.59 | -80.3 | 5.4
2001-08-15 | 23:16 | 17.18 | 64 | 54 | 90.3 | 8.7 | 62.1 | 0.194 | 7.57 | -78.9 | 5.4
2001-08-15 | 23:31 | 17.1 | 64 | 54 | 90.5 | 8.73 | 62.1 | 0.171 | 7.56 | -78.2 | 5.3
2001-08-15 | 23:46 | 17.02 | 64 | 54 | 90.5 | 8.74 | 62.1 | 0.123 | 7.55 | -77.6 | 5.3
2001-08-16 | 0:01 | 16.96 | 64 | 54 | 90.5 | 8.75 | 62.1 | 0.107 | 7.54 | -77 | 5.4
2001-08-16 | 0:16 | 16.89 | 64 | 54 | 90.4 | 8.76 | 62.1 | 0.131 | 7.53 | -76.4 | 5.3
2001-08-16 | 0:31 | 16.82 | 64 | 54 | 90.5 | 8.78 | 62.1 | 0.162 | 7.51 | -75.7 | 5.3
2001-08-16 | 0:46 | 16.75 | 64 | 54 | 90.6 | 8.8 | 62.1 | 0.182 | 7.5 | -75.2 | 5.3
2001-08-16 | 1:01 | 16.69 | 64 | 54 | 90.8 | 8.84 | 62.1 | 0.198 | 7.5 | -75 | 5.3
2001-08-16 | 1:16 | 16.62 | 64 | 53 | 91 | 8.87 | 62.1 | 0.19 | 7.5 | -74.8 | 5.3
2001-08-16 | 1:31 | 16.56 | 64 | 53 | 91 | 8.87 | 62.1 | 0.149 | 7.49 | -74.3 | 5.3
2001-08-16 | 1:46 | 16.51 | 64 | 53 | 90.9 | 8.88 | 62.1 | 0.123 | 7.48 | -74 | 5.3
2001-08-16 | 2:01 | 16.46 | 63 | 53 | 90.8 | 8.87 | 62.1 | 0.106 | 7.47 | -73.5 | 5.3
2001-08-16 | 2:16 | 16.41 | 64 | 53 | 90.8 | 8.88 | 62.1 | 0.137 | 7.47 | -73.5 | 5.3
2001-08-16 | 2:31 | 16.36 | 63 | 53 | 90.8 | 8.9 | 62.1 | 0.189 | 7.46 | -72.9 | 5.3
2001-08-16 | 2:46 | 16.32 | 63 | 53 | 90.9 | 8.92 | 62.1 | 0.156 | 7.46 | -72.8 | 5.3
2001-08-16 | 3:01 | 16.28 | 63 | 53 | 90.9 | 8.92 | 62.1 | 0.119 | 7.46 | -72.7 | 5.3
2001-08-16 | 3:16 | 16.24 | 63 | 53 | 90.9 | 8.93 | 60.9 | 0.125 | 7.46 | -72.6 | 5.3
2001-08-16 | 3:31 | 16.21 | 63 | 53 | 90.8 | 8.93 | 60.9 | 0.149 | 7.45 | -72 | 5.3
2001-08-16 | 3:46 | 16.18 | 63 | 53 | 91 | 8.95 | 60.9 | 0.168 | 7.44 | -71.7 | 5.3
2001-08-16 | 4:01 | 16.15 | 63 | 53 | 91.1 | 8.96 | 60.9 | 0.177 | 7.44 | -71.5 | 5.3
2001-08-16 | 4:16 | 16.13 | 63 | 52 | 91.2 | 8.98 | 60.9 | 0.191 | 7.44 | -71.5 | 5.3
2001-08-16 | 4:31 | 16.11 | 63 | 52 | 91.3 | 8.99 | 60.9 | 0.199 | 7.44 | -71.7 | 5.3
2001-08-16 | 4:46 | 16.09 | 63 | 52 | 91.3 | 9 | 60.9 | 0.204 | 7.44 | -71.6 | 5.3
2001-08-16 | 5:01 | 16.07 | 63 | 52 | 91.4 | 9.01 | 60.9 | 0.2 | 7.44 | -71.7 | 5.3
2001-08-16 | 5:16 | 16.06 | 63 | 52 | 91.4 | 9.01 | 60.9 | 0.198 | 7.44 | -71.5 | 5.3
2001-08-16 | 5:31 | 16.05 | 63 | 52 | 91.5 | 9.02 | 60.9 | 0.201 | 7.44 | -71.7 | 5.3
2001-08-16 | 5:46 | 16.03 | 63 | 52 | 91.4 | 9.02 | 60.9 | 0.206 | 7.44 | -71.5 | 5.3
2001-08-16 | 6:01 | 16.02 | 63 | 52 | 91.5 | 9.02 | 60.9 | 0.213 | 7.44 | -71.6 | 5.3
2001-08-16 | 6:16 | 16 | 63 | 52 | 91.6 | 9.04 | 60.9 | 0.215 | 7.44 | -71.4 | 5.3
2001-08-16 | 6:31 | 15.98 | 63 | 52 | 91.7 | 9.05 | 60.9 | 0.225 | 7.44 | -71.7 | 5.3
2001-08-16 | 6:46 | 15.96 | 63 | 52 | 91.9 | 9.08 | 60.9 | 0.227 | 7.45 | -71.9 | 5.3
2001-08-16 | 7:01 | 15.95 | 63 | 52 | 92.1 | 9.1 | 60.9 | 0.23 | 7.45 | -72.1 | 5.3
2001-08-16 | 7:16 | 15.94 | 63 | 52 | 92.5 | 9.14 | 60.9 | 0.227 | 7.46 | -72.8 | 5.3
2001-08-16 | 7:31 | 15.93 | 63 | 52 | 92.9 | 9.18 | 60.9 | 0.21 | 7.47 | -73.4 | 5.3
2001-08-16 | 7:46 | 15.93 | 63 | 52 | 93.2 | 9.22 | 60.9 | 0.198 | 7.48 | -73.9 | 5.3
2001-08-16 | 8:01 | 15.93 | 63 | 52 | 93.6 | 9.25 | 60.9 | 0.183 | 7.5 | -74.9 | 5.3
2001-08-16 | 8:16 | 15.94 | 63 | 52 | 94 | 9.29 | 60.9 | 0.179 | 7.52 | -75.7 | 5.3
2001-08-16 | 8:31 | 15.96 | 63 | 52 | 94.4 | 9.33 | 60.9 | 0.176 | 7.54 | -76.9 | 5.3
2001-08-16 | 8:46 | 16.01 | 63 | 52 | 95.2 | 9.39 | 60.9 | 0.194 | 7.56 | -78.1 | 5.3
2001-08-16 | 9:01 | 16.05 | 63 | 52 | 95.7 | 9.44 | 62.1 | 0.223 | 7.58 | -79.4 | 5.3
2001-08-16 | 9:16 | 16.11 | 63 | 52 | 96.3 | 9.48 | 62.1 | 0.184 | 7.61 | -81.2 | 5.3
2001-08-16 | 9:31 | 16.18 | 63 | 52 | 96.8 | 9.52 | 62.1 | 0.134 | 7.65 | -83.1 | 5.3
2001-08-16 | 9:46 | 16.24 | 63 | 52 | 97.2 | 9.55 | 62.1 | 0.148 | 7.67 | -84.6 | 5.4
2001-08-16 | 10:01 | 16.3 | 63 | 52 | 97.5 | 9.56 | 62.1 | 0.18 | 7.69 | -85.7 | 5.4
2001-08-16 | 10:16 | 16.38 | 63 | 52 | 97.8 | 9.57 | 62.1 | 0.17 | 7.72 | -87.3 | 5.3
2001-08-16 | 10:31 | 16.48 | 63 | 52 | 98.2 | 9.59 | 62.1 | 0.158 | 7.75 | -89 | 5.4
2001-08-16 | 10:46 | 16.58 | 63 | 52 | 98.5 | 9.6 | 62.1 | 0.148 | 7.78 | -90.9 | 5.4
2001-08-16 | 11:01 | 16.67 | 63 | 53 | 98.7 | 9.61 | 62.1 | 0.149 | 7.82 | -92.9 | 5.4
2001-08-16 | 11:16 | 16.76 | 63 | 53 | 99 | 9.62 | 62.1 | 0.149 | 7.85 | -94.5 | 5.4
2001-08-16 | 11:31 | 16.86 | 63 | 53 | 99.1 | 9.6 | 62.1 | 0.143 | 7.88 | -96.6 | 5.4
2001-08-16 | 11:46 | 16.95 | 63 | 53 | 99.3 | 9.61 | 62.1 | 0.134 | 7.91 | -98.2 | 5.4
2001-08-16 | 12:01 | 17.04 | 63 | 53 | 99.4 | 9.6 | 62.1 | 0.128 | 7.95 | -100.4 | 5.4
2001-08-16 | 12:16 | 17.14 | 63 | 53 | 99.5 | 9.58 | 62.1 | 0.116 | 7.98 | -102.2 | 5.4
2001-08-16 | 12:31 | 17.24 | 63 | 53 | 99.5 | 9.57 | 62.1 | 0.108 | 8.01 | -103.9 | 5.4
2001-08-16 | 12:46 | 17.32 | 63 | 53 | 99.6 | 9.56 | 62.1 | 0.102 | 8.04 | -105.7 | 5.4
2001-08-16 | 13:01 | 17.4 | 63 | 53 | 99.5 | 9.53 | 62.1 | 0.098 | 8.06 | -106.8 | 5.4
2001-08-16 | 13:16 | 17.49 | 63 | 54 | 99.5 | 9.51 | 62.1 | 0.089 | 8.09 | -108.4 | 5.4
2001-08-16 | 13:31 | 17.57 | 63 | 54 | 99.2 | 9.48 | 62.1 | 0.086 | 8.12 | -109.9 | 5.4
2001-08-16 | 13:46 | 17.65 | 63 | 54 | 99.1 | 9.45 | 62.1 | 0.083 | 8.14 | -111.1 | 5.4
2001-08-16 | 14:01 | 17.73 | 63 | 54 | 99 | 9.43 | 62.1 | 0.081 | 8.16 | -112.4 | 5.4
2001-08-16 | 14:16 | 17.8 | 63 | 54 | 99 | 9.41 | 62.1 | 0.076 | 8.18 | -113.4 | 5.4
2001-08-16 | 14:31 | 17.88 | 63 | 54 | 98.8 | 9.37 | 62.1 | 0.074 | 8.2 | -114.8 | 5.4
2001-08-16 | 14:46 | 17.95 | 63 | 54 | 98.6 | 9.34 | 62.1 | 0.113 | 8.21 | -115.5 | 5.4
2001-08-16 | 15:01 | 18.01 | 63 | 55 | 98 | 9.28 | 62.1 | 0.109 | 8.23 | -116.8 | 5.4
2001-08-16 | 15:16 | 18.08 | 63 | 55 | 97.8 | 9.24 | 62.1 | 0.061 | 8.25 | -117.4 | 5.4
2001-08-16 | 15:31 | 18.14 | 63 | 55 | 97.6 | 9.21 | 62.1 | 0.085 | 8.26 | -118.3 | 5.4
2001-08-16 | 15:46 | 18.18 | 63 | 55 | 97.4 | 9.19 | 62.1 | 0.084 | 8.27 | -118.7 | 5.4
2001-08-16 | 16:01 | 18.22 | 63 | 55 | 97.2 | 9.16 | 62.1 | 0.086 | 8.27 | -119 | 5.4
2001-08-16 | 16:16 | 18.25 | 63 | 55 | 96.8 | 9.12 | 62.1 | 0.088 | 8.27 | -118.7 | 5.4
2001-08-16 | 16:31 | 18.27 | 63 | 55 | 96.1 | 9.05 | 62.1 | 0.095 | 8.26 | -118.1 | 5.4
2001-08-16 | 16:46 | 18.3 | 63 | 55 | 95.6 | 8.99 | 62.1 | 0.067 | 8.26 | -118.3 | 5.4
2001-08-16 | 17:01 | 18.32 | 63 | 55 | 95.2 | 8.95 | 62.1 | 0.048 | 8.26 | -118.3 | 5.4
2001-08-16 | 17:16 | 18.33 | 63 | 55 | 94.7 | 8.91 | 62.1 | 0.085 | 8.25 | -117.8 | 5.4
2001-08-16 | 17:31 | 18.33 | 63 | 55 | 94.3 | 8.87 | 62.1 | 0.09 | 8.24 | -117.1 | 5.4
2001-08-16 | 17:46 | 18.33 | 63 | 55 | 94 | 8.84 | 60.9 | 0.09 | 8.23 | -116.5 | 5.4
2001-08-16 | 18:01 | 18.33 | 63 | 55 | 93.6 | 8.8 | 60.9 | 0.091 | 8.21 | -115.4 | 5.4
2001-08-16 | 18:16 | 18.32 | 63 | 55 | 93.2 | 8.76 | 60.9 | 0.126 | 8.19 | -114 | 5.4
2001-08-16 | 18:31 | 18.31 | 63 | 55 | 92.8 | 8.73 | 60.9 | 0.108 | 8.16 | -112.8 | 5.4
2001-08-16 | 18:46 | 18.3 | 63 | 55 | 92.3 | 8.69 | 60.9 | 0.071 | 8.15 | -111.9 | 5.4
2001-08-16 | 19:01 | 18.29 | 64 | 55 | 92 | 8.66 | 60.9 | 0.076 | 8.12 | -110.5 | 5.4
2001-08-16 | 19:16 | 18.27 | 64 | 55 | 91.8 | 8.65 | 60.9 | 0.118 | 8.1 | -109.2 | 5.4
2001-08-16 | 19:31 | 18.24 | 63 | 55 | 91.5 | 8.62 | 60.9 | 0.154 | 8.06 | -107.1 | 5.4
2001-08-16 | 19:46 | 18.21 | 64 | 55 | 91.4 | 8.61 | 60.9 | 0.173 | 8.02 | -104.8 | 5.3
2001-08-16 | 20:01 | 18.17 | 64 | 55 | 91.2 | 8.6 | 60.9 | 0.133 | 7.99 | -103 | 5.4
2001-08-16 | 20:16 | 18.13 | 64 | 55 | 90.9 | 8.58 | 60.9 | 0.102 | 7.95 | -100.8 | 5.4
2001-08-16 | 20:31 | 18.08 | 64 | 55 | 90.6 | 8.57 | 60.9 | 0.145 | 7.91 | -98.2 | 5.3
2001-08-16 | 20:46 | 18.01 | 64 | 55 | 90.5 | 8.56 | 60.9 | 0.189 | 7.86 | -95.3 | 5.3
2001-08-16 | 21:01 | 17.94 | 64 | 55 | 90.3 | 8.56 | 60.9 | 0.161 | 7.81 | -92.8 | 5.3
2001-08-16 | 21:16 | 17.88 | 64 | 55 | 90.2 | 8.56 | 60.9 | 0.132 | 7.77 | -90.6 | 5.3
2001-08-16 | 21:31 | 17.82 | 64 | 55 | 90.2 | 8.57 | 60.9 | 0.145 | 7.74 | -88.5 | 5.3
2001-08-16 | 21:46 | 17.75 | 64 | 55 | 90.2 | 8.58 | 60.9 | 0.176 | 7.7 | -86.6 | 5.3
2001-08-16 | 22:01 | 17.68 | 64 | 55 | 90.3 | 8.6 | 60.9 | 0.199 | 7.67 | -84.8 | 5.3
2001-08-16 | 22:16 | 17.61 | 64 | 55 | 90.5 | 8.63 | 60.9 | 0.209 | 7.64 | -83.2 | 5.3
2001-08-16 | 22:31 | 17.54 | 64 | 55 | 90.6 | 8.66 | 60.9 | 0.22 | 7.63 | -82.3 | 5.3
2001-08-16 | 22:46 | 17.47 | 64 | 55 | 90.7 | 8.68 | 60.4 | 0.223 | 7.61 | -81 | 5.3
2001-08-16 | 23:01 | 17.39 | 64 | 55 | 90.8 | 8.7 | 60.4 | 0.227 | 7.59 | -79.9 | 5.3
2001-08-16 | 23:16 | 17.32 | 64 | 54 | 90.8 | 8.72 | 60.4 | 0.233 | 7.57 | -79.2 | 5.3
2001-08-16 | 23:31 | 17.24 | 64 | 54 | 91 | 8.75 | 60.9 | 0.239 | 7.56 | -78.6 | 5.3
2001-08-16 | 23:46 | 17.17 | 64 | 54 | 91 | 8.77 | 60.4 | 0.24 | 7.54 | -77.5 | 5.3
2001-08-17 | 0:01 | 17.09 | 64 | 54 | 91.1 | 8.79 | 60.4 | 0.241 | 7.54 | -77.2 | 5.3
2001-08-17 | 0:16 | 17.02 | 64 | 54 | 91.2 | 8.81 | 60.4 | 0.246 | 7.53 | -76.6 | 5.3
2001-08-17 | 0:31 | 16.94 | 64 | 54 | 91.2 | 8.83 | 60.4 | 0.247 | 7.52 | -76.1 | 5.3
2001-08-17 | 0:46 | 16.87 | 64 | 54 | 91.3 | 8.84 | 60.4 | 0.244 | 7.51 | -75.5 | 5.3
2001-08-17 | 1:01 | 16.81 | 64 | 54 | 91.3 | 8.86 | 60.4 | 0.252 | 7.5 | -75 | 5.2
2001-08-17 | 1:16 | 16.74 | 64 | 54 | 91.3 | 8.87 | 60.4 | 0.248 | 7.5 | -74.8 | 5.3
2001-08-17 | 1:31 | 16.68 | 64 | 53 | 91.3 | 8.88 | 60.4 | 0.248 | 7.49 | -74.3 | 5.2
2001-08-17 | 1:46 | 16.62 | 64 | 53 | 91.4 | 8.9 | 60.4 | 0.247 | 7.48 | -74 | 5.2
2001-08-17 | 2:01 | 16.55 | 64 | 53 | 91.5 | 8.92 | 60.4 | 0.25 | 7.48 | -73.8 | 5.3
2001-08-17 | 2:16 | 16.49 | 64 | 53 | 91.5 | 8.94 | 60.4 | 0.249 | 7.47 | -73.5 | 5.3
2001-08-17 | 2:31 | 16.43 | 63 | 53 | 91.5 | 8.95 | 60.4 | 0.248 | 7.47 | -73.4 | 5.3
2001-08-17 | 2:46 | 16.38 | 63 | 53 | 91.5 | 8.96 | 60.4 | 0.249 | 7.47 | -73.4 | 5.2
2001-08-17 | 3:01 | 16.33 | 63 | 53 | 91.6 | 8.98 | 60.4 | 0.246 | 7.47 | -73.2 | 5.2
2001-08-17 | 3:16 | 16.28 | 63 | 53 | 91.7 | 9 | 60.4 | 0.242 | 7.47 | -73.2 | 5.2
2001-08-17 | 3:31 | 16.24 | 63 | 53 | 91.7 | 9 | 60.4 | 0.24 | 7.46 | -72.8 | 5.2
2001-08-17 | 3:46 | 16.2 | 63 | 53 | 91.7 | 9.02 | 60.4 | 0.248 | 7.46 | -72.9 | 5.2
2001-08-17 | 4:01 | 16.16 | 63 | 52 | 91.8 | 9.03 | 60.4 | 0.253 | 7.46 | -72.5 | 5.2
2001-08-17 | 4:16 | 16.12 | 63 | 52 | 91.8 | 9.04 | 60.4 | 0.241 | 7.46 | -72.5 | 5.3
2001-08-17 | 4:31 | 16.09 | 63 | 52 | 91.8 | 9.04 | 60.4 | 0.205 | 7.46 | -72.5 | 5.2
2001-08-17 | 4:46 | 16.06 | 63 | 52 | 91.6 | 9.03 | 60.4 | 0.166 | 7.46 | -72.6 | 5.2
2001-08-17 | 5:01 | 16.02 | 63 | 52 | 91.5 | 9.03 | 60.4 | 0.148 | 7.46 | -72.4 | 5.2
2001-08-17 | 5:16 | 16 | 63 | 52 | 91.5 | 9.03 | 60.4 | 0.185 | 7.44 | -71.8 | 5.2
2001-08-17 | 5:31 | 15.97 | 63 | 52 | 91.5 | 9.04 | 60.4 | 0.239 | 7.44 | -71.7 | 5.2
2001-08-17 | 5:46 | 15.93 | 63 | 52 | 91.7 | 9.06 | 60.4 | 0.227 | 7.44 | -71.4 | 5.2
2001-08-17 | 6:01 | 15.9 | 63 | 52 | 91.7 | 9.07 | 60.4 | 0.182 | 7.45 | -71.8 | 5.2
2001-08-17 | 6:16 | 15.86 | 63 | 52 | 91.7 | 9.08 | 60.4 | 0.169 | 7.44 | -71.7 | 5.2
2001-08-17 | 6:31 | 15.83 | 63 | 52 | 91.8 | 9.1 | 60.4 | 0.205 | 7.44 | -71.7 | 5.2
2001-08-17 | 6:46 | 15.8 | 63 | 52 | 92 | 9.12 | 60.4 | 0.24 | 7.44 | -71.6 | 5.2
2001-08-17 | 7:01 | 15.77 | 63 | 52 | 92.2 | 9.15 | 60.4 | 0.27 | 7.45 | -72.1 | 5.2
2001-08-17 | 7:16 | 15.75 | 63 | 52 | 92.6 | 9.19 | 60.4 | 0.275 | 7.46 | -72.4 | 5.2
2001-08-17 | 7:31 | 15.72 | 63 | 52 | 92.9 | 9.22 | 60.4 | 0.246 | 7.47 | -73.1 | 5.2
2001-08-17 | 7:46 | 15.71 | 63 | 52 | 93.1 | 9.25 | 60.4 | 0.221 | 7.48 | -73.5 | 5.2
2001-08-17 | 8:01 | 15.7 | 63 | 52 | 93.4 | 9.28 | 60.4 | 0.195 | 7.49 | -74.3 | 5.3
2001-08-17 | 8:16 | 15.7 | 63 | 52 | 93.8 | 9.32 | 60.4 | 0.182 | 7.5 | -74.9 | 5.2
2001-08-17 | 8:31 | 15.72 | 63 | 52 | 94.3 | 9.36 | 60.4 | 0.179 | 7.52 | -75.8 | 5.2
2001-08-17 | 8:46 | 15.76 | 63 | 52 | 95 | 9.43 | 60.4 | 0.17 | 7.54 | -77 | 5.3
2001-08-17 | 9:01 | 15.82 | 63 | 52 | 95.6 | 9.48 | 60.4 | 0.164 | 7.56 | -78.4 | 5.3
2001-08-17 | 9:16 | 15.88 | 63 | 52 | 96.3 | 9.53 | 60.4 | 0.16 | 7.59 | -80 | 5.3
2001-08-17 | 9:31 | 15.93 | 63 | 52 | 96.6 | 9.55 | 60.4 | 0.198 | 7.62 | -81.4 | 5.3
2001-08-17 | 9:46 | 15.98 | 63 | 52 | 97 | 9.58 | 60.4 | 0.24 | 7.63 | -82.4 | 5.3
2001-08-17 | 10:01 | 16.06 | 63 | 52 | 97.4 | 9.6 | 60.4 | 0.222 | 7.66 | -84 | 5.3
2001-08-17 | 10:16 | 16.17 | 63 | 52 | 97.9 | 9.63 | 60.4 | 0.186 | 7.7 | -85.9 | 5.3
2001-08-17 | 10:31 | 16.28 | 63 | 52 | 98.4 | 9.65 | 60.4 | 0.158 | 7.73 | -87.8 | 5.3
2001-08-17 | 10:46 | 16.38 | 63 | 52 | 98.7 | 9.66 | 60.4 | 0.152 | 7.77 | -89.9 | 5.3
2001-08-17 | 11:01 | 16.47 | 63 | 52 | 99 | 9.68 | 60.4 | 0.182 | 7.8 | -91.6 | 5.3
2001-08-17 | 11:16 | 16.56 | 63 | 53 | 99.1 | 9.67 | 60.4 | 0.213 | 7.82 | -93.1 | 5.3
2001-08-17 | 11:31 | 16.66 | 63 | 53 | 99.2 | 9.65 | 60.4 | 0.219 | 7.85 | -94.7 | 5.3
2001-08-17 | 11:46 | 16.77 | 63 | 53 | 99.2 | 9.63 | 60.4 | 0.164 | 7.89 | -96.9 | 5.3
2001-08-17 | 12:01 | 16.88 | 63 | 53 | 99.4 | 9.63 | 60.9 | 0.13 | 7.92 | -98.8 | 5.3
2001-08-17 | 12:16 | 16.96 | 63 | 53 | 99.5 | 9.63 | 60.4 | 0.176 | 7.95 | -100.5 | 5.3
2001-08-17 | 12:31 | 17.04 | 63 | 53 | 99.5 | 9.61 | 60.4 | 0.187 | 7.98 | -102 | 5.3
2001-08-17 | 12:46 | 17.13 | 63 | 53 | 99.3 | 9.58 | 60.4 | 0.166 | 8.01 | -103.7 | 5.3
2001-08-17 | 13:01 | 17.21 | 63 | 53 | 99.3 | 9.55 | 60.4 | 0.146 | 8.03 | -105.1 | 5.3
2001-08-17 | 13:16 | 17.28 | 63 | 54 | 99.2 | 9.54 | 60.4 | 0.136 | 8.06 | -106.6 | 5.3
2001-08-17 | 13:31 | 17.37 | 63 | 54 | 99.3 | 9.53 | 60.4 | 0.126 | 8.09 | -108.2 | 5.3
2001-08-17 | 13:46 | 17.43 | 63 | 54 | 99.1 | 9.5 | 60.9 | 0.116 | 8.11 | -109.3 | 5.3
2001-08-17 | 14:01 | 22.27 | 1 | 1 | 99.1 | 8.62 | 60.4 | 0.094 | 8.11 | -110.8 | 5.3
2001-08-17 | 14:16 | 29.64 | 3 | 4 | 96.1 | 7.31 | 60.9 | 0.024 | 7.81 | -94.8 | 5.3
2001-08-17 | 14:31 | 28.26 | 5 | 5 | 96.3 | 7.51 | 60.9 | 0.03 | 7.61 | -82.6 | 5.3
2001-08-17 | 14:46 | 27.43 | 6 | 6 | 96.2 | 7.61 | 60.9 | 0.024 | 7.45 | -73.1 | 5.3