## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | LDR 001.00
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Bridge
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2001-07-16 13:50 | 20.9 | 8.3 | 8 | 103 | 49 | 0
Sonde | 2001-07-16 13:46 | 20.86 | 8.26 | 8.84 | 113 | 46
Grade |  | A | A | B |  | A
Audit 2 | 2001-07-18 09:25 | 16.9 | 7.6 | 7.6 | 90 | 50 | 0
Sonde | 2001-07-18 09:31 | 17.11 | 7.43 | 7.79 | 93 | 47
Grade |  | A | A | A |  | A
Audit 3 | 2001-07-20 09:02 | 16.2 | 7.7 | 7.4 | 87 | 48 | 0
Sonde | 2001-07-20 09:01 | 16.56 | 7.21 | 7.7 | 91 | 44
Grade |  | A | B | A |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note: Audits 1-3 are macro (ed) into this main audit sheet and have audits that represent the duration of the deployment. Audits 4-5 represent the middle of the deployment and have the sonde values on the individual sheets. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | LDR 001.00
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-16 13:50 | 20.9 | 8.3 | 8 | 103 | 49
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | LDR 001.00
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-18 09:25 | 16.9 | 7.6 | 7.6 | 90 | 50
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | LDR 001.00
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-20 09:02 | 16.2 | 7.7 | 7.4 | 87 | 48
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | LDR 001.00
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-17 14:49 | 19.3 | 8.4 | 8.8 | 109 | 49
Sonde Values | 2001-07-17 18:46 | 19.62 | 8.49 | 9.1 | 114 | 46
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | LDR 001.00
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2001-07-19 13:25 | 18.9 | 8.4 | 8.8 | 108 | 47
Sonde Values | 2001-07-19 13:31 | 18.47 | 8.22 | 9.28 | 113 | 45
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | DO Sat
|  | °C |  | uS/cm | mg/L | %
2001-07-16 | 13:46 | 20.86 | 8.26 | 46 | 8.84 | 113.3%
2001-07-16 | 14:01 | 20.97 | 8.35 | 46 | 8.82 | 113.1%
2001-07-16 | 14:16 | 21.09 | 8.39 | 46 | 8.84 | 114.8%
2001-07-16 | 14:31 | 21.19 | 8.44 | 46 | 8.88 | 115.3%
2001-07-16 | 14:46 | 21.3 | 8.46 | 46 | 8.93 | 116%
2001-07-16 | 15:01 | 21.42 | 8.5 | 46 | 8.98 | 116.6%
2001-07-16 | 15:16 | 21.53 | 8.52 | 46 | 9 | 118.4%
2001-07-16 | 15:31 | 21.63 | 8.54 | 46 | 9.07 | 119.3%
2001-07-16 | 15:46 | 21.74 | 8.57 | 46 | 9.11 | 119.9%
2001-07-16 | 16:01 | 21.83 | 8.59 | 46 | 9.13 | 120.1%
2001-07-16 | 16:16 | 21.91 | 8.61 | 46 | 9.16 | 120.5%
2001-07-16 | 16:31 | 21.98 | 8.64 | 46 | 9.18 | 120.8%
2001-07-16 | 16:46 | 22.05 | 8.64 | 46 | 9.18 | 121.6%
2001-07-16 | 17:01 | 22.1 | 8.65 | 46 | 9.19 | 121.7%
2001-07-16 | 17:16 | 22.14 | 8.66 | 46 | 9.19 | 121.7%
2001-07-16 | 17:31 | 22.17 | 8.67 | 46 | 9.17 | 121.5%
2001-07-16 | 17:46 | 22.2 | 8.68 | 46 | 9.17 | 121.5%
2001-07-16 | 18:01 | 22.22 | 8.66 | 46 | 9.15 | 121.2%
2001-07-16 | 18:16 | 22.23 | 8.67 | 46 | 9.07 | 120.1%
2001-07-16 | 18:31 | 22.22 | 8.67 | 46 | 9.09 | 120.4%
2001-07-16 | 18:46 | 22.21 | 8.65 | 46 | 9.07 | 120.1%
2001-07-16 | 19:01 | 22.17 | 8.66 | 46 | 9.03 | 119.6%
2001-07-16 | 19:16 | 22.1 | 8.65 | 46 | 9.01 | 119.3%
2001-07-16 | 19:31 | 22.03 | 8.64 | 46 | 8.94 | 118.4%
2001-07-16 | 19:46 | 21.94 | 8.62 | 46 | 8.85 | 116.4%
2001-07-16 | 20:01 | 21.85 | 8.6 | 46 | 8.74 | 115%
2001-07-16 | 20:16 | 21.75 | 8.56 | 46 | 8.68 | 114.2%
2001-07-16 | 20:31 | 21.65 | 8.54 | 46 | 8.57 | 112.8%
2001-07-16 | 20:46 | 21.55 | 8.49 | 46 | 8.5 | 111.8%
2001-07-16 | 21:01 | 21.45 | 8.46 | 46 | 8.39 | 109%
2001-07-16 | 21:16 | 21.34 | 8.42 | 46 | 8.31 | 107.9%
2001-07-16 | 21:31 | 21.24 | 8.37 | 46 | 8.23 | 106.9%
2001-07-16 | 21:46 | 21.15 | 8.31 | 45 | 8.13 | 105.6%
2001-07-16 | 22:01 | 21.05 | 8.27 | 45 | 8.02 | 104.2%
2001-07-16 | 22:16 | 20.95 | 8.22 | 45 | 7.99 | 102.4%
2001-07-16 | 22:31 | 20.85 | 8.17 | 45 | 7.87 | 100.9%
2001-07-16 | 22:46 | 20.75 | 8.13 | 45 | 7.82 | 100.3%
2001-07-16 | 23:01 | 20.66 | 8.07 | 45 | 7.72 | 98.97%
2001-07-16 | 23:16 | 20.57 | 8.04 | 45 | 7.7 | 98.72%
2001-07-16 | 23:31 | 20.48 | 8 | 45 | 7.71 | 98.22%
2001-07-16 | 23:46 | 20.39 | 7.96 | 45 | 7.64 | 97.32%
2001-07-17 | 0:01 | 20.3 | 7.92 | 45 | 7.57 | 96.43%
2001-07-17 | 0:16 | 20.2 | 7.88 | 45 | 7.51 | 95.67%
2001-07-17 | 0:31 | 20.1 | 7.85 | 45 | 7.52 | 95.8%
2001-07-17 | 0:46 | 20 | 7.82 | 46 | 7.44 | 94.78%
2001-07-17 | 1:01 | 19.91 | 7.79 | 46 | 7.41 | 93.21%
2001-07-17 | 1:16 | 19.82 | 7.75 | 46 | 7.44 | 93.58%
2001-07-17 | 1:31 | 19.72 | 7.73 | 46 | 7.34 | 92.33%
2001-07-17 | 1:46 | 19.63 | 7.7 | 46 | 7.17 | 90.19%
2001-07-17 | 2:01 | 19.54 | 7.67 | 46 | 7.37 | 92.7%
2001-07-17 | 2:16 | 19.45 | 7.65 | 46 | 7.18 | 88.64%
2001-07-17 | 2:31 | 19.37 | 7.62 | 46 | 7.26 | 89.63%
2001-07-17 | 2:46 | 19.28 | 7.59 | 46 | 7.3 | 90.12%
2001-07-17 | 3:01 | 19.19 | 7.57 | 46 | 7.3 | 90.12%
2001-07-17 | 3:16 | 19.1 | 7.55 | 46 | 7.07 | 87.28%
2001-07-17 | 3:31 | 19.02 | 7.53 | 46 | 7.18 | 88.64%
2001-07-17 | 3:46 | 18.94 | 7.51 | 46 | 7.22 | 88.59%
2001-07-17 | 4:01 | 18.86 | 7.5 | 46 | 7.06 | 86.63%
2001-07-17 | 4:16 | 18.78 | 7.48 | 46 | 7.13 | 87.48%
2001-07-17 | 4:31 | 18.7 | 7.47 | 46 | 7.01 | 86.01%
2001-07-17 | 4:46 | 18.63 | 7.46 | 46 | 6.95 | 85.28%
2001-07-17 | 5:01 | 18.56 | 7.45 | 46 | 6.96 | 85.4%
2001-07-17 | 5:16 | 18.49 | 7.44 | 46 | 6.94 | 84.63%
2001-07-17 | 5:31 | 18.41 | 7.43 | 46 | 6.93 | 84.51%
2001-07-17 | 5:46 | 18.33 | 7.42 | 46 | 6.89 | 84.02%
2001-07-17 | 6:01 | 18.26 | 7.41 | 46 | 6.98 | 85.12%
2001-07-17 | 6:16 | 18.17 | 7.41 | 47 | 6.99 | 85.24%
2001-07-17 | 6:31 | 18.1 | 7.4 | 46 | 6.69 | 81.59%
2001-07-17 | 6:46 | 18.02 | 7.39 | 47 | 7.09 | 86.46%
2001-07-17 | 7:01 | 17.96 | 7.4 | 47 | 7.03 | 84.7%
2001-07-17 | 7:16 | 17.91 | 7.4 | 47 | 7.06 | 85.06%
2001-07-17 | 7:31 | 17.86 | 7.41 | 47 | 7.03 | 84.7%
2001-07-17 | 7:46 | 17.83 | 7.42 | 47 | 7.22 | 86.99%
2001-07-17 | 8:01 | 17.8 | 7.43 | 47 | 7.27 | 87.59%
2001-07-17 | 8:16 | 17.78 | 7.4 | 47 | 7.4 | 89.16%
2001-07-17 | 8:31 | 17.78 | 7.42 | 47 | 7.43 | 89.52%
2001-07-17 | 8:46 | 17.79 | 7.45 | 47 | 7.55 | 90.96%
2001-07-17 | 9:01 | 17.82 | 7.48 | 47 | 7.62 | 91.81%
2001-07-17 | 9:16 | 17.85 | 7.5 | 47 | 7.86 | 94.7%
2001-07-17 | 9:31 | 17.89 | 7.55 | 47 | 7.75 | 93.37%
2001-07-17 | 9:46 | 17.95 | 7.59 | 47 | 7.85 | 94.58%
2001-07-17 | 10:01 | 17.99 | 7.62 | 47 | 7.99 | 96.27%
2001-07-17 | 10:16 | 18.05 | 7.68 | 47 | 8.06 | 98.29%
2001-07-17 | 10:31 | 18.12 | 7.71 | 47 | 8.1 | 98.78%
2001-07-17 | 10:46 | 18.2 | 7.77 | 47 | 8.26 | 100.7%
2001-07-17 | 11:01 | 18.31 | 7.82 | 47 | 8.4 | 102.4%
2001-07-17 | 11:16 | 18.41 | 7.89 | 47 | 8.53 | 104%
2001-07-17 | 11:31 | 18.52 | 7.94 | 47 | 8.38 | 102.8%
2001-07-17 | 11:46 | 18.63 | 8.01 | 47 | 8.59 | 105.4%
2001-07-17 | 12:01 | 18.73 | 8.07 | 47 | 8.67 | 106.4%
2001-07-17 | 12:16 | 18.8 | 8.12 | 47 | 8.85 | 108.6%
2001-07-17 | 12:31 | 18.87 | 8.16 | 47 | 8.82 | 108.2%
2001-07-17 | 12:46 | 18.91 | 8.2 | 47 | 8.78 | 107.7%
2001-07-17 | 13:01 | 18.97 | 8.22 | 47 | 8.88 | 109%
2001-07-17 | 13:16 | 19.04 | 8.23 | 47 | 9 | 111.1%
2001-07-17 | 13:31 | 19.09 | 8.28 | 47 | 9.05 | 111.7%
2001-07-17 | 13:46 | 19.17 | 8.3 | 47 | 9.17 | 113.2%
2001-07-17 | 14:01 | 19.22 | 8.34 | 46 | 9.11 | 112.5%
2001-07-17 | 14:16 | 19.28 | 8.36 | 46 | 8.88 | 109.6%
2001-07-17 | 14:31 | 19.32 | 8.4 | 46 | 9.24 | 114.1%
2001-07-17 | 14:46 | 19.35 | 8.4 | 46 | 9.22 | 113.8%
2001-07-17 | 15:01 | 19.38 | 8.42 | 46 | 9.17 | 113.2%
2001-07-17 | 15:16 | 19.42 | 8.43 | 46 | 9.2 | 113.6%
2001-07-17 | 15:31 | 19.45 | 8.44 | 46 | 9.24 | 114.1%
2001-07-17 | 15:46 | 19.46 | 8.44 | 46 | 9.2 | 113.6%
2001-07-17 | 16:01 | 19.46 | 8.44 | 46 | 9.14 | 112.8%
2001-07-17 | 16:16 | 19.48 | 8.43 | 46 | 9.16 | 113.1%
2001-07-17 | 16:31 | 19.49 | 8.45 | 46 | 9.2 | 113.6%
2001-07-17 | 16:46 | 19.52 | 8.45 | 46 | 9.08 | 114.2%
2001-07-17 | 17:01 | 19.57 | 8.47 | 46 | 9.3 | 117%
2001-07-17 | 17:16 | 19.6 | 8.49 | 46 | 9.14 | 115%
2001-07-17 | 17:31 | 19.67 | 8.49 | 46 | 9.15 | 115.1%
2001-07-17 | 17:46 | 19.69 | 8.5 | 46 | 9.17 | 115.3%
2001-07-17 | 18:01 | 19.69 | 8.5 | 46 | 9.19 | 115.6%
2001-07-17 | 18:16 | 19.68 | 8.5 | 46 | 9.16 | 115.2%
2001-07-17 | 18:31 | 19.66 | 8.5 | 46 | 9.24 | 116.2%
2001-07-17 | 18:46 | 19.62 | 8.49 | 46 | 9.1 | 114.5%
2001-07-17 | 19:01 | 19.57 | 8.48 | 46 | 9.18 | 115.5%
2001-07-17 | 19:16 | 19.51 | 8.47 | 46 | 9.21 | 115.8%
2001-07-17 | 19:31 | 19.46 | 8.45 | 46 | 8.97 | 110.7%
2001-07-17 | 19:46 | 19.4 | 8.43 | 46 | 8.75 | 108%
2001-07-17 | 20:01 | 19.34 | 8.4 | 45 | 8.89 | 109.8%
2001-07-17 | 20:16 | 19.27 | 8.38 | 45 | 8.91 | 110%
2001-07-17 | 20:31 | 19.2 | 8.35 | 45 | 8.69 | 107.3%
2001-07-17 | 20:46 | 19.13 | 8.31 | 45 | 8.67 | 107%
2001-07-17 | 21:01 | 19.05 | 8.27 | 45 | 8.71 | 107.5%
2001-07-17 | 21:16 | 18.97 | 8.23 | 45 | 8.52 | 104.5%
2001-07-17 | 21:31 | 18.89 | 8.19 | 45 | 8.44 | 103.6%
2001-07-17 | 21:46 | 18.82 | 8.13 | 45 | 8.37 | 102.7%
2001-07-17 | 22:01 | 18.73 | 8.09 | 45 | 8.32 | 102.1%
2001-07-17 | 22:16 | 18.66 | 8.04 | 45 | 8.27 | 101.5%
2001-07-17 | 22:31 | 18.59 | 8.01 | 45 | 8.22 | 100.9%
2001-07-17 | 22:46 | 18.52 | 7.95 | 45 | 8.18 | 100.4%
2001-07-17 | 23:01 | 18.45 | 7.92 | 45 | 8.14 | 99.27%
2001-07-17 | 23:16 | 18.37 | 7.88 | 45 | 8.08 | 98.54%
2001-07-17 | 23:31 | 18.3 | 7.83 | 45 | 8.04 | 98.05%
2001-07-17 | 23:46 | 18.22 | 7.82 | 45 | 8.02 | 97.8%
2001-07-18 | 0:01 | 18.15 | 7.78 | 45 | 7.95 | 96.95%
2001-07-18 | 0:16 | 18.08 | 7.74 | 45 | 7.93 | 96.71%
2001-07-18 | 0:31 | 18.02 | 7.72 | 45 | 7.91 | 96.46%
2001-07-18 | 0:46 | 17.96 | 7.69 | 45 | 7.87 | 94.82%
2001-07-18 | 1:01 | 17.89 | 7.68 | 45 | 7.85 | 94.58%
2001-07-18 | 1:16 | 17.84 | 7.65 | 45 | 7.82 | 94.22%
2001-07-18 | 1:31 | 17.79 | 7.62 | 45 | 7.78 | 93.73%
2001-07-18 | 1:46 | 17.75 | 7.6 | 45 | 7.75 | 93.37%
2001-07-18 | 2:01 | 17.71 | 7.58 | 45 | 7.72 | 93.01%
2001-07-18 | 2:16 | 17.67 | 7.56 | 45 | 7.7 | 92.77%
2001-07-18 | 2:31 | 17.64 | 7.53 | 45 | 7.67 | 92.41%
2001-07-18 | 2:46 | 17.6 | 7.53 | 45 | 7.64 | 92.05%
2001-07-18 | 3:01 | 17.57 | 7.49 | 45 | 7.62 | 91.81%
2001-07-18 | 3:16 | 17.55 | 7.47 | 45 | 7.59 | 91.45%
2001-07-18 | 3:31 | 17.51 | 7.46 | 45 | 7.56 | 91.08%
2001-07-18 | 3:46 | 17.48 | 7.45 | 45 | 7.54 | 89.76%
2001-07-18 | 4:01 | 17.45 | 7.44 | 46 | 7.51 | 89.4%
2001-07-18 | 4:16 | 17.41 | 7.41 | 46 | 7.49 | 89.17%
2001-07-18 | 4:31 | 17.38 | 7.42 | 46 | 7.48 | 89.05%
2001-07-18 | 4:46 | 17.35 | 7.42 | 46 | 7.47 | 88.93%
2001-07-18 | 5:01 | 17.3 | 7.37 | 46 | 7.41 | 88.21%
2001-07-18 | 5:16 | 17.26 | 7.36 | 46 | 7.4 | 88.1%
2001-07-18 | 5:31 | 17.22 | 7.35 | 46 | 7.39 | 87.98%
2001-07-18 | 5:46 | 17.2 | 7.34 | 46 | 7.38 | 87.86%
2001-07-18 | 6:01 | 17.17 | 7.33 | 46 | 7.37 | 87.74%
2001-07-18 | 6:16 | 17.15 | 7.33 | 46 | 7.36 | 87.62%
2001-07-18 | 6:31 | 17.13 | 7.32 | 46 | 7.37 | 87.74%
2001-07-18 | 6:46 | 17.11 | 7.31 | 46 | 7.36 | 87.62%
2001-07-18 | 7:01 | 17.1 | 7.32 | 46 | 7.36 | 87.62%
2001-07-18 | 7:16 | 17.09 | 7.32 | 46 | 7.38 | 87.86%
2001-07-18 | 7:31 | 17.08 | 7.32 | 46 | 7.41 | 88.21%
2001-07-18 | 7:46 | 17.07 | 7.33 | 46 | 7.43 | 88.45%
2001-07-18 | 8:01 | 17.05 | 7.34 | 46 | 7.46 | 88.81%
2001-07-18 | 8:16 | 17.03 | 7.34 | 46 | 7.49 | 89.17%
2001-07-18 | 8:31 | 17.02 | 7.35 | 47 | 7.52 | 89.52%
2001-07-18 | 8:46 | 17.02 | 7.39 | 47 | 7.55 | 89.88%
2001-07-18 | 9:01 | 17.06 | 7.38 | 47 | 7.62 | 90.71%
2001-07-18 | 9:16 | 17.09 | 7.4 | 47 | 7.7 | 91.67%
2001-07-18 | 9:31 | 17.11 | 7.43 | 47 | 7.79 | 92.74%
2001-07-18 | 9:46 | 17.14 | 7.48 | 46 | 8.03 | 95.6%
2001-07-18 | 10:01 | 17.15 | 7.5 | 46 | 8.02 | 95.48%
2001-07-18 | 10:16 | 17.18 | 7.53 | 46 | 8.15 | 97.02%
2001-07-18 | 10:31 | 17.22 | 7.56 | 46 | 8.46 | 100.7%
2001-07-18 | 10:46 | 17.29 | 7.63 | 46 | 8.34 | 99.29%
2001-07-18 | 11:01 | 17.34 | 7.66 | 46 | 8.26 | 98.33%
2001-07-18 | 11:16 | 17.42 | 7.7 | 46 | 8.38 | 99.76%
2001-07-18 | 11:31 | 17.47 | 7.75 | 46 | 8.48 | 101%
2001-07-18 | 11:46 | 17.53 | 7.82 | 46 | 8.59 | 103.5%
2001-07-18 | 12:01 | 17.59 | 7.87 | 46 | 8.69 | 104.7%
2001-07-18 | 12:16 | 17.72 | 7.92 | 46 | 8.78 | 105.8%
2001-07-18 | 12:31 | 17.82 | 7.98 | 46 | 8.85 | 106.6%
2001-07-18 | 12:46 | 17.94 | 8.03 | 46 | 8.89 | 107.1%
2001-07-18 | 13:01 | 18.11 | 8.09 | 46 | 9.01 | 109.9%
2001-07-18 | 13:16 | 18.18 | 8.13 | 46 | 9.12 | 111.2%
2001-07-18 | 13:31 | 18.34 | 8.18 | 46 | 9.11 | 111.1%
2001-07-18 | 13:46 | 18.47 | 8.22 | 46 | 9.11 | 111.1%
2001-07-18 | 14:01 | 18.59 | 8.25 | 46 | 9.17 | 112.5%
2001-07-18 | 14:16 | 18.7 | 8.3 | 46 | 9.23 | 113.3%
2001-07-18 | 14:31 | 18.83 | 8.32 | 46 | 9.29 | 114%
2001-07-18 | 14:46 | 18.93 | 8.37 | 46 | 9.35 | 114.7%
2001-07-18 | 15:01 | 18.98 | 8.39 | 46 | 9.36 | 114.8%
2001-07-18 | 15:16 | 19.04 | 8.42 | 46 | 9.39 | 115.9%
2001-07-18 | 15:31 | 19.12 | 8.43 | 46 | 9.39 | 115.9%
2001-07-18 | 15:46 | 19.2 | 8.45 | 46 | 9.6 | 118.5%
2001-07-18 | 16:01 | 19.27 | 8.47 | 45 | 9.51 | 117.4%
2001-07-18 | 16:16 | 19.35 | 8.49 | 45 | 9.37 | 115.7%
2001-07-18 | 16:31 | 19.39 | 8.49 | 45 | 9.54 | 117.8%
2001-07-18 | 16:46 | 19.41 | 8.5 | 45 | 9.58 | 118.3%
2001-07-18 | 17:01 | 19.44 | 8.51 | 45 | 9.64 | 119%
2001-07-18 | 17:16 | 19.48 | 8.52 | 45 | 9.56 | 118%
2001-07-18 | 17:31 | 19.54 | 8.53 | 45 | 9.46 | 119%
2001-07-18 | 17:46 | 19.6 | 8.54 | 45 | 9.58 | 120.5%
2001-07-18 | 18:01 | 19.66 | 8.56 | 45 | 9.6 | 120.8%
2001-07-18 | 18:16 | 19.71 | 8.56 | 45 | 9.58 | 120.5%
2001-07-18 | 18:31 | 19.72 | 8.57 | 45 | 9.55 | 120.1%
2001-07-18 | 18:46 | 19.72 | 8.57 | 45 | 9.57 | 120.4%
2001-07-18 | 19:01 | 19.71 | 8.56 | 45 | 9.52 | 119.7%
2001-07-18 | 19:16 | 19.71 | 8.55 | 45 | 9.45 | 118.9%
2001-07-18 | 19:31 | 19.7 | 8.54 | 45 | 9.33 | 117.4%
2001-07-18 | 19:46 | 19.68 | 8.53 | 45 | 9.34 | 117.5%
2001-07-18 | 20:01 | 19.64 | 8.51 | 45 | 9.26 | 116.5%
2001-07-18 | 20:16 | 19.6 | 8.49 | 45 | 9.19 | 115.6%
2001-07-18 | 20:31 | 19.55 | 8.45 | 45 | 9.11 | 114.6%
2001-07-18 | 20:46 | 19.49 | 8.43 | 45 | 9.02 | 111.4%
2001-07-18 | 21:01 | 19.43 | 8.39 | 45 | 8.94 | 110.4%
2001-07-18 | 21:16 | 19.36 | 8.35 | 45 | 8.82 | 108.9%
2001-07-18 | 21:31 | 19.29 | 8.31 | 45 | 8.73 | 107.8%
2001-07-18 | 21:46 | 19.22 | 8.25 | 44 | 8.65 | 106.8%
2001-07-18 | 22:01 | 19.15 | 8.19 | 44 | 8.57 | 105.8%
2001-07-18 | 22:16 | 19.09 | 8.15 | 44 | 8.5 | 104.9%
2001-07-18 | 22:31 | 19.04 | 8.09 | 44 | 8.43 | 104.1%
2001-07-18 | 22:46 | 18.98 | 8.06 | 44 | 8.37 | 102.7%
2001-07-18 | 23:01 | 18.92 | 8.01 | 44 | 8.3 | 101.8%
2001-07-18 | 23:16 | 18.86 | 7.96 | 44 | 8.25 | 101.2%
2001-07-18 | 23:31 | 18.78 | 7.93 | 44 | 8.21 | 100.7%
2001-07-18 | 23:46 | 18.7 | 7.9 | 44 | 8.16 | 100.1%
2001-07-19 | 0:01 | 18.62 | 7.85 | 44 | 8.12 | 99.63%
2001-07-19 | 0:16 | 18.54 | 7.81 | 44 | 8.08 | 99.14%
2001-07-19 | 0:31 | 18.46 | 7.78 | 44 | 8.05 | 98.17%
2001-07-19 | 0:46 | 18.39 | 7.76 | 44 | 8.01 | 97.68%
2001-07-19 | 1:01 | 18.32 | 7.7 | 44 | 7.97 | 97.2%
2001-07-19 | 1:16 | 18.25 | 7.67 | 44 | 7.94 | 96.83%
2001-07-19 | 1:31 | 18.19 | 7.64 | 44 | 7.9 | 96.34%
2001-07-19 | 1:46 | 18.12 | 7.62 | 44 | 7.87 | 95.98%
2001-07-19 | 2:01 | 18.06 | 7.59 | 44 | 7.82 | 95.37%
2001-07-19 | 2:16 | 18 | 7.57 | 44 | 7.79 | 95%
2001-07-19 | 2:31 | 17.93 | 7.54 | 44 | 7.75 | 93.37%
2001-07-19 | 2:46 | 17.86 | 7.53 | 45 | 7.72 | 93.01%
2001-07-19 | 3:01 | 17.78 | 7.53 | 45 | 7.7 | 92.77%
2001-07-19 | 3:16 | 17.7 | 7.5 | 45 | 7.68 | 92.53%
2001-07-19 | 3:31 | 17.63 | 7.47 | 45 | 7.67 | 92.41%
2001-07-19 | 3:46 | 17.54 | 7.46 | 45 | 7.64 | 92.05%
2001-07-19 | 4:01 | 17.45 | 7.43 | 45 | 7.63 | 90.83%
2001-07-19 | 4:16 | 17.37 | 7.43 | 45 | 7.62 | 90.71%
2001-07-19 | 4:31 | 17.28 | 7.43 | 45 | 7.6 | 90.48%
2001-07-19 | 4:46 | 17.21 | 7.4 | 45 | 7.6 | 90.48%
2001-07-19 | 5:01 | 17.13 | 7.38 | 45 | 7.58 | 90.24%
2001-07-19 | 5:16 | 17.05 | 7.37 | 45 | 7.58 | 90.24%
2001-07-19 | 5:31 | 16.98 | 7.36 | 45 | 7.56 | 89.47%
2001-07-19 | 5:46 | 16.91 | 7.34 | 45 | 7.56 | 89.47%
2001-07-19 | 6:01 | 16.85 | 7.34 | 45 | 7.56 | 89.47%
2001-07-19 | 6:16 | 16.79 | 7.33 | 45 | 7.56 | 89.47%
2001-07-19 | 6:31 | 16.74 | 7.33 | 45 | 7.56 | 89.47%
2001-07-19 | 6:46 | 16.7 | 7.32 | 45 | 7.57 | 89.59%
2001-07-19 | 7:01 | 16.66 | 7.33 | 45 | 7.6 | 89.94%
2001-07-19 | 7:16 | 16.64 | 7.33 | 45 | 7.62 | 90.18%
2001-07-19 | 7:31 | 16.62 | 7.33 | 45 | 7.66 | 90.65%
2001-07-19 | 7:46 | 16.62 | 7.34 | 45 | 7.71 | 91.24%
2001-07-19 | 8:01 | 16.63 | 7.37 | 45 | 7.79 | 92.19%
2001-07-19 | 8:16 | 16.64 | 7.37 | 45 | 7.86 | 93.02%
2001-07-19 | 8:31 | 16.67 | 7.42 | 45 | 7.93 | 93.85%
2001-07-19 | 8:46 | 16.72 | 7.44 | 45 | 8.01 | 94.79%
2001-07-19 | 9:01 | 16.77 | 7.46 | 45 | 8.1 | 95.86%
2001-07-19 | 9:16 | 16.82 | 7.51 | 45 | 8.2 | 97.04%
2001-07-19 | 9:31 | 16.86 | 7.57 | 45 | 8.27 | 97.87%
2001-07-19 | 9:46 | 16.92 | 7.59 | 45 | 8.36 | 98.93%
2001-07-19 | 10:01 | 16.98 | 7.61 | 45 | 8.46 | 100.1%
2001-07-19 | 10:16 | 17.08 | 7.67 | 45 | 8.55 | 101.8%
2001-07-19 | 10:31 | 17.17 | 7.71 | 45 | 8.63 | 102.7%
2001-07-19 | 10:46 | 17.29 | 7.77 | 45 | 8.7 | 103.6%
2001-07-19 | 11:01 | 17.41 | 7.82 | 45 | 8.78 | 104.5%
2001-07-19 | 11:16 | 17.54 | 7.88 | 45 | 8.86 | 106.7%
2001-07-19 | 11:31 | 17.66 | 7.93 | 45 | 8.93 | 107.6%
2001-07-19 | 11:46 | 17.78 | 7.99 | 45 | 8.99 | 108.3%
2001-07-19 | 12:01 | 17.94 | 8.04 | 45 | 9.05 | 109%
2001-07-19 | 12:16 | 17.99 | 8.08 | 45 | 9.08 | 109.4%
2001-07-19 | 12:31 | 18.14 | 8.12 | 45 | 9.14 | 111.5%
2001-07-19 | 12:46 | 18.23 | 8.16 | 45 | 9.17 | 111.8%
2001-07-19 | 13:01 | 18.29 | 8.18 | 45 | 9.21 | 112.3%
2001-07-19 | 13:16 | 18.33 | 8.22 | 45 | 9.22 | 112.4%
2001-07-19 | 13:31 | 18.47 | 8.22 | 45 | 9.28 | 113.2%
2001-07-19 | 13:46 | 18.62 | 8.25 | 45 | 9.32 | 114.4%
2001-07-19 | 14:01 | 18.8 | 8.3 | 45 | 9.34 | 114.6%
2001-07-19 | 14:16 | 18.86 | 8.31 | 45 | 9.36 | 114.8%
2001-07-19 | 14:31 | 18.96 | 8.35 | 44 | 9.4 | 115.3%
2001-07-19 | 14:46 | 18.99 | 8.35 | 44 | 9.4 | 115.3%
2001-07-19 | 15:01 | 18.96 | 8.35 | 44 | 9.37 | 115%
2001-07-19 | 15:16 | 18.88 | 8.32 | 44 | 9.26 | 113.6%
2001-07-19 | 15:31 | 18.8 | 8.27 | 44 | 9.18 | 112.6%
2001-07-19 | 15:46 | 18.76 | 8.24 | 44 | 9.11 | 111.8%
2001-07-19 | 16:01 | 18.74 | 8.21 | 44 | 9.03 | 110.8%
2001-07-19 | 16:16 | 18.7 | 8.17 | 44 | 8.98 | 110.2%
2001-07-19 | 16:31 | 18.66 | 8.14 | 44 | 8.89 | 109.1%
2001-07-19 | 16:46 | 18.64 | 8.1 | 44 | 8.86 | 108.7%
2001-07-19 | 17:01 | 18.67 | 8.1 | 44 | 8.83 | 108.3%
2001-07-19 | 17:16 | 18.75 | 8.11 | 44 | 8.85 | 108.6%
2001-07-19 | 17:31 | 18.81 | 8.13 | 43 | 8.85 | 108.6%
2001-07-19 | 17:46 | 18.84 | 8.14 | 43 | 8.86 | 108.7%
2001-07-19 | 18:01 | 18.87 | 8.16 | 43 | 8.87 | 108.8%
2001-07-19 | 18:16 | 18.89 | 8.18 | 43 | 8.89 | 109.1%
2001-07-19 | 18:31 | 18.89 | 8.18 | 43 | 8.89 | 109.1%
2001-07-19 | 18:46 | 18.92 | 8.2 | 43 | 8.91 | 109.3%
2001-07-19 | 19:01 | 18.92 | 8.21 | 43 | 8.86 | 108.7%
2001-07-19 | 19:16 | 18.91 | 8.21 | 43 | 8.86 | 108.7%
2001-07-19 | 19:31 | 18.91 | 8.21 | 43 | 8.84 | 108.5%
2001-07-19 | 19:46 | 18.9 | 8.19 | 43 | 8.84 | 108.5%
2001-07-19 | 20:01 | 18.88 | 8.16 | 43 | 8.8 | 108%
2001-07-19 | 20:16 | 18.86 | 8.14 | 43 | 8.74 | 107.2%
2001-07-19 | 20:31 | 18.84 | 8.09 | 43 | 8.68 | 106.5%
2001-07-19 | 20:46 | 18.82 | 8.05 | 43 | 8.6 | 105.5%
2001-07-19 | 21:01 | 18.69 | 7.98 | 43 | 8.53 | 104.7%
2001-07-19 | 21:16 | 18.63 | 7.92 | 43 | 8.43 | 103.4%
2001-07-19 | 21:31 | 18.57 | 7.87 | 43 | 8.37 | 102.7%
2001-07-19 | 21:46 | 18.5 | 7.82 | 43 | 8.29 | 101.7%
2001-07-19 | 22:01 | 18.42 | 7.77 | 43 | 8.22 | 100.2%
2001-07-19 | 22:16 | 18.31 | 7.71 | 43 | 8.18 | 99.76%
2001-07-19 | 22:31 | 18.22 | 7.67 | 43 | 8.14 | 99.27%
2001-07-19 | 22:46 | 18.14 | 7.63 | 43 | 8.09 | 98.66%
2001-07-19 | 23:01 | 18.02 | 7.6 | 43 | 8.07 | 98.41%
2001-07-19 | 23:16 | 17.89 | 7.5 | 42 | 8.07 | 97.23%
2001-07-19 | 23:31 | 17.76 | 7.5 | 42 | 8.03 | 96.75%
2001-07-19 | 23:46 | 17.65 | 7.47 | 43 | 8.01 | 96.51%
2001-07-20 | 0:01 | 17.55 | 7.46 | 43 | 7.98 | 96.14%
2001-07-20 | 0:16 | 17.47 | 7.43 | 43 | 7.96 | 94.76%
2001-07-20 | 0:31 | 17.44 | 7.42 | 43 | 7.92 | 94.29%
2001-07-20 | 0:46 | 17.43 | 7.41 | 43 | 7.89 | 93.93%
2001-07-20 | 1:01 | 17.44 | 7.39 | 43 | 7.87 | 93.69%
2001-07-20 | 1:16 | 17.44 | 7.38 | 43 | 7.83 | 93.21%
2001-07-20 | 1:31 | 17.44 | 7.37 | 43 | 7.81 | 92.98%
2001-07-20 | 1:46 | 17.42 | 7.35 | 43 | 7.78 | 92.62%
2001-07-20 | 2:01 | 17.36 | 7.34 | 43 | 7.75 | 92.26%
2001-07-20 | 2:16 | 17.3 | 7.32 | 43 | 7.73 | 92.02%
2001-07-20 | 2:31 | 17.23 | 7.32 | 43 | 7.71 | 91.79%
2001-07-20 | 2:46 | 17.16 | 7.29 | 43 | 7.69 | 91.55%
2001-07-20 | 3:01 | 17.11 | 7.28 | 43 | 7.68 | 91.43%
2001-07-20 | 3:16 | 17.07 | 7.32 | 43 | 7.65 | 91.07%
2001-07-20 | 3:31 | 17.04 | 7.29 | 44 | 7.64 | 90.95%
2001-07-20 | 3:46 | 17.02 | 7.28 | 44 | 7.62 | 90.71%
2001-07-20 | 4:01 | 17 | 7.28 | 44 | 7.61 | 90.6%
2001-07-20 | 4:16 | 16.96 | 7.26 | 44 | 7.59 | 89.82%
2001-07-20 | 4:31 | 16.93 | 7.23 | 44 | 7.58 | 89.7%
2001-07-20 | 4:46 | 16.89 | 7.21 | 44 | 7.56 | 89.47%
2001-07-20 | 5:01 | 16.86 | 7.22 | 44 | 7.55 | 89.35%
2001-07-20 | 5:16 | 16.84 | 7.22 | 44 | 7.54 | 89.23%
2001-07-20 | 5:31 | 16.82 | 7.21 | 44 | 7.52 | 88.99%
2001-07-20 | 5:46 | 16.8 | 7.21 | 44 | 7.5 | 88.76%
2001-07-20 | 6:01 | 16.78 | 7.22 | 44 | 7.5 | 88.76%
2001-07-20 | 6:16 | 16.76 | 7.2 | 44 | 7.48 | 88.52%
2001-07-20 | 6:31 | 16.73 | 7.17 | 44 | 7.48 | 88.52%
2001-07-20 | 6:46 | 16.7 | 7.19 | 44 | 7.48 | 88.52%
2001-07-20 | 7:01 | 16.67 | 7.18 | 44 | 7.48 | 88.52%
2001-07-20 | 7:16 | 16.65 | 7.18 | 44 | 7.49 | 88.64%
2001-07-20 | 7:31 | 16.62 | 7.19 | 45 | 7.49 | 88.64%
2001-07-20 | 7:46 | 16.6 | 7.19 | 45 | 7.5 | 88.76%
2001-07-20 | 8:01 | 16.57 | 7.19 | 45 | 7.51 | 88.88%
2001-07-20 | 8:16 | 16.55 | 7.18 | 45 | 7.52 | 88.99%
2001-07-20 | 8:31 | 16.55 | 7.18 | 45 | 7.56 | 89.47%
2001-07-20 | 8:46 | 16.55 | 7.22 | 44 | 7.62 | 90.18%
2001-07-20 | 9:01 | 16.56 | 7.21 | 44 | 7.7 | 91.12%