## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | d/s Browns Crossing
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2002-04-29 14:09 | 11.2 | 7.8 | 9.2 | 97 | 43 | 0
Sonde | 2002-04-29 16:16 | 11.38 | 7.76 | 9.54 | 101 | 41
Grade |  | A | A | B |  | A
Audit 2 | 2002-05-01 10:25 | 11.6 | 7.8 | 9.8 | 105 | 45 | 0
Sonde | 2002-05-01 11:46 | 11.61 | 7.73 | 9.81 | 105 | 41
Grade |  | A | A | A |  | A
Audit 3 | 2002-05-03 09:26 | 11.9 | 7.9 | 9.5 | 102 | 42 | 0
Sonde | 2002-05-03 09:31 | 11.29 | 7.71 | 9.51 | 101 | 39
Grade |  | A | A | A |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note Audits 1, 2, & 3 are macro(ed) into this audit sheet and represent the beginning, middle and end of the deployment. Audits 4 and 5 were in the middle of the deployment and are not macro(ed) into this sheet. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-29 14:09 | 11.2 | 7.8 | 9.2 | 97 | 43
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-01 10:25 | 11.6 | 7.8 | 9.8 | 105 | 45
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-03 09:26 | 11.9 | 7.9 | 9.5 | 102 | 42
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-30 14:50 | 11.8 | 7.8 | 9 | 96 | 45
Sonde Values | 2002-04-30 15:01 | 11.92 | 7.76 | 9.8 | 105 | 41
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | DR 237.50
Program Name | Water Quality Monitoring Program |  | Description |  | d/s Browns Crossing
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 15:05 | 12.5 | 7.6 | 9 | 99 | 42
Sonde Values | 2002-05-02 15:01 | 12.64 | 7.75 | 9.88 | 109 | 40
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | DO
|  | °C |  | uS/cm | mg/L | % Sat
2002-04-29 | 14:16 | 11.31 | 7.74 | 41 | 9.9 | 104.8%
2002-04-29 | 14:31 | 11.34 | 7.75 | 41 | 9.95 | 105.3%
2002-04-29 | 14:46 | 11.29 | 7.75 | 41 | 9.76 | 103.3%
2002-04-29 | 15:01 | 11.36 | 7.76 | 41 | 9.82 | 103.9%
2002-04-29 | 15:16 | 11.37 | 7.75 | 41 | 9.75 | 103.2%
2002-04-29 | 15:31 | 11.4 | 7.76 | 41 | 9.68 | 102.4%
2002-04-29 | 15:46 | 11.44 | 7.76 | 41 | 9.65 | 102.1%
2002-04-29 | 16:01 | 11.4 | 7.76 | 41 | 9.62 | 101.8%
2002-04-29 | 16:16 | 11.38 | 7.76 | 41 | 9.54 | 101%
2002-04-29 | 16:31 | 11.37 | 7.75 | 41 | 9.49 | 100.4%
2002-04-29 | 16:46 | 11.4 | 7.76 | 41 | 9.45 | 100%
2002-04-29 | 17:01 | 11.39 | 7.76 | 41 | 9.39 | 99.37%
2002-04-29 | 17:16 | 11.41 | 7.76 | 41 | 9.33 | 98.73%
2002-04-29 | 17:31 | 11.42 | 7.76 | 41 | 9.31 | 98.52%
2002-04-29 | 17:46 | 11.41 | 7.74 | 41 | 9.28 | 98.2%
2002-04-29 | 18:01 | 11.39 | 7.74 | 41 | 9.27 | 98.1%
2002-04-29 | 18:16 | 11.4 | 7.73 | 41 | 9.24 | 97.78%
2002-04-29 | 18:31 | 11.43 | 7.73 | 41 | 9.21 | 97.46%
2002-04-29 | 18:46 | 11.47 | 7.73 | 41 | 9.18 | 97.14%
2002-04-29 | 19:01 | 11.5 | 7.72 | 41 | 9.15 | 97.86%
2002-04-29 | 19:16 | 11.5 | 7.72 | 41 | 9.12 | 97.54%
2002-04-29 | 19:31 | 11.52 | 7.72 | 41 | 9.09 | 97.22%
2002-04-29 | 19:46 | 11.55 | 7.71 | 41 | 9.06 | 96.9%
2002-04-29 | 20:01 | 11.59 | 7.72 | 41 | 9.03 | 96.58%
2002-04-29 | 20:16 | 11.63 | 7.71 | 41 | 9 | 96.26%
2002-04-29 | 20:31 | 11.68 | 7.71 | 41 | 8.97 | 95.94%
2002-04-29 | 20:46 | 11.72 | 7.71 | 41 | 8.94 | 95.61%
2002-04-29 | 21:01 | 11.77 | 7.71 | 41 | 8.92 | 95.4%
2002-04-29 | 21:16 | 11.8 | 7.71 | 41 | 8.91 | 95.29%
2002-04-29 | 21:31 | 11.79 | 7.71 | 41 | 8.9 | 95.19%
2002-04-29 | 21:46 | 11.79 | 7.71 | 41 | 8.9 | 95.19%
2002-04-29 | 22:01 | 11.79 | 7.71 | 41 | 8.89 | 95.08%
2002-04-29 | 22:16 | 11.77 | 7.71 | 41 | 8.89 | 95.08%
2002-04-29 | 22:31 | 11.75 | 7.71 | 41 | 8.89 | 95.08%
2002-04-29 | 22:46 | 11.74 | 7.71 | 41 | 8.89 | 95.08%
2002-04-29 | 23:01 | 11.72 | 7.71 | 41 | 8.89 | 95.08%
2002-04-29 | 23:16 | 11.7 | 7.71 | 41 | 8.89 | 95.08%
2002-04-29 | 23:31 | 11.69 | 7.71 | 41 | 8.89 | 95.08%
2002-04-29 | 23:46 | 11.68 | 7.71 | 41 | 8.89 | 95.08%
2002-04-30 | 0:01 | 11.66 | 7.71 | 41 | 8.89 | 95.08%
2002-04-30 | 0:16 | 11.63 | 7.71 | 41 | 8.89 | 95.08%
2002-04-30 | 0:31 | 11.62 | 7.71 | 41 | 8.9 | 95.19%
2002-04-30 | 0:46 | 11.61 | 7.71 | 41 | 8.9 | 95.19%
2002-04-30 | 1:01 | 11.6 | 7.71 | 41 | 8.9 | 95.19%
2002-04-30 | 1:16 | 11.58 | 7.71 | 41 | 8.9 | 95.19%
2002-04-30 | 1:31 | 11.55 | 7.71 | 41 | 8.9 | 95.19%
2002-04-30 | 1:46 | 11.52 | 7.71 | 41 | 8.91 | 95.29%
2002-04-30 | 2:01 | 11.48 | 7.71 | 41 | 8.91 | 94.29%
2002-04-30 | 2:16 | 11.44 | 7.71 | 41 | 8.92 | 94.39%
2002-04-30 | 2:31 | 11.4 | 7.71 | 41 | 8.92 | 94.39%
2002-04-30 | 2:46 | 11.37 | 7.71 | 41 | 8.93 | 94.5%
2002-04-30 | 3:01 | 11.35 | 7.71 | 41 | 8.94 | 94.6%
2002-04-30 | 3:16 | 11.32 | 7.71 | 41 | 8.94 | 94.6%
2002-04-30 | 3:31 | 11.29 | 7.71 | 41 | 8.94 | 94.6%
2002-04-30 | 3:46 | 11.27 | 7.71 | 41 | 8.95 | 94.71%
2002-04-30 | 4:01 | 11.25 | 7.71 | 41 | 8.95 | 94.71%
2002-04-30 | 4:16 | 11.22 | 7.71 | 41 | 8.95 | 94.71%
2002-04-30 | 4:31 | 11.2 | 7.71 | 40 | 8.95 | 94.71%
2002-04-30 | 4:46 | 11.18 | 7.71 | 41 | 8.96 | 94.81%
2002-04-30 | 5:01 | 11.16 | 7.71 | 41 | 8.97 | 94.92%
2002-04-30 | 5:16 | 11.13 | 7.71 | 41 | 8.97 | 94.92%
2002-04-30 | 5:31 | 11.11 | 7.7 | 41 | 8.97 | 94.92%
2002-04-30 | 5:46 | 11.09 | 7.7 | 41 | 8.98 | 95.03%
2002-04-30 | 6:01 | 11.07 | 7.7 | 41 | 8.98 | 95.03%
2002-04-30 | 6:16 | 11.06 | 7.7 | 41 | 8.99 | 95.13%
2002-04-30 | 6:31 | 11.05 | 7.7 | 41 | 9.01 | 95.34%
2002-04-30 | 6:46 | 11.05 | 7.71 | 41 | 9.03 | 95.56%
2002-04-30 | 7:01 | 11.05 | 7.71 | 41 | 9.06 | 95.87%
2002-04-30 | 7:16 | 11.05 | 7.71 | 41 | 9.09 | 96.19%
2002-04-30 | 7:31 | 11.06 | 7.71 | 41 | 9.13 | 96.61%
2002-04-30 | 7:46 | 11.06 | 7.73 | 41 | 9.16 | 96.93%
2002-04-30 | 8:01 | 11.07 | 7.73 | 41 | 9.2 | 97.35%
2002-04-30 | 8:16 | 11.07 | 7.73 | 41 | 9.24 | 97.78%
2002-04-30 | 8:31 | 11.05 | 7.74 | 41 | 9.27 | 98.1%
2002-04-30 | 8:46 | 11.03 | 7.73 | 41 | 9.3 | 98.41%
2002-04-30 | 9:01 | 11.01 | 7.73 | 41 | 9.32 | 98.62%
2002-04-30 | 9:16 | 11.02 | 7.74 | 41 | 9.33 | 98.73%
2002-04-30 | 9:31 | 11.06 | 7.74 | 41 | 9.35 | 98.94%
2002-04-30 | 9:46 | 11.09 | 7.74 | 41 | 9.36 | 99.05%
2002-04-30 | 10:01 | 11.13 | 7.75 | 41 | 9.42 | 99.68%
2002-04-30 | 10:16 | 11.24 | 7.76 | 41 | 9.44 | 99.89%
2002-04-30 | 10:31 | 11.33 | 7.76 | 41 | 9.44 | 99.89%
2002-04-30 | 10:46 | 11.32 | 7.76 | 41 | 9.46 | 100.1%
2002-04-30 | 11:01 | 11.35 | 7.76 | 41 | 9.47 | 100.2%
2002-04-30 | 11:16 | 11.42 | 7.76 | 41 | 9.48 | 100.3%
2002-04-30 | 11:31 | 11.4 | 7.77 | 41 | 9.53 | 100.8%
2002-04-30 | 11:46 | 11.46 | 7.77 | 41 | 9.61 | 101.7%
2002-04-30 | 12:01 | 11.5 | 7.76 | 41 | 9.6 | 102.7%
2002-04-30 | 12:16 | 11.65 | 7.77 | 41 | 9.69 | 103.6%
2002-04-30 | 12:31 | 11.82 | 7.77 | 41 | 9.75 | 104.3%
2002-04-30 | 12:46 | 11.86 | 7.76 | 41 | 9.78 | 104.6%
2002-04-30 | 13:01 | 11.89 | 7.77 | 41 | 9.81 | 104.9%
2002-04-30 | 13:16 | 11.95 | 7.77 | 41 | 9.79 | 104.7%
2002-04-30 | 13:31 | 11.86 | 7.76 | 41 | 9.75 | 104.3%
2002-04-30 | 13:46 | 11.9 | 7.77 | 41 | 9.8 | 104.8%
2002-04-30 | 14:01 | 11.85 | 7.76 | 39 | 9.79 | 104.7%
2002-04-30 | 14:16 | 11.79 | 7.76 | 41 | 9.75 | 104.3%
2002-04-30 | 14:31 | 11.9 | 7.76 | 41 | 9.81 | 104.9%
2002-04-30 | 14:46 | 11.9 | 7.76 | 41 | 9.8 | 104.8%
2002-04-30 | 15:01 | 11.92 | 7.76 | 41 | 9.8 | 104.8%
2002-04-30 | 15:16 | 11.9 | 7.76 | 39 | 9.78 | 104.6%
2002-04-30 | 15:31 | 11.88 | 7.76 | 41 | 9.75 | 104.3%
2002-04-30 | 15:46 | 11.83 | 7.76 | 41 | 9.71 | 103.9%
2002-04-30 | 16:01 | 11.8 | 7.76 | 41 | 9.66 | 103.3%
2002-04-30 | 16:16 | 11.76 | 7.76 | 41 | 9.57 | 102.4%
2002-04-30 | 16:31 | 11.76 | 7.76 | 41 | 9.52 | 101.8%
2002-04-30 | 16:46 | 11.75 | 7.76 | 41 | 9.47 | 101.3%
2002-04-30 | 17:01 | 11.74 | 7.76 | 41 | 9.43 | 100.9%
2002-04-30 | 17:16 | 11.73 | 7.75 | 41 | 9.39 | 100.4%
2002-04-30 | 17:31 | 11.72 | 7.75 | 41 | 9.37 | 100.2%
2002-04-30 | 17:46 | 11.73 | 7.75 | 41 | 9.34 | 99.89%
2002-04-30 | 18:01 | 11.72 | 7.75 | 41 | 9.3 | 99.47%
2002-04-30 | 18:16 | 11.72 | 7.75 | 41 | 9.28 | 99.25%
2002-04-30 | 18:31 | 11.7 | 7.75 | 41 | 9.25 | 98.93%
2002-04-30 | 18:46 | 11.68 | 7.74 | 41 | 9.23 | 98.72%
2002-04-30 | 19:01 | 11.68 | 7.74 | 41 | 9.2 | 98.4%
2002-04-30 | 19:16 | 11.68 | 7.74 | 41 | 9.17 | 98.07%
2002-04-30 | 19:31 | 11.67 | 7.74 | 41 | 9.14 | 97.75%
2002-04-30 | 19:46 | 11.64 | 7.74 | 41 | 9.11 | 97.43%
2002-04-30 | 20:01 | 11.62 | 7.73 | 41 | 9.09 | 97.22%
2002-04-30 | 20:16 | 11.61 | 7.73 | 41 | 9.07 | 97.01%
2002-04-30 | 20:31 | 11.61 | 7.73 | 41 | 9.06 | 96.9%
2002-04-30 | 20:46 | 11.62 | 7.73 | 41 | 9.05 | 96.79%
2002-04-30 | 21:01 | 11.62 | 7.73 | 41 | 9.04 | 96.68%
2002-04-30 | 21:16 | 11.61 | 7.73 | 41 | 9.03 | 96.58%
2002-04-30 | 21:31 | 11.61 | 7.73 | 41 | 9.03 | 96.58%
2002-04-30 | 21:46 | 11.62 | 7.72 | 41 | 9.01 | 96.36%
2002-04-30 | 22:01 | 11.62 | 7.72 | 41 | 9.01 | 96.36%
2002-04-30 | 22:16 | 11.62 | 7.72 | 41 | 9.01 | 96.36%
2002-04-30 | 22:31 | 11.6 | 7.72 | 41 | 9.01 | 96.36%
2002-04-30 | 22:46 | 11.59 | 7.72 | 41 | 9.01 | 96.36%
2002-04-30 | 23:01 | 11.58 | 7.72 | 41 | 9.01 | 96.36%
2002-04-30 | 23:16 | 11.58 | 7.72 | 41 | 9.01 | 96.36%
2002-04-30 | 23:31 | 11.56 | 7.73 | 41 | 9.01 | 96.36%
2002-04-30 | 23:46 | 11.55 | 7.72 | 41 | 9.01 | 96.36%
2002-05-01 | 0:01 | 11.54 | 7.72 | 41 | 9.01 | 96.36%
2002-05-01 | 0:16 | 11.51 | 7.72 | 41 | 9.01 | 96.36%
2002-05-01 | 0:31 | 11.51 | 7.72 | 41 | 9.01 | 96.36%
2002-05-01 | 0:46 | 11.49 | 7.72 | 41 | 9.01 | 95.34%
2002-05-01 | 1:01 | 11.48 | 7.72 | 41 | 9.02 | 95.45%
2002-05-01 | 1:16 | 11.46 | 7.72 | 41 | 9.02 | 95.45%
2002-05-01 | 1:31 | 11.45 | 7.72 | 41 | 9.02 | 95.45%
2002-05-01 | 1:46 | 11.45 | 7.71 | 41 | 9.02 | 95.45%
2002-05-01 | 2:01 | 11.44 | 7.72 | 41 | 9.02 | 95.45%
2002-05-01 | 2:16 | 11.42 | 7.72 | 41 | 9.03 | 95.56%
2002-05-01 | 2:31 | 11.42 | 7.71 | 41 | 9.02 | 95.45%
2002-05-01 | 2:46 | 11.41 | 7.71 | 41 | 9.03 | 95.56%
2002-05-01 | 3:01 | 11.39 | 7.71 | 41 | 9.03 | 95.56%
2002-05-01 | 3:16 | 11.38 | 7.71 | 41 | 9.03 | 95.56%
2002-05-01 | 3:31 | 11.35 | 7.71 | 41 | 9.03 | 95.56%
2002-05-01 | 3:46 | 11.33 | 7.71 | 41 | 9.04 | 95.66%
2002-05-01 | 4:01 | 11.31 | 7.71 | 41 | 9.05 | 95.77%
2002-05-01 | 4:16 | 11.27 | 7.71 | 41 | 9.06 | 95.87%
2002-05-01 | 4:31 | 11.23 | 7.71 | 41 | 9.06 | 95.87%
2002-05-01 | 4:46 | 11.21 | 7.7 | 41 | 9.07 | 95.98%
2002-05-01 | 5:01 | 11.17 | 7.7 | 41 | 9.08 | 96.08%
2002-05-01 | 5:16 | 11.15 | 7.71 | 41 | 9.08 | 96.08%
2002-05-01 | 5:31 | 11.12 | 7.7 | 41 | 9.08 | 96.08%
2002-05-01 | 5:46 | 11.09 | 7.7 | 41 | 9.09 | 96.19%
2002-05-01 | 6:01 | 11.07 | 7.7 | 41 | 9.1 | 96.3%
2002-05-01 | 6:16 | 11.05 | 7.7 | 41 | 9.11 | 96.4%
2002-05-01 | 6:31 | 11.03 | 7.7 | 41 | 9.12 | 96.51%
2002-05-01 | 6:46 | 11 | 7.7 | 41 | 9.15 | 96.83%
2002-05-01 | 7:01 | 10.97 | 7.71 | 41 | 9.18 | 96.63%
2002-05-01 | 7:16 | 10.94 | 7.7 | 41 | 9.23 | 97.16%
2002-05-01 | 7:31 | 10.91 | 7.71 | 41 | 9.26 | 97.47%
2002-05-01 | 7:46 | 10.9 | 7.71 | 39 | 9.31 | 98%
2002-05-01 | 8:01 | 10.91 | 7.71 | 41 | 9.34 | 98.32%
2002-05-01 | 8:16 | 10.9 | 7.72 | 41 | 9.37 | 98.63%
2002-05-01 | 8:31 | 10.88 | 7.71 | 41 | 9.39 | 98.84%
2002-05-01 | 8:46 | 10.85 | 7.71 | 41 | 9.41 | 99.05%
2002-05-01 | 9:01 | 10.86 | 7.71 | 41 | 9.43 | 99.26%
2002-05-01 | 9:16 | 10.89 | 7.71 | 41 | 9.44 | 99.37%
2002-05-01 | 9:31 | 10.93 | 7.71 | 41 | 9.45 | 99.47%
2002-05-01 | 9:46 | 10.97 | 7.71 | 41 | 9.47 | 99.68%
2002-05-01 | 10:01 | 11.03 | 7.71 | 41 | 9.54 | 101%
2002-05-01 | 10:16 | 11.09 | 7.72 | 41 | 9.53 | 100.8%
2002-05-01 | 10:31 | 11.1 | 7.72 | 41 | 9.51 | 100.6%
2002-05-01 | 10:46 | 11.08 | 7.72 | 41 | 9.53 | 100.8%
2002-05-01 | 11:01 | 11.26 | 7.73 | 41 | 9.52 | 100.7%
2002-05-01 | 11:16 | 11.4 | 7.73 | 41 | 9.53 | 100.8%
2002-05-01 | 11:31 | 11.55 | 7.74 | 41 | 9.64 | 103.1%
2002-05-01 | 11:46 | 11.61 | 7.73 | 41 | 9.81 | 104.9%
2002-05-01 | 12:01 | 11.71 | 7.74 | 41 | 9.89 | 105.8%
2002-05-01 | 12:16 | 11.8 | 7.74 | 41 | 9.93 | 106.2%
2002-05-01 | 12:31 | 11.89 | 7.74 | 41 | 9.96 | 106.5%
2002-05-01 | 12:46 | 11.95 | 7.74 | 41 | 9.98 | 106.7%
2002-05-01 | 13:01 | 11.94 | 7.74 | 41 | 10 | 107%
2002-05-01 | 13:16 | 12.06 | 7.74 | 41 | 9.97 | 108.4%
2002-05-01 | 13:31 | 12.12 | 7.74 | 41 | 9.99 | 108.6%
2002-05-01 | 13:46 | 12.17 | 7.75 | 41 | 10.01 | 108.8%
2002-05-01 | 14:01 | 12.17 | 7.75 | 41 | 9.98 | 108.5%
2002-05-01 | 14:16 | 12.18 | 7.74 | 41 | 9.97 | 108.4%
2002-05-01 | 14:31 | 12.19 | 7.74 | 41 | 9.98 | 108.5%
2002-05-01 | 14:46 | 12.25 | 7.74 | 41 | 9.96 | 108.3%
2002-05-01 | 15:01 | 12.23 | 7.74 | 41 | 9.93 | 107.9%
2002-05-01 | 15:16 | 12.23 | 7.75 | 41 | 9.92 | 107.8%
2002-05-01 | 15:31 | 12.17 | 7.74 | 41 | 9.88 | 107.4%
2002-05-01 | 15:46 | 12.16 | 7.74 | 41 | 9.86 | 107.2%
2002-05-01 | 16:01 | 12.13 | 7.74 | 41 | 9.8 | 106.5%
2002-05-01 | 16:16 | 12.1 | 7.74 | 41 | 9.66 | 105%
2002-05-01 | 16:31 | 12.08 | 7.74 | 41 | 9.59 | 104.2%
2002-05-01 | 16:46 | 12.07 | 7.74 | 41 | 9.51 | 103.4%
2002-05-01 | 17:01 | 12.05 | 7.74 | 41 | 9.45 | 102.7%
2002-05-01 | 17:16 | 12.04 | 7.73 | 41 | 9.41 | 102.3%
2002-05-01 | 17:31 | 12.03 | 7.74 | 41 | 9.37 | 101.8%
2002-05-01 | 17:46 | 12.02 | 7.74 | 41 | 9.35 | 101.6%
2002-05-01 | 18:01 | 12.03 | 7.73 | 41 | 9.32 | 101.3%
2002-05-01 | 18:16 | 12.03 | 7.73 | 41 | 9.3 | 101.1%
2002-05-01 | 18:31 | 12.03 | 7.73 | 41 | 9.27 | 100.8%
2002-05-01 | 18:46 | 12.02 | 7.73 | 41 | 9.25 | 100.5%
2002-05-01 | 19:01 | 12.02 | 7.72 | 41 | 9.22 | 100.2%
2002-05-01 | 19:16 | 12.03 | 7.72 | 41 | 9.2 | 100%
2002-05-01 | 19:31 | 12.02 | 7.72 | 41 | 9.17 | 99.67%
2002-05-01 | 19:46 | 12.01 | 7.72 | 41 | 9.14 | 99.35%
2002-05-01 | 20:01 | 12.01 | 7.72 | 41 | 9.12 | 99.13%
2002-05-01 | 20:16 | 12.01 | 7.72 | 41 | 9.1 | 98.91%
2002-05-01 | 20:31 | 12 | 7.72 | 41 | 9.08 | 98.7%
2002-05-01 | 20:46 | 12 | 7.72 | 41 | 9.07 | 98.59%
2002-05-01 | 21:01 | 12 | 7.72 | 41 | 9.06 | 98.48%
2002-05-01 | 21:16 | 12 | 7.71 | 41 | 9.05 | 98.37%
2002-05-01 | 21:31 | 12.01 | 7.72 | 41 | 9.04 | 98.26%
2002-05-01 | 21:46 | 12.01 | 7.72 | 41 | 9.03 | 98.15%
2002-05-01 | 22:01 | 12.01 | 7.72 | 41 | 9.03 | 98.15%
2002-05-01 | 22:16 | 12.02 | 7.72 | 39 | 9.02 | 98.04%
2002-05-01 | 22:31 | 12.03 | 7.72 | 41 | 9.02 | 98.04%
2002-05-01 | 22:46 | 12.03 | 7.72 | 40 | 9.01 | 97.93%
2002-05-01 | 23:01 | 12.03 | 7.72 | 41 | 9.01 | 97.93%
2002-05-01 | 23:16 | 12.04 | 7.72 | 41 | 9 | 97.83%
2002-05-01 | 23:31 | 12.04 | 7.72 | 40 | 9 | 97.83%
2002-05-01 | 23:46 | 12.04 | 7.72 | 40 | 9 | 97.83%
2002-05-02 | 0:01 | 12.04 | 7.72 | 40 | 9 | 97.83%
2002-05-02 | 0:16 | 12.04 | 7.72 | 40 | 8.99 | 97.72%
2002-05-02 | 0:31 | 12.03 | 7.72 | 40 | 8.99 | 97.72%
2002-05-02 | 0:46 | 12.02 | 7.72 | 40 | 8.99 | 97.72%
2002-05-02 | 1:01 | 12.01 | 7.72 | 40 | 8.99 | 97.72%
2002-05-02 | 1:16 | 11.99 | 7.71 | 40 | 8.99 | 96.15%
2002-05-02 | 1:31 | 11.97 | 7.71 | 40 | 8.99 | 96.15%
2002-05-02 | 1:46 | 11.96 | 7.71 | 40 | 9 | 96.26%
2002-05-02 | 2:01 | 11.95 | 7.71 | 40 | 9 | 96.26%
2002-05-02 | 2:16 | 11.94 | 7.71 | 40 | 9 | 96.26%
2002-05-02 | 2:31 | 11.92 | 7.71 | 40 | 9 | 96.26%
2002-05-02 | 2:46 | 11.9 | 7.71 | 40 | 9 | 96.26%
2002-05-02 | 3:01 | 11.89 | 7.71 | 40 | 9 | 96.26%
2002-05-02 | 3:16 | 11.88 | 7.71 | 40 | 9 | 96.26%
2002-05-02 | 3:31 | 11.88 | 7.71 | 40 | 9 | 96.26%
2002-05-02 | 3:46 | 11.87 | 7.71 | 40 | 9 | 96.26%
2002-05-02 | 4:01 | 11.85 | 7.7 | 40 | 9 | 96.26%
2002-05-02 | 4:16 | 11.83 | 7.7 | 40 | 9 | 96.26%
2002-05-02 | 4:31 | 11.82 | 7.7 | 40 | 9 | 96.26%
2002-05-02 | 4:46 | 11.8 | 7.7 | 40 | 9.01 | 96.36%
2002-05-02 | 5:01 | 11.77 | 7.7 | 40 | 9.01 | 96.36%
2002-05-02 | 5:16 | 11.75 | 7.7 | 40 | 9.01 | 96.36%
2002-05-02 | 5:31 | 11.73 | 7.7 | 40 | 9.02 | 96.47%
2002-05-02 | 5:46 | 11.71 | 7.7 | 38 | 9.02 | 96.47%
2002-05-02 | 6:01 | 11.69 | 7.7 | 40 | 9.02 | 96.47%
2002-05-02 | 6:16 | 11.67 | 7.7 | 40 | 9.04 | 96.68%
2002-05-02 | 6:31 | 11.66 | 7.7 | 40 | 9.06 | 96.9%
2002-05-02 | 6:46 | 11.64 | 7.7 | 40 | 9.09 | 97.22%
2002-05-02 | 7:01 | 11.62 | 7.7 | 40 | 9.11 | 97.43%
2002-05-02 | 7:16 | 11.61 | 7.7 | 40 | 9.15 | 97.86%
2002-05-02 | 7:31 | 11.59 | 7.7 | 40 | 9.18 | 98.18%
2002-05-02 | 7:46 | 11.56 | 7.7 | 40 | 9.21 | 98.5%
2002-05-02 | 8:01 | 11.55 | 7.7 | 40 | 9.24 | 98.82%
2002-05-02 | 8:16 | 11.56 | 7.7 | 40 | 9.26 | 99.04%
2002-05-02 | 8:31 | 11.57 | 7.71 | 40 | 9.28 | 99.25%
2002-05-02 | 8:46 | 11.59 | 7.71 | 40 | 9.29 | 99.36%
2002-05-02 | 9:01 | 11.61 | 7.72 | 40 | 9.31 | 99.57%
2002-05-02 | 9:16 | 11.63 | 7.72 | 40 | 9.32 | 99.68%
2002-05-02 | 9:31 | 11.67 | 7.72 | 40 | 9.33 | 99.79%
2002-05-02 | 9:46 | 11.69 | 7.73 | 40 | 9.34 | 99.89%
2002-05-02 | 10:01 | 11.74 | 7.73 | 40 | 9.44 | 101%
2002-05-02 | 10:16 | 11.81 | 7.73 | 40 | 9.42 | 100.7%
2002-05-02 | 10:31 | 11.87 | 7.74 | 40 | 9.38 | 100.3%
2002-05-02 | 10:46 | 11.96 | 7.74 | 40 | 9.38 | 100.3%
2002-05-02 | 11:01 | 12.06 | 7.75 | 40 | 9.36 | 101.7%
2002-05-02 | 11:16 | 12.13 | 7.75 | 40 | 9.4 | 102.2%
2002-05-02 | 11:31 | 12.23 | 7.75 | 40 | 9.53 | 103.6%
2002-05-02 | 11:46 | 12.3 | 7.75 | 40 | 9.71 | 105.5%
2002-05-02 | 12:01 | 12.36 | 7.75 | 40 | 9.8 | 106.5%
2002-05-02 | 12:16 | 12.43 | 7.76 | 40 | 9.84 | 107%
2002-05-02 | 12:31 | 12.48 | 7.76 | 40 | 9.86 | 107.2%
2002-05-02 | 12:46 | 12.55 | 7.76 | 40 | 9.89 | 108.7%
2002-05-02 | 13:01 | 12.58 | 7.76 | 40 | 9.9 | 108.8%
2002-05-02 | 13:16 | 12.63 | 7.76 | 40 | 9.91 | 108.9%
2002-05-02 | 13:31 | 12.67 | 7.76 | 40 | 9.91 | 108.9%
2002-05-02 | 13:46 | 12.69 | 7.76 | 40 | 9.93 | 109.1%
2002-05-02 | 14:01 | 12.67 | 7.76 | 40 | 9.92 | 109%
2002-05-02 | 14:16 | 12.69 | 7.76 | 40 | 9.92 | 109%
2002-05-02 | 14:31 | 12.69 | 7.76 | 40 | 9.92 | 109%
2002-05-02 | 14:46 | 12.67 | 7.76 | 40 | 9.9 | 108.8%
2002-05-02 | 15:01 | 12.64 | 7.75 | 40 | 9.88 | 108.6%
2002-05-02 | 15:16 | 12.62 | 7.75 | 40 | 9.89 | 108.7%
2002-05-02 | 15:31 | 12.56 | 7.75 | 40 | 9.9 | 108.8%
2002-05-02 | 15:46 | 12.53 | 7.75 | 40 | 9.86 | 108.4%
2002-05-02 | 16:01 | 12.48 | 7.75 | 40 | 9.79 | 106.4%
2002-05-02 | 16:16 | 12.45 | 7.75 | 40 | 9.64 | 104.8%
2002-05-02 | 16:31 | 12.41 | 7.74 | 40 | 9.58 | 104.1%
2002-05-02 | 16:46 | 12.37 | 7.74 | 40 | 9.49 | 103.2%
2002-05-02 | 17:01 | 12.36 | 7.74 | 40 | 9.42 | 102.4%
2002-05-02 | 17:16 | 12.35 | 7.74 | 40 | 9.38 | 102%
2002-05-02 | 17:31 | 12.34 | 7.74 | 40 | 9.35 | 101.6%
2002-05-02 | 17:46 | 12.34 | 7.73 | 40 | 9.32 | 101.3%
2002-05-02 | 18:01 | 12.34 | 7.74 | 40 | 9.29 | 101%
2002-05-02 | 18:16 | 12.34 | 7.73 | 40 | 9.26 | 100.7%
2002-05-02 | 18:31 | 12.36 | 7.73 | 40 | 9.23 | 100.3%
2002-05-02 | 18:46 | 12.36 | 7.73 | 39 | 9.21 | 100.1%
2002-05-02 | 19:01 | 12.38 | 7.74 | 40 | 9.18 | 99.78%
2002-05-02 | 19:16 | 12.39 | 7.74 | 39 | 9.15 | 99.46%
2002-05-02 | 19:31 | 12.38 | 7.73 | 39 | 9.13 | 99.24%
2002-05-02 | 19:46 | 12.37 | 7.73 | 39 | 9.1 | 98.91%
2002-05-02 | 20:01 | 12.37 | 7.73 | 39 | 9.08 | 98.7%
2002-05-02 | 20:16 | 12.36 | 7.73 | 39 | 9.05 | 98.37%
2002-05-02 | 20:31 | 12.35 | 7.73 | 39 | 9.03 | 98.15%
2002-05-02 | 20:46 | 12.34 | 7.72 | 39 | 9.02 | 98.04%
2002-05-02 | 21:01 | 12.33 | 7.73 | 39 | 9.01 | 97.93%
2002-05-02 | 21:16 | 12.33 | 7.72 | 39 | 9 | 97.83%
2002-05-02 | 21:31 | 12.32 | 7.72 | 39 | 8.99 | 97.72%
2002-05-02 | 21:46 | 12.31 | 7.72 | 39 | 8.99 | 97.72%
2002-05-02 | 22:01 | 12.3 | 7.72 | 37 | 8.99 | 97.72%
2002-05-02 | 22:16 | 12.26 | 7.72 | 39 | 8.99 | 97.72%
2002-05-02 | 22:31 | 12.25 | 7.72 | 39 | 8.99 | 97.72%
2002-05-02 | 22:46 | 12.22 | 7.72 | 39 | 8.99 | 97.72%
2002-05-02 | 23:01 | 12.21 | 7.72 | 39 | 9 | 97.83%
2002-05-02 | 23:16 | 12.19 | 7.72 | 39 | 8.99 | 97.72%
2002-05-02 | 23:31 | 12.19 | 7.72 | 39 | 8.99 | 97.72%
2002-05-02 | 23:46 | 12.18 | 7.72 | 39 | 8.99 | 97.72%
2002-05-03 | 0:01 | 12.17 | 7.72 | 39 | 8.99 | 97.72%
2002-05-03 | 0:16 | 12.15 | 7.72 | 39 | 9 | 97.83%
2002-05-03 | 0:31 | 12.11 | 7.72 | 39 | 9 | 97.83%
2002-05-03 | 0:46 | 12.07 | 7.72 | 39 | 9.01 | 97.93%
2002-05-03 | 1:01 | 12.02 | 7.71 | 39 | 9.02 | 98.04%
2002-05-03 | 1:16 | 12.01 | 7.72 | 39 | 9.03 | 98.15%
2002-05-03 | 1:31 | 11.99 | 7.71 | 39 | 9.03 | 96.58%
2002-05-03 | 1:46 | 11.98 | 7.72 | 39 | 9.03 | 96.58%
2002-05-03 | 2:01 | 11.95 | 7.71 | 39 | 9.03 | 96.58%
2002-05-03 | 2:16 | 11.94 | 7.71 | 39 | 9.03 | 96.58%
2002-05-03 | 2:31 | 11.91 | 7.71 | 39 | 9.03 | 96.58%
2002-05-03 | 2:46 | 11.87 | 7.71 | 39 | 9.04 | 96.68%
2002-05-03 | 3:01 | 11.82 | 7.71 | 39 | 9.05 | 96.79%
2002-05-03 | 3:16 | 11.8 | 7.71 | 39 | 9.05 | 96.79%
2002-05-03 | 3:31 | 11.75 | 7.71 | 39 | 9.06 | 96.9%
2002-05-03 | 3:46 | 11.73 | 7.7 | 39 | 9.06 | 96.9%
2002-05-03 | 4:01 | 11.69 | 7.71 | 39 | 9.07 | 97.01%
2002-05-03 | 4:16 | 11.69 | 7.7 | 39 | 9.08 | 97.11%
2002-05-03 | 4:31 | 11.68 | 7.71 | 39 | 9.08 | 97.11%
2002-05-03 | 4:46 | 11.67 | 7.7 | 39 | 9.08 | 97.11%
2002-05-03 | 5:01 | 11.64 | 7.7 | 39 | 9.08 | 97.11%
2002-05-03 | 5:16 | 11.6 | 7.7 | 39 | 9.1 | 97.33%
2002-05-03 | 5:31 | 11.56 | 7.7 | 39 | 9.1 | 97.33%
2002-05-03 | 5:46 | 11.51 | 7.7 | 39 | 9.11 | 97.43%
2002-05-03 | 6:01 | 11.48 | 7.7 | 39 | 9.12 | 96.51%
2002-05-03 | 6:16 | 11.45 | 7.7 | 39 | 9.14 | 96.72%
2002-05-03 | 6:31 | 11.44 | 7.7 | 39 | 9.16 | 96.93%
2002-05-03 | 6:46 | 11.41 | 7.7 | 39 | 9.19 | 97.25%
2002-05-03 | 7:01 | 11.39 | 7.7 | 39 | 9.23 | 97.67%
2002-05-03 | 7:16 | 11.36 | 7.7 | 39 | 9.27 | 98.1%
2002-05-03 | 7:31 | 11.34 | 7.7 | 39 | 9.31 | 98.52%
2002-05-03 | 7:46 | 11.32 | 7.7 | 39 | 9.34 | 98.84%
2002-05-03 | 8:01 | 11.28 | 7.7 | 39 | 9.38 | 99.26%
2002-05-03 | 8:16 | 11.28 | 7.7 | 39 | 9.41 | 99.58%
2002-05-03 | 8:31 | 11.29 | 7.71 | 39 | 9.44 | 99.89%
2002-05-03 | 8:46 | 11.3 | 7.71 | 39 | 9.46 | 100.1%
2002-05-03 | 9:01 | 11.29 | 7.71 | 39 | 9.48 | 100.3%
2002-05-03 | 9:16 | 11.28 | 7.7 | 39 | 9.5 | 100.5%
2002-05-03 | 9:31 | 11.29 | 7.71 | 39 | 9.51 | 100.6%