## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2002-04-29 17:05 | 11.9 | 7.2 | 9.2 | 97 | 55 | 0
Sonde | 2002-04-29 17:31 | 11.65 | 7.52 | 9.5 | 101 | 48
Grade |  | A | B | A |  | B
Audit 2 | 2002-05-01 15:25 | 12 | 7.7 | 9.2 | 98 | 55 | 0
Sonde | 2002-05-01 15:31 | 11.96 | 7.56 | 9.27 | 98 | 48
Grade |  | A | A | A |  | B
Audit 3 | 2002-05-03 11:20 | 11.6 | 8.1 | 8.8 | 93 | 56 | 0
Sonde | 2002-05-03 11:16 | 11.82 | 7.52 | 8.7 | 92 | 49
Grade |  | A | Fail | A |  | B
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note Audits 1, 2,& 3 are macro(ed) into this audit sheet and represent the beginning, middle and end of the deployment. Audits 4, 5 and 6 were in the middle of the deployment and are not macro(ed) into this sheet. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-29 17:05 | 11.9 | 7.2 | 9.2 | 97 | 55
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-01 15:25 | 12 | 7.7 | 9.2 | 98 | 55
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-03 11:20 | 11.6 | 8.1 | 8.8 | 93 | 56
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-30 10:07 | 10.3 | 7.9 | 8.7 | 90 | 56
Sonde Values | 2002-04-30 10:16 | 10.3 | 7.5 | 8.9 | 91 | 49
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 08:35 | 11.5 | 7.6 | 8.6 | 91 | 54
Sonde Values | 2002-05-02 08:31 | 11.4 | 7.5 | 8.5 | 90 | 49
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 6
Organization | ODEQ |  | Station ID |  | DR 191.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 16:40 | 12.8 | 7.6 | 9.2 | 99 | 54
Sonde Values | 2002-05-02 16:46 | 12.8 | 7.6 | 9.2 | 100 | 49
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | DO
|  | °C |  | uS/cm | mg/L | % Sat
2002-04-29 | 17:01 | 11.71 | 7.54 | 48 | 9.69 | 102.5%
2002-04-29 | 17:16 | 11.67 | 7.52 | 48 | 9.57 | 101.3%
2002-04-29 | 17:31 | 11.65 | 7.52 | 48 | 9.5 | 100.5%
2002-04-29 | 17:46 | 11.65 | 7.53 | 48 | 9.46 | 100.1%
2002-04-29 | 18:01 | 11.64 | 7.53 | 48 | 9.42 | 99.68%
2002-04-29 | 18:16 | 11.62 | 7.53 | 48 | 9.39 | 99.37%
2002-04-29 | 18:31 | 11.58 | 7.54 | 48 | 9.38 | 99.26%
2002-04-29 | 18:46 | 11.55 | 7.54 | 48 | 9.36 | 99.05%
2002-04-29 | 19:01 | 11.52 | 7.54 | 48 | 9.34 | 98.84%
2002-04-29 | 19:16 | 11.49 | 7.54 | 48 | 9.33 | 98.21%
2002-04-29 | 19:31 | 11.45 | 7.53 | 48 | 9.31 | 98%
2002-04-29 | 19:46 | 11.43 | 7.54 | 48 | 9.29 | 97.79%
2002-04-29 | 20:01 | 11.39 | 7.53 | 48 | 9.28 | 97.68%
2002-04-29 | 20:16 | 11.37 | 7.53 | 48 | 9.26 | 97.47%
2002-04-29 | 20:31 | 11.37 | 7.53 | 48 | 9.23 | 97.16%
2002-04-29 | 20:46 | 11.35 | 7.53 | 48 | 9.22 | 97.05%
2002-04-29 | 21:01 | 11.32 | 7.52 | 48 | 9.2 | 96.84%
2002-04-29 | 21:16 | 11.28 | 7.53 | 48 | 9.19 | 96.74%
2002-04-29 | 21:31 | 11.24 | 7.53 | 48 | 9.17 | 96.53%
2002-04-29 | 21:46 | 11.2 | 7.51 | 48 | 9.16 | 96.42%
2002-04-29 | 22:01 | 11.16 | 7.51 | 48 | 9.15 | 96.32%
2002-04-29 | 22:16 | 11.12 | 7.52 | 48 | 9.13 | 96.11%
2002-04-29 | 22:31 | 11.08 | 7.51 | 48 | 9.12 | 96%
2002-04-29 | 22:46 | 11.05 | 7.51 | 48 | 9.11 | 95.89%
2002-04-29 | 23:01 | 11.02 | 7.52 | 48 | 9.1 | 95.79%
2002-04-29 | 23:16 | 10.99 | 7.52 | 48 | 9.09 | 94.69%
2002-04-29 | 23:31 | 10.97 | 7.51 | 48 | 9.08 | 94.58%
2002-04-29 | 23:46 | 10.93 | 7.51 | 48 | 9.07 | 94.48%
2002-04-30 | 0:01 | 10.9 | 7.51 | 48 | 9.06 | 94.38%
2002-04-30 | 0:16 | 10.86 | 7.51 | 48 | 9.04 | 94.17%
2002-04-30 | 0:31 | 10.82 | 7.51 | 48 | 9.03 | 94.06%
2002-04-30 | 0:46 | 10.79 | 7.51 | 48 | 9.02 | 93.96%
2002-04-30 | 1:01 | 10.76 | 7.5 | 48 | 9 | 93.75%
2002-04-30 | 1:16 | 10.71 | 7.51 | 48 | 9 | 93.75%
2002-04-30 | 1:31 | 10.68 | 7.5 | 48 | 8.98 | 93.54%
2002-04-30 | 1:46 | 10.65 | 7.49 | 48 | 8.96 | 93.33%
2002-04-30 | 2:01 | 10.61 | 7.5 | 48 | 8.95 | 93.23%
2002-04-30 | 2:16 | 10.57 | 7.51 | 48 | 8.95 | 93.23%
2002-04-30 | 2:31 | 10.54 | 7.51 | 48 | 8.94 | 93.12%
2002-04-30 | 2:46 | 10.51 | 7.5 | 48 | 8.91 | 92.81%
2002-04-30 | 3:01 | 10.47 | 7.49 | 48 | 8.91 | 91.86%
2002-04-30 | 3:16 | 10.44 | 7.5 | 48 | 8.9 | 91.75%
2002-04-30 | 3:31 | 10.4 | 7.49 | 48 | 8.88 | 91.55%
2002-04-30 | 3:46 | 10.35 | 7.49 | 48 | 8.87 | 91.44%
2002-04-30 | 4:01 | 10.32 | 7.48 | 48 | 8.86 | 91.34%
2002-04-30 | 4:16 | 10.28 | 7.5 | 48 | 8.85 | 91.24%
2002-04-30 | 4:31 | 10.25 | 7.48 | 48 | 8.83 | 91.03%
2002-04-30 | 4:46 | 10.2 | 7.49 | 48 | 8.82 | 90.93%
2002-04-30 | 5:01 | 10.18 | 7.49 | 48 | 8.8 | 90.72%
2002-04-30 | 5:16 | 10.14 | 7.48 | 48 | 8.8 | 90.72%
2002-04-30 | 5:31 | 10.12 | 7.48 | 48 | 8.79 | 90.62%
2002-04-30 | 5:46 | 10.09 | 7.47 | 48 | 8.78 | 90.52%
2002-04-30 | 6:01 | 10.07 | 7.47 | 48 | 8.76 | 90.31%
2002-04-30 | 6:16 | 10.04 | 7.48 | 48 | 8.75 | 90.21%
2002-04-30 | 6:31 | 10.02 | 7.47 | 48 | 8.75 | 90.21%
2002-04-30 | 6:46 | 10 | 7.47 | 48 | 8.74 | 90.1%
2002-04-30 | 7:01 | 9.99 | 7.47 | 48 | 8.74 | 88.73%
2002-04-30 | 7:16 | 9.97 | 7.46 | 48 | 8.74 | 88.73%
2002-04-30 | 7:31 | 9.97 | 7.47 | 48 | 8.74 | 88.73%
2002-04-30 | 7:46 | 9.97 | 7.47 | 48 | 8.74 | 88.73%
2002-04-30 | 8:01 | 9.98 | 7.46 | 48 | 8.75 | 88.83%
2002-04-30 | 8:16 | 9.99 | 7.47 | 48 | 8.76 | 88.93%
2002-04-30 | 8:31 | 10 | 7.48 | 48 | 8.76 | 90.31%
2002-04-30 | 8:46 | 10.02 | 7.45 | 48 | 8.78 | 90.52%
2002-04-30 | 9:01 | 10.05 | 7.48 | 48 | 8.79 | 90.62%
2002-04-30 | 9:16 | 10.09 | 7.48 | 48 | 8.8 | 90.72%
2002-04-30 | 9:31 | 10.13 | 7.48 | 48 | 8.81 | 90.82%
2002-04-30 | 9:46 | 10.19 | 7.49 | 48 | 8.83 | 91.03%
2002-04-30 | 10:01 | 10.24 | 7.49 | 48 | 8.85 | 91.24%
2002-04-30 | 10:16 | 10.29 | 7.49 | 49 | 8.87 | 91.44%
2002-04-30 | 10:31 | 10.36 | 7.48 | 48 | 8.89 | 91.65%
2002-04-30 | 10:46 | 10.42 | 7.5 | 48 | 8.91 | 91.86%
2002-04-30 | 11:01 | 10.48 | 7.5 | 49 | 8.93 | 92.06%
2002-04-30 | 11:16 | 10.56 | 7.51 | 49 | 8.95 | 93.23%
2002-04-30 | 11:31 | 10.62 | 7.49 | 48 | 8.97 | 93.44%
2002-04-30 | 11:46 | 10.69 | 7.5 | 48 | 8.99 | 93.65%
2002-04-30 | 12:01 | 10.8 | 7.5 | 48 | 9.01 | 93.85%
2002-04-30 | 12:16 | 10.86 | 7.5 | 48 | 9.03 | 94.06%
2002-04-30 | 12:31 | 10.93 | 7.5 | 48 | 9.06 | 94.38%
2002-04-30 | 12:46 | 11.03 | 7.49 | 49 | 9.07 | 95.47%
2002-04-30 | 13:01 | 11.08 | 7.51 | 48 | 9.08 | 95.58%
2002-04-30 | 13:16 | 11.1 | 7.51 | 48 | 9.11 | 95.89%
2002-04-30 | 13:31 | 11.16 | 7.52 | 49 | 9.12 | 96%
2002-04-30 | 13:46 | 11.21 | 7.53 | 48 | 9.13 | 96.11%
2002-04-30 | 14:01 | 11.2 | 7.52 | 49 | 9.15 | 96.32%
2002-04-30 | 14:16 | 11.29 | 7.52 | 48 | 9.15 | 96.32%
2002-04-30 | 14:31 | 11.33 | 7.52 | 49 | 9.17 | 96.53%
2002-04-30 | 14:46 | 11.37 | 7.53 | 49 | 9.2 | 96.84%
2002-04-30 | 15:01 | 11.47 | 7.52 | 48 | 9.2 | 96.84%
2002-04-30 | 15:16 | 11.5 | 7.53 | 49 | 9.23 | 97.67%
2002-04-30 | 15:31 | 11.55 | 7.54 | 49 | 9.23 | 97.67%
2002-04-30 | 15:46 | 11.61 | 7.53 | 48 | 9.23 | 97.67%
2002-04-30 | 16:01 | 11.6 | 7.53 | 48 | 9.24 | 97.78%
2002-04-30 | 16:16 | 11.62 | 7.54 | 49 | 9.25 | 97.88%
2002-04-30 | 16:31 | 11.64 | 7.55 | 49 | 9.26 | 97.99%
2002-04-30 | 16:46 | 11.64 | 7.54 | 49 | 9.26 | 97.99%
2002-04-30 | 17:01 | 11.63 | 7.55 | 49 | 9.27 | 98.1%
2002-04-30 | 17:16 | 11.63 | 7.55 | 49 | 9.28 | 98.2%
2002-04-30 | 17:31 | 11.64 | 7.55 | 49 | 9.27 | 98.1%
2002-04-30 | 17:46 | 11.63 | 7.55 | 49 | 9.28 | 98.2%
2002-04-30 | 18:01 | 11.67 | 7.56 | 48 | 9.27 | 98.1%
2002-04-30 | 18:16 | 11.65 | 7.56 | 49 | 9.28 | 98.2%
2002-04-30 | 18:31 | 11.62 | 7.54 | 49 | 9.27 | 98.1%
2002-04-30 | 18:46 | 11.6 | 7.55 | 49 | 9.27 | 98.1%
2002-04-30 | 19:01 | 11.58 | 7.55 | 49 | 9.26 | 97.99%
2002-04-30 | 19:16 | 11.56 | 7.55 | 48 | 9.25 | 97.88%
2002-04-30 | 19:31 | 11.54 | 7.55 | 48 | 9.24 | 97.78%
2002-04-30 | 19:46 | 11.49 | 7.55 | 49 | 9.24 | 97.26%
2002-04-30 | 20:01 | 11.47 | 7.55 | 48 | 9.22 | 97.05%
2002-04-30 | 20:16 | 11.44 | 7.55 | 48 | 9.21 | 96.95%
2002-04-30 | 20:31 | 11.42 | 7.54 | 49 | 9.21 | 96.95%
2002-04-30 | 20:46 | 11.4 | 7.54 | 49 | 9.19 | 96.74%
2002-04-30 | 21:01 | 11.38 | 7.54 | 48 | 9.17 | 96.53%
2002-04-30 | 21:16 | 11.34 | 7.55 | 49 | 9.17 | 96.53%
2002-04-30 | 21:31 | 11.31 | 7.54 | 49 | 9.16 | 96.42%
2002-04-30 | 21:46 | 11.27 | 7.53 | 49 | 9.15 | 96.32%
2002-04-30 | 22:01 | 11.25 | 7.53 | 49 | 9.13 | 96.11%
2002-04-30 | 22:16 | 11.22 | 7.53 | 49 | 9.12 | 96%
2002-04-30 | 22:31 | 11.21 | 7.54 | 49 | 9.1 | 95.79%
2002-04-30 | 22:46 | 11.19 | 7.53 | 48 | 9.09 | 95.68%
2002-04-30 | 23:01 | 11.17 | 7.53 | 49 | 9.09 | 95.68%
2002-04-30 | 23:16 | 11.15 | 7.54 | 49 | 9.07 | 95.47%
2002-04-30 | 23:31 | 11.12 | 7.53 | 48 | 9.06 | 95.37%
2002-04-30 | 23:46 | 11.1 | 7.52 | 48 | 9.04 | 95.16%
2002-05-01 | 0:01 | 11.06 | 7.52 | 48 | 9.03 | 95.05%
2002-05-01 | 0:16 | 11.03 | 7.52 | 48 | 9.01 | 94.84%
2002-05-01 | 0:31 | 10.99 | 7.53 | 48 | 9.01 | 93.85%
2002-05-01 | 0:46 | 10.96 | 7.52 | 48 | 8.99 | 93.65%
2002-05-01 | 1:01 | 10.92 | 7.52 | 48 | 8.97 | 93.44%
2002-05-01 | 1:16 | 10.89 | 7.52 | 48 | 8.96 | 93.33%
2002-05-01 | 1:31 | 10.85 | 7.52 | 48 | 8.95 | 93.23%
2002-05-01 | 1:46 | 10.82 | 7.52 | 48 | 8.94 | 93.12%
2002-05-01 | 2:01 | 10.79 | 7.51 | 48 | 8.93 | 93.02%
2002-05-01 | 2:16 | 10.75 | 7.5 | 48 | 8.91 | 92.81%
2002-05-01 | 2:31 | 10.72 | 7.51 | 48 | 8.89 | 92.6%
2002-05-01 | 2:46 | 10.68 | 7.5 | 48 | 8.9 | 92.71%
2002-05-01 | 3:01 | 10.64 | 7.5 | 48 | 8.89 | 92.6%
2002-05-01 | 3:16 | 10.61 | 7.5 | 48 | 8.87 | 92.4%
2002-05-01 | 3:31 | 10.57 | 7.5 | 48 | 8.86 | 92.29%
2002-05-01 | 3:46 | 10.53 | 7.49 | 48 | 8.85 | 92.19%
2002-05-01 | 4:01 | 10.5 | 7.49 | 48 | 8.83 | 91.98%
2002-05-01 | 4:16 | 10.47 | 7.49 | 48 | 8.83 | 91.03%
2002-05-01 | 4:31 | 10.44 | 7.49 | 48 | 8.82 | 90.93%
2002-05-01 | 4:46 | 10.42 | 7.48 | 48 | 8.8 | 90.72%
2002-05-01 | 5:01 | 10.4 | 7.49 | 48 | 8.79 | 90.62%
2002-05-01 | 5:16 | 10.38 | 7.48 | 48 | 8.78 | 90.52%
2002-05-01 | 5:31 | 10.37 | 7.47 | 48 | 8.77 | 90.41%
2002-05-01 | 5:46 | 10.35 | 7.48 | 48 | 8.76 | 90.31%
2002-05-01 | 6:01 | 10.33 | 7.47 | 48 | 8.75 | 90.21%
2002-05-01 | 6:16 | 10.31 | 7.47 | 48 | 8.74 | 90.1%
2002-05-01 | 6:31 | 10.3 | 7.48 | 48 | 8.74 | 90.1%
2002-05-01 | 6:46 | 10.27 | 7.48 | 48 | 8.74 | 90.1%
2002-05-01 | 7:01 | 10.25 | 7.47 | 48 | 8.74 | 90.1%
2002-05-01 | 7:16 | 10.23 | 7.46 | 48 | 8.74 | 90.1%
2002-05-01 | 7:31 | 10.21 | 7.47 | 48 | 8.75 | 90.21%
2002-05-01 | 7:46 | 10.19 | 7.47 | 48 | 8.75 | 90.21%
2002-05-01 | 8:01 | 10.18 | 7.47 | 48 | 8.76 | 90.31%
2002-05-01 | 8:16 | 10.18 | 7.47 | 48 | 8.77 | 90.41%
2002-05-01 | 8:31 | 10.17 | 7.47 | 48 | 8.79 | 90.62%
2002-05-01 | 8:46 | 10.18 | 7.47 | 48 | 8.8 | 90.72%
2002-05-01 | 9:01 | 10.19 | 7.48 | 48 | 8.82 | 90.93%
2002-05-01 | 9:16 | 10.21 | 7.48 | 49 | 8.84 | 91.13%
2002-05-01 | 9:31 | 10.24 | 7.49 | 49 | 8.86 | 91.34%
2002-05-01 | 9:46 | 10.29 | 7.48 | 48 | 8.87 | 91.44%
2002-05-01 | 10:01 | 10.32 | 7.48 | 48 | 8.89 | 91.65%
2002-05-01 | 10:16 | 10.37 | 7.49 | 48 | 8.92 | 91.96%
2002-05-01 | 10:31 | 10.41 | 7.49 | 48 | 8.94 | 92.16%
2002-05-01 | 10:46 | 10.48 | 7.5 | 48 | 8.96 | 92.37%
2002-05-01 | 11:01 | 10.53 | 7.49 | 48 | 8.98 | 93.54%
2002-05-01 | 11:16 | 10.63 | 7.5 | 48 | 8.99 | 93.65%
2002-05-01 | 11:31 | 10.69 | 7.49 | 48 | 9.01 | 93.85%
2002-05-01 | 11:46 | 10.77 | 7.49 | 48 | 9.04 | 94.17%
2002-05-01 | 12:01 | 10.82 | 7.51 | 48 | 9.07 | 94.48%
2002-05-01 | 12:16 | 10.9 | 7.5 | 48 | 9.08 | 94.58%
2002-05-01 | 12:31 | 10.98 | 7.5 | 48 | 9.1 | 94.79%
2002-05-01 | 12:46 | 11.06 | 7.51 | 48 | 9.12 | 96%
2002-05-01 | 13:01 | 11.14 | 7.51 | 48 | 9.13 | 96.11%
2002-05-01 | 13:16 | 11.22 | 7.51 | 48 | 9.15 | 96.32%
2002-05-01 | 13:31 | 11.3 | 7.52 | 48 | 9.17 | 96.53%
2002-05-01 | 13:46 | 11.41 | 7.51 | 48 | 9.17 | 96.53%
2002-05-01 | 14:01 | 11.48 | 7.53 | 48 | 9.19 | 96.74%
2002-05-01 | 14:16 | 11.58 | 7.52 | 48 | 9.2 | 97.35%
2002-05-01 | 14:31 | 11.65 | 7.52 | 48 | 9.22 | 97.57%
2002-05-01 | 14:46 | 11.72 | 7.53 | 49 | 9.23 | 97.67%
2002-05-01 | 15:01 | 11.82 | 7.53 | 48 | 9.24 | 97.78%
2002-05-01 | 15:16 | 11.9 | 7.53 | 48 | 9.24 | 97.78%
2002-05-01 | 15:31 | 11.96 | 7.56 | 48 | 9.27 | 98.1%
2002-05-01 | 15:46 | 12.01 | 7.54 | 49 | 9.28 | 98.72%
2002-05-01 | 16:01 | 12.09 | 7.55 | 48 | 9.29 | 98.83%
2002-05-01 | 16:16 | 12.15 | 7.53 | 48 | 9.29 | 98.83%
2002-05-01 | 16:31 | 12.2 | 7.54 | 48 | 9.3 | 98.94%
2002-05-01 | 16:46 | 12.29 | 7.55 | 48 | 9.29 | 98.83%
2002-05-01 | 17:01 | 12.32 | 7.55 | 48 | 9.3 | 98.94%
2002-05-01 | 17:16 | 12.37 | 7.55 | 48 | 9.3 | 98.94%
2002-05-01 | 17:31 | 12.42 | 7.55 | 48 | 9.31 | 99.04%
2002-05-01 | 17:46 | 12.45 | 7.56 | 48 | 9.29 | 98.83%
2002-05-01 | 18:01 | 12.49 | 7.56 | 48 | 9.3 | 98.94%
2002-05-01 | 18:16 | 12.53 | 7.55 | 48 | 9.29 | 100.4%
2002-05-01 | 18:31 | 12.54 | 7.55 | 48 | 9.28 | 100.3%
2002-05-01 | 18:46 | 12.57 | 7.56 | 48 | 9.28 | 100.3%
2002-05-01 | 19:01 | 12.58 | 7.55 | 48 | 9.26 | 100.1%
2002-05-01 | 19:16 | 12.59 | 7.55 | 48 | 9.24 | 99.89%
2002-05-01 | 19:31 | 12.59 | 7.55 | 48 | 9.23 | 99.78%
2002-05-01 | 19:46 | 12.6 | 7.56 | 48 | 9.22 | 99.68%
2002-05-01 | 20:01 | 12.61 | 7.55 | 48 | 9.2 | 99.46%
2002-05-01 | 20:16 | 12.61 | 7.55 | 48 | 9.19 | 99.35%
2002-05-01 | 20:31 | 12.61 | 7.55 | 48 | 9.18 | 99.24%
2002-05-01 | 20:46 | 12.61 | 7.55 | 48 | 9.16 | 99.03%
2002-05-01 | 21:01 | 12.6 | 7.55 | 48 | 9.14 | 98.81%
2002-05-01 | 21:16 | 12.59 | 7.54 | 48 | 9.12 | 98.59%
2002-05-01 | 21:31 | 12.58 | 7.54 | 48 | 9.11 | 98.49%
2002-05-01 | 21:46 | 12.57 | 7.54 | 48 | 9.09 | 98.27%
2002-05-01 | 22:01 | 12.55 | 7.53 | 48 | 9.07 | 98.05%
2002-05-01 | 22:16 | 12.53 | 7.53 | 48 | 9.05 | 97.84%
2002-05-01 | 22:31 | 12.51 | 7.53 | 48 | 9.04 | 97.73%
2002-05-01 | 22:46 | 12.49 | 7.53 | 48 | 9.01 | 95.85%
2002-05-01 | 23:01 | 12.47 | 7.53 | 48 | 8.99 | 95.64%
2002-05-01 | 23:16 | 12.46 | 7.52 | 48 | 8.98 | 95.53%
2002-05-01 | 23:31 | 12.44 | 7.52 | 48 | 8.96 | 95.32%
2002-05-01 | 23:46 | 12.41 | 7.52 | 48 | 8.95 | 95.21%
2002-05-02 | 0:01 | 12.39 | 7.52 | 48 | 8.93 | 95%
2002-05-02 | 0:16 | 12.36 | 7.52 | 48 | 8.9 | 94.68%
2002-05-02 | 0:31 | 12.33 | 7.52 | 48 | 8.88 | 94.47%
2002-05-02 | 0:46 | 12.31 | 7.52 | 48 | 8.86 | 94.26%
2002-05-02 | 1:01 | 12.29 | 7.52 | 48 | 8.84 | 94.04%
2002-05-02 | 1:16 | 12.26 | 7.52 | 48 | 8.83 | 93.94%
2002-05-02 | 1:31 | 12.24 | 7.51 | 48 | 8.81 | 93.72%
2002-05-02 | 1:46 | 12.21 | 7.51 | 48 | 8.79 | 93.51%
2002-05-02 | 2:01 | 12.18 | 7.51 | 48 | 8.77 | 93.3%
2002-05-02 | 2:16 | 12.15 | 7.51 | 48 | 8.75 | 93.09%
2002-05-02 | 2:31 | 12.12 | 7.5 | 48 | 8.74 | 92.98%
2002-05-02 | 2:46 | 12.08 | 7.5 | 48 | 8.71 | 92.66%
2002-05-02 | 3:01 | 12.03 | 7.5 | 48 | 8.7 | 92.55%
2002-05-02 | 3:16 | 11.99 | 7.5 | 48 | 8.69 | 91.96%
2002-05-02 | 3:31 | 11.94 | 7.5 | 48 | 8.67 | 91.75%
2002-05-02 | 3:46 | 11.9 | 7.5 | 48 | 8.66 | 91.64%
2002-05-02 | 4:01 | 11.85 | 7.5 | 48 | 8.64 | 91.43%
2002-05-02 | 4:16 | 11.81 | 7.49 | 48 | 8.63 | 91.32%
2002-05-02 | 4:31 | 11.76 | 7.49 | 48 | 8.61 | 91.11%
2002-05-02 | 4:46 | 11.73 | 7.49 | 48 | 8.6 | 91.01%
2002-05-02 | 5:01 | 11.69 | 7.49 | 48 | 8.58 | 90.79%
2002-05-02 | 5:16 | 11.65 | 7.48 | 48 | 8.57 | 90.69%
2002-05-02 | 5:31 | 11.61 | 7.48 | 48 | 8.56 | 90.58%
2002-05-02 | 5:46 | 11.57 | 7.47 | 48 | 8.55 | 90.48%
2002-05-02 | 6:01 | 11.53 | 7.47 | 48 | 8.54 | 90.37%
2002-05-02 | 6:16 | 11.5 | 7.47 | 48 | 8.53 | 90.26%
2002-05-02 | 6:31 | 11.47 | 7.47 | 48 | 8.52 | 89.68%
2002-05-02 | 6:46 | 11.45 | 7.46 | 48 | 8.51 | 89.58%
2002-05-02 | 7:01 | 11.43 | 7.47 | 48 | 8.51 | 89.58%
2002-05-02 | 7:16 | 11.42 | 7.47 | 48 | 8.51 | 89.58%
2002-05-02 | 7:31 | 11.41 | 7.47 | 48 | 8.52 | 89.68%
2002-05-02 | 7:46 | 11.4 | 7.47 | 48 | 8.52 | 89.68%
2002-05-02 | 8:01 | 11.4 | 7.47 | 48 | 8.52 | 89.68%
2002-05-02 | 8:16 | 11.41 | 7.47 | 48 | 8.53 | 89.79%
2002-05-02 | 8:31 | 11.42 | 7.48 | 49 | 8.54 | 89.89%
2002-05-02 | 8:46 | 11.43 | 7.48 | 49 | 8.55 | 90%
2002-05-02 | 9:01 | 11.46 | 7.48 | 49 | 8.56 | 90.11%
2002-05-02 | 9:16 | 11.48 | 7.49 | 49 | 8.58 | 90.32%
2002-05-02 | 9:31 | 11.51 | 7.49 | 49 | 8.59 | 90.9%
2002-05-02 | 9:46 | 11.53 | 7.48 | 49 | 8.62 | 91.22%
2002-05-02 | 10:01 | 11.56 | 7.49 | 49 | 8.64 | 91.43%
2002-05-02 | 10:16 | 11.59 | 7.49 | 49 | 8.66 | 91.64%
2002-05-02 | 10:31 | 11.62 | 7.51 | 49 | 8.69 | 91.96%
2002-05-02 | 10:46 | 11.65 | 7.5 | 49 | 8.72 | 92.28%
2002-05-02 | 11:01 | 11.68 | 7.5 | 49 | 8.74 | 92.49%
2002-05-02 | 11:16 | 11.72 | 7.51 | 49 | 8.77 | 92.8%
2002-05-02 | 11:31 | 11.76 | 7.51 | 49 | 8.8 | 93.12%
2002-05-02 | 11:46 | 11.82 | 7.51 | 49 | 8.83 | 93.44%
2002-05-02 | 12:01 | 11.85 | 7.52 | 49 | 8.86 | 93.76%
2002-05-02 | 12:16 | 11.94 | 7.52 | 49 | 8.88 | 93.97%
2002-05-02 | 12:31 | 11.95 | 7.52 | 49 | 8.91 | 94.29%
2002-05-02 | 12:46 | 12 | 7.53 | 49 | 8.94 | 95.11%
2002-05-02 | 13:01 | 12.03 | 7.53 | 49 | 8.97 | 95.43%
2002-05-02 | 13:16 | 12.09 | 7.53 | 49 | 9 | 95.74%
2002-05-02 | 13:31 | 12.17 | 7.53 | 49 | 9.02 | 95.96%
2002-05-02 | 13:46 | 12.2 | 7.54 | 49 | 9.05 | 96.28%
2002-05-02 | 14:01 | 12.26 | 7.55 | 49 | 9.07 | 96.49%
2002-05-02 | 14:16 | 12.33 | 7.55 | 49 | 9.09 | 96.7%
2002-05-02 | 14:31 | 12.35 | 7.55 | 49 | 9.12 | 97.02%
2002-05-02 | 14:46 | 12.43 | 7.55 | 49 | 9.13 | 97.13%
2002-05-02 | 15:01 | 12.47 | 7.56 | 49 | 9.15 | 97.34%
2002-05-02 | 15:16 | 12.52 | 7.56 | 49 | 9.16 | 99.03%
2002-05-02 | 15:31 | 12.56 | 7.57 | 49 | 9.18 | 99.24%
2002-05-02 | 15:46 | 12.64 | 7.57 | 49 | 9.19 | 99.35%
2002-05-02 | 16:01 | 12.66 | 7.57 | 49 | 9.2 | 99.46%
2002-05-02 | 16:16 | 12.72 | 7.58 | 49 | 9.21 | 99.57%
2002-05-02 | 16:31 | 12.76 | 7.57 | 49 | 9.22 | 99.68%
2002-05-02 | 16:46 | 12.8 | 7.58 | 49 | 9.23 | 99.78%
2002-05-02 | 17:01 | 12.85 | 7.58 | 49 | 9.23 | 99.78%
2002-05-02 | 17:16 | 12.87 | 7.58 | 49 | 9.24 | 99.89%
2002-05-02 | 17:31 | 12.9 | 7.58 | 49 | 9.24 | 99.89%
2002-05-02 | 17:46 | 12.93 | 7.58 | 49 | 9.23 | 99.78%
2002-05-02 | 18:01 | 12.94 | 7.58 | 49 | 9.23 | 99.78%
2002-05-02 | 18:16 | 12.94 | 7.59 | 49 | 9.22 | 99.68%
2002-05-02 | 18:31 | 12.95 | 7.58 | 49 | 9.21 | 99.57%
2002-05-02 | 18:46 | 12.96 | 7.58 | 49 | 9.2 | 99.46%
2002-05-02 | 19:01 | 12.94 | 7.58 | 49 | 9.19 | 99.35%
2002-05-02 | 19:16 | 12.92 | 7.58 | 49 | 9.19 | 99.35%
2002-05-02 | 19:31 | 12.91 | 7.58 | 49 | 9.18 | 99.24%
2002-05-02 | 19:46 | 12.87 | 7.58 | 49 | 9.16 | 99.03%
2002-05-02 | 20:01 | 12.86 | 7.58 | 49 | 9.14 | 98.81%
2002-05-02 | 20:16 | 12.83 | 7.57 | 49 | 9.13 | 98.7%
2002-05-02 | 20:31 | 12.81 | 7.57 | 49 | 9.11 | 98.49%
2002-05-02 | 20:46 | 12.78 | 7.58 | 49 | 9.1 | 98.38%
2002-05-02 | 21:01 | 12.78 | 7.56 | 49 | 9.08 | 98.16%
2002-05-02 | 21:16 | 12.76 | 7.56 | 49 | 9.06 | 97.95%
2002-05-02 | 21:31 | 12.73 | 7.57 | 49 | 9.05 | 97.84%
2002-05-02 | 21:46 | 12.72 | 7.56 | 49 | 9.02 | 97.51%
2002-05-02 | 22:01 | 12.68 | 7.56 | 49 | 9.01 | 97.41%
2002-05-02 | 22:16 | 12.64 | 7.55 | 49 | 9 | 97.3%
2002-05-02 | 22:31 | 12.61 | 7.57 | 49 | 8.99 | 97.19%
2002-05-02 | 22:46 | 12.58 | 7.56 | 49 | 8.96 | 96.86%
2002-05-02 | 23:01 | 12.56 | 7.56 | 49 | 8.94 | 96.65%
2002-05-02 | 23:16 | 12.52 | 7.55 | 49 | 8.93 | 96.54%
2002-05-02 | 23:31 | 12.5 | 7.55 | 49 | 8.92 | 96.43%
2002-05-02 | 23:46 | 12.46 | 7.55 | 49 | 8.9 | 94.68%
2002-05-03 | 0:01 | 12.44 | 7.55 | 49 | 8.88 | 94.47%
2002-05-03 | 0:16 | 12.43 | 7.54 | 49 | 8.86 | 94.26%
2002-05-03 | 0:31 | 12.39 | 7.54 | 49 | 8.84 | 94.04%
2002-05-03 | 0:46 | 12.37 | 7.54 | 49 | 8.83 | 93.94%
2002-05-03 | 1:01 | 12.33 | 7.53 | 49 | 8.81 | 93.72%
2002-05-03 | 1:16 | 12.3 | 7.54 | 49 | 8.79 | 93.51%
2002-05-03 | 1:31 | 12.25 | 7.54 | 49 | 8.78 | 93.4%
2002-05-03 | 1:46 | 12.22 | 7.53 | 49 | 8.76 | 93.19%
2002-05-03 | 2:01 | 12.18 | 7.53 | 49 | 8.73 | 92.87%
2002-05-03 | 2:16 | 12.13 | 7.53 | 49 | 8.72 | 92.77%
2002-05-03 | 2:31 | 12.09 | 7.53 | 49 | 8.71 | 92.66%
2002-05-03 | 2:46 | 12.06 | 7.54 | 49 | 8.69 | 92.45%
2002-05-03 | 3:01 | 12.02 | 7.53 | 49 | 8.68 | 92.34%
2002-05-03 | 3:16 | 11.98 | 7.53 | 49 | 8.66 | 91.64%
2002-05-03 | 3:31 | 11.92 | 7.53 | 49 | 8.65 | 91.53%
2002-05-03 | 3:46 | 11.87 | 7.53 | 49 | 8.64 | 91.43%
2002-05-03 | 4:01 | 11.85 | 7.53 | 49 | 8.61 | 91.11%
2002-05-03 | 4:16 | 11.8 | 7.52 | 49 | 8.6 | 91.01%
2002-05-03 | 4:31 | 11.78 | 7.53 | 49 | 8.59 | 90.9%
2002-05-03 | 4:46 | 11.73 | 7.51 | 49 | 8.58 | 90.79%
2002-05-03 | 5:01 | 11.69 | 7.51 | 49 | 8.56 | 90.58%
2002-05-03 | 5:16 | 11.66 | 7.51 | 49 | 8.53 | 90.26%
2002-05-03 | 5:31 | 11.61 | 7.52 | 49 | 8.54 | 90.37%
2002-05-03 | 5:46 | 11.58 | 7.5 | 49 | 8.52 | 90.16%
2002-05-03 | 6:01 | 11.56 | 7.51 | 49 | 8.51 | 90.05%
2002-05-03 | 6:16 | 11.54 | 7.51 | 49 | 8.49 | 89.84%
2002-05-03 | 6:31 | 11.51 | 7.49 | 49 | 8.48 | 89.74%
2002-05-03 | 6:46 | 11.49 | 7.5 | 49 | 8.48 | 89.26%
2002-05-03 | 7:01 | 11.47 | 7.5 | 49 | 8.46 | 89.05%
2002-05-03 | 7:16 | 11.45 | 7.49 | 49 | 8.46 | 89.05%
2002-05-03 | 7:31 | 11.43 | 7.5 | 49 | 8.47 | 89.16%
2002-05-03 | 7:46 | 11.43 | 7.5 | 49 | 8.48 | 89.26%
2002-05-03 | 8:01 | 11.42 | 7.48 | 49 | 8.47 | 89.16%
2002-05-03 | 8:16 | 11.44 | 7.5 | 49 | 8.48 | 89.26%
2002-05-03 | 8:31 | 11.44 | 7.49 | 49 | 8.5 | 89.47%
2002-05-03 | 8:46 | 11.45 | 7.51 | 49 | 8.51 | 89.58%
2002-05-03 | 9:01 | 11.47 | 7.51 | 49 | 8.53 | 89.79%
2002-05-03 | 9:16 | 11.49 | 7.5 | 49 | 8.54 | 89.89%
2002-05-03 | 9:31 | 11.51 | 7.5 | 49 | 8.55 | 90.48%
2002-05-03 | 9:46 | 11.56 | 7.51 | 49 | 8.57 | 90.69%
2002-05-03 | 10:01 | 11.57 | 7.51 | 49 | 8.6 | 91.01%
2002-05-03 | 10:16 | 11.63 | 7.51 | 49 | 8.61 | 91.11%
2002-05-03 | 10:31 | 11.67 | 7.51 | 49 | 8.63 | 91.32%
2002-05-03 | 10:46 | 11.72 | 7.51 | 49 | 8.66 | 91.64%
2002-05-03 | 11:01 | 11.76 | 7.52 | 49 | 8.69 | 91.96%
2002-05-03 | 11:16 | 11.86 | 7.52 | 49 | 8.7 | 92.06%
2002-05-03 | 11:31 | 11.88 | 7.53 | 49 | 8.76 | 92.7%