## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2002-04-29 13:13 | 11.6 | 8.36 | 9.8 | 103 | 63 | 0
Sonde | 2002-04-29 14:01 | 11.94 | 8.02 | 9.52 | 100 | 58
Grade |  | A | B | A |  | A
Audit 2 | 2002-05-01 13:10 | 12.3 | 8.1 | 9.9 | 105 | 63 | 0
Sonde | 2002-05-01 15:46 | 13 | 8.1 | 10.3 | 109 | 59
Grade |  | A | A | B |  | A
Audit 3 | 2002-05-03 08:30 | 10.8 | 7.8 | 9.5 | 96 | 61 | 0
Sonde | 2002-05-03 08:16 | 10.77 | 7.88 | 7.49 | 76 | 58
Grade |  | A | A | 0 |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note Audits 1, 2, & 3 are macro(ed) into this audit sheet and represent the beginning, middle and end of the deployment. Audits 4 and 5 were in the middle of the deployment and are not macro(ed) into this sheet. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-29 13:13 | 11.6 | 8.36 | 9.8 | 103 | 63
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-01 13:10 | 12.3 | 8.1 | 9.9 | 105 | 63
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-03 08:30 | 10.8 | 7.8 | 9.5 | 96 | 61
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-30 08:11 | 10.2 | 7.97 | 9.8 | 98 | 63
Sonde Values | 2002-04-30 09:16 | 10.5 | 7.8 | 9.5 | 97 | 58
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | DR 164.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, u/s Riverhouse
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 11:00 | 11.5 | 7.9 | 9.7 | 102 | 62
Sonde Values | 2002-05-02 11:01 | 11.6 | 8 | 10.6 | 111 | 58
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | DO
|  | °C |  | uS/cm | mg/L | % Sat
2002-04-29 | 13:16 | 11.65 | 7.82 | 58 | 9.62 | 100.7%
2002-04-29 | 13:31 | 11.74 | 7.98 | 58 | 9.58 | 100.3%
2002-04-29 | 13:46 | 11.84 | 8.01 | 58 | 9.56 | 100.1%
2002-04-29 | 14:01 | 11.94 | 8.02 | 58 | 9.52 | 99.69%
2002-04-29 | 14:16 | 12.02 | 8.03 | 58 | 9.48 | 100.3%
2002-04-29 | 14:31 | 12.13 | 8.04 | 58 | 9.45 | 100%
2002-04-29 | 14:46 | 12.23 | 8.05 | 58 | 9.43 | 99.79%
2002-04-29 | 15:01 | 12.25 | 8.05 | 58 | 9.39 | 99.37%
2002-04-29 | 15:16 | 12.29 | 8.05 | 58 | 9.34 | 98.84%
2002-04-29 | 15:31 | 12.32 | 8.05 | 58 | 9.29 | 98.31%
2002-04-29 | 15:46 | 12.35 | 8.04 | 58 | 9.25 | 97.88%
2002-04-29 | 16:01 | 12.39 | 8.04 | 58 | 9.24 | 97.78%
2002-04-29 | 16:16 | 12.42 | 8.04 | 58 | 9.23 | 97.67%
2002-04-29 | 16:31 | 12.45 | 8.04 | 58 | 9.2 | 97.35%
2002-04-29 | 16:46 | 12.47 | 8.04 | 58 | 9.18 | 97.14%
2002-04-29 | 17:01 | 12.51 | 8.04 | 58 | 9.14 | 97.23%
2002-04-29 | 17:16 | 12.52 | 8.04 | 58 | 9.11 | 96.91%
2002-04-29 | 17:31 | 12.52 | 8.03 | 59 | 9.08 | 96.6%
2002-04-29 | 17:46 | 12.49 | 8.01 | 59 | 9 | 95.24%
2002-04-29 | 18:01 | 12.49 | 7.99 | 59 | 8.95 | 94.71%
2002-04-29 | 18:16 | 12.49 | 7.98 | 59 | 8.94 | 94.6%
2002-04-29 | 18:31 | 12.46 | 7.97 | 59 | 8.9 | 94.18%
2002-04-29 | 18:46 | 12.43 | 7.95 | 59 | 8.86 | 93.76%
2002-04-29 | 19:01 | 12.43 | 7.93 | 59 | 8.85 | 93.65%
2002-04-29 | 19:16 | 12.41 | 7.92 | 59 | 8.82 | 93.33%
2002-04-29 | 19:31 | 12.39 | 7.9 | 59 | 8.8 | 93.12%
2002-04-29 | 19:46 | 12.35 | 7.88 | 59 | 8.77 | 92.8%
2002-04-29 | 20:01 | 12.3 | 7.87 | 59 | 8.75 | 92.59%
2002-04-29 | 20:16 | 12.26 | 7.85 | 59 | 8.73 | 92.38%
2002-04-29 | 20:31 | 12.21 | 7.83 | 59 | 8.71 | 92.17%
2002-04-29 | 20:46 | 12.16 | 7.81 | 59 | 8.71 | 92.17%
2002-04-29 | 21:01 | 12.11 | 7.8 | 59 | 8.71 | 92.17%
2002-04-29 | 21:16 | 12.05 | 7.78 | 59 | 8.73 | 92.38%
2002-04-29 | 21:31 | 11.99 | 7.77 | 59 | 8.74 | 91.52%
2002-04-29 | 21:46 | 11.92 | 7.77 | 59 | 8.75 | 91.62%
2002-04-29 | 22:01 | 11.86 | 7.76 | 59 | 8.77 | 91.83%
2002-04-29 | 22:16 | 11.77 | 7.75 | 59 | 8.79 | 92.04%
2002-04-29 | 22:31 | 11.67 | 7.74 | 59 | 8.81 | 92.25%
2002-04-29 | 22:46 | 11.58 | 7.73 | 59 | 8.83 | 92.46%
2002-04-29 | 23:01 | 11.5 | 7.73 | 59 | 8.85 | 92.67%
2002-04-29 | 23:16 | 11.42 | 7.72 | 59 | 8.87 | 91.44%
2002-04-29 | 23:31 | 11.32 | 7.72 | 59 | 8.9 | 91.75%
2002-04-29 | 23:46 | 11.23 | 7.71 | 59 | 8.92 | 91.96%
2002-04-30 | 0:01 | 11.15 | 7.71 | 59 | 8.94 | 92.16%
2002-04-30 | 0:16 | 11.08 | 7.71 | 59 | 8.95 | 92.27%
2002-04-30 | 0:31 | 11.02 | 7.7 | 59 | 8.95 | 92.27%
2002-04-30 | 0:46 | 10.96 | 7.69 | 59 | 8.97 | 91.07%
2002-04-30 | 1:01 | 10.89 | 7.69 | 59 | 8.98 | 91.17%
2002-04-30 | 1:16 | 10.82 | 7.68 | 59 | 9.01 | 91.47%
2002-04-30 | 1:31 | 10.76 | 7.69 | 59 | 9.02 | 91.57%
2002-04-30 | 1:46 | 10.7 | 7.68 | 59 | 9.03 | 91.68%
2002-04-30 | 2:01 | 10.65 | 7.68 | 59 | 9.04 | 91.78%
2002-04-30 | 2:16 | 10.6 | 7.68 | 59 | 9.06 | 91.98%
2002-04-30 | 2:31 | 10.54 | 7.67 | 59 | 9.07 | 92.08%
2002-04-30 | 2:46 | 10.5 | 7.67 | 59 | 9.08 | 92.18%
2002-04-30 | 3:01 | 10.45 | 7.67 | 59 | 9.08 | 90.8%
2002-04-30 | 3:16 | 10.41 | 7.67 | 59 | 9.08 | 90.8%
2002-04-30 | 3:31 | 10.38 | 7.66 | 59 | 9.09 | 90.9%
2002-04-30 | 3:46 | 10.35 | 7.66 | 59 | 9.09 | 90.9%
2002-04-30 | 4:01 | 10.32 | 7.66 | 59 | 9.1 | 91%
2002-04-30 | 4:16 | 10.29 | 7.66 | 59 | 9.11 | 91.1%
2002-04-30 | 4:31 | 10.28 | 7.66 | 59 | 9.11 | 91.1%
2002-04-30 | 4:46 | 10.26 | 7.66 | 59 | 9.1 | 91%
2002-04-30 | 5:01 | 10.24 | 7.65 | 59 | 9.1 | 91%
2002-04-30 | 5:16 | 10.21 | 7.65 | 59 | 9.1 | 91%
2002-04-30 | 5:31 | 10.19 | 7.65 | 59 | 9.1 | 91%
2002-04-30 | 5:46 | 10.18 | 7.64 | 59 | 9.09 | 90.9%
2002-04-30 | 6:01 | 10.16 | 7.64 | 59 | 9.1 | 91%
2002-04-30 | 6:16 | 10.15 | 7.64 | 59 | 9.12 | 91.2%
2002-04-30 | 6:31 | 10.15 | 7.65 | 59 | 9.15 | 91.5%
2002-04-30 | 6:46 | 10.15 | 7.65 | 59 | 9.18 | 91.8%
2002-04-30 | 7:01 | 10.15 | 7.66 | 59 | 9.22 | 92.2%
2002-04-30 | 7:16 | 10.16 | 7.67 | 59 | 9.25 | 92.5%
2002-04-30 | 7:31 | 10.18 | 7.68 | 58 | 9.29 | 92.9%
2002-04-30 | 7:46 | 10.2 | 7.7 | 58 | 9.34 | 93.4%
2002-04-30 | 8:01 | 10.21 | 7.71 | 58 | 9.38 | 93.8%
2002-04-30 | 8:16 | 10.23 | 7.73 | 58 | 9.41 | 94.1%
2002-04-30 | 8:31 | 10.28 | 7.74 | 58 | 9.43 | 94.3%
2002-04-30 | 8:46 | 10.36 | 7.75 | 58 | 9.48 | 94.8%
2002-04-30 | 9:01 | 10.45 | 7.77 | 58 | 9.52 | 95.2%
2002-04-30 | 9:16 | 10.52 | 7.78 | 58 | 9.54 | 96.85%
2002-04-30 | 9:31 | 10.61 | 7.8 | 58 | 9.55 | 96.95%
2002-04-30 | 9:46 | 10.76 | 7.83 | 58 | 9.59 | 97.36%
2002-04-30 | 10:01 | 10.92 | 7.86 | 58 | 9.63 | 97.77%
2002-04-30 | 10:16 | 10.94 | 7.86 | 58 | 9.67 | 98.17%
2002-04-30 | 10:31 | 10.98 | 7.85 | 58 | 9.61 | 97.56%
2002-04-30 | 10:46 | 11.09 | 7.86 | 58 | 9.57 | 98.66%
2002-04-30 | 11:01 | 11.16 | 7.89 | 58 | 9.59 | 98.87%
2002-04-30 | 11:16 | 11.23 | 7.91 | 58 | 9.6 | 98.97%
2002-04-30 | 11:31 | 11.34 | 7.93 | 58 | 9.6 | 98.97%
2002-04-30 | 11:46 | 11.47 | 7.94 | 58 | 9.6 | 98.97%
2002-04-30 | 12:01 | 11.52 | 7.96 | 58 | 9.58 | 100.3%
2002-04-30 | 12:16 | 11.64 | 7.97 | 58 | 9.57 | 100.2%
2002-04-30 | 12:31 | 11.75 | 7.99 | 58 | 9.55 | 100%
2002-04-30 | 12:46 | 11.71 | 7.99 | 58 | 9.5 | 99.48%
2002-04-30 | 13:01 | 11.79 | 8.01 | 58 | 9.48 | 99.27%
2002-04-30 | 13:16 | 11.84 | 8.01 | 58 | 9.46 | 99.06%
2002-04-30 | 13:31 | 11.88 | 8.02 | 58 | 9.43 | 98.74%
2002-04-30 | 13:46 | 11.92 | 8.04 | 58 | 9.41 | 98.53%
2002-04-30 | 14:01 | 11.99 | 8.04 | 59 | 9.41 | 98.53%
2002-04-30 | 14:16 | 12.07 | 8.05 | 58 | 9.41 | 99.58%
2002-04-30 | 14:31 | 12.12 | 8.05 | 58 | 9.37 | 99.15%
2002-04-30 | 14:46 | 12.18 | 8.06 | 58 | 9.34 | 98.84%
2002-04-30 | 15:01 | 12.21 | 8.06 | 58 | 9.3 | 98.41%
2002-04-30 | 15:16 | 12.26 | 8.05 | 58 | 9.28 | 98.2%
2002-04-30 | 15:31 | 12.28 | 8.07 | 58 | 9.25 | 97.88%
2002-04-30 | 15:46 | 12.32 | 8.06 | 58 | 9.23 | 97.67%
2002-04-30 | 16:01 | 12.33 | 8.05 | 58 | 9.2 | 97.35%
2002-04-30 | 16:16 | 12.35 | 8.05 | 59 | 9.17 | 97.04%
2002-04-30 | 16:31 | 12.36 | 8.05 | 59 | 9.14 | 96.72%
2002-04-30 | 16:46 | 12.37 | 8.05 | 59 | 9.11 | 96.4%
2002-04-30 | 17:01 | 12.35 | 8.03 | 59 | 9.05 | 95.77%
2002-04-30 | 17:16 | 12.33 | 8.02 | 59 | 9.02 | 95.45%
2002-04-30 | 17:31 | 12.33 | 8 | 59 | 8.99 | 95.13%
2002-04-30 | 17:46 | 12.31 | 7.98 | 59 | 8.96 | 94.81%
2002-04-30 | 18:01 | 12.3 | 7.97 | 59 | 8.94 | 94.6%
2002-04-30 | 18:16 | 12.29 | 7.96 | 59 | 8.93 | 94.5%
2002-04-30 | 18:31 | 12.26 | 7.95 | 59 | 8.9 | 94.18%
2002-04-30 | 18:46 | 12.22 | 7.93 | 59 | 8.87 | 93.86%
2002-04-30 | 19:01 | 12.18 | 7.91 | 59 | 8.83 | 93.44%
2002-04-30 | 19:16 | 12.14 | 7.89 | 59 | 8.82 | 93.33%
2002-04-30 | 19:31 | 12.08 | 7.87 | 59 | 8.81 | 93.23%
2002-04-30 | 19:46 | 12.02 | 7.85 | 59 | 8.79 | 93.02%
2002-04-30 | 20:01 | 11.96 | 7.84 | 59 | 8.78 | 91.94%
2002-04-30 | 20:16 | 11.9 | 7.82 | 59 | 8.77 | 91.83%
2002-04-30 | 20:31 | 11.82 | 7.81 | 59 | 8.77 | 91.83%
2002-04-30 | 20:46 | 11.75 | 7.79 | 59 | 8.76 | 91.73%
2002-04-30 | 21:01 | 11.7 | 7.78 | 59 | 8.77 | 91.83%
2002-04-30 | 21:16 | 11.65 | 7.76 | 59 | 8.79 | 92.04%
2002-04-30 | 21:31 | 11.59 | 7.75 | 59 | 8.81 | 92.25%
2002-04-30 | 21:46 | 11.53 | 7.75 | 59 | 8.83 | 92.46%
2002-04-30 | 22:01 | 11.48 | 7.74 | 59 | 8.84 | 91.13%
2002-04-30 | 22:16 | 11.42 | 7.74 | 59 | 8.85 | 91.24%
2002-04-30 | 22:31 | 11.37 | 7.73 | 59 | 8.86 | 91.34%
2002-04-30 | 22:46 | 11.32 | 7.72 | 59 | 8.88 | 91.55%
2002-04-30 | 23:01 | 11.24 | 7.71 | 59 | 8.89 | 91.65%
2002-04-30 | 23:16 | 11.17 | 7.71 | 59 | 8.91 | 91.86%
2002-04-30 | 23:31 | 11.1 | 7.7 | 59 | 8.93 | 92.06%
2002-04-30 | 23:46 | 11.03 | 7.7 | 59 | 8.95 | 92.27%
2002-05-01 | 0:01 | 10.95 | 7.69 | 59 | 8.97 | 91.07%
2002-05-01 | 0:16 | 10.87 | 7.69 | 59 | 8.99 | 91.27%
2002-05-01 | 0:31 | 10.8 | 7.68 | 59 | 9 | 91.37%
2002-05-01 | 0:46 | 10.73 | 7.72 | 59 | 9.03 | 91.68%
2002-05-01 | 1:01 | 10.67 | 7.71 | 59 | 9.05 | 91.88%
2002-05-01 | 1:16 | 10.61 | 7.7 | 59 | 9.09 | 92.28%
2002-05-01 | 1:31 | 10.55 | 7.7 | 59 | 9.14 | 92.79%
2002-05-01 | 1:46 | 10.49 | 7.69 | 59 | 9.19 | 91.9%
2002-05-01 | 2:01 | 10.44 | 7.7 | 59 | 9.24 | 92.4%
2002-05-01 | 2:16 | 10.4 | 7.7 | 59 | 9.28 | 92.8%
2002-05-01 | 2:31 | 10.35 | 7.7 | 59 | 9.33 | 93.3%
2002-05-01 | 2:46 | 10.31 | 7.69 | 59 | 9.38 | 93.8%
2002-05-01 | 3:01 | 10.27 | 7.69 | 59 | 9.43 | 94.3%
2002-05-01 | 3:16 | 10.22 | 7.69 | 59 | 9.47 | 94.7%
2002-05-01 | 3:31 | 10.2 | 7.69 | 59 | 9.51 | 95.1%
2002-05-01 | 3:46 | 10.18 | 7.69 | 59 | 9.55 | 95.5%
2002-05-01 | 4:01 | 10.16 | 7.69 | 59 | 9.59 | 95.9%
2002-05-01 | 4:16 | 10.14 | 7.7 | 59 | 9.62 | 96.2%
2002-05-01 | 4:31 | 10.14 | 7.7 | 59 | 9.66 | 96.6%
2002-05-01 | 4:46 | 10.13 | 7.7 | 59 | 9.69 | 96.9%
2002-05-01 | 5:01 | 10.12 | 7.7 | 59 | 9.73 | 97.3%
2002-05-01 | 5:16 | 10.11 | 7.7 | 59 | 9.76 | 97.6%
2002-05-01 | 5:31 | 10.12 | 7.7 | 59 | 9.8 | 98%
2002-05-01 | 5:46 | 10.11 | 7.7 | 59 | 9.82 | 98.2%
2002-05-01 | 6:01 | 10.12 | 7.7 | 59 | 9.86 | 98.6%
2002-05-01 | 6:16 | 10.12 | 7.7 | 59 | 9.9 | 99%
2002-05-01 | 6:31 | 10.12 | 7.71 | 58 | 9.93 | 99.3%
2002-05-01 | 6:46 | 10.14 | 7.71 | 58 | 9.95 | 99.5%
2002-05-01 | 7:01 | 10.15 | 7.71 | 58 | 9.98 | 99.8%
2002-05-01 | 7:16 | 10.16 | 7.71 | 58 | 10.04 | 100.4%
2002-05-01 | 7:31 | 10.18 | 7.72 | 58 | 10.08 | 100.8%
2002-05-01 | 7:46 | 10.2 | 7.73 | 58 | 10.13 | 101.3%
2002-05-01 | 8:01 | 10.22 | 7.73 | 58 | 10.2 | 102%
2002-05-01 | 8:16 | 10.26 | 7.75 | 58 | 10.27 | 102.7%
2002-05-01 | 8:31 | 10.33 | 7.77 | 58 | 10.35 | 103.5%
2002-05-01 | 8:46 | 10.43 | 7.78 | 58 | 10.43 | 104.3%
2002-05-01 | 9:01 | 10.49 | 7.8 | 58 | 10.51 | 105.1%
2002-05-01 | 9:16 | 10.57 | 7.82 | 58 | 10.55 | 107.1%
2002-05-01 | 9:31 | 10.67 | 7.84 | 58 | 10.59 | 107.5%
2002-05-01 | 9:46 | 10.79 | 7.86 | 58 | 10.64 | 108%
2002-05-01 | 10:01 | 10.89 | 7.89 | 58 | 10.68 | 108.4%
2002-05-01 | 10:16 | 10.99 | 7.91 | 58 | 10.71 | 108.7%
2002-05-01 | 10:31 | 11.1 | 7.92 | 58 | 10.74 | 110.7%
2002-05-01 | 10:46 | 11.21 | 7.94 | 58 | 10.73 | 110.6%
2002-05-01 | 11:01 | 11.31 | 7.95 | 58 | 10.74 | 110.7%
2002-05-01 | 11:16 | 11.42 | 7.97 | 58 | 10.74 | 110.7%
2002-05-01 | 11:31 | 11.55 | 7.99 | 58 | 10.74 | 112.5%
2002-05-01 | 11:46 | 11.68 | 8.01 | 58 | 10.75 | 112.6%
2002-05-01 | 12:01 | 11.79 | 8.02 | 58 | 10.74 | 112.5%
2002-05-01 | 12:16 | 11.9 | 8.04 | 58 | 10.74 | 112.5%
2002-05-01 | 12:31 | 11.99 | 8.05 | 58 | 10.71 | 112.1%
2002-05-01 | 12:46 | 12.07 | 8.05 | 58 | 10.67 | 112.9%
2002-05-01 | 13:01 | 12.16 | 8.06 | 58 | 10.64 | 112.6%
2002-05-01 | 13:16 | 12.27 | 8.07 | 58 | 10.6 | 112.2%
2002-05-01 | 13:31 | 12.36 | 8.08 | 58 | 10.57 | 111.9%
2002-05-01 | 13:46 | 12.45 | 8.09 | 58 | 10.55 | 111.6%
2002-05-01 | 14:01 | 12.55 | 8.1 | 58 | 10.54 | 112.1%
2002-05-01 | 14:16 | 12.63 | 8.1 | 58 | 10.51 | 111.8%
2002-05-01 | 14:31 | 12.7 | 8.1 | 58 | 10.5 | 111.7%
2002-05-01 | 14:46 | 12.76 | 8.11 | 59 | 10.43 | 111%
2002-05-01 | 15:01 | 12.82 | 8.11 | 59 | 10.4 | 110.6%
2002-05-01 | 15:16 | 12.89 | 8.11 | 59 | 10.38 | 110.4%
2002-05-01 | 15:31 | 12.93 | 8.11 | 58 | 10.34 | 110%
2002-05-01 | 15:46 | 12.97 | 8.11 | 59 | 10.29 | 109.5%
2002-05-01 | 16:01 | 12.99 | 8.11 | 59 | 10.25 | 109%
2002-05-01 | 16:16 | 13.02 | 8.09 | 59 | 10.23 | 110%
2002-05-01 | 16:31 | 13.03 | 8.09 | 59 | 10.15 | 109.1%
2002-05-01 | 16:46 | 13.04 | 8.07 | 59 | 10.11 | 108.7%
2002-05-01 | 17:01 | 13.07 | 8.07 | 59 | 10.09 | 108.5%
2002-05-01 | 17:16 | 13.08 | 8.05 | 59 | 10.04 | 108%
2002-05-01 | 17:31 | 13.09 | 8.04 | 59 | 9.99 | 107.4%
2002-05-01 | 17:46 | 13.09 | 8.03 | 59 | 9.97 | 107.2%
2002-05-01 | 18:01 | 13.09 | 8.02 | 59 | 9.93 | 106.8%
2002-05-01 | 18:16 | 13.08 | 8.01 | 59 | 9.9 | 106.5%
2002-05-01 | 18:31 | 13.1 | 8 | 59 | 9.89 | 106.3%
2002-05-01 | 18:46 | 13.11 | 7.99 | 59 | 9.87 | 106.1%
2002-05-01 | 19:01 | 13.09 | 7.98 | 59 | 9.86 | 106%
2002-05-01 | 19:16 | 13.06 | 7.97 | 59 | 9.84 | 105.8%
2002-05-01 | 19:31 | 13.03 | 7.96 | 59 | 9.82 | 105.6%
2002-05-01 | 19:46 | 13 | 7.95 | 59 | 9.82 | 105.6%
2002-05-01 | 20:01 | 12.95 | 7.93 | 59 | 9.78 | 104%
2002-05-01 | 20:16 | 12.89 | 7.91 | 59 | 9.78 | 104%
2002-05-01 | 20:31 | 12.82 | 7.89 | 59 | 9.78 | 104%
2002-05-01 | 20:46 | 12.74 | 7.88 | 59 | 9.82 | 104.5%
2002-05-01 | 21:01 | 12.66 | 7.87 | 59 | 9.81 | 104.4%
2002-05-01 | 21:16 | 12.59 | 7.86 | 59 | 9.81 | 104.4%
2002-05-01 | 21:31 | 12.51 | 7.85 | 59 | 9.83 | 104.6%
2002-05-01 | 21:46 | 12.42 | 7.84 | 59 | 9.87 | 104.4%
2002-05-01 | 22:01 | 12.34 | 7.83 | 59 | 9.88 | 104.6%
2002-05-01 | 22:16 | 12.25 | 7.82 | 59 | 9.89 | 104.7%
2002-05-01 | 22:31 | 12.15 | 7.81 | 59 | 9.94 | 105.2%
2002-05-01 | 22:46 | 12.06 | 7.81 | 59 | 9.93 | 105.1%
2002-05-01 | 23:01 | 11.96 | 7.8 | 59 | 9.97 | 104.4%
2002-05-01 | 23:16 | 11.87 | 7.79 | 59 | 9.99 | 104.6%
2002-05-01 | 23:31 | 11.76 | 7.78 | 59 | 10.02 | 104.9%
2002-05-01 | 23:46 | 11.67 | 7.78 | 59 | 10.06 | 105.3%
2002-05-02 | 0:01 | 11.57 | 7.78 | 59 | 10.09 | 105.7%
2002-05-02 | 0:16 | 11.48 | 7.77 | 59 | 10.11 | 104.2%
2002-05-02 | 0:31 | 11.4 | 7.77 | 59 | 10.12 | 104.3%
2002-05-02 | 0:46 | 11.32 | 7.77 | 59 | 10.16 | 104.7%
2002-05-02 | 1:01 | 11.24 | 7.77 | 59 | 10.18 | 104.9%
2002-05-02 | 1:16 | 11.17 | 7.77 | 59 | 10.2 | 105.2%
2002-05-02 | 1:31 | 11.11 | 7.76 | 59 | 10.22 | 105.4%
2002-05-02 | 1:46 | 11.06 | 7.76 | 59 | 10.21 | 105.3%
2002-05-02 | 2:01 | 11.01 | 7.76 | 59 | 10.23 | 105.5%
2002-05-02 | 2:16 | 10.96 | 7.76 | 59 | 10.24 | 104%
2002-05-02 | 2:31 | 10.93 | 7.76 | 59 | 10.21 | 103.7%
2002-05-02 | 2:46 | 10.89 | 7.76 | 58 | 10.25 | 104.1%
2002-05-02 | 3:01 | 10.85 | 7.76 | 59 | 10.24 | 104%
2002-05-02 | 3:16 | 10.81 | 7.76 | 59 | 10.27 | 104.3%
2002-05-02 | 3:31 | 10.79 | 7.76 | 58 | 10.23 | 103.9%
2002-05-02 | 3:46 | 10.76 | 7.75 | 58 | 10.25 | 104.1%
2002-05-02 | 4:01 | 10.73 | 7.76 | 58 | 10.26 | 104.2%
2002-05-02 | 4:16 | 10.7 | 7.75 | 58 | 10.26 | 104.2%
2002-05-02 | 4:31 | 10.66 | 7.75 | 58 | 10.28 | 104.4%
2002-05-02 | 4:46 | 10.64 | 7.75 | 58 | 10.26 | 104.2%
2002-05-02 | 5:01 | 10.61 | 7.75 | 58 | 10.27 | 104.3%
2002-05-02 | 5:16 | 10.59 | 7.76 | 58 | 10.24 | 104%
2002-05-02 | 5:31 | 10.57 | 7.75 | 58 | 10.28 | 104.4%
2002-05-02 | 5:46 | 10.56 | 7.76 | 58 | 10.24 | 104%
2002-05-02 | 6:01 | 10.54 | 7.75 | 58 | 10.25 | 104.1%
2002-05-02 | 6:16 | 10.53 | 7.75 | 58 | 10.25 | 104.1%
2002-05-02 | 6:31 | 10.52 | 7.76 | 58 | 10.22 | 103.8%
2002-05-02 | 6:46 | 10.51 | 7.76 | 58 | 10.27 | 104.3%
2002-05-02 | 7:01 | 10.51 | 7.77 | 58 | 10.25 | 104.1%
2002-05-02 | 7:16 | 10.51 | 7.78 | 58 | 10.27 | 104.3%
2002-05-02 | 7:31 | 10.5 | 7.78 | 58 | 10.29 | 104.5%
2002-05-02 | 7:46 | 10.52 | 7.8 | 58 | 10.34 | 105%
2002-05-02 | 8:01 | 10.54 | 7.81 | 58 | 10.32 | 104.8%
2002-05-02 | 8:16 | 10.57 | 7.82 | 58 | 10.39 | 105.5%
2002-05-02 | 8:31 | 10.65 | 7.85 | 58 | 10.43 | 105.9%
2002-05-02 | 8:46 | 10.72 | 7.87 | 58 | 10.46 | 106.2%
2002-05-02 | 9:01 | 10.79 | 7.88 | 58 | 10.56 | 107.2%
2002-05-02 | 9:16 | 10.86 | 7.9 | 58 | 10.55 | 107.1%
2002-05-02 | 9:31 | 10.94 | 7.92 | 58 | 10.55 | 107.1%
2002-05-02 | 9:46 | 11.05 | 7.94 | 58 | 10.66 | 109.9%
2002-05-02 | 10:01 | 11.15 | 7.96 | 58 | 10.62 | 109.5%
2002-05-02 | 10:16 | 11.26 | 7.98 | 58 | 10.58 | 109.1%
2002-05-02 | 10:31 | 11.37 | 8 | 58 | 10.67 | 110%
2002-05-02 | 10:46 | 11.5 | 8.01 | 58 | 10.52 | 110.2%
2002-05-02 | 11:01 | 11.63 | 8.04 | 58 | 10.58 | 110.8%
2002-05-02 | 11:16 | 11.76 | 8.05 | 58 | 10.62 | 111.2%
2002-05-02 | 11:31 | 11.9 | 8.07 | 58 | 10.46 | 109.5%
2002-05-02 | 11:46 | 12.05 | 8.09 | 58 | 10.47 | 110.8%
2002-05-02 | 12:01 | 12.19 | 8.1 | 58 | 10.44 | 110.5%
2002-05-02 | 12:16 | 12.27 | 8.11 | 58 | 10.48 | 110.9%
2002-05-02 | 12:31 | 12.39 | 8.11 | 58 | 10.33 | 109.3%
2002-05-02 | 12:46 | 12.48 | 8.12 | 58 | 10.3 | 109%
2002-05-02 | 13:01 | 12.59 | 8.13 | 58 | 10.21 | 108.6%
2002-05-02 | 13:16 | 12.7 | 8.14 | 58 | 10.19 | 108.4%
2002-05-02 | 13:31 | 12.83 | 8.15 | 58 | 10.16 | 108.1%
2002-05-02 | 13:46 | 12.93 | 8.16 | 58 | 10.01 | 106.5%
2002-05-02 | 14:01 | 13.02 | 8.16 | 58 | 9.93 | 106.8%
2002-05-02 | 14:16 | 13.1 | 8.17 | 58 | 10.02 | 107.7%
2002-05-02 | 14:31 | 13.18 | 8.18 | 58 | 9.85 | 105.9%
2002-05-02 | 14:46 | 13.28 | 8.18 | 58 | 9.88 | 106.2%
2002-05-02 | 15:01 | 13.35 | 8.17 | 58 | 9.69 | 104.2%
2002-05-02 | 15:16 | 13.44 | 8.17 | 58 | 9.65 | 103.8%
2002-05-02 | 15:31 | 13.52 | 8.17 | 58 | 9.61 | 104.5%
2002-05-02 | 15:46 | 13.59 | 8.18 | 58 | 9.51 | 103.4%
2002-05-02 | 16:01 | 13.64 | 8.17 | 58 | 9.54 | 103.7%
2002-05-02 | 16:16 | 13.67 | 8.15 | 58 | 9.36 | 101.7%
2002-05-02 | 16:31 | 13.72 | 8.15 | 59 | 9.4 | 102.2%
2002-05-02 | 16:46 | 13.77 | 8.13 | 58 | 9.18 | 99.78%
2002-05-02 | 17:01 | 13.79 | 8.12 | 59 | 9.17 | 99.67%
2002-05-02 | 17:16 | 13.81 | 8.11 | 59 | 9.21 | 100.1%
2002-05-02 | 17:31 | 13.81 | 8.1 | 59 | 9.04 | 98.26%
2002-05-02 | 17:46 | 13.8 | 8.08 | 59 | 9.01 | 97.93%
2002-05-02 | 18:01 | 13.78 | 8.07 | 59 | 9.11 | 99.02%
2002-05-02 | 18:16 | 13.75 | 8.05 | 59 | 8.85 | 96.2%
2002-05-02 | 18:31 | 13.72 | 8.04 | 59 | 8.86 | 96.3%
2002-05-02 | 18:46 | 13.68 | 8.02 | 59 | 8.86 | 96.3%
2002-05-02 | 19:01 | 13.66 | 8.01 | 59 | 8.73 | 94.89%
2002-05-02 | 19:16 | 13.63 | 7.99 | 59 | 8.85 | 96.2%
2002-05-02 | 19:31 | 13.59 | 7.98 | 59 | 8.68 | 94.35%
2002-05-02 | 19:46 | 13.55 | 7.97 | 59 | 8.75 | 95.11%
2002-05-02 | 20:01 | 13.51 | 7.96 | 59 | 8.6 | 93.48%
2002-05-02 | 20:16 | 13.47 | 7.94 | 59 | 8.54 | 91.83%
2002-05-02 | 20:31 | 13.4 | 7.93 | 59 | 8.61 | 92.58%
2002-05-02 | 20:46 | 13.3 | 7.91 | 59 | 8.55 | 91.94%
2002-05-02 | 21:01 | 13.23 | 7.9 | 59 | 8.58 | 92.26%
2002-05-02 | 21:16 | 13.17 | 7.89 | 59 | 8.44 | 90.75%
2002-05-02 | 21:31 | 13.09 | 7.88 | 59 | 8.57 | 92.15%
2002-05-02 | 21:46 | 12.99 | 7.87 | 59 | 8.46 | 90%
2002-05-02 | 22:01 | 12.89 | 7.86 | 59 | 8.45 | 89.89%
2002-05-02 | 22:16 | 12.81 | 7.85 | 59 | 8.48 | 90.21%
2002-05-02 | 22:31 | 12.72 | 7.85 | 59 | 8.39 | 89.26%
2002-05-02 | 22:46 | 12.61 | 7.85 | 59 | 8.38 | 89.15%
2002-05-02 | 23:01 | 12.49 | 7.84 | 59 | 8.41 | 88.99%
2002-05-02 | 23:16 | 12.38 | 7.83 | 59 | 8.41 | 88.99%
2002-05-02 | 23:31 | 12.26 | 7.83 | 59 | 8.41 | 88.99%
2002-05-02 | 23:46 | 12.18 | 7.82 | 59 | 8.39 | 88.78%
2002-05-03 | 0:01 | 12.06 | 7.82 | 59 | 8.32 | 88.04%
2002-05-03 | 0:16 | 11.95 | 7.81 | 59 | 8.22 | 86.07%
2002-05-03 | 0:31 | 11.85 | 7.81 | 59 | 8.22 | 86.07%
2002-05-03 | 0:46 | 11.76 | 7.81 | 59 | 8.4 | 87.96%
2002-05-03 | 1:01 | 11.67 | 7.81 | 59 | 8.3 | 86.91%
2002-05-03 | 1:16 | 11.61 | 7.8 | 59 | 8.18 | 85.65%
2002-05-03 | 1:31 | 11.56 | 7.8 | 59 | 8.1 | 84.82%
2002-05-03 | 1:46 | 11.49 | 7.8 | 59 | 8.14 | 83.92%
2002-05-03 | 2:01 | 11.42 | 7.8 | 59 | 8.16 | 84.12%
2002-05-03 | 2:16 | 11.35 | 7.8 | 59 | 8.12 | 83.71%
2002-05-03 | 2:31 | 11.3 | 7.8 | 59 | 8.13 | 83.81%
2002-05-03 | 2:46 | 11.23 | 7.79 | 59 | 8.01 | 82.58%
2002-05-03 | 3:01 | 11.16 | 7.79 | 59 | 8.01 | 82.58%
2002-05-03 | 3:16 | 11.12 | 7.79 | 59 | 7.95 | 81.96%
2002-05-03 | 3:31 | 11.07 | 7.79 | 59 | 7.99 | 82.37%
2002-05-03 | 3:46 | 11.02 | 7.79 | 59 | 7.98 | 82.27%
2002-05-03 | 4:01 | 10.99 | 7.78 | 59 | 7.84 | 79.59%
2002-05-03 | 4:16 | 10.95 | 7.78 | 59 | 8.02 | 81.42%
2002-05-03 | 4:31 | 10.93 | 7.78 | 59 | 7.88 | 80%
2002-05-03 | 4:46 | 10.91 | 7.78 | 59 | 7.85 | 79.7%
2002-05-03 | 5:01 | 10.87 | 7.78 | 59 | 7.84 | 79.59%
2002-05-03 | 5:16 | 10.87 | 7.78 | 59 | 7.96 | 80.81%
2002-05-03 | 5:31 | 10.86 | 7.78 | 59 | 7.89 | 80.1%
2002-05-03 | 5:46 | 10.83 | 7.78 | 59 | 7.62 | 77.36%
2002-05-03 | 6:01 | 10.8 | 7.78 | 59 | 7.66 | 77.77%
2002-05-03 | 6:16 | 10.8 | 7.79 | 58 | 7.55 | 76.65%
2002-05-03 | 6:31 | 10.8 | 7.79 | 58 | 7.68 | 77.97%
2002-05-03 | 6:46 | 10.79 | 7.8 | 58 | 7.69 | 78.07%
2002-05-03 | 7:01 | 10.77 | 7.81 | 58 | 7.56 | 76.75%
2002-05-03 | 7:16 | 10.75 | 7.82 | 58 | 7.63 | 77.46%
2002-05-03 | 7:31 | 10.75 | 7.83 | 58 | 7.53 | 76.45%
2002-05-03 | 7:46 | 10.77 | 7.85 | 58 | 7.63 | 77.46%
2002-05-03 | 8:01 | 10.76 | 7.86 | 58 | 7.59 | 77.06%
2002-05-03 | 8:16 | 10.77 | 7.88 | 58 | 7.49 | 76.04%