## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | DR 226.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s of USGS Gauging Station d/s Wickiup
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2002-04-29 15:31 | 10 | 7.05 | 10.2 | 105 | 58 | 0
Sonde | 2002-04-29 15:31 | 10.01 | 7.56 | 11.02 | 114 | 55
Grade |  | A | Fail | B |  | A
Audit 2 | 2002-05-01 11:08 | 11 | 7.9 | 10.5 | 111 | 58 | 0
Sonde | 2002-05-01 11:16 | 10.49 | 7.83 | 11.01 | 114 | 51
Grade |  | A | A | B |  | B
Audit 3 | 2002-05-03 10:20 | 10.5 | 7.9 | 10.2 | 106 | 60 | 0
Sonde | 2002-05-03 10:16 | 10.8 | 7.82 | 11.17 | 116 | 53
Grade |  | A | A | B |  | B
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note Audits 1, 2, & 3 are macro(ed) into this audit sheet and represent the beginning, middle and end of the deployment. Audits 4 and 5 were in the middle of the deployment and are not macro(ed) into this sheet. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | DR 226.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s of USGS Gauging Station d/s Wickiup
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-29 15:31 | 10 | 7.05 | 10.2 | 105 | 58
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | DR 226.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s of USGS Gauging Station d/s Wickiup
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-01 11:08 | 11 | 7.9 | 10.5 | 111 | 58
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | DR 226.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s of USGS Gauging Station d/s Wickiup
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-03 10:20 | 10.5 | 7.9 | 10.2 | 106 | 60
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | DR 226.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s of USGS Gauging Station d/s Wickiup
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-30 13:35 | 10.1 | 8 | 10.2 | 105 | 59
Sonde Values | 2002-04-30 13:31 | 10 | 7.8 | 11.2 | 115 | 55
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | DR 226.75
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s of USGS Gauging Station d/s Wickiup
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 15:49 | 10.7 | 7.9 | 10.2 | 106 | 58
Sonde Values | 2002-05-02 15:46 | 10.9 | 7.9 | 11 | 115 | 54
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | DO
|  | °C |  | uS/cm | mg/L | % Sat
2002-04-29 | 15:31 | 10.1 | 7.56 | 55 | 11.02 | 113.6%
2002-04-29 | 15:46 | 10.11 | 7.8 | 55 | 11.08 | 114.2%
2002-04-29 | 16:01 | 10.11 | 7.83 | 55 | 11.01 | 113.5%
2002-04-29 | 16:16 | 10.1 | 7.85 | 55 | 10.94 | 112.8%
2002-04-29 | 16:31 | 10.09 | 7.81 | 55 | 10.92 | 112.6%
2002-04-29 | 16:46 | 10.07 | 7.83 | 55 | 10.88 | 112.2%
2002-04-29 | 17:01 | 10.04 | 7.83 | 55 | 10.86 | 112%
2002-04-29 | 17:16 | 10.04 | 7.82 | 55 | 10.82 | 111.5%
2002-04-29 | 17:31 | 10.02 | 7.81 | 55 | 10.88 | 112.2%
2002-04-29 | 17:46 | 10.02 | 7.78 | 55 | 10.98 | 113.2%
2002-04-29 | 18:01 | 10.01 | 7.79 | 55 | 10.98 | 113.2%
2002-04-29 | 18:16 | 10.02 | 7.78 | 55 | 10.98 | 113.2%
2002-04-29 | 18:31 | 10.01 | 7.79 | 55 | 10.98 | 113.2%
2002-04-29 | 18:46 | 10 | 7.81 | 55 | 10.93 | 112.7%
2002-04-29 | 19:01 | 9.99 | 7.82 | 55 | 10.92 | 110.9%
2002-04-29 | 19:16 | 9.99 | 7.77 | 55 | 10.94 | 111.1%
2002-04-29 | 19:31 | 9.97 | 7.78 | 55 | 10.95 | 111.2%
2002-04-29 | 19:46 | 9.96 | 7.8 | 55 | 10.93 | 111%
2002-04-29 | 20:01 | 9.96 | 7.78 | 55 | 10.93 | 111%
2002-04-29 | 20:16 | 9.95 | 7.75 | 55 | 10.95 | 111.2%
2002-04-29 | 20:31 | 9.94 | 7.78 | 55 | 10.92 | 110.9%
2002-04-29 | 20:46 | 9.93 | 7.76 | 55 | 10.95 | 111.2%
2002-04-29 | 21:01 | 9.93 | 7.75 | 55 | 10.92 | 110.9%
2002-04-29 | 21:16 | 9.92 | 7.76 | 55 | 10.91 | 110.8%
2002-04-29 | 21:31 | 9.92 | 7.76 | 55 | 10.91 | 110.8%
2002-04-29 | 21:46 | 9.9 | 7.77 | 56 | 10.9 | 110.7%
2002-04-29 | 22:01 | 9.89 | 7.75 | 55 | 10.91 | 110.8%
2002-04-29 | 22:16 | 9.89 | 7.74 | 55 | 10.92 | 110.9%
2002-04-29 | 22:31 | 9.89 | 7.75 | 55 | 10.91 | 110.8%
2002-04-29 | 22:46 | 9.89 | 7.74 | 55 | 10.91 | 110.8%
2002-04-29 | 23:01 | 9.89 | 7.73 | 55 | 10.9 | 110.7%
2002-04-29 | 23:16 | 9.89 | 7.75 | 55 | 10.91 | 110.8%
2002-04-29 | 23:31 | 9.88 | 7.76 | 55 | 10.89 | 110.6%
2002-04-29 | 23:46 | 9.88 | 7.73 | 55 | 10.9 | 110.7%
2002-04-30 | 0:01 | 9.88 | 7.73 | 55 | 10.92 | 110.9%
2002-04-30 | 0:16 | 9.88 | 7.78 | 55 | 10.88 | 110.5%
2002-04-30 | 0:31 | 9.88 | 7.74 | 55 | 10.9 | 110.7%
2002-04-30 | 0:46 | 9.87 | 7.73 | 55 | 10.92 | 110.9%
2002-04-30 | 1:01 | 9.87 | 7.76 | 55 | 10.91 | 110.8%
2002-04-30 | 1:16 | 9.87 | 7.74 | 55 | 10.89 | 110.6%
2002-04-30 | 1:31 | 9.87 | 7.75 | 55 | 10.91 | 110.8%
2002-04-30 | 1:46 | 9.87 | 7.75 | 55 | 10.91 | 110.8%
2002-04-30 | 2:01 | 9.86 | 7.75 | 55 | 10.9 | 110.7%
2002-04-30 | 2:16 | 9.85 | 7.74 | 55 | 10.91 | 110.8%
2002-04-30 | 2:31 | 9.85 | 7.76 | 55 | 10.91 | 110.8%
2002-04-30 | 2:46 | 9.84 | 7.74 | 55 | 10.93 | 111%
2002-04-30 | 3:01 | 9.84 | 7.75 | 55 | 10.91 | 110.8%
2002-04-30 | 3:16 | 9.83 | 7.74 | 55 | 10.93 | 111%
2002-04-30 | 3:31 | 9.83 | 7.76 | 55 | 10.91 | 110.8%
2002-04-30 | 3:46 | 9.82 | 7.73 | 55 | 10.91 | 110.8%
2002-04-30 | 4:01 | 9.82 | 7.76 | 55 | 10.91 | 110.8%
2002-04-30 | 4:16 | 9.81 | 7.74 | 55 | 10.91 | 110.8%
2002-04-30 | 4:31 | 9.81 | 7.75 | 55 | 10.89 | 110.6%
2002-04-30 | 4:46 | 9.8 | 7.74 | 55 | 10.92 | 110.9%
2002-04-30 | 5:01 | 9.8 | 7.75 | 55 | 10.91 | 110.8%
2002-04-30 | 5:16 | 9.8 | 7.72 | 55 | 10.92 | 110.9%
2002-04-30 | 5:31 | 9.8 | 7.75 | 55 | 10.92 | 110.9%
2002-04-30 | 5:46 | 9.79 | 7.75 | 55 | 10.91 | 110.8%
2002-04-30 | 6:01 | 9.79 | 7.73 | 55 | 10.93 | 111%
2002-04-30 | 6:16 | 9.78 | 7.73 | 55 | 10.95 | 111.2%
2002-04-30 | 6:31 | 9.77 | 7.75 | 55 | 10.91 | 110.8%
2002-04-30 | 6:46 | 9.77 | 7.73 | 55 | 10.92 | 110.9%
2002-04-30 | 7:01 | 9.77 | 7.73 | 55 | 10.94 | 111.1%
2002-04-30 | 7:16 | 9.77 | 7.75 | 55 | 10.93 | 111%
2002-04-30 | 7:31 | 9.78 | 7.74 | 55 | 10.97 | 111.4%
2002-04-30 | 7:46 | 9.77 | 7.74 | 55 | 10.96 | 111.3%
2002-04-30 | 8:01 | 9.75 | 7.75 | 55 | 10.97 | 111.4%
2002-04-30 | 8:16 | 9.75 | 7.73 | 55 | 10.99 | 111.6%
2002-04-30 | 8:31 | 9.75 | 7.72 | 55 | 11 | 111.7%
2002-04-30 | 8:46 | 9.75 | 7.73 | 55 | 11.01 | 111.8%
2002-04-30 | 9:01 | 9.75 | 7.75 | 55 | 11.03 | 112%
2002-04-30 | 9:16 | 9.77 | 7.76 | 55 | 11.01 | 111.8%
2002-04-30 | 9:31 | 9.78 | 7.75 | 55 | 11.05 | 112.2%
2002-04-30 | 9:46 | 9.8 | 7.77 | 55 | 11.06 | 112.3%
2002-04-30 | 10:01 | 9.82 | 7.73 | 55 | 11.08 | 112.5%
2002-04-30 | 10:16 | 9.83 | 7.76 | 55 | 11.18 | 113.5%
2002-04-30 | 10:31 | 9.87 | 7.77 | 55 | 11.16 | 113.3%
2002-04-30 | 10:46 | 9.88 | 7.75 | 55 | 11.19 | 113.6%
2002-04-30 | 11:01 | 9.88 | 7.76 | 55 | 11.2 | 113.7%
2002-04-30 | 11:16 | 9.89 | 7.77 | 55 | 11.2 | 113.7%
2002-04-30 | 11:31 | 9.91 | 7.74 | 55 | 11.2 | 113.7%
2002-04-30 | 11:46 | 9.91 | 7.77 | 55 | 11.18 | 113.5%
2002-04-30 | 12:01 | 9.91 | 7.76 | 55 | 11.19 | 113.6%
2002-04-30 | 12:16 | 9.91 | 7.75 | 55 | 11.2 | 113.7%
2002-04-30 | 12:31 | 9.92 | 7.77 | 55 | 11.2 | 113.7%
2002-04-30 | 12:46 | 9.97 | 7.74 | 55 | 11.19 | 113.6%
2002-04-30 | 13:01 | 10.01 | 7.76 | 55 | 11.19 | 115.4%
2002-04-30 | 13:16 | 10.04 | 7.78 | 55 | 11.19 | 115.4%
2002-04-30 | 13:31 | 10.02 | 7.78 | 55 | 11.17 | 115.2%
2002-04-30 | 13:46 | 10.14 | 7.78 | 55 | 11.15 | 114.9%
2002-04-30 | 14:01 | 10.18 | 7.79 | 55 | 11.14 | 114.8%
2002-04-30 | 14:16 | 10.23 | 7.85 | 55 | 11 | 113.4%
2002-04-30 | 14:31 | 10.29 | 7.86 | 55 | 10.95 | 112.9%
2002-04-30 | 14:46 | 10.36 | 7.84 | 55 | 10.98 | 113.2%
2002-04-30 | 15:01 | 10.43 | 7.86 | 54 | 10.91 | 112.5%
2002-04-30 | 15:16 | 10.41 | 7.84 | 54 | 11.01 | 113.5%
2002-04-30 | 15:31 | 10.44 | 7.86 | 54 | 10.95 | 112.9%
2002-04-30 | 15:46 | 10.44 | 7.87 | 53 | 10.96 | 113%
2002-04-30 | 16:01 | 10.44 | 7.86 | 53 | 10.9 | 112.4%
2002-04-30 | 16:16 | 10.49 | 7.85 | 53 | 10.85 | 111.9%
2002-04-30 | 16:31 | 10.51 | 7.87 | 52 | 10.86 | 113.1%
2002-04-30 | 16:46 | 10.52 | 7.86 | 52 | 10.84 | 112.9%
2002-04-30 | 17:01 | 10.49 | 7.86 | 52 | 10.91 | 112.5%
2002-04-30 | 17:16 | 10.43 | 7.84 | 51 | 10.94 | 112.8%
2002-04-30 | 17:31 | 10.39 | 7.86 | 51 | 10.87 | 112.1%
2002-04-30 | 17:46 | 10.37 | 7.84 | 51 | 10.88 | 112.2%
2002-04-30 | 18:01 | 10.41 | 7.87 | 51 | 10.81 | 111.4%
2002-04-30 | 18:16 | 10.49 | 7.87 | 51 | 10.85 | 111.9%
2002-04-30 | 18:31 | 10.51 | 7.85 | 51 | 10.87 | 113.2%
2002-04-30 | 18:46 | 10.53 | 7.88 | 52 | 10.84 | 112.9%
2002-04-30 | 19:01 | 10.55 | 7.85 | 52 | 10.81 | 112.6%
2002-04-30 | 19:16 | 10.56 | 7.86 | 52 | 10.8 | 112.5%
2002-04-30 | 19:31 | 10.59 | 7.88 | 51 | 10.79 | 112.4%
2002-04-30 | 19:46 | 10.55 | 7.87 | 51 | 10.8 | 112.5%
2002-04-30 | 20:01 | 10.58 | 7.88 | 52 | 10.81 | 112.6%
2002-04-30 | 20:16 | 10.6 | 7.9 | 52 | 10.78 | 112.3%
2002-04-30 | 20:31 | 10.6 | 7.89 | 52 | 10.72 | 111.7%
2002-04-30 | 20:46 | 10.6 | 7.9 | 51 | 10.74 | 111.9%
2002-04-30 | 21:01 | 10.62 | 7.89 | 51 | 10.72 | 111.7%
2002-04-30 | 21:16 | 10.63 | 7.89 | 51 | 10.78 | 112.3%
2002-04-30 | 21:31 | 10.64 | 7.9 | 51 | 10.77 | 112.2%
2002-04-30 | 21:46 | 10.66 | 7.91 | 51 | 10.72 | 111.7%
2002-04-30 | 22:01 | 10.67 | 7.89 | 50 | 10.74 | 111.9%
2002-04-30 | 22:16 | 10.66 | 7.92 | 51 | 10.7 | 111.5%
2002-04-30 | 22:31 | 10.58 | 7.92 | 50 | 10.82 | 112.7%
2002-04-30 | 22:46 | 10.49 | 7.87 | 50 | 10.83 | 111.6%
2002-04-30 | 23:01 | 10.44 | 7.89 | 51 | 10.79 | 111.2%
2002-04-30 | 23:16 | 10.4 | 7.87 | 51 | 10.77 | 111%
2002-04-30 | 23:31 | 10.4 | 7.86 | 51 | 10.81 | 111.4%
2002-04-30 | 23:46 | 10.4 | 7.86 | 52 | 10.78 | 111.1%
2002-05-01 | 0:01 | 10.4 | 7.87 | 51 | 10.78 | 111.1%
2002-05-01 | 0:16 | 10.41 | 7.87 | 51 | 10.78 | 111.1%
2002-05-01 | 0:31 | 10.4 | 7.87 | 50 | 10.77 | 111%
2002-05-01 | 0:46 | 10.4 | 7.88 | 50 | 10.79 | 111.2%
2002-05-01 | 1:01 | 10.42 | 7.86 | 50 | 10.81 | 111.4%
2002-05-01 | 1:16 | 10.43 | 7.87 | 50 | 10.83 | 111.6%
2002-05-01 | 1:31 | 10.42 | 7.84 | 52 | 10.84 | 111.8%
2002-05-01 | 1:46 | 10.43 | 7.84 | 52 | 10.79 | 111.2%
2002-05-01 | 2:01 | 10.44 | 7.86 | 52 | 10.81 | 111.4%
2002-05-01 | 2:16 | 10.43 | 7.86 | 52 | 10.81 | 111.4%
2002-05-01 | 2:31 | 10.42 | 7.87 | 52 | 10.84 | 111.8%
2002-05-01 | 2:46 | 10.41 | 7.88 | 52 | 10.81 | 111.4%
2002-05-01 | 3:01 | 10.41 | 7.87 | 51 | 10.82 | 111.5%
2002-05-01 | 3:16 | 10.41 | 7.87 | 51 | 10.8 | 111.3%
2002-05-01 | 3:31 | 10.41 | 7.87 | 51 | 10.83 | 111.6%
2002-05-01 | 3:46 | 10.4 | 7.87 | 51 | 10.79 | 111.2%
2002-05-01 | 4:01 | 10.39 | 7.85 | 51 | 10.82 | 111.5%
2002-05-01 | 4:16 | 10.39 | 7.87 | 50 | 10.83 | 111.6%
2002-05-01 | 4:31 | 10.39 | 7.88 | 51 | 10.8 | 111.3%
2002-05-01 | 4:46 | 10.4 | 7.87 | 50 | 10.82 | 111.5%
2002-05-01 | 5:01 | 10.4 | 7.88 | 50 | 10.77 | 111%
2002-05-01 | 5:16 | 10.4 | 7.85 | 50 | 10.79 | 111.2%
2002-05-01 | 5:31 | 10.41 | 7.88 | 51 | 10.79 | 111.2%
2002-05-01 | 5:46 | 10.41 | 7.85 | 51 | 10.82 | 111.5%
2002-05-01 | 6:01 | 10.42 | 7.85 | 51 | 10.8 | 111.3%
2002-05-01 | 6:16 | 10.41 | 7.86 | 51 | 10.81 | 111.4%
2002-05-01 | 6:31 | 10.42 | 7.85 | 52 | 10.81 | 111.4%
2002-05-01 | 6:46 | 10.42 | 7.86 | 52 | 10.79 | 111.2%
2002-05-01 | 7:01 | 10.43 | 7.88 | 52 | 10.8 | 111.3%
2002-05-01 | 7:16 | 10.43 | 7.85 | 52 | 10.82 | 111.5%
2002-05-01 | 7:31 | 10.43 | 7.87 | 51 | 10.8 | 111.3%
2002-05-01 | 7:46 | 10.42 | 7.87 | 51 | 10.82 | 111.5%
2002-05-01 | 8:01 | 10.42 | 7.84 | 53 | 10.84 | 111.8%
2002-05-01 | 8:16 | 10.42 | 7.83 | 53 | 10.84 | 111.8%
2002-05-01 | 8:31 | 10.41 | 7.86 | 53 | 10.83 | 111.6%
2002-05-01 | 8:46 | 10.42 | 7.85 | 52 | 10.83 | 111.6%
2002-05-01 | 9:01 | 10.43 | 7.85 | 52 | 10.86 | 112%
2002-05-01 | 9:16 | 10.43 | 7.86 | 52 | 10.96 | 113%
2002-05-01 | 9:31 | 10.45 | 7.86 | 52 | 10.98 | 113.2%
2002-05-01 | 9:46 | 10.45 | 7.83 | 52 | 11 | 113.4%
2002-05-01 | 10:01 | 10.46 | 7.84 | 52 | 11.01 | 113.5%
2002-05-01 | 10:16 | 10.48 | 7.85 | 52 | 10.98 | 113.2%
2002-05-01 | 10:31 | 10.49 | 7.87 | 51 | 10.97 | 113.1%
2002-05-01 | 10:46 | 10.51 | 7.88 | 51 | 10.98 | 114.4%
2002-05-01 | 11:01 | 10.5 | 7.86 | 52 | 11 | 114.6%
2002-05-01 | 11:16 | 10.49 | 7.83 | 51 | 11.01 | 113.5%
2002-05-01 | 11:31 | 10.48 | 7.83 | 51 | 11.1 | 114.4%
2002-05-01 | 11:46 | 10.49 | 7.86 | 51 | 11.12 | 114.6%
2002-05-01 | 12:01 | 10.48 | 7.86 | 51 | 11.13 | 114.7%
2002-05-01 | 12:16 | 10.47 | 7.85 | 51 | 11.14 | 114.8%
2002-05-01 | 12:31 | 10.46 | 7.85 | 51 | 11.14 | 114.8%
2002-05-01 | 12:46 | 10.44 | 7.85 | 52 | 11.13 | 114.7%
2002-05-01 | 13:01 | 10.43 | 7.84 | 52 | 11.15 | 114.9%
2002-05-01 | 13:16 | 10.43 | 7.85 | 52 | 11.09 | 114.3%
2002-05-01 | 13:31 | 10.44 | 7.84 | 52 | 11.12 | 114.6%
2002-05-01 | 13:46 | 10.45 | 7.84 | 52 | 11.13 | 114.7%
2002-05-01 | 14:01 | 10.45 | 7.85 | 51 | 11.13 | 114.7%
2002-05-01 | 14:16 | 10.46 | 7.83 | 51 | 11.14 | 114.8%
2002-05-01 | 14:31 | 10.45 | 7.86 | 51 | 11.14 | 114.8%
2002-05-01 | 14:46 | 10.44 | 7.85 | 51 | 11.14 | 114.8%
2002-05-01 | 15:01 | 10.44 | 7.86 | 50 | 11.14 | 114.8%
2002-05-01 | 15:16 | 10.41 | 7.84 | 50 | 11.13 | 114.7%
2002-05-01 | 15:31 | 10.37 | 7.86 | 50 | 11.11 | 114.5%
2002-05-01 | 15:46 | 10.37 | 7.84 | 50 | 11.09 | 114.3%
2002-05-01 | 16:01 | 10.4 | 7.88 | 51 | 11.02 | 113.6%
2002-05-01 | 16:16 | 10.39 | 7.84 | 51 | 11.07 | 114.1%
2002-05-01 | 16:31 | 10.39 | 7.86 | 52 | 11.04 | 113.8%
2002-05-01 | 16:46 | 10.38 | 7.88 | 52 | 11.01 | 113.5%
2002-05-01 | 17:01 | 10.37 | 7.86 | 52 | 11.1 | 114.4%
2002-05-01 | 17:16 | 10.4 | 7.88 | 54 | 11.14 | 114.8%
2002-05-01 | 17:31 | 10.36 | 7.88 | 53 | 11.1 | 114.4%
2002-05-01 | 17:46 | 10.36 | 7.88 | 53 | 11.06 | 114%
2002-05-01 | 18:01 | 10.36 | 7.89 | 53 | 11.02 | 113.6%
2002-05-01 | 18:16 | 10.4 | 7.9 | 53 | 10.98 | 113.2%
2002-05-01 | 18:31 | 10.46 | 7.89 | 53 | 11 | 113.4%
2002-05-01 | 18:46 | 10.49 | 7.92 | 53 | 10.95 | 112.9%
2002-05-01 | 19:01 | 10.5 | 7.93 | 53 | 10.92 | 113.8%
2002-05-01 | 19:16 | 10.48 | 7.91 | 53 | 10.95 | 112.9%
2002-05-01 | 19:31 | 10.46 | 7.89 | 53 | 10.97 | 113.1%
2002-05-01 | 19:46 | 10.42 | 7.91 | 53 | 10.95 | 112.9%
2002-05-01 | 20:01 | 10.4 | 7.9 | 52 | 10.94 | 112.8%
2002-05-01 | 20:16 | 10.37 | 7.89 | 52 | 10.94 | 112.8%
2002-05-01 | 20:31 | 10.38 | 7.88 | 52 | 10.96 | 113%
2002-05-01 | 20:46 | 10.39 | 7.87 | 52 | 10.95 | 112.9%
2002-05-01 | 21:01 | 10.41 | 7.89 | 52 | 10.92 | 112.6%
2002-05-01 | 21:16 | 10.39 | 7.89 | 52 | 10.93 | 112.7%
2002-05-01 | 21:31 | 10.4 | 7.88 | 52 | 10.94 | 112.8%
2002-05-01 | 21:46 | 10.39 | 7.88 | 52 | 10.91 | 112.5%
2002-05-01 | 22:01 | 10.4 | 7.86 | 51 | 10.94 | 112.8%
2002-05-01 | 22:16 | 10.39 | 7.87 | 52 | 10.93 | 112.7%
2002-05-01 | 22:31 | 10.39 | 7.88 | 51 | 10.9 | 112.4%
2002-05-01 | 22:46 | 10.38 | 7.88 | 51 | 11.01 | 113.5%
2002-05-01 | 23:01 | 10.38 | 7.87 | 52 | 10.97 | 113.1%
2002-05-01 | 23:16 | 10.38 | 7.87 | 53 | 11.02 | 113.6%
2002-05-01 | 23:31 | 10.38 | 7.87 | 53 | 11 | 113.4%
2002-05-01 | 23:46 | 10.4 | 7.88 | 53 | 10.95 | 112.9%
2002-05-02 | 0:01 | 10.41 | 7.88 | 53 | 10.97 | 113.1%
2002-05-02 | 0:16 | 10.42 | 7.89 | 52 | 10.91 | 112.5%
2002-05-02 | 0:31 | 10.41 | 7.87 | 53 | 10.93 | 112.7%
2002-05-02 | 0:46 | 10.42 | 7.88 | 53 | 10.91 | 112.5%
2002-05-02 | 1:01 | 10.41 | 7.87 | 53 | 10.92 | 112.6%
2002-05-02 | 1:16 | 10.41 | 7.87 | 53 | 10.89 | 112.3%
2002-05-02 | 1:31 | 10.44 | 7.89 | 53 | 10.88 | 112.2%
2002-05-02 | 1:46 | 10.42 | 7.88 | 52 | 10.91 | 112.5%
2002-05-02 | 2:01 | 10.4 | 7.87 | 52 | 10.91 | 112.5%
2002-05-02 | 2:16 | 10.38 | 7.86 | 52 | 10.97 | 113.1%
2002-05-02 | 2:31 | 10.4 | 7.84 | 52 | 10.93 | 112.7%
2002-05-02 | 2:46 | 10.42 | 7.86 | 52 | 10.92 | 112.6%
2002-05-02 | 3:01 | 10.46 | 7.84 | 52 | 10.89 | 112.3%
2002-05-02 | 3:16 | 10.47 | 7.85 | 52 | 10.92 | 112.6%
2002-05-02 | 3:31 | 10.5 | 7.87 | 53 | 10.96 | 114.2%
2002-05-02 | 3:46 | 10.53 | 7.86 | 53 | 10.94 | 114%
2002-05-02 | 4:01 | 10.52 | 7.88 | 53 | 10.89 | 113.4%
2002-05-02 | 4:16 | 10.5 | 7.88 | 53 | 10.85 | 113%
2002-05-02 | 4:31 | 10.47 | 7.87 | 54 | 10.88 | 112.2%
2002-05-02 | 4:46 | 10.47 | 7.87 | 54 | 10.85 | 111.9%
2002-05-02 | 5:01 | 10.47 | 7.86 | 54 | 10.87 | 112.1%
2002-05-02 | 5:16 | 10.47 | 7.86 | 54 | 10.86 | 112%
2002-05-02 | 5:31 | 10.47 | 7.87 | 54 | 10.87 | 112.1%
2002-05-02 | 5:46 | 10.48 | 7.87 | 54 | 10.85 | 111.9%
2002-05-02 | 6:01 | 10.51 | 7.85 | 53 | 10.85 | 113%
2002-05-02 | 6:16 | 10.52 | 7.85 | 53 | 10.95 | 114.1%
2002-05-02 | 6:31 | 10.53 | 7.85 | 53 | 10.96 | 114.2%
2002-05-02 | 6:46 | 10.55 | 7.87 | 53 | 10.9 | 113.5%
2002-05-02 | 7:01 | 10.57 | 7.87 | 54 | 10.91 | 113.6%
2002-05-02 | 7:16 | 10.58 | 7.87 | 54 | 10.87 | 113.2%
2002-05-02 | 7:31 | 10.58 | 7.87 | 53 | 10.87 | 113.2%
2002-05-02 | 7:46 | 10.6 | 7.86 | 53 | 10.86 | 113.1%
2002-05-02 | 8:01 | 10.61 | 7.86 | 53 | 10.87 | 113.2%
2002-05-02 | 8:16 | 10.62 | 7.88 | 53 | 10.86 | 113.1%
2002-05-02 | 8:31 | 10.63 | 7.87 | 54 | 10.94 | 114%
2002-05-02 | 8:46 | 10.64 | 7.86 | 53 | 10.9 | 113.5%
2002-05-02 | 9:01 | 10.64 | 7.85 | 53 | 10.96 | 114.2%
2002-05-02 | 9:16 | 10.64 | 7.85 | 53 | 11.04 | 115%
2002-05-02 | 9:31 | 10.65 | 7.85 | 53 | 11.03 | 114.9%
2002-05-02 | 9:46 | 10.66 | 7.85 | 53 | 11.02 | 114.8%
2002-05-02 | 10:01 | 10.68 | 7.85 | 53 | 11.04 | 115%
2002-05-02 | 10:16 | 10.7 | 7.85 | 53 | 10.94 | 114%
2002-05-02 | 10:31 | 10.72 | 7.87 | 53 | 10.96 | 114.2%
2002-05-02 | 10:46 | 10.75 | 7.85 | 54 | 11.01 | 114.7%
2002-05-02 | 11:01 | 10.79 | 7.85 | 54 | 11 | 114.6%
2002-05-02 | 11:16 | 10.83 | 7.87 | 54 | 10.97 | 114.3%
2002-05-02 | 11:31 | 10.85 | 7.88 | 54 | 10.94 | 114%
2002-05-02 | 11:46 | 10.88 | 7.88 | 53 | 10.96 | 114.2%
2002-05-02 | 12:01 | 10.9 | 7.87 | 54 | 11.05 | 115.1%
2002-05-02 | 12:16 | 10.89 | 7.87 | 54 | 11.08 | 115.4%
2002-05-02 | 12:31 | 10.89 | 7.88 | 54 | 11.08 | 115.4%
2002-05-02 | 12:46 | 10.88 | 7.85 | 54 | 11.12 | 115.8%
2002-05-02 | 13:01 | 10.88 | 7.86 | 54 | 11.1 | 115.6%
2002-05-02 | 13:16 | 10.9 | 7.88 | 54 | 11.08 | 115.4%
2002-05-02 | 13:31 | 10.9 | 7.89 | 54 | 11.08 | 115.4%
2002-05-02 | 13:46 | 10.91 | 7.87 | 54 | 11.05 | 115.1%
2002-05-02 | 14:01 | 10.91 | 7.88 | 54 | 11.04 | 115%
2002-05-02 | 14:16 | 10.93 | 7.87 | 53 | 11.03 | 114.9%
2002-05-02 | 14:31 | 10.92 | 7.87 | 53 | 11.02 | 114.8%
2002-05-02 | 14:46 | 10.87 | 7.87 | 53 | 11.03 | 114.9%
2002-05-02 | 15:01 | 10.84 | 7.87 | 53 | 11 | 114.6%
2002-05-02 | 15:16 | 10.84 | 7.85 | 53 | 11.07 | 115.3%
2002-05-02 | 15:31 | 10.85 | 7.84 | 53 | 11.02 | 114.8%
2002-05-02 | 15:46 | 10.88 | 7.86 | 54 | 11 | 114.6%
2002-05-02 | 16:01 | 10.89 | 7.85 | 54 | 11 | 114.6%
2002-05-02 | 16:16 | 10.91 | 7.87 | 53 | 10.93 | 113.9%
2002-05-02 | 16:31 | 10.93 | 7.86 | 53 | 10.95 | 114.1%
2002-05-02 | 16:46 | 10.9 | 7.89 | 53 | 10.9 | 113.5%
2002-05-02 | 17:01 | 10.9 | 7.85 | 53 | 10.94 | 114%
2002-05-02 | 17:16 | 10.89 | 7.85 | 53 | 10.93 | 113.9%
2002-05-02 | 17:31 | 10.85 | 7.86 | 55 | 10.88 | 113.3%
2002-05-02 | 17:46 | 10.84 | 7.84 | 55 | 10.91 | 113.6%
2002-05-02 | 18:01 | 10.85 | 7.86 | 55 | 10.88 | 113.3%
2002-05-02 | 18:16 | 10.82 | 7.86 | 55 | 10.89 | 113.4%
2002-05-02 | 18:31 | 10.81 | 7.85 | 55 | 10.89 | 113.4%
2002-05-02 | 18:46 | 10.83 | 7.83 | 55 | 10.91 | 113.6%
2002-05-02 | 19:01 | 10.85 | 7.86 | 54 | 10.85 | 113%
2002-05-02 | 19:16 | 10.92 | 7.86 | 54 | 10.86 | 113.1%
2002-05-02 | 19:31 | 10.98 | 7.88 | 54 | 10.8 | 112.5%
2002-05-02 | 19:46 | 10.98 | 7.87 | 54 | 10.82 | 112.7%
2002-05-02 | 20:01 | 11.07 | 7.87 | 54 | 10.79 | 113.6%
2002-05-02 | 20:16 | 11.12 | 7.91 | 54 | 10.74 | 113.1%
2002-05-02 | 20:31 | 11.18 | 7.88 | 54 | 10.73 | 112.9%
2002-05-02 | 20:46 | 11.22 | 7.91 | 54 | 10.69 | 112.5%
2002-05-02 | 21:01 | 11.22 | 7.89 | 54 | 10.72 | 112.8%
2002-05-02 | 21:16 | 11.2 | 7.88 | 54 | 10.73 | 112.9%
2002-05-02 | 21:31 | 11.25 | 7.89 | 54 | 10.73 | 112.9%
2002-05-02 | 21:46 | 11.29 | 7.87 | 54 | 10.69 | 112.5%
2002-05-02 | 22:01 | 11.1 | 7.87 | 54 | 10.71 | 112.7%
2002-05-02 | 22:16 | 10.99 | 7.86 | 53 | 10.74 | 111.9%
2002-05-02 | 22:31 | 10.96 | 7.84 | 53 | 10.76 | 112.1%
2002-05-02 | 22:46 | 10.95 | 7.82 | 53 | 10.77 | 112.2%
2002-05-02 | 23:01 | 10.97 | 7.85 | 54 | 10.81 | 112.6%
2002-05-02 | 23:16 | 10.95 | 7.86 | 54 | 10.75 | 112%
2002-05-02 | 23:31 | 10.94 | 7.87 | 54 | 10.78 | 112.3%
2002-05-02 | 23:46 | 10.94 | 7.87 | 54 | 10.79 | 112.4%
2002-05-03 | 0:01 | 10.92 | 7.86 | 55 | 10.83 | 112.8%
2002-05-03 | 0:16 | 10.88 | 7.86 | 55 | 10.81 | 112.6%
2002-05-03 | 0:31 | 10.9 | 7.85 | 55 | 10.81 | 112.6%
2002-05-03 | 0:46 | 10.93 | 7.88 | 55 | 10.79 | 112.4%
2002-05-03 | 1:01 | 10.92 | 7.88 | 55 | 10.8 | 112.5%
2002-05-03 | 1:16 | 10.92 | 7.89 | 54 | 10.78 | 112.3%
2002-05-03 | 1:31 | 10.94 | 7.87 | 54 | 10.78 | 112.3%
2002-05-03 | 1:46 | 10.98 | 7.87 | 54 | 10.76 | 112.1%
2002-05-03 | 2:01 | 10.99 | 7.87 | 54 | 10.78 | 112.3%
2002-05-03 | 2:16 | 11.06 | 7.9 | 54 | 10.74 | 113.1%
2002-05-03 | 2:31 | 11.06 | 7.9 | 54 | 10.75 | 113.2%
2002-05-03 | 2:46 | 11.05 | 7.88 | 54 | 10.78 | 113.5%
2002-05-03 | 3:01 | 11.02 | 7.88 | 54 | 10.78 | 113.5%
2002-05-03 | 3:16 | 11.04 | 7.89 | 54 | 10.8 | 113.7%
2002-05-03 | 3:31 | 11.04 | 7.9 | 54 | 10.75 | 113.2%
2002-05-03 | 3:46 | 10.91 | 7.88 | 54 | 10.79 | 112.4%
2002-05-03 | 4:01 | 10.96 | 7.87 | 54 | 10.83 | 112.8%
2002-05-03 | 4:16 | 11 | 7.87 | 54 | 10.77 | 113.4%
2002-05-03 | 4:31 | 11 | 7.89 | 54 | 10.77 | 113.4%
2002-05-03 | 4:46 | 11.01 | 7.9 | 54 | 10.75 | 113.2%
2002-05-03 | 5:01 | 10.98 | 7.87 | 54 | 10.78 | 112.3%
2002-05-03 | 5:16 | 10.99 | 7.89 | 54 | 10.75 | 112%
2002-05-03 | 5:31 | 11 | 7.9 | 54 | 10.79 | 113.6%
2002-05-03 | 5:46 | 11 | 7.87 | 54 | 10.77 | 113.4%
2002-05-03 | 6:01 | 11 | 7.89 | 54 | 10.79 | 113.6%
2002-05-03 | 6:16 | 10.98 | 7.87 | 54 | 10.82 | 112.7%
2002-05-03 | 6:31 | 10.95 | 7.87 | 54 | 10.8 | 112.5%
2002-05-03 | 6:46 | 10.93 | 7.88 | 54 | 10.81 | 112.6%
2002-05-03 | 7:01 | 10.88 | 7.87 | 55 | 10.85 | 113%
2002-05-03 | 7:16 | 10.84 | 7.84 | 55 | 10.88 | 113.3%
2002-05-03 | 7:31 | 10.8 | 7.86 | 54 | 10.89 | 113.4%
2002-05-03 | 7:46 | 10.77 | 7.85 | 54 | 10.91 | 113.6%
2002-05-03 | 8:01 | 10.76 | 7.84 | 54 | 10.93 | 113.9%
2002-05-03 | 8:16 | 10.74 | 7.83 | 54 | 10.93 | 113.9%
2002-05-03 | 8:31 | 10.74 | 7.82 | 54 | 10.97 | 114.3%
2002-05-03 | 8:46 | 10.77 | 7.82 | 54 | 10.95 | 114.1%
2002-05-03 | 9:01 | 10.78 | 7.82 | 54 | 10.99 | 114.5%
2002-05-03 | 9:16 | 10.79 | 7.84 | 54 | 11.06 | 115.2%
2002-05-03 | 9:31 | 10.81 | 7.85 | 54 | 11.05 | 115.1%
2002-05-03 | 9:46 | 10.81 | 7.85 | 53 | 11.11 | 115.7%
2002-05-03 | 10:01 | 10.81 | 7.85 | 53 | 11.13 | 115.9%
2002-05-03 | 10:16 | 10.8 | 7.82 | 53 | 11.17 | 116.4%