## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | DR 191.75??
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Golf Course?? Station ID=DR @ Harper Bridge
Contact Name |  |  | Latitude | [5 empty] | Note: Elevation was 4170
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2002-04-29 15:00 | 12.6 | 7.3 | 8.6 | 93 | 44 | 0
Sonde | 2002-04-29 15:01 | 12.44 | 7.02 | 8.69 | 92 | 41
Grade |  | A | A | A |  | A
Audit 2 | 2002-05-01 09:20 | 10.2 | 7.4 | 8.7 | 90 | 43 | 0
Sonde | 2002-05-01 09:16 | 10.11 | 6.97 | 8.73 | 90 | 41
Grade |  | A | B | A |  | A
Audit 3 | 2002-05-03 09:34 | 12.1 | 7.4 | 8.2 | 87 | 44 | 0
Sonde | 2002-05-03 09:31 | 12.15 | 6.99 | 8.47 | 90 | 42
Grade |  | A | B | A |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note Audits 1, 2,& 3 are macro(ed) into this audit sheet and represent the beginning, middle and end of the deployment. Audits 4, 5, 6 & 7 were in the middle of the deployment and are not macro(ed) into this sheet. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | DR 191.75??
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Golf Course?? Station ID=DR @ Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-29 15:00 | 12.6 | 7.3 | 8.6 | 93 | 44
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | DR 191.75??
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Golf Course?? Station ID=DR @ Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-01 09:20 | 10.2 | 7.4 | 8.7 | 90 | 43
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | DR 191.75??
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Golf Course?? Station ID=DR @ Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-03 09:34 | 12.1 | 7.4 | 8.2 | 87 | 44
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | DR 191.75??
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Golf Course?? Station ID=DR @ Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-30 13:46 | 11.8 | 7.3 | 8.8 | 93 | 44
Sonde Value | 2002-04-30 13:46 | 11.55 | 7.01 | 8.79 | 93 | 41
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | DR 191.75??
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Golf Course?? Station ID=DR @ Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-01 17:05 | 12.9 | 7.4 | 9 | 97 | 42
Sonde Values | 2002-05-01 17:01 | 12.7 | 7 | 9 | 97 | 40
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 6
Organization | ODEQ |  | Station ID |  | DR 191.75??
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Golf Course?? Station ID=DR @ Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 08:55 | 11.4 | 7.4 | 8.9 | 94 | 43
Sonde Values | 2007-05-02 09:01 | 11.4 | 7 | 8.6 | 90 | 41
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 7
Organization | ODEQ |  | Station ID |  | DR 191.75??
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, Crosswater Golf Course?? Station ID=DR @ Harper Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 16:20 | 14.2 | 7.4 | 8.6 | 97 | 43
Sonde Values | 2002-05-02 16:16 | 14 | 7.1 | 8.9 | 98 | 41
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | DO
|  | °C |  | uS/cm | mg/L | % Sat
2002-04-29 | 15:01 | 12.44 | 7.02 | 41 | 8.69 | 92.45%
2002-04-29 | 15:16 | 12.48 | 7.09 | 41 | 8.65 | 92.02%
2002-04-29 | 15:31 | 12.49 | 7.08 | 41 | 8.63 | 91.81%
2002-04-29 | 15:46 | 12.48 | 7.08 | 41 | 8.61 | 91.6%
2002-04-29 | 16:01 | 12.49 | 7.08 | 41 | 8.63 | 91.81%
2002-04-29 | 16:16 | 12.5 | 7.06 | 41 | 8.63 | 93.3%
2002-04-29 | 16:31 | 12.49 | 7.06 | 41 | 8.63 | 91.81%
2002-04-29 | 16:46 | 12.48 | 7.05 | 41 | 8.63 | 91.81%
2002-04-29 | 17:01 | 12.46 | 7.04 | 41 | 8.73 | 92.87%
2002-04-29 | 17:16 | 12.44 | 7.03 | 41 | 8.67 | 92.23%
2002-04-29 | 17:31 | 12.44 | 7.03 | 41 | 8.64 | 91.91%
2002-04-29 | 17:46 | 12.45 | 7.03 | 41 | 8.67 | 92.23%
2002-04-29 | 18:01 | 12.44 | 7.03 | 41 | 8.62 | 91.7%
2002-04-29 | 18:16 | 12.42 | 7.02 | 41 | 8.6 | 91.49%
2002-04-29 | 18:31 | 12.4 | 7.03 | 41 | 8.59 | 91.38%
2002-04-29 | 18:46 | 12.37 | 7.04 | 41 | 8.59 | 91.38%
2002-04-29 | 19:01 | 12.33 | 7.03 | 41 | 8.58 | 91.28%
2002-04-29 | 19:16 | 12.28 | 7.02 | 41 | 8.57 | 91.17%
2002-04-29 | 19:31 | 12.21 | 7.01 | 41 | 8.56 | 91.06%
2002-04-29 | 19:46 | 12.16 | 7.01 | 41 | 8.55 | 90.96%
2002-04-29 | 20:01 | 12.09 | 7.01 | 41 | 8.54 | 90.85%
2002-04-29 | 20:16 | 12.04 | 7.01 | 41 | 8.53 | 90.74%
2002-04-29 | 20:31 | 12 | 7.01 | 41 | 8.52 | 90.64%
2002-04-29 | 20:46 | 11.96 | 7.01 | 41 | 8.51 | 90.05%
2002-04-29 | 21:01 | 11.92 | 7 | 41 | 8.49 | 89.84%
2002-04-29 | 21:16 | 11.86 | 7 | 41 | 8.48 | 89.74%
2002-04-29 | 21:31 | 11.81 | 6.99 | 41 | 8.48 | 89.74%
2002-04-29 | 21:46 | 11.76 | 7 | 41 | 8.47 | 89.63%
2002-04-29 | 22:01 | 11.72 | 6.99 | 41 | 8.46 | 89.52%
2002-04-29 | 22:16 | 11.68 | 6.99 | 41 | 8.45 | 89.42%
2002-04-29 | 22:31 | 11.65 | 7 | 40 | 8.44 | 89.31%
2002-04-29 | 22:46 | 11.62 | 6.99 | 40 | 8.44 | 89.31%
2002-04-29 | 23:01 | 11.59 | 6.98 | 40 | 8.44 | 89.31%
2002-04-29 | 23:16 | 11.56 | 6.98 | 40 | 8.43 | 89.21%
2002-04-29 | 23:31 | 11.51 | 6.98 | 40 | 8.42 | 89.1%
2002-04-29 | 23:46 | 11.47 | 6.98 | 40 | 8.42 | 88.63%
2002-04-30 | 0:01 | 11.41 | 6.98 | 40 | 8.43 | 88.74%
2002-04-30 | 0:16 | 11.35 | 6.98 | 40 | 8.43 | 88.74%
2002-04-30 | 0:31 | 11.29 | 6.98 | 40 | 8.43 | 88.74%
2002-04-30 | 0:46 | 11.22 | 6.97 | 40 | 8.43 | 88.74%
2002-04-30 | 1:01 | 11.15 | 6.97 | 40 | 8.43 | 88.74%
2002-04-30 | 1:16 | 11.08 | 6.97 | 40 | 8.43 | 88.74%
2002-04-30 | 1:31 | 11.02 | 6.97 | 40 | 8.44 | 88.84%
2002-04-30 | 1:46 | 10.95 | 6.96 | 40 | 8.44 | 87.92%
2002-04-30 | 2:01 | 10.88 | 6.97 | 40 | 8.45 | 88.02%
2002-04-30 | 2:16 | 10.82 | 6.97 | 40 | 8.45 | 88.02%
2002-04-30 | 2:31 | 10.77 | 6.96 | 40 | 8.44 | 87.92%
2002-04-30 | 2:46 | 10.71 | 6.96 | 40 | 8.45 | 88.02%
2002-04-30 | 3:01 | 10.66 | 6.95 | 40 | 8.45 | 88.02%
2002-04-30 | 3:16 | 10.62 | 6.95 | 40 | 8.45 | 88.02%
2002-04-30 | 3:31 | 10.57 | 6.96 | 40 | 8.45 | 88.02%
2002-04-30 | 3:46 | 10.53 | 6.96 | 40 | 8.46 | 88.12%
2002-04-30 | 4:01 | 10.48 | 6.95 | 41 | 8.45 | 87.11%
2002-04-30 | 4:16 | 10.44 | 6.96 | 41 | 8.44 | 87.01%
2002-04-30 | 4:31 | 10.39 | 6.95 | 40 | 8.45 | 87.11%
2002-04-30 | 4:46 | 10.34 | 6.96 | 40 | 8.45 | 87.11%
2002-04-30 | 5:01 | 10.29 | 6.95 | 41 | 8.45 | 87.11%
2002-04-30 | 5:16 | 10.23 | 6.95 | 41 | 8.46 | 87.22%
2002-04-30 | 5:31 | 10.18 | 6.95 | 40 | 8.47 | 87.32%
2002-04-30 | 5:46 | 10.13 | 6.95 | 41 | 8.48 | 87.42%
2002-04-30 | 6:01 | 10.07 | 6.95 | 41 | 8.49 | 87.53%
2002-04-30 | 6:16 | 10.02 | 6.95 | 41 | 8.49 | 87.53%
2002-04-30 | 6:31 | 9.97 | 6.94 | 41 | 8.49 | 86.19%
2002-04-30 | 6:46 | 9.92 | 6.95 | 41 | 8.5 | 86.29%
2002-04-30 | 7:01 | 9.89 | 6.95 | 41 | 8.52 | 86.5%
2002-04-30 | 7:16 | 9.85 | 6.95 | 41 | 8.51 | 86.4%
2002-04-30 | 7:31 | 9.83 | 6.96 | 41 | 8.53 | 86.6%
2002-04-30 | 7:46 | 9.81 | 6.96 | 41 | 8.54 | 86.7%
2002-04-30 | 8:01 | 9.8 | 6.96 | 41 | 8.55 | 86.8%
2002-04-30 | 8:16 | 9.8 | 6.96 | 41 | 8.57 | 87.01%
2002-04-30 | 8:31 | 9.81 | 6.95 | 41 | 8.57 | 87.01%
2002-04-30 | 8:46 | 9.83 | 6.96 | 41 | 8.59 | 87.21%
2002-04-30 | 9:01 | 9.87 | 6.96 | 41 | 8.59 | 87.21%
2002-04-30 | 9:16 | 9.93 | 6.96 | 41 | 8.61 | 87.41%
2002-04-30 | 9:31 | 9.99 | 6.96 | 41 | 8.62 | 87.51%
2002-04-30 | 9:46 | 10.06 | 6.97 | 41 | 8.63 | 88.97%
2002-04-30 | 10:01 | 10.15 | 6.96 | 41 | 8.64 | 89.07%
2002-04-30 | 10:16 | 10.25 | 6.97 | 41 | 8.66 | 89.28%
2002-04-30 | 10:31 | 10.36 | 6.97 | 41 | 8.67 | 89.38%
2002-04-30 | 10:46 | 10.47 | 6.98 | 41 | 8.69 | 89.59%
2002-04-30 | 11:01 | 10.57 | 6.99 | 41 | 8.69 | 90.52%
2002-04-30 | 11:16 | 10.68 | 6.99 | 41 | 8.71 | 90.73%
2002-04-30 | 11:31 | 10.8 | 6.99 | 41 | 8.71 | 90.73%
2002-04-30 | 11:46 | 10.9 | 6.98 | 41 | 8.72 | 90.83%
2002-04-30 | 12:01 | 11.01 | 6.99 | 41 | 8.74 | 92%
2002-04-30 | 12:16 | 11.12 | 7 | 41 | 8.74 | 92%
2002-04-30 | 12:31 | 11.19 | 7 | 41 | 8.75 | 92.11%
2002-04-30 | 12:46 | 11.29 | 7 | 41 | 8.76 | 92.21%
2002-04-30 | 13:01 | 11.4 | 7 | 41 | 8.76 | 92.21%
2002-04-30 | 13:16 | 11.43 | 7.01 | 41 | 8.78 | 92.42%
2002-04-30 | 13:31 | 11.46 | 7 | 41 | 8.78 | 92.42%
2002-04-30 | 13:46 | 11.55 | 7.01 | 41 | 8.79 | 93.02%
2002-04-30 | 14:01 | 11.58 | 7.01 | 41 | 8.79 | 93.02%
2002-04-30 | 14:16 | 11.63 | 7.01 | 41 | 8.79 | 93.02%
2002-04-30 | 14:31 | 11.67 | 7.01 | 41 | 8.8 | 93.12%
2002-04-30 | 14:46 | 11.74 | 7.01 | 41 | 8.8 | 93.12%
2002-04-30 | 15:01 | 11.78 | 7.02 | 41 | 8.81 | 93.23%
2002-04-30 | 15:16 | 11.83 | 7.02 | 41 | 8.81 | 93.23%
2002-04-30 | 15:31 | 11.89 | 7.02 | 41 | 8.82 | 93.33%
2002-04-30 | 15:46 | 11.95 | 7.02 | 41 | 8.82 | 93.33%
2002-04-30 | 16:01 | 11.94 | 7.02 | 41 | 8.82 | 93.33%
2002-04-30 | 16:16 | 11.96 | 7.03 | 41 | 8.82 | 93.33%
2002-04-30 | 16:31 | 11.99 | 7.02 | 41 | 8.82 | 93.33%
2002-04-30 | 16:46 | 12.01 | 7.03 | 41 | 8.82 | 93.83%
2002-04-30 | 17:01 | 12.03 | 7.03 | 41 | 8.82 | 93.83%
2002-04-30 | 17:16 | 12.01 | 7.03 | 41 | 8.82 | 93.83%
2002-04-30 | 17:31 | 12.01 | 7.02 | 41 | 8.82 | 93.83%
2002-04-30 | 17:46 | 12.02 | 7.02 | 41 | 8.82 | 93.83%
2002-04-30 | 18:01 | 12.04 | 7.03 | 41 | 8.81 | 93.72%
2002-04-30 | 18:16 | 12.03 | 7.03 | 41 | 8.81 | 93.72%
2002-04-30 | 18:31 | 12.01 | 7.02 | 41 | 8.8 | 93.62%
2002-04-30 | 18:46 | 11.97 | 7.02 | 41 | 8.8 | 93.12%
2002-04-30 | 19:01 | 11.92 | 7.02 | 41 | 8.79 | 93.02%
2002-04-30 | 19:16 | 11.87 | 7.02 | 41 | 8.79 | 93.02%
2002-04-30 | 19:31 | 11.83 | 7.02 | 41 | 8.77 | 92.8%
2002-04-30 | 19:46 | 11.77 | 7.02 | 41 | 8.76 | 92.7%
2002-04-30 | 20:01 | 11.72 | 7.01 | 41 | 8.75 | 92.59%
2002-04-30 | 20:16 | 11.67 | 7.01 | 41 | 8.73 | 92.38%
2002-04-30 | 20:31 | 11.64 | 7.01 | 40 | 8.73 | 92.38%
2002-04-30 | 20:46 | 11.6 | 7 | 41 | 8.73 | 92.38%
2002-04-30 | 21:01 | 11.55 | 7.01 | 40 | 8.71 | 92.17%
2002-04-30 | 21:16 | 11.5 | 7 | 40 | 8.71 | 92.17%
2002-04-30 | 21:31 | 11.46 | 7 | 40 | 8.68 | 91.37%
2002-04-30 | 21:46 | 11.42 | 6.99 | 40 | 8.68 | 91.37%
2002-04-30 | 22:01 | 11.39 | 7 | 40 | 8.66 | 91.16%
2002-04-30 | 22:16 | 11.36 | 7 | 40 | 8.65 | 91.05%
2002-04-30 | 22:31 | 11.34 | 6.99 | 40 | 8.65 | 91.05%
2002-04-30 | 22:46 | 11.31 | 6.99 | 40 | 8.64 | 90.95%
2002-04-30 | 23:01 | 11.29 | 6.99 | 40 | 8.65 | 91.05%
2002-04-30 | 23:16 | 11.26 | 6.99 | 40 | 8.63 | 90.84%
2002-04-30 | 23:31 | 11.23 | 6.98 | 40 | 8.63 | 90.84%
2002-04-30 | 23:46 | 11.2 | 6.99 | 40 | 8.63 | 90.84%
2002-05-01 | 0:01 | 11.16 | 6.99 | 40 | 8.62 | 90.74%
2002-05-01 | 0:16 | 11.12 | 6.98 | 40 | 8.61 | 90.63%
2002-05-01 | 0:31 | 11.08 | 6.97 | 40 | 8.61 | 90.63%
2002-05-01 | 0:46 | 11.03 | 6.98 | 40 | 8.61 | 90.63%
2002-05-01 | 1:01 | 10.98 | 6.96 | 40 | 8.61 | 89.69%
2002-05-01 | 1:16 | 10.94 | 6.97 | 40 | 8.61 | 89.69%
2002-05-01 | 1:31 | 10.89 | 6.97 | 40 | 8.61 | 89.69%
2002-05-01 | 1:46 | 10.84 | 6.96 | 40 | 8.6 | 89.58%
2002-05-01 | 2:01 | 10.79 | 6.96 | 40 | 8.6 | 89.58%
2002-05-01 | 2:16 | 10.74 | 6.97 | 40 | 8.6 | 89.58%
2002-05-01 | 2:31 | 10.7 | 6.97 | 40 | 8.61 | 89.69%
2002-05-01 | 2:46 | 10.65 | 6.96 | 40 | 8.61 | 89.69%
2002-05-01 | 3:01 | 10.61 | 6.96 | 40 | 8.61 | 89.69%
2002-05-01 | 3:16 | 10.57 | 6.97 | 40 | 8.61 | 89.69%
2002-05-01 | 3:31 | 10.53 | 6.96 | 40 | 8.61 | 89.69%
2002-05-01 | 3:46 | 10.49 | 6.96 | 40 | 8.6 | 88.66%
2002-05-01 | 4:01 | 10.45 | 6.95 | 40 | 8.59 | 88.56%
2002-05-01 | 4:16 | 10.41 | 6.96 | 40 | 8.6 | 88.66%
2002-05-01 | 4:31 | 10.38 | 6.96 | 40 | 8.59 | 88.56%
2002-05-01 | 4:46 | 10.34 | 6.95 | 40 | 8.6 | 88.66%
2002-05-01 | 5:01 | 10.31 | 6.95 | 40 | 8.6 | 88.66%
2002-05-01 | 5:16 | 10.28 | 6.95 | 40 | 8.6 | 88.66%
2002-05-01 | 5:31 | 10.24 | 6.96 | 40 | 8.6 | 88.66%
2002-05-01 | 5:46 | 10.21 | 6.95 | 40 | 8.6 | 88.66%
2002-05-01 | 6:01 | 10.17 | 6.96 | 40 | 8.61 | 88.76%
2002-05-01 | 6:16 | 10.14 | 6.95 | 40 | 8.6 | 88.66%
2002-05-01 | 6:31 | 10.1 | 6.95 | 40 | 8.61 | 88.76%
2002-05-01 | 6:46 | 10.06 | 6.95 | 40 | 8.61 | 88.76%
2002-05-01 | 7:01 | 10.03 | 6.95 | 40 | 8.62 | 88.87%
2002-05-01 | 7:16 | 10 | 6.96 | 40 | 8.62 | 88.87%
2002-05-01 | 7:31 | 9.98 | 6.95 | 40 | 8.63 | 87.61%
2002-05-01 | 7:46 | 9.96 | 6.96 | 40 | 8.64 | 87.72%
2002-05-01 | 8:01 | 9.95 | 6.95 | 40 | 8.66 | 87.92%
2002-05-01 | 8:16 | 9.96 | 6.95 | 40 | 8.67 | 88.02%
2002-05-01 | 8:31 | 9.98 | 6.96 | 40 | 8.68 | 88.12%
2002-05-01 | 8:46 | 10.01 | 6.96 | 40 | 8.7 | 89.69%
2002-05-01 | 9:01 | 10.05 | 6.96 | 40 | 8.72 | 89.9%
2002-05-01 | 9:16 | 10.11 | 6.97 | 41 | 8.73 | 90%
2002-05-01 | 9:31 | 10.17 | 6.96 | 40 | 8.75 | 90.21%
2002-05-01 | 9:46 | 10.24 | 6.97 | 40 | 8.77 | 90.41%
2002-05-01 | 10:01 | 10.31 | 6.96 | 41 | 8.78 | 90.52%
2002-05-01 | 10:16 | 10.39 | 6.97 | 40 | 8.8 | 90.72%
2002-05-01 | 10:31 | 10.46 | 6.97 | 41 | 8.81 | 90.82%
2002-05-01 | 10:46 | 10.54 | 6.98 | 41 | 8.83 | 91.98%
2002-05-01 | 11:01 | 10.63 | 6.98 | 40 | 8.85 | 92.19%
2002-05-01 | 11:16 | 10.71 | 6.99 | 41 | 8.87 | 92.4%
2002-05-01 | 11:31 | 10.81 | 6.99 | 41 | 8.89 | 92.6%
2002-05-01 | 11:46 | 10.9 | 7 | 41 | 8.9 | 92.71%
2002-05-01 | 12:01 | 11.01 | 6.98 | 40 | 8.92 | 93.89%
2002-05-01 | 12:16 | 11.11 | 6.99 | 41 | 8.92 | 93.89%
2002-05-01 | 12:31 | 11.21 | 7 | 41 | 8.93 | 94%
2002-05-01 | 12:46 | 11.33 | 7.01 | 41 | 8.94 | 94.11%
2002-05-01 | 13:01 | 11.44 | 7.01 | 41 | 8.96 | 94.32%
2002-05-01 | 13:16 | 11.55 | 7 | 41 | 8.96 | 94.81%
2002-05-01 | 13:31 | 11.65 | 7.01 | 40 | 8.97 | 94.92%
2002-05-01 | 13:46 | 11.76 | 7.01 | 41 | 8.98 | 95.03%
2002-05-01 | 14:01 | 11.86 | 7.02 | 41 | 8.98 | 95.03%
2002-05-01 | 14:16 | 11.96 | 7.02 | 40 | 9 | 95.24%
2002-05-01 | 14:31 | 12.06 | 7.03 | 41 | 9.01 | 95.85%
2002-05-01 | 14:46 | 12.16 | 7.03 | 41 | 9.01 | 95.85%
2002-05-01 | 15:01 | 12.25 | 7.02 | 40 | 9.02 | 95.96%
2002-05-01 | 15:16 | 12.33 | 7.03 | 41 | 9.02 | 95.96%
2002-05-01 | 15:31 | 12.41 | 7.03 | 41 | 9.01 | 95.85%
2002-05-01 | 15:46 | 12.47 | 7.03 | 40 | 9.02 | 95.96%
2002-05-01 | 16:01 | 12.53 | 7.03 | 40 | 9.03 | 97.62%
2002-05-01 | 16:16 | 12.59 | 7.04 | 40 | 9.02 | 97.51%
2002-05-01 | 16:31 | 12.64 | 7.05 | 40 | 9.01 | 97.41%
2002-05-01 | 16:46 | 12.69 | 7.04 | 40 | 9.01 | 97.41%
2002-05-01 | 17:01 | 12.73 | 7.04 | 40 | 9.01 | 97.41%
2002-05-01 | 17:16 | 12.76 | 7.04 | 40 | 9.01 | 97.41%
2002-05-01 | 17:31 | 12.79 | 7.05 | 40 | 9.01 | 97.41%
2002-05-01 | 17:46 | 12.8 | 7.04 | 40 | 8.99 | 97.19%
2002-05-01 | 18:01 | 12.79 | 7.04 | 40 | 8.98 | 97.08%
2002-05-01 | 18:16 | 12.79 | 7.04 | 40 | 8.97 | 96.97%
2002-05-01 | 18:31 | 12.77 | 7.04 | 40 | 8.97 | 96.97%
2002-05-01 | 18:46 | 12.74 | 7.04 | 40 | 8.97 | 96.97%
2002-05-01 | 19:01 | 12.72 | 7.04 | 40 | 8.94 | 96.65%
2002-05-01 | 19:16 | 12.69 | 7.03 | 40 | 8.93 | 96.54%
2002-05-01 | 19:31 | 12.65 | 7.03 | 40 | 8.91 | 96.32%
2002-05-01 | 19:46 | 12.62 | 7.03 | 40 | 8.9 | 96.22%
2002-05-01 | 20:01 | 12.61 | 7.03 | 40 | 8.88 | 96%
2002-05-01 | 20:16 | 12.6 | 7.03 | 40 | 8.86 | 95.78%
2002-05-01 | 20:31 | 12.58 | 7.02 | 40 | 8.84 | 95.57%
2002-05-01 | 20:46 | 12.56 | 7.02 | 40 | 8.83 | 95.46%
2002-05-01 | 21:01 | 12.55 | 7.02 | 40 | 8.8 | 95.14%
2002-05-01 | 21:16 | 12.54 | 7.02 | 40 | 8.79 | 95.03%
2002-05-01 | 21:31 | 12.53 | 7.02 | 40 | 8.78 | 94.92%
2002-05-01 | 21:46 | 12.51 | 7.01 | 40 | 8.75 | 94.59%
2002-05-01 | 22:01 | 12.5 | 7.01 | 40 | 8.73 | 94.38%
2002-05-01 | 22:16 | 12.49 | 7.01 | 40 | 8.72 | 92.77%
2002-05-01 | 22:31 | 12.47 | 7.02 | 40 | 8.7 | 92.55%
2002-05-01 | 22:46 | 12.46 | 7 | 40 | 8.68 | 92.34%
2002-05-01 | 23:01 | 12.44 | 7.01 | 40 | 8.67 | 92.23%
2002-05-01 | 23:16 | 12.41 | 7 | 40 | 8.65 | 92.02%
2002-05-01 | 23:31 | 12.38 | 7 | 40 | 8.64 | 91.91%
2002-05-01 | 23:46 | 12.35 | 6.99 | 40 | 8.63 | 91.81%
2002-05-02 | 0:01 | 12.31 | 6.99 | 40 | 8.61 | 91.6%
2002-05-02 | 0:16 | 12.28 | 6.98 | 40 | 8.6 | 91.49%
2002-05-02 | 0:31 | 12.24 | 7 | 40 | 8.6 | 91.49%
2002-05-02 | 0:46 | 12.2 | 6.99 | 40 | 8.59 | 91.38%
2002-05-02 | 1:01 | 12.16 | 6.98 | 40 | 8.58 | 91.28%
2002-05-02 | 1:16 | 12.12 | 6.98 | 40 | 8.57 | 91.17%
2002-05-02 | 1:31 | 12.08 | 6.98 | 40 | 8.56 | 91.06%
2002-05-02 | 1:46 | 12.04 | 6.97 | 40 | 8.55 | 90.96%
2002-05-02 | 2:01 | 12.01 | 6.98 | 40 | 8.54 | 90.85%
2002-05-02 | 2:16 | 11.97 | 6.96 | 40 | 8.55 | 90.48%
2002-05-02 | 2:31 | 11.94 | 6.97 | 40 | 8.54 | 90.37%
2002-05-02 | 2:46 | 11.9 | 6.96 | 40 | 8.53 | 90.26%
2002-05-02 | 3:01 | 11.86 | 6.98 | 40 | 8.53 | 90.26%
2002-05-02 | 3:16 | 11.83 | 6.96 | 40 | 8.52 | 90.16%
2002-05-02 | 3:31 | 11.79 | 6.97 | 40 | 8.51 | 90.05%
2002-05-02 | 3:46 | 11.75 | 6.96 | 40 | 8.51 | 90.05%
2002-05-02 | 4:01 | 11.71 | 6.96 | 40 | 8.5 | 89.95%
2002-05-02 | 4:16 | 11.67 | 6.96 | 40 | 8.5 | 89.95%
2002-05-02 | 4:31 | 11.63 | 6.97 | 40 | 8.49 | 89.84%
2002-05-02 | 4:46 | 11.6 | 6.96 | 40 | 8.49 | 89.84%
2002-05-02 | 5:01 | 11.56 | 6.95 | 40 | 8.48 | 89.74%
2002-05-02 | 5:16 | 11.52 | 6.96 | 40 | 8.49 | 89.84%
2002-05-02 | 5:31 | 11.48 | 6.96 | 40 | 8.48 | 89.26%
2002-05-02 | 5:46 | 11.44 | 6.95 | 40 | 8.48 | 89.26%
2002-05-02 | 6:01 | 11.41 | 6.96 | 40 | 8.48 | 89.26%
2002-05-02 | 6:16 | 11.38 | 6.96 | 40 | 8.47 | 89.16%
2002-05-02 | 6:31 | 11.35 | 6.95 | 40 | 8.47 | 89.16%
2002-05-02 | 6:46 | 11.32 | 6.95 | 40 | 8.47 | 89.16%
2002-05-02 | 7:01 | 11.3 | 6.95 | 40 | 8.48 | 89.26%
2002-05-02 | 7:16 | 11.28 | 6.96 | 41 | 8.49 | 89.37%
2002-05-02 | 7:31 | 11.27 | 6.97 | 41 | 8.49 | 89.37%
2002-05-02 | 7:46 | 11.27 | 6.96 | 40 | 8.5 | 89.47%
2002-05-02 | 8:01 | 11.27 | 6.97 | 40 | 8.51 | 89.58%
2002-05-02 | 8:16 | 11.29 | 6.97 | 41 | 8.53 | 89.79%
2002-05-02 | 8:31 | 11.31 | 6.97 | 40 | 8.54 | 89.89%
2002-05-02 | 8:46 | 11.34 | 6.97 | 41 | 8.56 | 90.11%
2002-05-02 | 9:01 | 11.39 | 6.97 | 41 | 8.56 | 90.11%
2002-05-02 | 9:16 | 11.44 | 6.97 | 41 | 8.57 | 90.21%
2002-05-02 | 9:31 | 11.5 | 6.98 | 41 | 8.59 | 90.9%
2002-05-02 | 9:46 | 11.57 | 6.97 | 41 | 8.6 | 91.01%
2002-05-02 | 10:01 | 11.63 | 6.98 | 41 | 8.62 | 91.22%
2002-05-02 | 10:16 | 11.7 | 6.98 | 41 | 8.64 | 91.43%
2002-05-02 | 10:31 | 11.78 | 6.98 | 41 | 8.66 | 91.64%
2002-05-02 | 10:46 | 11.86 | 6.99 | 41 | 8.67 | 91.75%
2002-05-02 | 11:01 | 11.95 | 7 | 41 | 8.69 | 91.96%
2002-05-02 | 11:16 | 12.05 | 7 | 41 | 8.7 | 92.55%
2002-05-02 | 11:31 | 12.16 | 7.01 | 41 | 8.72 | 92.77%
2002-05-02 | 11:46 | 12.27 | 7.01 | 41 | 8.74 | 92.98%
2002-05-02 | 12:01 | 12.38 | 7.02 | 41 | 8.75 | 93.09%
2002-05-02 | 12:16 | 12.49 | 7.01 | 41 | 8.78 | 93.4%
2002-05-02 | 12:31 | 12.6 | 7.02 | 41 | 8.77 | 94.81%
2002-05-02 | 12:46 | 12.71 | 7.02 | 41 | 8.78 | 94.92%
2002-05-02 | 13:01 | 12.81 | 7.03 | 41 | 8.79 | 95.03%
2002-05-02 | 13:16 | 12.92 | 7.02 | 41 | 8.8 | 95.14%
2002-05-02 | 13:31 | 13.02 | 7.02 | 41 | 8.82 | 96.92%
2002-05-02 | 13:46 | 13.12 | 7.02 | 41 | 8.85 | 97.25%
2002-05-02 | 14:01 | 13.22 | 7.03 | 41 | 8.84 | 97.14%
2002-05-02 | 14:16 | 13.33 | 7.03 | 41 | 8.84 | 97.14%
2002-05-02 | 14:31 | 13.42 | 7.04 | 41 | 8.86 | 97.36%
2002-05-02 | 14:46 | 13.51 | 7.04 | 41 | 8.84 | 98.22%
2002-05-02 | 15:01 | 13.6 | 7.05 | 41 | 8.85 | 98.33%
2002-05-02 | 15:16 | 13.69 | 7.05 | 41 | 8.85 | 98.33%
2002-05-02 | 15:31 | 13.77 | 7.05 | 41 | 8.86 | 98.44%
2002-05-02 | 15:46 | 13.84 | 7.05 | 41 | 8.87 | 98.56%
2002-05-02 | 16:01 | 13.91 | 7.06 | 41 | 8.85 | 98.33%
2002-05-02 | 16:16 | 13.97 | 7.06 | 41 | 8.86 | 98.44%
2002-05-02 | 16:31 | 14.01 | 7.07 | 41 | 8.85 | 99.44%
2002-05-02 | 16:46 | 14.05 | 7.06 | 41 | 8.85 | 99.44%
2002-05-02 | 17:01 | 14.08 | 7.06 | 41 | 8.85 | 99.44%
2002-05-02 | 17:16 | 14.1 | 7.07 | 41 | 8.86 | 99.55%
2002-05-02 | 17:31 | 14.11 | 7.06 | 41 | 8.83 | 99.21%
2002-05-02 | 17:46 | 14.12 | 7.07 | 41 | 8.84 | 99.33%
2002-05-02 | 18:01 | 14.12 | 7.07 | 41 | 8.84 | 99.33%
2002-05-02 | 18:16 | 14.1 | 7.07 | 41 | 8.81 | 98.99%
2002-05-02 | 18:31 | 14.08 | 7.06 | 41 | 8.81 | 98.99%
2002-05-02 | 18:46 | 14.04 | 7.06 | 41 | 8.8 | 98.88%
2002-05-02 | 19:01 | 13.99 | 7.06 | 41 | 8.78 | 97.56%
2002-05-02 | 19:16 | 13.93 | 7.06 | 41 | 8.76 | 97.33%
2002-05-02 | 19:31 | 13.89 | 7.06 | 41 | 8.75 | 97.22%
2002-05-02 | 19:46 | 13.84 | 7.06 | 41 | 8.73 | 97%
2002-05-02 | 20:01 | 13.77 | 7.06 | 41 | 8.71 | 96.78%
2002-05-02 | 20:16 | 13.71 | 7.05 | 41 | 8.7 | 96.67%
2002-05-02 | 20:31 | 13.65 | 7.04 | 41 | 8.68 | 96.44%
2002-05-02 | 20:46 | 13.6 | 7.05 | 41 | 8.66 | 96.22%
2002-05-02 | 21:01 | 13.55 | 7.04 | 41 | 8.65 | 96.11%
2002-05-02 | 21:16 | 13.51 | 7.04 | 41 | 8.64 | 96%
2002-05-02 | 21:31 | 13.46 | 7.04 | 40 | 8.62 | 94.73%
2002-05-02 | 21:46 | 13.42 | 7.03 | 41 | 8.61 | 94.62%
2002-05-02 | 22:01 | 13.37 | 7.03 | 41 | 8.6 | 94.51%
2002-05-02 | 22:16 | 13.32 | 7.03 | 41 | 8.59 | 94.4%
2002-05-02 | 22:31 | 13.26 | 7.03 | 40 | 8.58 | 94.29%
2002-05-02 | 22:46 | 13.22 | 7.02 | 41 | 8.57 | 94.18%
2002-05-02 | 23:01 | 13.18 | 7.02 | 40 | 8.56 | 94.07%
2002-05-02 | 23:16 | 13.14 | 7.02 | 41 | 8.54 | 93.85%
2002-05-02 | 23:31 | 13.1 | 7.02 | 41 | 8.53 | 93.74%
2002-05-02 | 23:46 | 13.07 | 7.02 | 41 | 8.52 | 93.63%
2002-05-03 | 0:01 | 13.02 | 7.01 | 41 | 8.51 | 93.52%
2002-05-03 | 0:16 | 12.99 | 7 | 41 | 8.5 | 91.89%
2002-05-03 | 0:31 | 12.96 | 7 | 41 | 8.5 | 91.89%
2002-05-03 | 0:46 | 12.92 | 7 | 41 | 8.48 | 91.68%
2002-05-03 | 1:01 | 12.88 | 7 | 41 | 8.46 | 91.46%
2002-05-03 | 1:16 | 12.85 | 7 | 41 | 8.44 | 91.24%
2002-05-03 | 1:31 | 12.82 | 6.99 | 41 | 8.45 | 91.35%
2002-05-03 | 1:46 | 12.79 | 7 | 41 | 8.45 | 91.35%
2002-05-03 | 2:01 | 12.76 | 7 | 41 | 8.41 | 90.92%
2002-05-03 | 2:16 | 12.72 | 7 | 41 | 8.41 | 90.92%
2002-05-03 | 2:31 | 12.69 | 6.99 | 41 | 8.42 | 91.03%
2002-05-03 | 2:46 | 12.66 | 6.99 | 41 | 8.41 | 90.92%
2002-05-03 | 3:01 | 12.62 | 6.99 | 41 | 8.4 | 90.81%
2002-05-03 | 3:16 | 12.57 | 6.98 | 41 | 8.38 | 90.59%
2002-05-03 | 3:31 | 12.52 | 6.99 | 41 | 8.37 | 90.49%
2002-05-03 | 3:46 | 12.48 | 6.99 | 41 | 8.36 | 88.94%
2002-05-03 | 4:01 | 12.44 | 6.98 | 41 | 8.36 | 88.94%
2002-05-03 | 4:16 | 12.42 | 6.99 | 41 | 8.36 | 88.94%
2002-05-03 | 4:31 | 12.39 | 6.98 | 41 | 8.36 | 88.94%
2002-05-03 | 4:46 | 12.36 | 6.99 | 41 | 8.35 | 88.83%
2002-05-03 | 5:01 | 12.33 | 6.98 | 41 | 8.34 | 88.72%
2002-05-03 | 5:16 | 12.29 | 6.97 | 41 | 8.33 | 88.62%
2002-05-03 | 5:31 | 12.26 | 6.98 | 41 | 8.34 | 88.72%
2002-05-03 | 5:46 | 12.23 | 6.99 | 41 | 8.32 | 88.51%
2002-05-03 | 6:01 | 12.19 | 6.97 | 41 | 8.33 | 88.62%
2002-05-03 | 6:16 | 12.16 | 6.99 | 41 | 8.32 | 88.51%
2002-05-03 | 6:31 | 12.13 | 6.97 | 41 | 8.34 | 88.72%
2002-05-03 | 6:46 | 12.09 | 6.98 | 41 | 8.34 | 88.72%
2002-05-03 | 7:01 | 12.06 | 6.98 | 41 | 8.34 | 88.72%
2002-05-03 | 7:16 | 12.02 | 6.97 | 41 | 8.35 | 88.83%
2002-05-03 | 7:31 | 11.99 | 6.97 | 41 | 8.34 | 88.25%
2002-05-03 | 7:46 | 11.97 | 6.99 | 41 | 8.36 | 88.47%
2002-05-03 | 8:01 | 11.96 | 6.99 | 42 | 8.37 | 88.57%
2002-05-03 | 8:16 | 11.97 | 6.98 | 41 | 8.38 | 88.68%
2002-05-03 | 8:31 | 11.98 | 6.99 | 41 | 8.4 | 88.89%
2002-05-03 | 8:46 | 12.01 | 6.99 | 42 | 8.41 | 89.47%
2002-05-03 | 9:01 | 12.04 | 7 | 42 | 8.44 | 89.79%
2002-05-03 | 9:16 | 12.09 | 6.99 | 42 | 8.45 | 89.89%
2002-05-03 | 9:31 | 12.15 | 6.99 | 42 | 8.47 | 90.11%