## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2002-04-29 13:50 | 8.7 | 6.9 | 9.5 | 95 | 40 | 0
Sonde | 2002-04-29 15:16 | 8.74 | 7.29 | 9.78 | 98 | 37
Grade |  | A | B | A |  | A
Audit 2 | 2002-05-01 12:45 | 10.2 | 7.3 | 9.5 | 99 | 42 | 0
Sonde | 2002-05-01 12:46 | 10.07 | 7.23 | 9.69 | 101 | 38
Grade |  | A | A | A |  | A
Audit 3 | 2002-05-03 10:30 | 9.6 | 7.3 | 9.2 | 95 | 41 | 0
Sonde | 2002-05-03 10:31 | 9.61 | 7.15 | 9.42 | 97 | 39
Grade |  | A | A | A |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note Audits 1, 2, & 3 are macro(ed) into this audit sheet and represent the beginning, middle and end of the deployment. Audits 4, 5 and 6 were in the middle of the deployment and are not macro(ed) into this sheet. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-29 13:50 | 8.7 | 6.9 | 9.5 | 95 | 40
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-01 12:45 | 10.2 | 7.3 | 9.5 | 99 | 42
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-03 10:30 | 9.6 | 7.3 | 9.2 | 95 | 41
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-30 10:10 | 7.8 | 7.2 | 9.7 | 95 | 42
Sonde Values | 2002-04-30 10:16 | 7.35 | 7.18 | 9.84 | 95 | 38
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 10:50 | 9.9 | 7.4 | 9.2 | 95 | 42
Sonde Values | 2002-05-02 10:46 | 9.4 | 9.2 | 9.5 | 96 | 39
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 6
Organization | ODEQ |  | Station ID |  | LDR 064.75
Program Name | Water Quality Monitoring Program |  | Description |  | Little Deschutes River, u/s of Gilchrist Mill Pond
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 15:35 | 12.6 | 7.4 | 9 | 99 | 41
Sonde Values | 2002-05-02 15:31 | 12.2 | 7.3 | 9.3 | 101 | 39
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO Conc | DO
|  | °C |  | uS/cm | mg/L | % Sat
2002-04-29 | 14:16 | 8.65 | 7.37 | 37 | 9.86 | 98.6%
2002-04-29 | 14:31 | 8.66 | 7.34 | 37 | 9.83 | 98.3%
2002-04-29 | 14:46 | 8.71 | 7.32 | 37 | 9.81 | 98.1%
2002-04-29 | 15:01 | 8.76 | 7.3 | 37 | 9.8 | 98%
2002-04-29 | 15:16 | 8.74 | 7.29 | 37 | 9.78 | 97.8%
2002-04-29 | 15:31 | 8.71 | 7.28 | 37 | 9.77 | 97.7%
2002-04-29 | 15:46 | 8.67 | 7.27 | 37 | 9.76 | 97.6%
2002-04-29 | 16:01 | 8.61 | 7.27 | 37 | 9.75 | 97.5%
2002-04-29 | 16:16 | 8.57 | 7.26 | 37 | 9.75 | 97.5%
2002-04-29 | 16:31 | 8.56 | 7.26 | 37 | 9.75 | 97.5%
2002-04-29 | 16:46 | 8.55 | 7.26 | 37 | 9.76 | 97.6%
2002-04-29 | 17:01 | 8.53 | 7.26 | 37 | 9.76 | 97.6%
2002-04-29 | 17:16 | 8.51 | 7.25 | 37 | 9.76 | 97.6%
2002-04-29 | 17:31 | 8.47 | 7.25 | 37 | 9.76 | 96.63%
2002-04-29 | 17:46 | 8.41 | 7.25 | 37 | 9.76 | 96.63%
2002-04-29 | 18:01 | 8.41 | 7.25 | 37 | 9.77 | 96.73%
2002-04-29 | 18:16 | 8.36 | 7.24 | 37 | 9.77 | 96.73%
2002-04-29 | 18:31 | 8.29 | 7.24 | 37 | 9.78 | 96.83%
2002-04-29 | 18:46 | 8.23 | 7.25 | 37 | 9.77 | 96.73%
2002-04-29 | 19:01 | 8.17 | 7.24 | 37 | 9.77 | 96.73%
2002-04-29 | 19:16 | 8.09 | 7.23 | 37 | 9.76 | 96.63%
2002-04-29 | 19:31 | 8.02 | 7.23 | 37 | 9.76 | 96.63%
2002-04-29 | 19:46 | 7.96 | 7.23 | 37 | 9.75 | 95.12%
2002-04-29 | 20:01 | 7.89 | 7.23 | 37 | 9.74 | 95.02%
2002-04-29 | 20:16 | 7.83 | 7.22 | 37 | 9.74 | 95.02%
2002-04-29 | 20:31 | 7.77 | 7.22 | 37 | 9.73 | 94.93%
2002-04-29 | 20:46 | 7.72 | 7.22 | 37 | 9.71 | 94.73%
2002-04-29 | 21:01 | 7.68 | 7.21 | 37 | 9.7 | 94.63%
2002-04-29 | 21:16 | 7.64 | 7.21 | 37 | 9.69 | 94.54%
2002-04-29 | 21:31 | 7.6 | 7.2 | 37 | 9.69 | 94.54%
2002-04-29 | 21:46 | 7.55 | 7.2 | 37 | 9.68 | 94.44%
2002-04-29 | 22:01 | 7.51 | 7.2 | 37 | 9.67 | 94.34%
2002-04-29 | 22:16 | 7.48 | 7.19 | 37 | 9.66 | 92.88%
2002-04-29 | 22:31 | 7.45 | 7.19 | 37 | 9.65 | 92.79%
2002-04-29 | 22:46 | 7.43 | 7.19 | 37 | 9.65 | 92.79%
2002-04-29 | 23:01 | 7.41 | 7.19 | 37 | 9.64 | 92.69%
2002-04-29 | 23:16 | 7.38 | 7.18 | 37 | 9.64 | 92.69%
2002-04-29 | 23:31 | 7.36 | 7.19 | 37 | 9.63 | 92.6%
2002-04-29 | 23:46 | 7.34 | 7.18 | 37 | 9.62 | 92.5%
2002-04-30 | 0:01 | 7.33 | 7.18 | 37 | 9.62 | 92.5%
2002-04-30 | 0:16 | 7.31 | 7.18 | 37 | 9.61 | 92.4%
2002-04-30 | 0:31 | 7.3 | 7.18 | 37 | 9.61 | 92.4%
2002-04-30 | 0:46 | 7.29 | 7.18 | 37 | 9.6 | 92.31%
2002-04-30 | 1:01 | 7.28 | 7.17 | 37 | 9.6 | 92.31%
2002-04-30 | 1:16 | 7.27 | 7.17 | 37 | 9.59 | 92.21%
2002-04-30 | 1:31 | 7.25 | 7.17 | 37 | 9.59 | 92.21%
2002-04-30 | 1:46 | 7.24 | 7.17 | 37 | 9.59 | 92.21%
2002-04-30 | 2:01 | 7.23 | 7.17 | 37 | 9.59 | 92.21%
2002-04-30 | 2:16 | 7.21 | 7.16 | 37 | 9.58 | 92.12%
2002-04-30 | 2:31 | 7.2 | 7.16 | 37 | 9.58 | 92.12%
2002-04-30 | 2:46 | 7.18 | 7.17 | 37 | 9.58 | 92.12%
2002-04-30 | 3:01 | 7.16 | 7.16 | 38 | 9.58 | 92.12%
2002-04-30 | 3:16 | 7.13 | 7.16 | 38 | 9.58 | 92.12%
2002-04-30 | 3:31 | 7.09 | 7.15 | 37 | 9.58 | 92.12%
2002-04-30 | 3:46 | 7.06 | 7.15 | 37 | 9.58 | 92.12%
2002-04-30 | 4:01 | 7.03 | 7.15 | 37 | 9.58 | 92.12%
2002-04-30 | 4:16 | 7 | 7.15 | 37 | 9.58 | 92.12%
2002-04-30 | 4:31 | 6.96 | 7.15 | 38 | 9.58 | 91.24%
2002-04-30 | 4:46 | 6.92 | 7.14 | 38 | 9.59 | 91.33%
2002-04-30 | 5:01 | 6.88 | 7.14 | 38 | 9.59 | 91.33%
2002-04-30 | 5:16 | 6.85 | 7.14 | 38 | 9.6 | 91.43%
2002-04-30 | 5:31 | 6.81 | 7.13 | 38 | 9.6 | 91.43%
2002-04-30 | 5:46 | 6.76 | 7.13 | 38 | 9.61 | 91.52%
2002-04-30 | 6:01 | 6.71 | 7.13 | 38 | 9.61 | 91.52%
2002-04-30 | 6:16 | 6.66 | 7.13 | 38 | 9.62 | 91.62%
2002-04-30 | 6:31 | 6.61 | 7.14 | 38 | 9.64 | 91.81%
2002-04-30 | 6:46 | 6.58 | 7.14 | 38 | 9.65 | 91.9%
2002-04-30 | 7:01 | 6.56 | 7.14 | 38 | 9.66 | 92%
2002-04-30 | 7:16 | 6.56 | 7.13 | 38 | 9.68 | 92.19%
2002-04-30 | 7:31 | 6.58 | 7.14 | 38 | 9.7 | 92.38%
2002-04-30 | 7:46 | 6.62 | 7.14 | 38 | 9.72 | 92.57%
2002-04-30 | 8:01 | 6.68 | 7.15 | 38 | 9.73 | 92.67%
2002-04-30 | 8:16 | 6.74 | 7.15 | 38 | 9.75 | 92.86%
2002-04-30 | 8:31 | 6.8 | 7.15 | 38 | 9.77 | 93.05%
2002-04-30 | 8:46 | 6.85 | 7.16 | 38 | 9.78 | 93.14%
2002-04-30 | 9:01 | 6.93 | 7.16 | 38 | 9.79 | 93.24%
2002-04-30 | 9:16 | 6.99 | 7.17 | 38 | 9.81 | 93.43%
2002-04-30 | 9:31 | 7.06 | 7.17 | 38 | 9.82 | 94.42%
2002-04-30 | 9:46 | 7.15 | 7.18 | 38 | 9.83 | 94.52%
2002-04-30 | 10:01 | 7.24 | 7.18 | 38 | 9.83 | 94.52%
2002-04-30 | 10:16 | 7.35 | 7.18 | 38 | 9.84 | 94.62%
2002-04-30 | 10:31 | 7.52 | 7.19 | 38 | 9.85 | 96.1%
2002-04-30 | 10:46 | 7.65 | 7.2 | 38 | 9.84 | 96%
2002-04-30 | 11:01 | 7.82 | 7.2 | 38 | 9.84 | 96%
2002-04-30 | 11:16 | 8.04 | 7.21 | 38 | 9.84 | 97.43%
2002-04-30 | 11:31 | 8.17 | 7.21 | 38 | 9.84 | 97.43%
2002-04-30 | 11:46 | 8.32 | 7.21 | 38 | 9.83 | 97.33%
2002-04-30 | 12:01 | 8.44 | 7.22 | 38 | 9.82 | 97.23%
2002-04-30 | 12:16 | 8.63 | 7.22 | 38 | 9.82 | 98.2%
2002-04-30 | 12:31 | 8.75 | 7.23 | 38 | 9.81 | 98.1%
2002-04-30 | 12:46 | 8.95 | 7.23 | 38 | 9.8 | 98%
2002-04-30 | 13:01 | 9.11 | 7.24 | 38 | 9.79 | 99.39%
2002-04-30 | 13:16 | 9.29 | 7.24 | 38 | 9.78 | 99.29%
2002-04-30 | 13:31 | 9.43 | 7.24 | 38 | 9.77 | 99.19%
2002-04-30 | 13:46 | 9.67 | 7.25 | 38 | 9.75 | 100.5%
2002-04-30 | 14:01 | 9.89 | 7.25 | 38 | 9.74 | 100.4%
2002-04-30 | 14:16 | 10.05 | 7.26 | 38 | 9.72 | 101.2%
2002-04-30 | 14:31 | 10.19 | 7.26 | 38 | 9.71 | 101.1%
2002-04-30 | 14:46 | 10.37 | 7.26 | 38 | 9.69 | 100.9%
2002-04-30 | 15:01 | 10.51 | 7.27 | 38 | 9.67 | 101.8%
2002-04-30 | 15:16 | 10.64 | 7.27 | 38 | 9.66 | 101.7%
2002-04-30 | 15:31 | 10.68 | 7.27 | 38 | 9.64 | 101.5%
2002-04-30 | 15:46 | 10.77 | 7.27 | 38 | 9.62 | 101.3%
2002-04-30 | 16:01 | 10.78 | 7.28 | 38 | 9.6 | 101.1%
2002-04-30 | 16:16 | 10.82 | 7.27 | 38 | 9.59 | 100.9%
2002-04-30 | 16:31 | 10.73 | 7.28 | 38 | 9.57 | 100.7%
2002-04-30 | 16:46 | 10.66 | 7.28 | 38 | 9.55 | 100.5%
2002-04-30 | 17:01 | 10.62 | 7.28 | 38 | 9.53 | 100.3%
2002-04-30 | 17:16 | 10.53 | 7.28 | 38 | 9.52 | 100.2%
2002-04-30 | 17:31 | 10.41 | 7.28 | 38 | 9.5 | 98.96%
2002-04-30 | 17:46 | 10.29 | 7.27 | 38 | 9.49 | 98.85%
2002-04-30 | 18:01 | 10.2 | 7.27 | 38 | 9.47 | 98.65%
2002-04-30 | 18:16 | 10.13 | 7.26 | 38 | 9.47 | 98.65%
2002-04-30 | 18:31 | 10.05 | 7.27 | 38 | 9.47 | 98.65%
2002-04-30 | 18:46 | 9.94 | 7.26 | 38 | 9.47 | 97.63%
2002-04-30 | 19:01 | 9.84 | 7.26 | 38 | 9.47 | 97.63%
2002-04-30 | 19:16 | 9.72 | 7.27 | 38 | 9.46 | 97.53%
2002-04-30 | 19:31 | 9.59 | 7.26 | 38 | 9.45 | 97.42%
2002-04-30 | 19:46 | 9.46 | 7.25 | 38 | 9.44 | 95.84%
2002-04-30 | 20:01 | 9.35 | 7.25 | 38 | 9.44 | 95.84%
2002-04-30 | 20:16 | 9.23 | 7.25 | 38 | 9.43 | 95.74%
2002-04-30 | 20:31 | 9.12 | 7.24 | 38 | 9.42 | 95.63%
2002-04-30 | 20:46 | 9.02 | 7.24 | 38 | 9.42 | 95.63%
2002-04-30 | 21:01 | 8.93 | 7.23 | 38 | 9.41 | 94.1%
2002-04-30 | 21:16 | 8.84 | 7.23 | 38 | 9.4 | 94%
2002-04-30 | 21:31 | 8.77 | 7.22 | 38 | 9.4 | 94%
2002-04-30 | 21:46 | 8.71 | 7.22 | 38 | 9.39 | 93.9%
2002-04-30 | 22:01 | 8.64 | 7.22 | 38 | 9.39 | 93.9%
2002-04-30 | 22:16 | 8.58 | 7.21 | 38 | 9.38 | 93.8%
2002-04-30 | 22:31 | 8.52 | 7.21 | 38 | 9.38 | 93.8%
2002-04-30 | 22:46 | 8.46 | 7.2 | 38 | 9.38 | 92.87%
2002-04-30 | 23:01 | 8.4 | 7.2 | 38 | 9.38 | 92.87%
2002-04-30 | 23:16 | 8.33 | 7.2 | 38 | 9.37 | 92.77%
2002-04-30 | 23:31 | 8.26 | 7.19 | 38 | 9.38 | 92.87%
2002-04-30 | 23:46 | 8.2 | 7.19 | 38 | 9.38 | 92.87%
2002-05-01 | 0:01 | 8.14 | 7.19 | 38 | 9.38 | 92.87%
2002-05-01 | 0:16 | 8.09 | 7.18 | 38 | 9.38 | 92.87%
2002-05-01 | 0:31 | 8.02 | 7.18 | 38 | 9.38 | 92.87%
2002-05-01 | 0:46 | 7.96 | 7.18 | 38 | 9.38 | 91.51%
2002-05-01 | 1:01 | 7.91 | 7.17 | 38 | 9.39 | 91.61%
2002-05-01 | 1:16 | 7.85 | 7.17 | 38 | 9.39 | 91.61%
2002-05-01 | 1:31 | 7.81 | 7.17 | 38 | 9.39 | 91.61%
2002-05-01 | 1:46 | 7.76 | 7.17 | 38 | 9.39 | 91.61%
2002-05-01 | 2:01 | 7.72 | 7.17 | 38 | 9.39 | 91.61%
2002-05-01 | 2:16 | 7.67 | 7.16 | 38 | 9.39 | 91.61%
2002-05-01 | 2:31 | 7.63 | 7.16 | 38 | 9.4 | 91.71%
2002-05-01 | 2:46 | 7.59 | 7.16 | 38 | 9.4 | 91.71%
2002-05-01 | 3:01 | 7.55 | 7.16 | 38 | 9.4 | 91.71%
2002-05-01 | 3:16 | 7.5 | 7.16 | 38 | 9.41 | 91.8%
2002-05-01 | 3:31 | 7.47 | 7.15 | 38 | 9.41 | 90.48%
2002-05-01 | 3:46 | 7.42 | 7.15 | 38 | 9.41 | 90.48%
2002-05-01 | 4:01 | 7.38 | 7.15 | 38 | 9.42 | 90.58%
2002-05-01 | 4:16 | 7.34 | 7.15 | 38 | 9.42 | 90.58%
2002-05-01 | 4:31 | 7.3 | 7.15 | 38 | 9.43 | 90.67%
2002-05-01 | 4:46 | 7.26 | 7.15 | 38 | 9.43 | 90.67%
2002-05-01 | 5:01 | 7.22 | 7.15 | 38 | 9.44 | 90.77%
2002-05-01 | 5:16 | 7.18 | 7.14 | 38 | 9.45 | 90.87%
2002-05-01 | 5:31 | 7.13 | 7.15 | 38 | 9.46 | 90.96%
2002-05-01 | 5:46 | 7.09 | 7.15 | 38 | 9.46 | 90.96%
2002-05-01 | 6:01 | 7.05 | 7.14 | 38 | 9.47 | 91.06%
2002-05-01 | 6:16 | 7 | 7.14 | 38 | 9.49 | 91.25%
2002-05-01 | 6:31 | 6.96 | 7.14 | 38 | 9.5 | 90.48%
2002-05-01 | 6:46 | 6.93 | 7.14 | 38 | 9.51 | 90.57%
2002-05-01 | 7:01 | 6.91 | 7.14 | 38 | 9.53 | 90.76%
2002-05-01 | 7:16 | 6.91 | 7.14 | 38 | 9.55 | 90.95%
2002-05-01 | 7:31 | 6.91 | 7.14 | 38 | 9.57 | 91.14%
2002-05-01 | 7:46 | 6.94 | 7.15 | 38 | 9.59 | 91.33%
2002-05-01 | 8:01 | 6.98 | 7.15 | 38 | 9.61 | 91.52%
2002-05-01 | 8:16 | 7.04 | 7.16 | 38 | 9.63 | 92.6%
2002-05-01 | 8:31 | 7.1 | 7.16 | 38 | 9.66 | 92.88%
2002-05-01 | 8:46 | 7.2 | 7.16 | 38 | 9.68 | 93.08%
2002-05-01 | 9:01 | 7.31 | 7.16 | 38 | 9.69 | 93.17%
2002-05-01 | 9:16 | 7.42 | 7.17 | 38 | 9.71 | 93.37%
2002-05-01 | 9:31 | 7.55 | 7.18 | 38 | 9.72 | 94.83%
2002-05-01 | 9:46 | 7.7 | 7.18 | 38 | 9.73 | 94.93%
2002-05-01 | 10:01 | 7.86 | 7.18 | 38 | 9.74 | 95.02%
2002-05-01 | 10:16 | 8.02 | 7.19 | 38 | 9.75 | 96.53%
2002-05-01 | 10:31 | 8.2 | 7.2 | 38 | 9.75 | 96.53%
2002-05-01 | 10:46 | 8.38 | 7.2 | 38 | 9.75 | 96.53%
2002-05-01 | 11:01 | 8.59 | 7.21 | 38 | 9.75 | 97.5%
2002-05-01 | 11:16 | 8.8 | 7.21 | 38 | 9.75 | 97.5%
2002-05-01 | 11:31 | 9.02 | 7.21 | 38 | 9.75 | 98.98%
2002-05-01 | 11:46 | 9.24 | 7.22 | 38 | 9.73 | 98.78%
2002-05-01 | 12:01 | 9.5 | 7.23 | 38 | 9.73 | 100.3%
2002-05-01 | 12:16 | 9.7 | 7.23 | 38 | 9.71 | 100.1%
2002-05-01 | 12:31 | 9.9 | 7.23 | 38 | 9.7 | 100%
2002-05-01 | 12:46 | 10.07 | 7.23 | 38 | 9.69 | 100.9%
2002-05-01 | 13:01 | 10.24 | 7.24 | 38 | 9.67 | 100.7%
2002-05-01 | 13:16 | 10.4 | 7.24 | 38 | 9.65 | 100.5%
2002-05-01 | 13:31 | 10.57 | 7.25 | 38 | 9.64 | 101.5%
2002-05-01 | 13:46 | 10.71 | 7.25 | 38 | 9.62 | 101.3%
2002-05-01 | 14:01 | 10.83 | 7.25 | 38 | 9.61 | 101.2%
2002-05-01 | 14:16 | 10.98 | 7.26 | 38 | 9.59 | 100.9%
2002-05-01 | 14:31 | 11.14 | 7.26 | 38 | 9.57 | 101.3%
2002-05-01 | 14:46 | 11.31 | 7.27 | 38 | 9.55 | 101.1%
2002-05-01 | 15:01 | 11.44 | 7.27 | 38 | 9.53 | 100.8%
2002-05-01 | 15:16 | 11.55 | 7.27 | 38 | 9.51 | 101.7%
2002-05-01 | 15:31 | 11.67 | 7.27 | 38 | 9.48 | 101.4%
2002-05-01 | 15:46 | 11.79 | 7.27 | 38 | 9.46 | 101.2%
2002-05-01 | 16:01 | 11.9 | 7.27 | 38 | 9.44 | 101%
2002-05-01 | 16:16 | 11.95 | 7.28 | 38 | 9.42 | 100.7%
2002-05-01 | 16:31 | 11.81 | 7.27 | 38 | 9.4 | 100.5%
2002-05-01 | 16:46 | 11.9 | 7.28 | 38 | 9.38 | 100.3%
2002-05-01 | 17:01 | 11.95 | 7.28 | 38 | 9.36 | 100.1%
2002-05-01 | 17:16 | 11.93 | 7.28 | 38 | 9.34 | 99.89%
2002-05-01 | 17:31 | 11.78 | 7.28 | 38 | 9.32 | 99.68%
2002-05-01 | 17:46 | 11.7 | 7.28 | 38 | 9.3 | 99.47%
2002-05-01 | 18:01 | 11.71 | 7.27 | 38 | 9.29 | 99.36%
2002-05-01 | 18:16 | 11.7 | 7.28 | 38 | 9.28 | 99.25%
2002-05-01 | 18:31 | 11.63 | 7.28 | 38 | 9.27 | 99.14%
2002-05-01 | 18:46 | 11.55 | 7.28 | 38 | 9.26 | 99.04%
2002-05-01 | 19:01 | 11.47 | 7.27 | 38 | 9.24 | 97.78%
2002-05-01 | 19:16 | 11.39 | 7.27 | 38 | 9.22 | 97.57%
2002-05-01 | 19:31 | 11.31 | 7.27 | 38 | 9.21 | 97.46%
2002-05-01 | 19:46 | 11.22 | 7.26 | 38 | 9.19 | 97.25%
2002-05-01 | 20:01 | 11.11 | 7.26 | 38 | 9.17 | 97.04%
2002-05-01 | 20:16 | 11 | 7.25 | 38 | 9.15 | 96.83%
2002-05-01 | 20:31 | 10.89 | 7.25 | 38 | 9.14 | 96.21%
2002-05-01 | 20:46 | 10.76 | 7.24 | 38 | 9.12 | 96%
2002-05-01 | 21:01 | 10.65 | 7.24 | 38 | 9.11 | 95.89%
2002-05-01 | 21:16 | 10.55 | 7.23 | 38 | 9.1 | 95.79%
2002-05-01 | 21:31 | 10.45 | 7.23 | 38 | 9.09 | 94.69%
2002-05-01 | 21:46 | 10.37 | 7.22 | 38 | 9.08 | 94.58%
2002-05-01 | 22:01 | 10.29 | 7.21 | 38 | 9.07 | 94.48%
2002-05-01 | 22:16 | 10.21 | 7.21 | 38 | 9.06 | 94.38%
2002-05-01 | 22:31 | 10.14 | 7.19 | 38 | 9.06 | 94.38%
2002-05-01 | 22:46 | 10.07 | 7.2 | 38 | 9.05 | 94.27%
2002-05-01 | 23:01 | 10 | 7.2 | 38 | 9.04 | 94.17%
2002-05-01 | 23:16 | 9.94 | 7.19 | 38 | 9.04 | 93.2%
2002-05-01 | 23:31 | 9.88 | 7.19 | 38 | 9.04 | 93.2%
2002-05-01 | 23:46 | 9.82 | 7.18 | 38 | 9.03 | 93.09%
2002-05-02 | 0:01 | 9.76 | 7.18 | 38 | 9.03 | 93.09%
2002-05-02 | 0:16 | 9.7 | 7.17 | 38 | 9.03 | 93.09%
2002-05-02 | 0:31 | 9.63 | 7.17 | 38 | 9.03 | 93.09%
2002-05-02 | 0:46 | 9.57 | 7.16 | 38 | 9.03 | 93.09%
2002-05-02 | 1:01 | 9.5 | 7.16 | 38 | 9.03 | 93.09%
2002-05-02 | 1:16 | 9.43 | 7.16 | 38 | 9.04 | 91.78%
2002-05-02 | 1:31 | 9.36 | 7.16 | 38 | 9.04 | 91.78%
2002-05-02 | 1:46 | 9.29 | 7.16 | 38 | 9.05 | 91.88%
2002-05-02 | 2:01 | 9.23 | 7.16 | 38 | 9.05 | 91.88%
2002-05-02 | 2:16 | 9.16 | 7.15 | 38 | 9.06 | 91.98%
2002-05-02 | 2:31 | 9.1 | 7.16 | 38 | 9.06 | 91.98%
2002-05-02 | 2:46 | 9.03 | 7.16 | 38 | 9.07 | 92.08%
2002-05-02 | 3:01 | 8.97 | 7.15 | 38 | 9.07 | 90.7%
2002-05-02 | 3:16 | 8.91 | 7.15 | 38 | 9.08 | 90.8%
2002-05-02 | 3:31 | 8.85 | 7.15 | 38 | 9.09 | 90.9%
2002-05-02 | 3:46 | 8.79 | 7.15 | 38 | 9.1 | 91%
2002-05-02 | 4:01 | 8.73 | 7.15 | 38 | 9.1 | 91%
2002-05-02 | 4:16 | 8.67 | 7.15 | 38 | 9.11 | 91.1%
2002-05-02 | 4:31 | 8.61 | 7.14 | 38 | 9.12 | 91.2%
2002-05-02 | 4:46 | 8.54 | 7.14 | 38 | 9.13 | 91.3%
2002-05-02 | 5:01 | 8.48 | 7.15 | 38 | 9.14 | 90.5%
2002-05-02 | 5:16 | 8.41 | 7.14 | 38 | 9.16 | 90.69%
2002-05-02 | 5:31 | 8.34 | 7.14 | 38 | 9.17 | 90.79%
2002-05-02 | 5:46 | 8.28 | 7.14 | 38 | 9.18 | 90.89%
2002-05-02 | 6:01 | 8.22 | 7.13 | 38 | 9.19 | 90.99%
2002-05-02 | 6:16 | 8.15 | 7.14 | 38 | 9.21 | 91.19%
2002-05-02 | 6:31 | 8.1 | 7.13 | 39 | 9.22 | 91.29%
2002-05-02 | 6:46 | 8.06 | 7.14 | 38 | 9.24 | 91.49%
2002-05-02 | 7:01 | 8.02 | 7.14 | 38 | 9.27 | 91.78%
2002-05-02 | 7:16 | 8.01 | 7.14 | 38 | 9.29 | 91.98%
2002-05-02 | 7:31 | 8.01 | 7.14 | 39 | 9.31 | 92.18%
2002-05-02 | 7:46 | 8.02 | 7.15 | 38 | 9.34 | 92.48%
2002-05-02 | 8:01 | 8.06 | 7.15 | 38 | 9.36 | 92.67%
2002-05-02 | 8:16 | 8.12 | 7.16 | 38 | 9.38 | 92.87%
2002-05-02 | 8:31 | 8.2 | 7.16 | 38 | 9.4 | 93.07%
2002-05-02 | 8:46 | 8.28 | 7.17 | 39 | 9.42 | 93.27%
2002-05-02 | 9:01 | 8.39 | 7.17 | 39 | 9.44 | 93.47%
2002-05-02 | 9:16 | 8.5 | 7.18 | 39 | 9.46 | 94.6%
2002-05-02 | 9:31 | 8.63 | 7.19 | 38 | 9.47 | 94.7%
2002-05-02 | 9:46 | 8.77 | 7.19 | 39 | 9.48 | 94.8%
2002-05-02 | 10:01 | 8.92 | 7.2 | 38 | 9.49 | 94.9%
2002-05-02 | 10:16 | 9.08 | 7.2 | 39 | 9.5 | 96.45%
2002-05-02 | 10:31 | 9.25 | 7.2 | 39 | 9.5 | 96.45%
2002-05-02 | 10:46 | 9.42 | 7.21 | 39 | 9.5 | 96.45%
2002-05-02 | 11:01 | 9.62 | 7.21 | 39 | 9.5 | 97.94%
2002-05-02 | 11:16 | 9.81 | 7.21 | 39 | 9.5 | 97.94%
2002-05-02 | 11:31 | 10.02 | 7.22 | 39 | 9.49 | 98.85%
2002-05-02 | 11:46 | 10.24 | 7.23 | 39 | 9.48 | 98.75%
2002-05-02 | 12:01 | 10.45 | 7.23 | 38 | 9.48 | 98.75%
2002-05-02 | 12:16 | 10.66 | 7.24 | 38 | 9.47 | 99.68%
2002-05-02 | 12:31 | 10.87 | 7.25 | 39 | 9.45 | 99.47%
2002-05-02 | 12:46 | 11.08 | 7.25 | 39 | 9.44 | 99.89%
2002-05-02 | 13:01 | 11.28 | 7.25 | 39 | 9.42 | 99.68%
2002-05-02 | 13:16 | 11.49 | 7.25 | 39 | 9.41 | 99.58%
2002-05-02 | 13:31 | 11.7 | 7.26 | 39 | 9.39 | 100.4%
2002-05-02 | 13:46 | 11.91 | 7.26 | 39 | 9.37 | 100.2%
2002-05-02 | 14:01 | 12.09 | 7.27 | 39 | 9.36 | 101.7%
2002-05-02 | 14:16 | 12.22 | 7.27 | 39 | 9.34 | 101.5%
2002-05-02 | 14:31 | 12.33 | 7.27 | 39 | 9.32 | 101.3%
2002-05-02 | 14:46 | 12.39 | 7.28 | 39 | 9.3 | 101.1%
2002-05-02 | 15:01 | 12.4 | 7.28 | 39 | 9.28 | 100.9%
2002-05-02 | 15:16 | 12.23 | 7.27 | 39 | 9.26 | 100.7%
2002-05-02 | 15:31 | 12.2 | 7.28 | 39 | 9.25 | 100.5%
2002-05-02 | 15:46 | 12.16 | 7.28 | 39 | 9.24 | 100.4%
2002-05-02 | 16:01 | 12.08 | 7.28 | 39 | 9.23 | 100.3%
2002-05-02 | 16:16 | 12.06 | 7.28 | 39 | 9.22 | 100.2%
2002-05-02 | 16:31 | 12.09 | 7.28 | 39 | 9.22 | 100.2%
2002-05-02 | 16:46 | 12.18 | 7.29 | 39 | 9.21 | 100.1%
2002-05-02 | 17:01 | 12.25 | 7.28 | 39 | 9.21 | 100.1%
2002-05-02 | 17:16 | 12.29 | 7.29 | 39 | 9.2 | 100%
2002-05-02 | 17:31 | 12.31 | 7.28 | 39 | 9.19 | 99.89%
2002-05-02 | 17:46 | 12.34 | 7.29 | 39 | 9.19 | 99.89%
2002-05-02 | 18:01 | 12.32 | 7.29 | 39 | 9.18 | 99.78%
2002-05-02 | 18:16 | 12.28 | 7.29 | 39 | 9.17 | 99.67%
2002-05-02 | 18:31 | 12.23 | 7.29 | 39 | 9.14 | 99.35%
2002-05-02 | 18:46 | 12.17 | 7.29 | 39 | 9.13 | 99.24%
2002-05-02 | 19:01 | 12.1 | 7.29 | 39 | 9.11 | 99.02%
2002-05-02 | 19:16 | 12.02 | 7.28 | 39 | 9.09 | 98.8%
2002-05-02 | 19:31 | 11.93 | 7.28 | 39 | 9.07 | 97.01%
2002-05-02 | 19:46 | 11.83 | 7.28 | 39 | 9.05 | 96.79%
2002-05-02 | 20:01 | 11.73 | 7.27 | 39 | 9.02 | 96.47%
2002-05-02 | 20:16 | 11.62 | 7.27 | 39 | 9.01 | 96.36%
2002-05-02 | 20:31 | 11.51 | 7.26 | 39 | 8.99 | 96.15%
2002-05-02 | 20:46 | 11.41 | 7.25 | 39 | 8.97 | 94.92%
2002-05-02 | 21:01 | 11.3 | 7.25 | 39 | 8.95 | 94.71%
2002-05-02 | 21:16 | 11.19 | 7.24 | 39 | 8.94 | 94.6%
2002-05-02 | 21:31 | 11.09 | 7.23 | 39 | 8.93 | 94.5%
2002-05-02 | 21:46 | 10.99 | 7.23 | 39 | 8.92 | 93.89%
2002-05-02 | 22:01 | 10.88 | 7.22 | 39 | 8.92 | 93.89%
2002-05-02 | 22:16 | 10.76 | 7.22 | 39 | 8.91 | 93.79%
2002-05-02 | 22:31 | 10.65 | 7.22 | 39 | 8.91 | 93.79%
2002-05-02 | 22:46 | 10.56 | 7.21 | 39 | 8.91 | 93.79%
2002-05-02 | 23:01 | 10.47 | 7.21 | 39 | 8.9 | 92.71%
2002-05-02 | 23:16 | 10.4 | 7.21 | 39 | 8.9 | 92.71%
2002-05-02 | 23:31 | 10.32 | 7.2 | 39 | 8.9 | 92.71%
2002-05-02 | 23:46 | 10.26 | 7.2 | 39 | 8.9 | 92.71%
2002-05-03 | 0:01 | 10.21 | 7.19 | 39 | 8.9 | 92.71%
2002-05-03 | 0:16 | 10.16 | 7.19 | 39 | 8.9 | 92.71%
2002-05-03 | 0:31 | 10.1 | 7.19 | 39 | 8.9 | 92.71%
2002-05-03 | 0:46 | 10.07 | 7.19 | 39 | 8.9 | 92.71%
2002-05-03 | 1:01 | 10.06 | 7.18 | 39 | 8.9 | 92.71%
2002-05-03 | 1:16 | 10.04 | 7.19 | 39 | 8.89 | 92.6%
2002-05-03 | 1:31 | 10.02 | 7.18 | 39 | 8.89 | 92.6%
2002-05-03 | 1:46 | 10 | 7.19 | 39 | 8.89 | 92.6%
2002-05-03 | 2:01 | 9.96 | 7.18 | 39 | 8.89 | 91.65%
2002-05-03 | 2:16 | 9.91 | 7.18 | 39 | 8.9 | 91.75%
2002-05-03 | 2:31 | 9.86 | 7.17 | 39 | 8.9 | 91.75%
2002-05-03 | 2:46 | 9.8 | 7.17 | 39 | 8.9 | 91.75%
2002-05-03 | 3:01 | 9.74 | 7.17 | 39 | 8.9 | 91.75%
2002-05-03 | 3:16 | 9.68 | 7.18 | 39 | 8.91 | 91.86%
2002-05-03 | 3:31 | 9.62 | 7.17 | 39 | 8.91 | 91.86%
2002-05-03 | 3:46 | 9.56 | 7.17 | 39 | 8.93 | 92.06%
2002-05-03 | 4:01 | 9.5 | 7.17 | 39 | 8.93 | 92.06%
2002-05-03 | 4:16 | 9.44 | 7.17 | 39 | 8.94 | 90.76%
2002-05-03 | 4:31 | 9.38 | 7.17 | 39 | 8.95 | 90.86%
2002-05-03 | 4:46 | 9.31 | 7.17 | 39 | 8.96 | 90.96%
2002-05-03 | 5:01 | 9.23 | 7.16 | 39 | 8.97 | 91.07%
2002-05-03 | 5:16 | 9.16 | 7.17 | 39 | 8.98 | 91.17%
2002-05-03 | 5:31 | 9.08 | 7.16 | 39 | 8.99 | 91.27%
2002-05-03 | 5:46 | 9.02 | 7.16 | 39 | 9.01 | 91.47%
2002-05-03 | 6:01 | 8.95 | 7.16 | 39 | 9.02 | 90.2%
2002-05-03 | 6:16 | 8.89 | 7.16 | 39 | 9.04 | 90.4%
2002-05-03 | 6:31 | 8.83 | 7.16 | 39 | 9.05 | 90.5%
2002-05-03 | 6:46 | 8.78 | 7.16 | 39 | 9.07 | 90.7%
2002-05-03 | 7:01 | 8.75 | 7.16 | 39 | 9.1 | 91%
2002-05-03 | 7:16 | 8.73 | 7.17 | 39 | 9.12 | 91.2%
2002-05-03 | 7:31 | 8.73 | 7.16 | 39 | 9.15 | 91.5%
2002-05-03 | 7:46 | 8.72 | 7.17 | 39 | 9.17 | 91.7%
2002-05-03 | 8:01 | 8.72 | 7.17 | 39 | 9.2 | 92%
2002-05-03 | 8:16 | 8.73 | 7.17 | 39 | 9.23 | 92.3%
2002-05-03 | 8:31 | 8.78 | 7.18 | 39 | 9.25 | 92.5%
2002-05-03 | 8:46 | 8.84 | 7.19 | 39 | 9.28 | 92.8%
2002-05-03 | 9:01 | 8.9 | 7.19 | 39 | 9.31 | 93.1%
2002-05-03 | 9:16 | 8.97 | 7.2 | 39 | 9.33 | 93.3%
2002-05-03 | 9:31 | 9.07 | 7.2 | 39 | 9.35 | 94.92%
2002-05-03 | 9:46 | 9.19 | 7.21 | 39 | 9.37 | 95.13%
2002-05-03 | 10:01 | 9.32 | 7.21 | 39 | 9.39 | 95.33%
2002-05-03 | 10:16 | 9.46 | 7.21 | 39 | 9.4 | 95.43%
2002-05-03 | 10:31 | 9.61 | 7.15 | 39 | 9.42 | 97.11%