## cont multiparameter audit
Organization | ODEQ |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name |  |  | Latitude
Number |  |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2002-04-29 14:30 | 16.9 | 9.2 | 11.2 | 125 | 83 | 0
Sonde | 2002-04-29 14:31 | 17.04 | 9.2 | 11.02 | 125 | 80
Grade |  | A | A | A |  | A
Audit 2 | 2002-05-01 10:01 | 12.7 | 8.5 | 11 | 113 | 93 | 0
Sonde | 2002-05-01 10:01 | 12.77 | 8.5 | 10.66 | 110 | 89
Grade |  | A | A | B |  | A
Audit 3 | 2002-05-03 13:05 | 16 | 9.3 | 11.3 | 124 | 88 | 0
Sonde | 2002-05-03 13:01 | 16.13 | 9.12 | 11.35 | 125 | 86
Grade |  | A | A | A |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note Audits 1, 2, & 3 are macro(ed) into this audit sheet and represent the beginning, middle and end of the deployment. Audits 4 and 5 were in the middle of the deployment and are not macro(ed) into this sheet. |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | ODEQ |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-29 14:30 | 16.9 | 9.2 | 11.2 | 125 | 83
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | ODEQ |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-01 10:01 | 12.7 | 8.5 | 11 | 113 | 93
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | ODEQ |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-03 13:05 | 16 | 9.3 | 11.3 | 124 | 88
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | ODEQ |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-04-30 18:05 | 15.4 | 9.3 | 10.1 | 109 | 87
Sonde Values | 2002-04-30 17:01 | 15.8 | 9.4 | 10.5 | 114 | 80
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | ODEQ |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | 0 |  | Latitude |  | 0
Number | 0 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2002-05-02 08:55 | 12.5 | 8.1 | 10.1 | 104 | 92
Sonde Values | 2002-05-02 09:01 | 12.6 | 8.2 | 10.1 | 104 | 91
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | SpCond | DO | DO
|  | °C |  | uS/cm | mg/L | % Sat
2002-04-29 | 14:31 | 17.04 | 9.2 | 80 | 11.02 | 125.2%
2002-04-29 | 14:46 | 17.32 | 9.23 | 80 | 10.92 | 124.1%
2002-04-29 | 15:01 | 17.36 | 9.25 | 80 | 10.87 | 123.5%
2002-04-29 | 15:16 | 17.48 | 9.27 | 79 | 10.77 | 122.4%
2002-04-29 | 15:31 | 17.67 | 9.29 | 79 | 10.7 | 122.3%
2002-04-29 | 15:46 | 17.81 | 9.31 | 79 | 10.66 | 121.8%
2002-04-29 | 16:01 | 18.01 | 9.33 | 78 | 10.58 | 121.6%
2002-04-29 | 16:16 | 18.14 | 9.34 | 77 | 10.52 | 120.9%
2002-04-29 | 16:31 | 18.22 | 9.36 | 77 | 10.44 | 120%
2002-04-29 | 16:46 | 18.25 | 9.37 | 77 | 10.37 | 119.2%
2002-04-29 | 17:01 | 18.24 | 9.38 | 76 | 10.33 | 118.7%
2002-04-29 | 17:16 | 18.22 | 9.39 | 76 | 10.27 | 118%
2002-04-29 | 17:31 | 18.19 | 9.39 | 75 | 10.21 | 117.4%
2002-04-29 | 17:46 | 18.08 | 9.39 | 74 | 10.14 | 116.6%
2002-04-29 | 18:01 | 17.94 | 9.39 | 74 | 10.06 | 115%
2002-04-29 | 18:16 | 17.78 | 9.38 | 74 | 9.96 | 113.8%
2002-04-29 | 18:31 | 17.61 | 9.38 | 74 | 9.84 | 112.5%
2002-04-29 | 18:46 | 17.43 | 9.36 | 74 | 9.72 | 110.5%
2002-04-29 | 19:01 | 17.23 | 9.35 | 74 | 9.59 | 109%
2002-04-29 | 19:16 | 17.03 | 9.32 | 74 | 9.42 | 107%
2002-04-29 | 19:31 | 16.8 | 9.29 | 75 | 9.23 | 103.1%
2002-04-29 | 19:46 | 16.63 | 9.26 | 76 | 9.03 | 100.9%
2002-04-29 | 20:01 | 16.43 | 9.22 | 77 | 8.85 | 97.25%
2002-04-29 | 20:16 | 16.27 | 9.19 | 77 | 8.7 | 95.6%
2002-04-29 | 20:31 | 16.11 | 9.17 | 78 | 8.58 | 94.29%
2002-04-29 | 20:46 | 15.95 | 9.13 | 78 | 8.46 | 91.96%
2002-04-29 | 21:01 | 15.79 | 9.1 | 79 | 8.41 | 91.41%
2002-04-29 | 21:16 | 15.63 | 9.07 | 80 | 8.33 | 90.54%
2002-04-29 | 21:31 | 15.47 | 9.04 | 80 | 8.28 | 89.03%
2002-04-29 | 21:46 | 15.35 | 9.01 | 80 | 8.22 | 88.39%
2002-04-29 | 22:01 | 15.23 | 8.97 | 81 | 8.17 | 87.85%
2002-04-29 | 22:16 | 15.08 | 8.93 | 81 | 8.11 | 87.2%
2002-04-29 | 22:31 | 14.96 | 8.89 | 82 | 8.07 | 85.85%
2002-04-29 | 22:46 | 14.82 | 8.84 | 83 | 8.02 | 85.32%
2002-04-29 | 23:01 | 14.68 | 8.79 | 83 | 7.97 | 84.79%
2002-04-29 | 23:16 | 14.51 | 8.72 | 84 | 7.95 | 84.57%
2002-04-29 | 23:31 | 14.36 | 8.66 | 84 | 7.94 | 84.02%
2002-04-29 | 23:46 | 14.22 | 8.59 | 85 | 7.94 | 84.02%
2002-04-30 | 0:01 | 14.07 | 8.51 | 85 | 7.93 | 83.92%
2002-04-30 | 0:16 | 13.93 | 8.43 | 85 | 7.94 | 83.58%
2002-04-30 | 0:31 | 13.79 | 8.35 | 86 | 7.94 | 83.58%
2002-04-30 | 0:46 | 13.65 | 8.27 | 86 | 7.95 | 83.68%
2002-04-30 | 1:01 | 13.5 | 8.18 | 87 | 7.95 | 83.68%
2002-04-30 | 1:16 | 13.37 | 8.11 | 87 | 7.97 | 83.02%
2002-04-30 | 1:31 | 13.23 | 8.04 | 87 | 7.98 | 83.12%
2002-04-30 | 1:46 | 13.08 | 7.97 | 88 | 7.99 | 83.23%
2002-04-30 | 2:01 | 12.94 | 7.91 | 88 | 8.01 | 82.58%
2002-04-30 | 2:16 | 12.81 | 7.86 | 89 | 8.02 | 82.68%
2002-04-30 | 2:31 | 12.67 | 7.83 | 89 | 8.05 | 82.99%
2002-04-30 | 2:46 | 12.53 | 7.79 | 89 | 8.06 | 83.09%
2002-04-30 | 3:01 | 12.41 | 7.76 | 89 | 8.08 | 82.03%
2002-04-30 | 3:16 | 12.28 | 7.73 | 89 | 8.1 | 82.23%
2002-04-30 | 3:31 | 12.16 | 7.71 | 90 | 8.13 | 82.54%
2002-04-30 | 3:46 | 12.02 | 7.69 | 90 | 8.15 | 82.74%
2002-04-30 | 4:01 | 11.91 | 7.67 | 90 | 8.17 | 81.7%
2002-04-30 | 4:16 | 11.78 | 7.66 | 90 | 8.19 | 81.9%
2002-04-30 | 4:31 | 11.68 | 7.65 | 90 | 8.22 | 82.2%
2002-04-30 | 4:46 | 11.56 | 7.64 | 90 | 8.25 | 82.5%
2002-04-30 | 5:01 | 11.45 | 7.62 | 90 | 8.26 | 81.78%
2002-04-30 | 5:16 | 11.37 | 7.62 | 90 | 8.29 | 82.08%
2002-04-30 | 5:31 | 11.27 | 7.61 | 90 | 8.32 | 82.38%
2002-04-30 | 5:46 | 11.17 | 7.6 | 90 | 8.34 | 82.57%
2002-04-30 | 6:01 | 11.07 | 7.6 | 90 | 8.37 | 82.87%
2002-04-30 | 6:16 | 10.97 | 7.6 | 90 | 8.42 | 82.55%
2002-04-30 | 6:31 | 10.87 | 7.61 | 90 | 8.51 | 83.43%
2002-04-30 | 6:46 | 10.78 | 7.63 | 90 | 8.64 | 84.71%
2002-04-30 | 7:01 | 10.7 | 7.66 | 90 | 8.79 | 86.18%
2002-04-30 | 7:16 | 10.64 | 7.69 | 90 | 8.96 | 87.84%
2002-04-30 | 7:31 | 10.62 | 7.74 | 89 | 9.15 | 89.71%
2002-04-30 | 7:46 | 10.66 | 7.8 | 89 | 9.37 | 91.86%
2002-04-30 | 8:01 | 10.74 | 7.88 | 89 | 9.62 | 94.31%
2002-04-30 | 8:16 | 10.89 | 7.97 | 89 | 9.84 | 96.47%
2002-04-30 | 8:31 | 11.09 | 8.06 | 89 | 10.03 | 99.31%
2002-04-30 | 8:46 | 11.23 | 8.14 | 89 | 10.18 | 100.8%
2002-04-30 | 9:01 | 11.38 | 8.21 | 88 | 10.3 | 102%
2002-04-30 | 9:16 | 11.48 | 8.28 | 88 | 10.41 | 103.1%
2002-04-30 | 9:31 | 11.64 | 8.35 | 88 | 10.52 | 105.2%
2002-04-30 | 9:46 | 11.86 | 8.42 | 88 | 10.63 | 106.3%
2002-04-30 | 10:01 | 12.11 | 8.49 | 88 | 10.72 | 108.8%
2002-04-30 | 10:16 | 12.35 | 8.56 | 88 | 10.81 | 109.7%
2002-04-30 | 10:31 | 12.54 | 8.62 | 87 | 10.88 | 112.2%
2002-04-30 | 10:46 | 12.76 | 8.69 | 87 | 10.97 | 113.1%
2002-04-30 | 11:01 | 13.15 | 8.75 | 87 | 11.02 | 114.8%
2002-04-30 | 11:16 | 13.43 | 8.8 | 87 | 11.1 | 115.6%
2002-04-30 | 11:31 | 13.82 | 8.85 | 86 | 11.12 | 117.1%
2002-04-30 | 11:46 | 14.17 | 8.9 | 86 | 11.16 | 118.1%
2002-04-30 | 12:01 | 14.62 | 8.94 | 86 | 11.21 | 119.3%
2002-04-30 | 12:16 | 14.69 | 8.98 | 86 | 11.2 | 119.1%
2002-04-30 | 12:31 | 14.53 | 9 | 85 | 11.12 | 118.3%
2002-04-30 | 12:46 | 14.47 | 9.03 | 85 | 11.15 | 118%
2002-04-30 | 13:01 | 14.4 | 9.06 | 84 | 11.1 | 117.5%
2002-04-30 | 13:16 | 14.52 | 9.09 | 84 | 11.1 | 118.1%
2002-04-30 | 13:31 | 14.9 | 9.13 | 84 | 11.15 | 118.6%
2002-04-30 | 13:46 | 15.13 | 9.17 | 83 | 11.16 | 120%
2002-04-30 | 14:01 | 15.29 | 9.2 | 83 | 11.12 | 119.6%
2002-04-30 | 14:16 | 15.38 | 9.22 | 82 | 11.09 | 119.2%
2002-04-30 | 14:31 | 15.38 | 9.24 | 82 | 11.02 | 118.5%
2002-04-30 | 14:46 | 15.36 | 9.25 | 82 | 10.94 | 117.6%
2002-04-30 | 15:01 | 15.3 | 9.27 | 81 | 10.85 | 116.7%
2002-04-30 | 15:16 | 15.23 | 9.26 | 81 | 10.7 | 115.1%
2002-04-30 | 15:31 | 15.18 | 9.26 | 80 | 10.55 | 113.4%
2002-04-30 | 15:46 | 15.18 | 9.27 | 81 | 10.43 | 112.2%
2002-04-30 | 16:01 | 15.22 | 9.28 | 80 | 10.39 | 111.7%
2002-04-30 | 16:16 | 15.26 | 9.29 | 80 | 10.35 | 111.3%
2002-04-30 | 16:31 | 15.38 | 9.31 | 80 | 10.39 | 111.7%
2002-04-30 | 16:46 | 15.65 | 9.33 | 80 | 10.47 | 113.8%
2002-04-30 | 17:01 | 15.77 | 9.35 | 80 | 10.47 | 113.8%
2002-04-30 | 17:16 | 15.77 | 9.36 | 79 | 10.42 | 113.3%
2002-04-30 | 17:31 | 15.68 | 9.35 | 79 | 10.27 | 111.6%
2002-04-30 | 17:46 | 15.56 | 9.33 | 80 | 10.11 | 109.9%
2002-04-30 | 18:01 | 15.38 | 9.32 | 80 | 10 | 107.5%
2002-04-30 | 18:16 | 15.26 | 9.32 | 80 | 9.93 | 106.8%
2002-04-30 | 18:31 | 15.15 | 9.31 | 80 | 9.84 | 105.8%
2002-04-30 | 18:46 | 15.12 | 9.3 | 80 | 9.76 | 104.9%
2002-04-30 | 19:01 | 15.08 | 9.28 | 80 | 9.55 | 102.7%
2002-04-30 | 19:16 | 15.01 | 9.26 | 80 | 9.43 | 101.4%
2002-04-30 | 19:31 | 14.97 | 9.23 | 80 | 9.25 | 98.4%
2002-04-30 | 19:46 | 14.9 | 9.19 | 80 | 9.07 | 96.49%
2002-04-30 | 20:01 | 14.8 | 9.15 | 81 | 8.89 | 94.57%
2002-04-30 | 20:16 | 14.75 | 9.11 | 81 | 8.75 | 93.09%
2002-04-30 | 20:31 | 14.66 | 9.07 | 82 | 8.64 | 91.91%
2002-04-30 | 20:46 | 14.53 | 9.04 | 82 | 8.56 | 91.06%
2002-04-30 | 21:01 | 14.43 | 9.01 | 82 | 8.5 | 89.95%
2002-04-30 | 21:16 | 14.33 | 8.98 | 83 | 8.43 | 89.21%
2002-04-30 | 21:31 | 14.23 | 8.94 | 83 | 8.38 | 88.68%
2002-04-30 | 21:46 | 14.15 | 8.9 | 84 | 8.33 | 88.15%
2002-04-30 | 22:01 | 14.08 | 8.87 | 84 | 8.29 | 87.72%
2002-04-30 | 22:16 | 14 | 8.82 | 84 | 8.24 | 87.2%
2002-04-30 | 22:31 | 13.92 | 8.77 | 84 | 8.2 | 86.32%
2002-04-30 | 22:46 | 13.83 | 8.71 | 84 | 8.16 | 85.89%
2002-04-30 | 23:01 | 13.77 | 8.65 | 85 | 8.13 | 85.58%
2002-04-30 | 23:16 | 13.68 | 8.59 | 85 | 8.11 | 85.37%
2002-04-30 | 23:31 | 13.61 | 8.52 | 85 | 8.09 | 85.16%
2002-04-30 | 23:46 | 13.54 | 8.46 | 85 | 8.08 | 85.05%
2002-05-01 | 0:01 | 13.45 | 8.38 | 86 | 8.07 | 84.06%
2002-05-01 | 0:16 | 13.37 | 8.31 | 86 | 8.05 | 83.85%
2002-05-01 | 0:31 | 13.28 | 8.23 | 87 | 8.05 | 83.85%
2002-05-01 | 0:46 | 13.2 | 8.16 | 87 | 8.05 | 83.85%
2002-05-01 | 1:01 | 13.12 | 8.09 | 87 | 8.04 | 83.75%
2002-05-01 | 1:16 | 13.02 | 8.02 | 88 | 8.04 | 83.75%
2002-05-01 | 1:31 | 12.96 | 7.96 | 88 | 8.04 | 82.89%
2002-05-01 | 1:46 | 12.85 | 7.91 | 88 | 8.05 | 82.99%
2002-05-01 | 2:01 | 12.75 | 7.86 | 89 | 8.05 | 82.99%
2002-05-01 | 2:16 | 12.66 | 7.83 | 89 | 8.07 | 83.2%
2002-05-01 | 2:31 | 12.59 | 7.8 | 89 | 8.08 | 83.3%
2002-05-01 | 2:46 | 12.47 | 7.77 | 89 | 8.09 | 82.13%
2002-05-01 | 3:01 | 12.38 | 7.74 | 90 | 8.11 | 82.34%
2002-05-01 | 3:16 | 12.29 | 7.72 | 90 | 8.14 | 82.64%
2002-05-01 | 3:31 | 12.19 | 7.7 | 91 | 8.16 | 82.84%
2002-05-01 | 3:46 | 12.1 | 7.69 | 91 | 8.19 | 83.15%
2002-05-01 | 4:01 | 12.03 | 7.67 | 91 | 8.21 | 83.35%
2002-05-01 | 4:16 | 11.96 | 7.66 | 91 | 8.23 | 82.3%
2002-05-01 | 4:31 | 11.89 | 7.65 | 92 | 8.26 | 82.6%
2002-05-01 | 4:46 | 11.81 | 7.64 | 92 | 8.27 | 82.7%
2002-05-01 | 5:01 | 11.74 | 7.63 | 92 | 8.29 | 82.9%
2002-05-01 | 5:16 | 11.67 | 7.62 | 92 | 8.31 | 83.1%
2002-05-01 | 5:31 | 11.61 | 7.62 | 92 | 8.34 | 83.4%
2002-05-01 | 5:46 | 11.53 | 7.61 | 92 | 8.36 | 83.6%
2002-05-01 | 6:01 | 11.47 | 7.61 | 92 | 8.39 | 83.07%
2002-05-01 | 6:16 | 11.41 | 7.61 | 93 | 8.42 | 83.37%
2002-05-01 | 6:31 | 11.36 | 7.61 | 93 | 8.49 | 84.06%
2002-05-01 | 6:46 | 11.32 | 7.63 | 92 | 8.58 | 84.95%
2002-05-01 | 7:01 | 11.29 | 7.65 | 92 | 8.7 | 86.14%
2002-05-01 | 7:16 | 11.27 | 7.68 | 92 | 8.84 | 87.52%
2002-05-01 | 7:31 | 11.27 | 7.73 | 92 | 9.04 | 89.5%
2002-05-01 | 7:46 | 11.33 | 7.8 | 92 | 9.3 | 92.08%
2002-05-01 | 8:01 | 11.42 | 7.88 | 92 | 9.52 | 94.26%
2002-05-01 | 8:16 | 11.54 | 7.95 | 91 | 9.73 | 97.3%
2002-05-01 | 8:31 | 11.67 | 8.03 | 91 | 9.9 | 99%
2002-05-01 | 8:46 | 11.83 | 8.11 | 91 | 10.04 | 100.4%
2002-05-01 | 9:01 | 11.99 | 8.18 | 91 | 10.17 | 101.7%
2002-05-01 | 9:16 | 12.15 | 8.25 | 91 | 10.29 | 104.5%
2002-05-01 | 9:31 | 12.37 | 8.32 | 90 | 10.4 | 105.6%
2002-05-01 | 9:46 | 12.57 | 8.41 | 90 | 10.53 | 108.6%
2002-05-01 | 10:01 | 12.77 | 8.5 | 89 | 10.66 | 109.9%
2002-05-01 | 10:16 | 13.07 | 8.56 | 89 | 10.75 | 112%
2002-05-01 | 10:31 | 13.32 | 8.62 | 89 | 10.83 | 112.8%
2002-05-01 | 10:46 | 13.59 | 8.69 | 89 | 10.92 | 114.9%
2002-05-01 | 11:01 | 13.9 | 8.74 | 88 | 10.98 | 115.6%
2002-05-01 | 11:16 | 14.18 | 8.8 | 88 | 11.05 | 116.9%
2002-05-01 | 11:31 | 14.51 | 8.85 | 88 | 11.11 | 118.2%
2002-05-01 | 11:46 | 14.8 | 8.9 | 87 | 11.15 | 118.6%
2002-05-01 | 12:01 | 15.14 | 8.95 | 87 | 11.18 | 120.2%
2002-05-01 | 12:16 | 15.5 | 8.99 | 87 | 11.21 | 121.8%
2002-05-01 | 12:31 | 15.78 | 9.03 | 86 | 11.23 | 122.1%
2002-05-01 | 12:46 | 16.01 | 9.07 | 86 | 11.26 | 123.7%
2002-05-01 | 13:01 | 16.36 | 9.1 | 85 | 11.24 | 123.5%
2002-05-01 | 13:16 | 16.65 | 9.14 | 85 | 11.22 | 125.4%
2002-05-01 | 13:31 | 16.93 | 9.17 | 84 | 11.21 | 125.3%
2002-05-01 | 13:46 | 17.2 | 9.2 | 83 | 11.15 | 126.7%
2002-05-01 | 14:01 | 17.5 | 9.23 | 82 | 11.08 | 126.6%
2002-05-01 | 14:16 | 17.73 | 9.26 | 81 | 11.03 | 126.1%
2002-05-01 | 14:31 | 17.96 | 9.28 | 80 | 10.96 | 125.3%
2002-05-01 | 14:46 | 18.2 | 9.3 | 79 | 10.95 | 125.9%
2002-05-01 | 15:01 | 18.38 | 9.32 | 78 | 10.9 | 125.3%
2002-05-01 | 15:16 | 18.59 | 9.34 | 78 | 10.83 | 125.9%
2002-05-01 | 15:31 | 18.75 | 9.36 | 77 | 10.74 | 124.9%
2002-05-01 | 15:46 | 18.88 | 9.37 | 76 | 10.67 | 124.1%
2002-05-01 | 16:01 | 18.98 | 9.39 | 78 | 10.62 | 123.5%
2002-05-01 | 16:16 | 19.11 | 9.41 | 77 | 10.57 | 124.4%
2002-05-01 | 16:31 | 19.2 | 9.41 | 76 | 10.51 | 123.6%
2002-05-01 | 16:46 | 19.18 | 9.42 | 76 | 10.48 | 123.3%
2002-05-01 | 17:01 | 19.25 | 9.43 | 76 | 10.4 | 122.4%
2002-05-01 | 17:16 | 19.34 | 9.44 | 76 | 10.31 | 121.3%
2002-05-01 | 17:31 | 19.32 | 9.45 | 76 | 10.23 | 120.4%
2002-05-01 | 17:46 | 19.22 | 9.45 | 75 | 10.16 | 119.5%
2002-05-01 | 18:01 | 19.16 | 9.45 | 74 | 10.05 | 118.2%
2002-05-01 | 18:16 | 19.02 | 9.45 | 74 | 9.93 | 116.8%
2002-05-01 | 18:31 | 18.89 | 9.44 | 74 | 9.78 | 113.7%
2002-05-01 | 18:46 | 18.71 | 9.43 | 73 | 9.63 | 112%
2002-05-01 | 19:01 | 18.5 | 9.42 | 77 | 9.47 | 110.1%
2002-05-01 | 19:16 | 18.28 | 9.39 | 79 | 9.27 | 106.6%
2002-05-01 | 19:31 | 18.06 | 9.36 | 79 | 9.02 | 103.7%
2002-05-01 | 19:46 | 17.84 | 9.32 | 78 | 8.8 | 100.6%
2002-05-01 | 20:01 | 17.67 | 9.29 | 78 | 8.63 | 98.63%
2002-05-01 | 20:16 | 17.51 | 9.27 | 80 | 8.46 | 96.69%
2002-05-01 | 20:31 | 17.38 | 9.24 | 80 | 8.32 | 94.55%
2002-05-01 | 20:46 | 17.23 | 9.2 | 80 | 8.19 | 93.07%
2002-05-01 | 21:01 | 17.07 | 9.17 | 80 | 8.08 | 91.82%
2002-05-01 | 21:16 | 16.93 | 9.14 | 81 | 8.01 | 89.5%
2002-05-01 | 21:31 | 16.78 | 9.11 | 81 | 7.94 | 88.72%
2002-05-01 | 21:46 | 16.64 | 9.08 | 82 | 7.89 | 88.16%
2002-05-01 | 22:01 | 16.54 | 9.05 | 83 | 7.85 | 87.71%
2002-05-01 | 22:16 | 16.4 | 9.01 | 83 | 7.8 | 85.71%
2002-05-01 | 22:31 | 16.26 | 8.97 | 83 | 7.76 | 85.27%
2002-05-01 | 22:46 | 16.12 | 8.93 | 84 | 7.72 | 84.84%
2002-05-01 | 23:01 | 15.97 | 8.88 | 84 | 7.7 | 83.7%
2002-05-01 | 23:16 | 15.82 | 8.83 | 85 | 7.69 | 83.59%
2002-05-01 | 23:31 | 15.66 | 8.77 | 85 | 7.67 | 83.37%
2002-05-01 | 23:46 | 15.51 | 8.71 | 85 | 7.66 | 83.26%
2002-05-02 | 0:01 | 15.37 | 8.65 | 86 | 7.67 | 82.47%
2002-05-02 | 0:16 | 15.22 | 8.58 | 86 | 7.67 | 82.47%
2002-05-02 | 0:31 | 15.07 | 8.5 | 86 | 7.67 | 82.47%
2002-05-02 | 0:46 | 14.92 | 8.42 | 87 | 7.68 | 81.7%
2002-05-02 | 1:01 | 14.79 | 8.34 | 87 | 7.7 | 81.91%
2002-05-02 | 1:16 | 14.64 | 8.25 | 88 | 7.72 | 82.13%
2002-05-02 | 1:31 | 14.5 | 8.17 | 88 | 7.73 | 82.23%
2002-05-02 | 1:46 | 14.37 | 8.09 | 89 | 7.74 | 81.9%
2002-05-02 | 2:01 | 14.23 | 8.02 | 89 | 7.77 | 82.22%
2002-05-02 | 2:16 | 14.1 | 7.95 | 90 | 7.79 | 82.43%
2002-05-02 | 2:31 | 13.97 | 7.9 | 90 | 7.8 | 82.11%
2002-05-02 | 2:46 | 13.84 | 7.85 | 91 | 7.82 | 82.32%
2002-05-02 | 3:01 | 13.72 | 7.81 | 91 | 7.85 | 82.63%
2002-05-02 | 3:16 | 13.6 | 7.78 | 92 | 7.87 | 82.84%
2002-05-02 | 3:31 | 13.47 | 7.75 | 92 | 7.89 | 82.19%
2002-05-02 | 3:46 | 13.36 | 7.73 | 92 | 7.9 | 82.29%
2002-05-02 | 4:01 | 13.24 | 7.71 | 93 | 7.93 | 82.6%
2002-05-02 | 4:16 | 13.13 | 7.69 | 93 | 7.95 | 82.81%
2002-05-02 | 4:31 | 13 | 7.67 | 93 | 7.97 | 83.02%
2002-05-02 | 4:46 | 12.9 | 7.66 | 93 | 8 | 82.47%
2002-05-02 | 5:01 | 12.8 | 7.65 | 93 | 8.03 | 82.78%
2002-05-02 | 5:16 | 12.68 | 7.64 | 94 | 8.05 | 82.99%
2002-05-02 | 5:31 | 12.57 | 7.63 | 94 | 8.08 | 83.3%
2002-05-02 | 5:46 | 12.48 | 7.62 | 94 | 8.11 | 82.34%
2002-05-02 | 6:01 | 12.36 | 7.62 | 94 | 8.14 | 82.64%
2002-05-02 | 6:16 | 12.27 | 7.62 | 94 | 8.2 | 83.25%
2002-05-02 | 6:31 | 12.17 | 7.63 | 94 | 8.3 | 84.26%
2002-05-02 | 6:46 | 12.09 | 7.66 | 94 | 8.46 | 85.89%
2002-05-02 | 7:01 | 12.04 | 7.69 | 94 | 8.64 | 87.72%
2002-05-02 | 7:16 | 12.03 | 7.74 | 94 | 8.84 | 89.75%
2002-05-02 | 7:31 | 12.05 | 7.79 | 93 | 9.04 | 91.78%
2002-05-02 | 7:46 | 12.1 | 7.85 | 93 | 9.24 | 93.81%
2002-05-02 | 8:01 | 12.16 | 7.92 | 93 | 9.43 | 95.74%
2002-05-02 | 8:16 | 12.24 | 7.99 | 93 | 9.6 | 97.46%
2002-05-02 | 8:31 | 12.35 | 8.06 | 92 | 9.76 | 99.09%
2002-05-02 | 8:46 | 12.48 | 8.13 | 92 | 9.91 | 100.6%
2002-05-02 | 9:01 | 12.63 | 8.21 | 91 | 10.06 | 103.7%
2002-05-02 | 9:16 | 12.77 | 8.29 | 91 | 10.2 | 105.2%
2002-05-02 | 9:31 | 12.93 | 8.37 | 91 | 10.33 | 106.5%
2002-05-02 | 9:46 | 13.11 | 8.45 | 90 | 10.44 | 108.7%
2002-05-02 | 10:01 | 13.32 | 8.52 | 90 | 10.54 | 109.8%
2002-05-02 | 10:16 | 13.53 | 8.59 | 89 | 10.64 | 112%
2002-05-02 | 10:31 | 13.8 | 8.65 | 89 | 10.73 | 112.9%
2002-05-02 | 10:46 | 14.1 | 8.72 | 89 | 10.81 | 114.4%
2002-05-02 | 11:01 | 14.37 | 8.78 | 88 | 10.89 | 115.2%
2002-05-02 | 11:16 | 14.66 | 8.83 | 88 | 10.95 | 116.5%
2002-05-02 | 11:31 | 14.93 | 8.88 | 87 | 11.01 | 117.1%
2002-05-02 | 11:46 | 15.28 | 8.93 | 87 | 11.05 | 118.8%
2002-05-02 | 12:01 | 15.52 | 8.97 | 87 | 11.08 | 120.4%
2002-05-02 | 12:16 | 15.77 | 9.02 | 87 | 11.12 | 120.9%
2002-05-02 | 12:31 | 16.1 | 9.06 | 86 | 11.14 | 122.4%
2002-05-02 | 12:46 | 16.39 | 9.09 | 86 | 11.2 | 123.1%
2002-05-02 | 13:01 | 16.69 | 9.13 | 86 | 11.2 | 125.1%
2002-05-02 | 13:16 | 16.98 | 9.17 | 85 | 11.22 | 125.4%
2002-05-02 | 13:31 | 17.28 | 9.19 | 85 | 11.16 | 126.8%
2002-05-02 | 13:46 | 17.51 | 9.22 | 85 | 11.1 | 126.9%
2002-05-02 | 14:01 | 17.73 | 9.25 | 84 | 11.07 | 126.5%
2002-05-02 | 14:16 | 17.99 | 9.27 | 83 | 11.04 | 126.2%
2002-05-02 | 14:31 | 18.2 | 9.29 | 83 | 10.98 | 126.2%
2002-05-02 | 14:46 | 18.33 | 9.32 | 82 | 10.95 | 125.9%
2002-05-02 | 15:01 | 18.53 | 9.35 | 82 | 10.9 | 126.7%
2002-05-02 | 15:16 | 18.7 | 9.36 | 81 | 10.84 | 126%
2002-05-02 | 15:31 | 18.79 | 9.38 | 81 | 10.78 | 125.3%
2002-05-02 | 15:46 | 18.93 | 9.41 | 80 | 10.71 | 124.5%
2002-05-02 | 16:01 | 18.98 | 9.42 | 81 | 10.69 | 124.3%
2002-05-02 | 16:16 | 18.95 | 9.43 | 80 | 10.61 | 123.4%
2002-05-02 | 16:31 | 18.97 | 9.44 | 80 | 10.52 | 122.3%
2002-05-02 | 16:46 | 18.95 | 9.46 | 80 | 10.47 | 121.7%
2002-05-02 | 17:01 | 18.93 | 9.47 | 79 | 10.38 | 120.7%
2002-05-02 | 17:16 | 18.89 | 9.48 | 79 | 10.31 | 119.9%
2002-05-02 | 17:31 | 18.83 | 9.48 | 79 | 10.19 | 118.5%
2002-05-02 | 17:46 | 18.77 | 9.48 | 79 | 10.11 | 117.6%
2002-05-02 | 18:01 | 18.64 | 9.48 | 78 | 10.01 | 116.4%
2002-05-02 | 18:16 | 18.51 | 9.48 | 78 | 9.9 | 115.1%
2002-05-02 | 18:31 | 18.35 | 9.48 | 78 | 9.8 | 112.6%
2002-05-02 | 18:46 | 18.16 | 9.46 | 78 | 9.64 | 110.8%
2002-05-02 | 19:01 | 17.91 | 9.45 | 79 | 9.43 | 107.8%
2002-05-02 | 19:16 | 17.67 | 9.42 | 79 | 9.21 | 105.3%
2002-05-02 | 19:31 | 17.46 | 9.4 | 80 | 9.02 | 102.5%
2002-05-02 | 19:46 | 17.28 | 9.37 | 81 | 8.83 | 100.3%
2002-05-02 | 20:01 | 17.13 | 9.34 | 82 | 8.66 | 98.41%
2002-05-02 | 20:16 | 16.9 | 9.31 | 82 | 8.49 | 94.86%
2002-05-02 | 20:31 | 16.75 | 9.28 | 83 | 8.35 | 93.3%
2002-05-02 | 20:46 | 16.58 | 9.25 | 84 | 8.22 | 91.84%
2002-05-02 | 21:01 | 16.36 | 9.22 | 85 | 8.12 | 89.23%
2002-05-02 | 21:16 | 16.12 | 9.19 | 85 | 8.05 | 88.46%
2002-05-02 | 21:31 | 15.87 | 9.16 | 86 | 8 | 86.96%
2002-05-02 | 21:46 | 15.63 | 9.13 | 87 | 7.97 | 86.63%
2002-05-02 | 22:01 | 15.43 | 9.1 | 88 | 7.93 | 85.27%
2002-05-02 | 22:16 | 15.31 | 9.08 | 88 | 7.9 | 84.95%
2002-05-02 | 22:31 | 15.15 | 9.04 | 89 | 7.86 | 84.52%
2002-05-02 | 22:46 | 14.91 | 9.01 | 89 | 7.85 | 83.51%
2002-05-02 | 23:01 | 14.68 | 8.98 | 92 | 7.85 | 83.51%
2002-05-02 | 23:16 | 14.53 | 8.95 | 90 | 7.85 | 83.51%
2002-05-02 | 23:31 | 14.36 | 8.9 | 91 | 7.83 | 82.86%
2002-05-02 | 23:46 | 14.21 | 8.86 | 90 | 7.83 | 82.86%
2002-05-03 | 0:01 | 14.12 | 8.82 | 91 | 7.85 | 83.07%
2002-05-03 | 0:16 | 13.98 | 8.78 | 91 | 7.85 | 82.63%
2002-05-03 | 0:31 | 13.75 | 8.73 | 92 | 7.86 | 82.74%
2002-05-03 | 0:46 | 13.62 | 8.66 | 92 | 7.88 | 82.95%
2002-05-03 | 1:01 | 13.6 | 8.62 | 92 | 7.91 | 83.26%
2002-05-03 | 1:16 | 13.53 | 8.55 | 93 | 7.92 | 83.37%
2002-05-03 | 1:31 | 13.48 | 8.47 | 93 | 7.93 | 82.6%
2002-05-03 | 1:46 | 13.41 | 8.39 | 94 | 7.94 | 82.71%
2002-05-03 | 2:01 | 13.28 | 8.31 | 95 | 7.97 | 83.02%
2002-05-03 | 2:16 | 13.16 | 8.24 | 95 | 7.99 | 83.23%
2002-05-03 | 2:31 | 12.99 | 8.15 | 95 | 8.01 | 82.58%
2002-05-03 | 2:46 | 12.77 | 8.07 | 96 | 8.05 | 82.99%
2002-05-03 | 3:01 | 12.68 | 8 | 97 | 8.06 | 83.09%
2002-05-03 | 3:16 | 12.63 | 7.94 | 97 | 8.07 | 83.2%
2002-05-03 | 3:31 | 12.55 | 7.9 | 97 | 8.08 | 83.3%
2002-05-03 | 3:46 | 12.48 | 7.86 | 98 | 8.1 | 82.23%
2002-05-03 | 4:01 | 12.39 | 7.83 | 98 | 8.11 | 82.34%
2002-05-03 | 4:16 | 12.32 | 7.8 | 98 | 8.13 | 82.54%
2002-05-03 | 4:31 | 12.2 | 7.77 | 98 | 8.16 | 82.84%
2002-05-03 | 4:46 | 12.05 | 7.75 | 98 | 8.19 | 83.15%
2002-05-03 | 5:01 | 11.93 | 7.73 | 99 | 8.21 | 82.1%
2002-05-03 | 5:16 | 11.88 | 7.72 | 99 | 8.23 | 82.3%
2002-05-03 | 5:31 | 11.87 | 7.7 | 99 | 8.24 | 82.4%
2002-05-03 | 5:46 | 11.82 | 7.69 | 99 | 8.26 | 82.6%
2002-05-03 | 6:01 | 11.76 | 7.68 | 99 | 8.29 | 82.9%
2002-05-03 | 6:16 | 11.69 | 7.68 | 99 | 8.36 | 83.6%
2002-05-03 | 6:31 | 11.62 | 7.7 | 98 | 8.49 | 84.9%
2002-05-03 | 6:46 | 11.57 | 7.73 | 99 | 8.67 | 86.7%
2002-05-03 | 7:01 | 11.55 | 7.77 | 98 | 8.86 | 88.6%
2002-05-03 | 7:16 | 11.55 | 7.82 | 98 | 9.06 | 90.6%
2002-05-03 | 7:31 | 11.53 | 7.87 | 97 | 9.24 | 92.4%
2002-05-03 | 7:46 | 11.51 | 7.92 | 97 | 9.43 | 94.3%
2002-05-03 | 8:01 | 11.55 | 7.98 | 96 | 9.59 | 95.9%
2002-05-03 | 8:16 | 11.61 | 8.04 | 96 | 9.77 | 97.7%
2002-05-03 | 8:31 | 11.71 | 8.1 | 95 | 9.92 | 99.2%
2002-05-03 | 8:46 | 11.84 | 8.17 | 94 | 10.07 | 100.7%
2002-05-03 | 9:01 | 11.97 | 8.25 | 93 | 10.22 | 102.2%
2002-05-03 | 9:16 | 12.14 | 8.33 | 93 | 10.36 | 105.2%
2002-05-03 | 9:31 | 12.29 | 8.4 | 92 | 10.47 | 106.3%
2002-05-03 | 9:46 | 12.47 | 8.47 | 91 | 10.58 | 107.4%
2002-05-03 | 10:01 | 12.6 | 8.52 | 91 | 10.65 | 109.8%
2002-05-03 | 10:16 | 12.83 | 8.58 | 90 | 10.74 | 110.7%
2002-05-03 | 10:31 | 13.12 | 8.65 | 90 | 10.83 | 112.8%
2002-05-03 | 10:46 | 13.4 | 8.71 | 89 | 10.93 | 113.9%
2002-05-03 | 11:01 | 13.71 | 8.77 | 89 | 11 | 115.8%
2002-05-03 | 11:16 | 13.96 | 8.83 | 89 | 11.07 | 116.5%
2002-05-03 | 11:31 | 14.28 | 8.87 | 88 | 11.12 | 117.7%
2002-05-03 | 11:46 | 14.56 | 8.92 | 88 | 11.17 | 118.8%
2002-05-03 | 12:01 | 14.88 | 8.96 | 87 | 11.21 | 119.3%
2002-05-03 | 12:16 | 15.18 | 9.01 | 87 | 11.25 | 121%
2002-05-03 | 12:31 | 15.48 | 9.05 | 87 | 11.3 | 121.5%
2002-05-03 | 12:46 | 15.81 | 9.09 | 86 | 11.33 | 123.2%
2002-05-03 | 13:01 | 16.13 | 9.12 | 86 | 11.35 | 124.7%