## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 164.75
Program Name | City of Bend Stormwater Management Plan: Ambient Monitoring Project |  | Description |  | Deschutes River, u/s Riverhouse Hotel
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 6920 #04G11769AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2005-09-14 |  | 2005-09-14 |  | A
Time | Computer set to atomic time |  | 15:42 |  | 15:42 |  | A
Temp | VWR NIST traceable surface water |  | 14.9 |  | 14.9 |  | A
DO | Winkler titration of surface water |  | 10.6 |  | 11.1 |  | B
SpC | 148 µS/mL |  | 63 |  | 60 |  | A
Ph | pH 7 standard solution |  | 8.3 |  | 8.7 |  | B
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2005-09-15 13:30 | 14.4 | 7.7 | 9.1 | 0 | 62 | 0
Sonde |  | 14.4 | 8.2 | 10.1 |  | 63
Grade |  | A | B | B |  | A
Audit 2 | 2005-09-20 12:31 | 13.3 | 7.9 | 9.3 | 0 | 62 | 0
Sonde |  | 13.3 | 8.1 | 10 |  | 63
Grade |  | A | A | B |  | A
Audit 3 | 2005-09-22 11:01 | 12.1 | 7.6 | 9.2 | 0 | 62 | 0
Sonde |  | 12.1 | 7.9 | 10.1 |  | 64
Grade |  | A | A | B |  | A
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 164.75
Program Name | City of Bend Stormwater Management Plan: Ambient Monitoring Project |  | Description |  | Deschutes River, u/s Riverhouse Hotel
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 #04C1624AX HACH 2100P #040300035008
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2005-09-15 13:30 | 14.4 | 7.7 | 9.1 |  | 62
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 164.75
Program Name | City of Bend Stormwater Management Plan: Ambient Monitoring Project |  | Description |  | Deschutes River, u/s Riverhouse Hotel
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 #04C1624AX HACH 2100P #040300035008
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2005-09-20 12:31 | 13.3 | 7.9 | 9.3 |  | 62
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 164.75
Program Name | City of Bend Stormwater Management Plan: Ambient Monitoring Project |  | Description |  | Deschutes River, u/s Riverhouse Hotel
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 #04C1624AX HACH 2100P #040300035008
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2005-09-22 11:01 | 12.1 | 7.6 | 9.2 |  | 62
3
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | SpCond | DO Conc | pH | Turbidity
|  | °C | mS/mL | mg/L |  | NTU+
2005-09-15 | 13:30 | 14.43 | 0.063 | 10.16 | 8.2 | 2.1
2005-09-15 | 13:45 | 14.51 | 0.063 | 9.99 | 8.24 | 1.7
2005-09-15 | 14:00 | 14.58 | 0.063 | 9.82 | 8.27 | 1.7
2005-09-15 | 14:15 | 14.65 | 0.064 | 9.72 | 8.3 | 1.7
2005-09-15 | 14:30 | 14.72 | 0.064 | 9.66 | 8.32 | 1.8
2005-09-15 | 14:45 | 14.78 | 0.064 | 9.61 | 8.35 | 1.7
2005-09-15 | 15:00 | 14.84 | 0.064 | 9.57 | 8.37 | 1.8
2005-09-15 | 15:15 | 14.89 | 0.064 | 9.52 | 8.38 | 1.7
2005-09-15 | 15:30 | 14.94 | 0.064 | 9.49 | 8.4 | 1.8
2005-09-15 | 15:45 | 14.99 | 0.064 | 9.47 | 8.4 | 1.8
2005-09-15 | 16:00 | 15.02 | 0.064 | 9.45 | 8.42 | 1.7
2005-09-15 | 16:15 | 15.06 | 0.064 | 9.42 | 8.42 | 1.8
2005-09-15 | 16:30 | 15.09 | 0.064 | 9.4 | 8.43 | 1.8
2005-09-15 | 16:45 | 15.11 | 0.064 | 9.39 | 8.43 | 1.7
2005-09-15 | 17:00 | 15.12 | 0.064 | 9.37 | 8.42 | 1.7
2005-09-15 | 17:15 | 15.13 | 0.064 | 9.34 | 8.42 | 1.7
2005-09-15 | 17:30 | 15.13 | 0.064 | 9.32 | 8.41 | 1.7
2005-09-15 | 17:45 | 15.13 | 0.064 | 9.32 | 8.4 | 1.7
2005-09-15 | 18:00 | 15.12 | 0.064 | 9.32 | 8.38 | 1.8
2005-09-15 | 18:15 | 15.1 | 0.064 | 9.32 | 8.37 | 1.7
2005-09-15 | 18:30 | 15.07 | 0.064 | 9.32 | 8.35 | 1.7
2005-09-15 | 18:45 | 15.03 | 0.064 | 9.33 | 8.32 | 1.8
2005-09-15 | 19:00 | 14.98 | 0.064 | 9.33 | 8.28 | 2.3
2005-09-15 | 19:15 | 14.93 | 0.064 | 9.34 | 8.24 | 1.7
2005-09-15 | 19:30 | 14.88 | 0.064 | 9.36 | 8.19 | 1.7
2005-09-15 | 19:45 | 14.83 | 0.064 | 9.37 | 8.15 | 1.7
2005-09-15 | 20:00 | 14.78 | 0.064 | 9.38 | 8.11 | 1.9
2005-09-15 | 20:15 | 14.73 | 0.065 | 9.4 | 8.08 | 1.7
2005-09-15 | 20:30 | 14.68 | 0.065 | 9.41 | 8.04 | 1.7
2005-09-15 | 20:45 | 14.62 | 0.065 | 9.43 | 8.01 | 1.8
2005-09-15 | 21:00 | 14.56 | 0.065 | 9.44 | 7.98 | 1.8
2005-09-15 | 21:15 | 14.49 | 0.065 | 9.46 | 7.95 | 1.8
2005-09-15 | 21:30 | 14.41 | 0.065 | 9.48 | 7.92 | 1.7
2005-09-15 | 21:45 | 14.33 | 0.065 | 9.5 | 7.9 | 1.8
2005-09-15 | 22:00 | 14.25 | 0.065 | 9.53 | 7.88 | 1.7
2005-09-15 | 22:15 | 14.17 | 0.065 | 9.55 | 7.86 | 1.7
2005-09-15 | 22:30 | 14.09 | 0.065 | 9.57 | 7.84 | 1.7
2005-09-15 | 22:45 | 14.02 | 0.065 | 9.59 | 7.83 | 1.7
2005-09-15 | 23:00 | 13.94 | 0.065 | 9.6 | 7.81 | 1.7
2005-09-15 | 23:15 | 13.87 | 0.065 | 9.62 | 7.8 | 1.7
2005-09-15 | 23:30 | 13.8 | 0.065 | 9.64 | 7.79 | 1.6
2005-09-15 | 23:45 | 13.74 | 0.065 | 9.65 | 7.78 | 1.7
2005-09-16 | 24:00 | 13.67 | 0.065 | 9.67 | 7.77 | 1.7
2005-09-16 | 24:15 | 13.61 | 0.065 | 9.68 | 7.77 | 1.7
2005-09-16 | 24:30 | 13.55 | 0.065 | 9.7 | 7.76 | 1.7
2005-09-16 | 24:45 | 13.5 | 0.065 | 9.71 | 7.76 | 1.6
2005-09-16 | 25:00 | 13.44 | 0.065 | 9.73 | 7.75 | 1.7
2005-09-16 | 25:15 | 13.39 | 0.065 | 9.74 | 7.75 | 1.6
2005-09-16 | 25:30 | 13.35 | 0.065 | 9.75 | 7.74 | 1.7
2005-09-16 | 25:45 | 13.3 | 0.065 | 9.77 | 7.74 | 1.8
2005-09-16 | 26:00 | 13.25 | 0.065 | 9.78 | 7.73 | 1.8
2005-09-16 | 26:15 | 13.22 | 0.065 | 9.79 | 7.73 | 1.7
2005-09-16 | 26:30 | 13.18 | 0.065 | 9.8 | 7.73 | 1.8
2005-09-16 | 26:45 | 13.15 | 0.064 | 9.81 | 7.73 | 1.8
2005-09-16 | 27:00 | 13.12 | 0.064 | 9.82 | 7.73 | 1.7
2005-09-16 | 27:15 | 13.09 | 0.064 | 9.82 | 7.73 | 1.8
2005-09-16 | 27:30 | 13.06 | 0.064 | 9.83 | 7.72 | 1.7
2005-09-16 | 27:45 | 13.04 | 0.064 | 9.83 | 7.72 | 1.7
2005-09-16 | 28:00 | 13.02 | 0.064 | 9.85 | 7.72 | 1.8
2005-09-16 | 28:15 | 13 | 0.064 | 9.85 | 7.72 | 1.6
2005-09-16 | 28:30 | 12.98 | 0.064 | 9.85 | 7.72 | 1.9
2005-09-16 | 28:45 | 12.97 | 0.064 | 9.86 | 7.72 | 1.7
2005-09-16 | 29:00 | 12.96 | 0.064 | 9.86 | 7.72 | 1.8
2005-09-16 | 29:15 | 12.94 | 0.064 | 9.86 | 7.72 | 1.8
2005-09-16 | 29:30 | 12.93 | 0.064 | 9.87 | 7.73 | 1.8
2005-09-16 | 29:45 | 12.92 | 0.064 | 9.87 | 7.73 | 1.7
2005-09-16 | 30:00 | 12.91 | 0.064 | 9.87 | 7.72 | 1.8
2005-09-16 | 30:15 | 12.91 | 0.064 | 9.88 | 7.73 | 1.8
2005-09-16 | 30:30 | 12.9 | 0.064 | 9.88 | 7.72 | 1.8
2005-09-16 | 30:45 | 12.9 | 0.064 | 9.89 | 7.73 | 1.8
2005-09-16 | 31:00 | 12.9 | 0.064 | 9.89 | 7.73 | 1.8
2005-09-16 | 31:15 | 12.89 | 0.064 | 9.9 | 7.73 | 1.8
2005-09-16 | 31:30 | 12.9 | 0.064 | 9.93 | 7.74 | 1.9
2005-09-16 | 31:45 | 12.91 | 0.064 | 9.93 | 7.75 | 2
2005-09-16 | 32:00 | 12.92 | 0.064 | 9.94 | 7.76 | 2
2005-09-16 | 32:15 | 12.93 | 0.064 | 9.93 | 7.76 | 2.1
2005-09-16 | 32:30 | 12.94 | 0.064 | 9.94 | 7.77 | 2
2005-09-16 | 32:45 | 12.95 | 0.064 | 9.93 | 7.77 | 1.9
2005-09-16 | 33:00 | 12.96 | 0.064 | 9.93 | 7.77 | 1.8
2005-09-16 | 33:15 | 12.98 | 0.064 | 9.96 | 7.78 | 1.9
2005-09-16 | 33:30 | 13 | 0.064 | 9.95 | 7.8 | 1.9
2005-09-16 | 33:45 | 13.02 | 0.064 | 9.96 | 7.81 | 2
2005-09-16 | 34:00 | 13.05 | 0.064 | 9.97 | 7.82 | 1.9
2005-09-16 | 34:15 | 13.08 | 0.064 | 9.98 | 7.84 | 1.8
2005-09-16 | 34:30 | 13.15 | 0.064 | 9.99 | 7.87 | 2
2005-09-16 | 34:45 | 13.19 | 0.064 | 9.97 | 7.88 | 1.9
2005-09-16 | 35:00 | 13.22 | 0.064 | 9.95 | 7.89 | 2
2005-09-16 | 35:15 | 13.29 | 0.064 | 9.97 | 7.92 | 1.9
2005-09-16 | 35:30 | 13.41 | 0.064 | 9.96 | 7.95 | 1.9
2005-09-16 | 35:45 | 13.54 | 0.063 | 9.94 | 7.98 | 1.9
2005-09-16 | 36:00 | 13.66 | 0.064 | 9.93 | 8.01 | 2
2005-09-16 | 36:15 | 13.66 | 0.064 | 9.87 | 8.01 | 1.8
2005-09-16 | 36:30 | 13.64 | 0.064 | 9.85 | 8.01 | 1.8
2005-09-16 | 36:45 | 13.64 | 0.064 | 9.86 | 8.03 | 1.8
2005-09-16 | 37:00 | 13.64 | 0.064 | 9.84 | 8.03 | 2.2
2005-09-16 | 37:15 | 13.7 | 0.064 | 9.85 | 8.05 | 1.8
2005-09-16 | 37:30 | 13.72 | 0.064 | 9.86 | 8.07 | 1.8
2005-09-16 | 37:45 | 13.72 | 0.064 | 9.81 | 8.08 | 1.8
2005-09-16 | 38:00 | 13.73 | 0.064 | 9.81 | 8.1 | 1.8
2005-09-16 | 38:15 | 13.79 | 0.064 | 9.84 | 8.13 | 2.1
2005-09-16 | 38:30 | 13.85 | 0.064 | 9.83 | 8.16 | 2
2005-09-16 | 38:45 | 13.9 | 0.064 | 9.81 | 8.18 | 1.9
2005-09-16 | 39:00 | 13.91 | 0.064 | 9.78 | 8.19 | 2
2005-09-16 | 39:15 | 13.9 | 0.064 | 9.75 | 8.2 | 1.9
2005-09-16 | 39:30 | 13.9 | 0.064 | 9.76 | 8.21 | 1.9
2005-09-16 | 39:45 | 13.91 | 0.064 | 9.77 | 8.22 | 1.8
2005-09-16 | 40:00 | 13.93 | 0.064 | 9.74 | 8.23 | 2
2005-09-16 | 40:15 | 13.91 | 0.064 | 9.72 | 8.23 | 2
2005-09-16 | 40:30 | 13.9 | 0.064 | 9.71 | 8.23 | 1.8
2005-09-16 | 40:45 | 13.9 | 0.064 | 9.73 | 8.23 | 1.9
2005-09-16 | 41:00 | 13.9 | 0.064 | 9.72 | 8.23 | 2
2005-09-16 | 41:15 | 13.89 | 0.064 | 9.72 | 8.23 | 2
2005-09-16 | 41:30 | 13.88 | 0.064 | 9.7 | 8.22 | 1.9
2005-09-16 | 41:45 | 13.87 | 0.064 | 9.68 | 8.19 | 2
2005-09-16 | 42:00 | 13.86 | 0.064 | 9.66 | 8.16 | 1.9
2005-09-16 | 42:15 | 13.84 | 0.064 | 9.67 | 8.13 | 1.9
2005-09-16 | 42:30 | 13.81 | 0.064 | 9.67 | 8.1 | 1.9
2005-09-16 | 42:45 | 13.78 | 0.064 | 9.68 | 8.07 | 2
2005-09-16 | 43:00 | 13.75 | 0.064 | 9.69 | 8.04 | 2.1
2005-09-16 | 43:15 | 13.71 | 0.064 | 9.69 | 8.02 | 2.2
2005-09-16 | 43:30 | 13.68 | 0.064 | 9.7 | 7.99 | 2.1
2005-09-16 | 43:45 | 13.65 | 0.065 | 9.7 | 7.97 | 2.1
2005-09-16 | 44:00 | 13.61 | 0.065 | 9.71 | 7.95 | 2.1
2005-09-16 | 44:15 | 13.57 | 0.065 | 9.73 | 7.92 | 2.2
2005-09-16 | 44:30 | 13.53 | 0.065 | 9.73 | 7.9 | 2.3
2005-09-16 | 44:45 | 13.5 | 0.065 | 9.74 | 7.88 | 2
2005-09-16 | 45:00 | 13.46 | 0.065 | 9.75 | 7.87 | 2
2005-09-16 | 45:15 | 13.43 | 0.065 | 9.76 | 7.85 | 2.1
2005-09-16 | 45:30 | 13.4 | 0.065 | 9.78 | 7.83 | 2.1
2005-09-16 | 45:45 | 13.36 | 0.065 | 9.77 | 7.82 | 2
2005-09-16 | 46:00 | 13.34 | 0.065 | 9.79 | 7.81 | 2
2005-09-16 | 46:15 | 13.32 | 0.065 | 9.79 | 7.8 | 2
2005-09-16 | 46:30 | 13.3 | 0.065 | 9.8 | 7.79 | 2
2005-09-16 | 46:45 | 13.29 | 0.065 | 9.8 | 7.78 | 2.1
2005-09-16 | 47:00 | 13.27 | 0.065 | 9.8 | 7.78 | 2
2005-09-16 | 47:15 | 13.24 | 0.065 | 9.81 | 7.77 | 2
2005-09-16 | 47:30 | 13.22 | 0.065 | 9.82 | 7.76 | 2
2005-09-16 | 47:45 | 13.2 | 0.065 | 9.82 | 7.75 | 2
2005-09-17 | 48:00 | 13.17 | 0.065 | 9.83 | 7.75 | 2
2005-09-17 | 48:15 | 13.14 | 0.065 | 9.84 | 7.74 | 1.9
2005-09-17 | 48:30 | 13.11 | 0.065 | 9.85 | 7.74 | 2
2005-09-17 | 48:45 | 13.07 | 0.065 | 9.86 | 7.73 | 2
2005-09-17 | 49:00 | 13.04 | 0.065 | 9.86 | 7.73 | 2.5
2005-09-17 | 49:15 | 13 | 0.065 | 9.87 | 7.72 | 2
2005-09-17 | 49:30 | 12.97 | 0.065 | 9.88 | 7.72 | 2
2005-09-17 | 49:45 | 12.94 | 0.065 | 9.89 | 7.72 | 2
2005-09-17 | 50:00 | 12.92 | 0.065 | 9.9 | 7.72 | 2.6
2005-09-17 | 50:15 | 12.89 | 0.065 | 9.9 | 7.71 | 2.1
2005-09-17 | 50:30 | 12.87 | 0.065 | 9.91 | 7.72 | 2.1
2005-09-17 | 50:45 | 12.85 | 0.065 | 9.92 | 7.72 | 2.1
2005-09-17 | 51:00 | 12.83 | 0.065 | 9.93 | 7.72 | 2.3
2005-09-17 | 51:15 | 12.82 | 0.065 | 9.93 | 7.72 | 2.1
2005-09-17 | 51:30 | 12.8 | 0.065 | 9.93 | 7.71 | 2.1
2005-09-17 | 51:45 | 12.79 | 0.065 | 9.93 | 7.71 | 2.1
2005-09-17 | 52:00 | 12.77 | 0.065 | 9.94 | 7.71 | 2.1
2005-09-17 | 52:15 | 12.75 | 0.065 | 9.95 | 7.71 | 2.1
2005-09-17 | 52:30 | 12.73 | 0.065 | 9.95 | 7.71 | 2.1
2005-09-17 | 52:45 | 12.71 | 0.065 | 9.95 | 7.71 | 2.1
2005-09-17 | 53:00 | 12.69 | 0.065 | 9.96 | 7.71 | 2.1
2005-09-17 | 53:15 | 12.67 | 0.065 | 9.97 | 7.71 | 2.2
2005-09-17 | 53:30 | 12.65 | 0.065 | 9.98 | 7.71 | 2.2
2005-09-17 | 53:45 | 12.64 | 0.065 | 9.99 | 7.71 | 2.2
2005-09-17 | 54:00 | 12.63 | 0.065 | 10 | 7.7 | 2.1
2005-09-17 | 54:15 | 12.62 | 0.065 | 9.99 | 7.7 | 2.1
2005-09-17 | 54:30 | 12.62 | 0.065 | 9.99 | 7.7 | 2.3
2005-09-17 | 54:45 | 12.61 | 0.065 | 10 | 7.7 | 2.3
2005-09-17 | 55:00 | 12.59 | 0.065 | 10.01 | 7.71 | 2.4
2005-09-17 | 55:15 | 12.58 | 0.065 | 10.01 | 7.71 | 2.2
2005-09-17 | 55:30 | 12.56 | 0.065 | 10.02 | 7.71 | 2.2
2005-09-17 | 55:45 | 12.55 | 0.065 | 10.03 | 7.71 | 2.2
2005-09-17 | 56:00 | 12.53 | 0.065 | 10.04 | 7.71 | 2.3
2005-09-17 | 56:15 | 12.52 | 0.065 | 10.05 | 7.72 | 2.2
2005-09-17 | 56:30 | 12.54 | 0.065 | 10.07 | 7.74 | 2.4
2005-09-17 | 56:45 | 12.57 | 0.065 | 10.08 | 7.76 | 2.4
2005-09-17 | 57:00 | 12.61 | 0.064 | 10.13 | 7.78 | 2.5
2005-09-17 | 57:15 | 12.63 | 0.064 | 10.12 | 7.8 | 2.5
2005-09-17 | 57:30 | 12.65 | 0.064 | 10.13 | 7.81 | 2.2
2005-09-17 | 57:45 | 12.65 | 0.064 | 10.14 | 7.82 | 2.3
2005-09-17 | 58:00 | 12.66 | 0.064 | 10.14 | 7.84 | 2.1
2005-09-17 | 58:15 | 12.71 | 0.064 | 10.16 | 7.86 | 2.2
2005-09-17 | 58:30 | 12.78 | 0.064 | 10.15 | 7.87 | 2.3
2005-09-17 | 58:45 | 12.87 | 0.064 | 10.14 | 7.9 | 2.5
2005-09-17 | 59:00 | 12.95 | 0.064 | 10.12 | 7.92 | 2.3
2005-09-17 | 59:15 | 13.05 | 0.064 | 10.13 | 7.95 | 2.2
2005-09-17 | 59:30 | 13.15 | 0.064 | 10.12 | 7.97 | 2.1
2005-09-17 | 59:45 | 13.26 | 0.064 | 10.11 | 8.01 | 2.1
2005-09-17 | 60:00 | 13.39 | 0.064 | 10.1 | 8.03 | 2.2
2005-09-17 | 60:15 | 13.51 | 0.064 | 10.04 | 8.06 | 2.2
2005-09-17 | 60:30 | 13.61 | 0.064 | 10.01 | 8.09 | 2.1
2005-09-17 | 60:45 | 13.67 | 0.064 | 9.96 | 8.11 | 2.2
2005-09-17 | 61:00 | 13.73 | 0.064 | 9.93 | 8.13 | 2.1
2005-09-17 | 61:15 | 13.81 | 0.064 | 9.93 | 8.16 | 2.1
2005-09-17 | 61:30 | 13.88 | 0.064 | 9.92 | 8.19 | 2.2
2005-09-17 | 61:45 | 13.91 | 0.064 | 9.91 | 8.21 | 2.2
2005-09-17 | 62:00 | 14.01 | 0.064 | 9.9 | 8.24 | 2.3
2005-09-17 | 62:15 | 14.11 | 0.064 | 9.88 | 8.28 | 2.3
2005-09-17 | 62:30 | 14.12 | 0.064 | 9.82 | 8.28 | 2.2
2005-09-17 | 62:45 | 14.21 | 0.064 | 9.8 | 8.31 | 2.3
2005-09-17 | 63:00 | 14.29 | 0.064 | 9.8 | 8.34 | 2.2
2005-09-17 | 63:15 | 14.31 | 0.064 | 9.76 | 8.36 | 2.2
2005-09-17 | 63:30 | 14.36 | 0.064 | 9.74 | 8.37 | 2.1
2005-09-17 | 63:45 | 14.36 | 0.064 | 9.7 | 8.38 | 2.1
2005-09-17 | 64:00 | 14.41 | 0.064 | 9.7 | 8.39 | 2.1
2005-09-17 | 64:15 | 14.41 | 0.064 | 9.71 | 8.39 | 2.1
2005-09-17 | 64:30 | 14.4 | 0.065 | 9.71 | 8.39 | 2.1
2005-09-17 | 64:45 | 14.37 | 0.065 | 9.68 | 8.39 | 2.1
2005-09-17 | 65:00 | 14.35 | 0.065 | 9.67 | 8.38 | 2.1
2005-09-17 | 65:15 | 14.34 | 0.065 | 9.66 | 8.39 | 2.2
2005-09-17 | 65:30 | 14.34 | 0.065 | 9.65 | 8.38 | 2.3
2005-09-17 | 65:45 | 14.33 | 0.065 | 9.64 | 8.38 | 2.1
2005-09-17 | 66:00 | 14.31 | 0.065 | 9.63 | 8.37 | 2.1
2005-09-17 | 66:15 | 14.29 | 0.065 | 9.61 | 8.35 | 2.3
2005-09-17 | 66:30 | 14.27 | 0.065 | 9.61 | 8.33 | 2.2
2005-09-17 | 66:45 | 14.24 | 0.065 | 9.6 | 8.3 | 2.2
2005-09-17 | 67:00 | 14.19 | 0.065 | 9.64 | 8.26 | 2.1
2005-09-17 | 67:15 | 14.13 | 0.065 | 9.62 | 8.22 | 2.1
2005-09-17 | 67:30 | 14.06 | 0.065 | 9.63 | 8.18 | 2.1
2005-09-17 | 67:45 | 13.98 | 0.065 | 9.64 | 8.14 | 2.2
2005-09-17 | 68:00 | 13.9 | 0.065 | 9.66 | 8.09 | 2.2
2005-09-17 | 68:15 | 13.82 | 0.065 | 9.69 | 8.06 | 2.3
2005-09-17 | 68:30 | 13.74 | 0.065 | 9.71 | 8.02 | 2.1
2005-09-17 | 68:45 | 13.66 | 0.065 | 9.73 | 7.99 | 2.4
2005-09-17 | 69:00 | 13.58 | 0.065 | 9.75 | 7.97 | 2.1
2005-09-17 | 69:15 | 13.51 | 0.065 | 9.76 | 7.94 | 2.3
2005-09-17 | 69:30 | 13.44 | 0.065 | 9.78 | 7.92 | 2.1
2005-09-17 | 69:45 | 13.38 | 0.065 | 9.8 | 7.91 | 2.3
2005-09-17 | 70:00 | 13.32 | 0.065 | 9.82 | 7.89 | 2.1
2005-09-17 | 70:15 | 13.26 | 0.065 | 9.83 | 7.87 | 2.1
2005-09-17 | 70:30 | 13.19 | 0.065 | 9.85 | 7.86 | 2.2
2005-09-17 | 70:45 | 13.11 | 0.065 | 9.87 | 7.84 | 2.2
2005-09-17 | 71:00 | 13.04 | 0.065 | 9.9 | 7.83 | 2.1
2005-09-17 | 71:15 | 12.96 | 0.066 | 9.91 | 7.81 | 2.2
2005-09-17 | 71:30 | 12.9 | 0.065 | 9.93 | 7.8 | 2.5
2005-09-17 | 71:45 | 12.83 | 0.065 | 9.94 | 7.79 | 2.2
2005-09-18 | 72:00 | 12.77 | 0.065 | 9.96 | 7.78 | 2.3
2005-09-18 | 72:15 | 12.7 | 0.065 | 9.98 | 7.77 | 2.1
2005-09-18 | 72:30 | 12.64 | 0.065 | 10 | 7.77 | 2.3
2005-09-18 | 72:45 | 12.59 | 0.065 | 10.01 | 7.76 | 2.1
2005-09-18 | 73:00 | 12.53 | 0.065 | 10.02 | 7.76 | 2.1
2005-09-18 | 73:15 | 12.47 | 0.065 | 10.03 | 7.75 | 2.3
2005-09-18 | 73:30 | 12.41 | 0.065 | 10.05 | 7.75 | 2.2
2005-09-18 | 73:45 | 12.36 | 0.065 | 10.1 | 7.74 | 2.1
2005-09-18 | 74:00 | 12.3 | 0.065 | 10.1 | 7.74 | 2.1
2005-09-18 | 74:15 | 12.24 | 0.065 | 10.1 | 7.74 | 2.1
2005-09-18 | 74:30 | 12.18 | 0.065 | 10.11 | 7.73 | 2.2
2005-09-18 | 74:45 | 12.13 | 0.065 | 10.12 | 7.73 | 2.3
2005-09-18 | 75:00 | 12.07 | 0.065 | 10.15 | 7.73 | 2.1
2005-09-18 | 75:15 | 12.02 | 0.065 | 10.15 | 7.73 | 2.2
2005-09-18 | 75:30 | 11.97 | 0.065 | 10.17 | 7.73 | 2.2
2005-09-18 | 75:45 | 11.92 | 0.065 | 10.18 | 7.73 | 2.3
2005-09-18 | 76:00 | 11.88 | 0.065 | 10.19 | 7.72 | 2.3
2005-09-18 | 76:15 | 11.83 | 0.065 | 10.21 | 7.72 | 2.3
2005-09-18 | 76:30 | 11.79 | 0.065 | 10.21 | 7.72 | 2.3
2005-09-18 | 76:45 | 11.75 | 0.065 | 10.22 | 7.72 | 2.3
2005-09-18 | 77:00 | 11.72 | 0.065 | 10.23 | 7.72 | 2.3
2005-09-18 | 77:15 | 11.69 | 0.065 | 10.24 | 7.72 | 2.3
2005-09-18 | 77:30 | 11.66 | 0.065 | 10.25 | 7.72 | 2.3
2005-09-18 | 77:45 | 11.64 | 0.065 | 10.25 | 7.72 | 2.3
2005-09-18 | 78:00 | 11.62 | 0.065 | 10.26 | 7.72 | 2.5
2005-09-18 | 78:15 | 11.6 | 0.065 | 10.27 | 7.72 | 2.3
2005-09-18 | 78:30 | 11.58 | 0.065 | 10.27 | 7.72 | 2.4
2005-09-18 | 78:45 | 11.57 | 0.065 | 10.28 | 7.71 | 2.3
2005-09-18 | 79:00 | 11.56 | 0.065 | 10.29 | 7.72 | 2.4
2005-09-18 | 79:15 | 11.55 | 0.065 | 10.29 | 7.72 | 2.5
2005-09-18 | 79:30 | 11.55 | 0.065 | 10.3 | 7.72 | 2.5
2005-09-18 | 79:45 | 11.55 | 0.064 | 10.32 | 7.73 | 2.5
2005-09-18 | 80:00 | 11.56 | 0.064 | 10.32 | 7.73 | 2.5
2005-09-18 | 80:15 | 11.57 | 0.064 | 10.32 | 7.74 | 2.5
2005-09-18 | 80:30 | 11.59 | 0.064 | 10.32 | 7.75 | 2.5
2005-09-18 | 80:45 | 11.62 | 0.064 | 10.33 | 7.76 | 2.5
2005-09-18 | 81:00 | 11.66 | 0.064 | 10.33 | 7.78 | 2.5
2005-09-18 | 81:15 | 11.7 | 0.064 | 10.33 | 7.79 | 2.6
2005-09-18 | 81:30 | 11.75 | 0.064 | 10.34 | 7.8 | 2.4
2005-09-18 | 81:45 | 11.8 | 0.064 | 10.34 | 7.82 | 2.5
2005-09-18 | 82:00 | 11.87 | 0.064 | 10.34 | 7.83 | 2.5
2005-09-18 | 82:15 | 11.94 | 0.064 | 10.35 | 7.85 | 2.5
2005-09-18 | 82:30 | 12.02 | 0.064 | 10.32 | 7.88 | 2.4
2005-09-18 | 82:45 | 12.11 | 0.064 | 10.32 | 7.89 | 2.5
2005-09-18 | 83:00 | 12.2 | 0.064 | 10.3 | 7.92 | 2.5
2005-09-18 | 83:15 | 12.29 | 0.064 | 10.29 | 7.94 | 2.5
2005-09-18 | 83:30 | 12.38 | 0.064 | 10.28 | 7.97 | 2.5
2005-09-18 | 83:45 | 12.47 | 0.064 | 10.26 | 7.99 | 2.4
2005-09-18 | 84:00 | 12.57 | 0.064 | 10.25 | 8.02 | 2.5
2005-09-18 | 84:15 | 12.66 | 0.064 | 10.22 | 8.04 | 2.4
2005-09-18 | 84:30 | 12.75 | 0.064 | 10.19 | 8.07 | 2.4
2005-09-18 | 84:45 | 12.85 | 0.064 | 10.18 | 8.1 | 2.3
2005-09-18 | 85:00 | 12.95 | 0.064 | 10.15 | 8.13 | 2.5
2005-09-18 | 85:15 | 13.05 | 0.064 | 10.13 | 8.16 | 2.3
2005-09-18 | 85:30 | 13.14 | 0.064 | 10.1 | 8.19 | 2.3
2005-09-18 | 85:45 | 13.24 | 0.064 | 10.07 | 8.22 | 2.4
2005-09-18 | 86:00 | 13.32 | 0.064 | 10.03 | 8.25 | 2.3
2005-09-18 | 86:15 | 13.41 | 0.064 | 10 | 8.27 | 2.3
2005-09-18 | 86:30 | 13.49 | 0.064 | 9.98 | 8.3 | 2.3
2005-09-18 | 86:45 | 13.57 | 0.064 | 9.96 | 8.32 | 2.3
2005-09-18 | 87:00 | 13.64 | 0.064 | 9.92 | 8.34 | 2.3
2005-09-18 | 87:15 | 13.71 | 0.064 | 9.89 | 8.35 | 2.3
2005-09-18 | 87:30 | 13.78 | 0.064 | 9.86 | 8.36 | 3.6
2005-09-18 | 87:45 | 13.84 | 0.064 | 9.83 | 8.37 | 2.4
2005-09-18 | 88:00 | 13.9 | 0.064 | 9.8 | 8.38 | 2.5
2005-09-18 | 88:15 | 13.96 | 0.064 | 9.77 | 8.38 | 2.5
2005-09-18 | 88:30 | 14.01 | 0.064 | 9.74 | 8.39 | 2.5
2005-09-18 | 88:45 | 14.05 | 0.064 | 9.71 | 8.39 | 2.5
2005-09-18 | 89:00 | 14.08 | 0.064 | 9.68 | 8.39 | 2.7
2005-09-18 | 89:15 | 14.11 | 0.064 | 9.65 | 8.38 | 2.5
2005-09-18 | 89:30 | 14.14 | 0.064 | 9.63 | 8.38 | 2.4
2005-09-18 | 89:45 | 14.17 | 0.064 | 9.61 | 8.37 | 2.6
2005-09-18 | 90:00 | 14.19 | 0.064 | 9.6 | 8.36 | 2.4
2005-09-18 | 90:15 | 14.2 | 0.064 | 9.59 | 8.35 | 2.5
2005-09-18 | 90:30 | 14.21 | 0.064 | 9.58 | 8.33 | 2.5
2005-09-18 | 90:45 | 14.21 | 0.064 | 9.58 | 8.31 | 2.6
2005-09-18 | 91:00 | 14.19 | 0.064 | 9.57 | 8.28 | 2.5
2005-09-18 | 91:15 | 14.17 | 0.065 | 9.57 | 8.23 | 2.4
2005-09-18 | 91:30 | 14.14 | 0.065 | 9.58 | 8.19 | 2.5
2005-09-18 | 91:45 | 14.1 | 0.065 | 9.59 | 8.15 | 2.5
2005-09-18 | 92:00 | 14.06 | 0.065 | 9.6 | 8.11 | 2.8
2005-09-18 | 92:15 | 14.02 | 0.065 | 9.62 | 8.08 | 2.5
2005-09-18 | 92:30 | 13.99 | 0.065 | 9.64 | 8.05 | 2.4
2005-09-18 | 92:45 | 13.94 | 0.065 | 9.64 | 8.02 | 2.3
2005-09-18 | 93:00 | 13.89 | 0.065 | 9.65 | 7.98 | 2.3
2005-09-18 | 93:15 | 13.83 | 0.065 | 9.66 | 7.95 | 2.3
2005-09-18 | 93:30 | 13.77 | 0.065 | 9.69 | 7.93 | 2.3
2005-09-18 | 93:45 | 13.71 | 0.065 | 9.7 | 7.91 | 2.5
2005-09-18 | 94:00 | 13.63 | 0.065 | 9.72 | 7.88 | 2.3
2005-09-18 | 94:15 | 13.56 | 0.065 | 9.74 | 7.86 | 2.4
2005-09-18 | 94:30 | 13.48 | 0.065 | 9.77 | 7.84 | 2.3
2005-09-18 | 94:45 | 13.39 | 0.065 | 9.79 | 7.82 | 2.3
2005-09-18 | 95:00 | 13.31 | 0.065 | 9.81 | 7.81 | 2.4
2005-09-18 | 95:15 | 13.23 | 0.065 | 9.83 | 7.8 | 2.3
2005-09-18 | 95:30 | 13.15 | 0.065 | 9.85 | 7.79 | 2.2
2005-09-18 | 95:45 | 13.07 | 0.065 | 9.87 | 7.78 | 2.2
2005-09-19 | 96:00 | 13 | 0.065 | 9.89 | 7.77 | 2.2
2005-09-19 | 96:15 | 12.93 | 0.065 | 9.9 | 7.76 | 2.3
2005-09-19 | 96:30 | 12.86 | 0.065 | 9.92 | 7.75 | 2.3
2005-09-19 | 96:45 | 12.79 | 0.065 | 9.94 | 7.75 | 2.3
2005-09-19 | 97:00 | 12.73 | 0.065 | 9.96 | 7.75 | 2.2
2005-09-19 | 97:15 | 12.66 | 0.065 | 9.98 | 7.75 | 2.2
2005-09-19 | 97:30 | 12.6 | 0.065 | 9.99 | 7.74 | 2.2
2005-09-19 | 97:45 | 12.54 | 0.065 | 10 | 7.74 | 2.3
2005-09-19 | 98:00 | 12.49 | 0.065 | 10.02 | 7.73 | 2.2
2005-09-19 | 98:15 | 12.44 | 0.065 | 10.04 | 7.73 | 2.3
2005-09-19 | 98:30 | 12.39 | 0.065 | 10.05 | 7.73 | 2.3
2005-09-19 | 98:45 | 12.34 | 0.065 | 10.06 | 7.72 | 2.2
2005-09-19 | 99:00 | 12.3 | 0.065 | 10.07 | 7.73 | 2.3
2005-09-19 | 99:15 | 12.25 | 0.065 | 10.08 | 7.72 | 2.2
2005-09-19 | 99:30 | 12.21 | 0.065 | 10.09 | 7.72 | 2.3
2005-09-19 | 99:45 | 12.17 | 0.065 | 10.1 | 7.72 | 2.2
2005-09-19 | 100:00 | 12.12 | 0.065 | 10.11 | 7.72 | 2.2
2005-09-19 | 100:15 | 12.08 | 0.065 | 10.12 | 7.72 | 2.3
2005-09-19 | 100:30 | 12.05 | 0.065 | 10.14 | 7.72 | 2.3
2005-09-19 | 100:45 | 12.01 | 0.065 | 10.15 | 7.71 | 2.3
2005-09-19 | 101:00 | 11.97 | 0.065 | 10.15 | 7.71 | 2.3
2005-09-19 | 101:15 | 11.95 | 0.065 | 10.16 | 7.71 | 2.3
2005-09-19 | 101:30 | 11.92 | 0.065 | 10.17 | 7.71 | 2.4
2005-09-19 | 101:45 | 11.89 | 0.065 | 10.18 | 7.71 | 2.3
2005-09-19 | 102:00 | 11.87 | 0.065 | 10.18 | 7.71 | 2.4
2005-09-19 | 102:15 | 11.84 | 0.064 | 10.19 | 7.71 | 2.4
2005-09-19 | 102:30 | 11.82 | 0.064 | 10.2 | 7.71 | 2.3
2005-09-19 | 102:45 | 11.79 | 0.064 | 10.2 | 7.71 | 2.3
2005-09-19 | 103:00 | 11.77 | 0.064 | 10.21 | 7.71 | 2.3
2005-09-19 | 103:15 | 11.75 | 0.064 | 10.23 | 7.71 | 2.5
2005-09-19 | 103:30 | 11.73 | 0.064 | 10.24 | 7.72 | 2.5
2005-09-19 | 103:45 | 11.73 | 0.064 | 10.26 | 7.72 | 2.6
2005-09-19 | 104:00 | 11.72 | 0.064 | 10.27 | 7.73 | 2.6
2005-09-19 | 104:15 | 11.72 | 0.064 | 10.28 | 7.73 | 2.6
2005-09-19 | 104:30 | 11.73 | 0.064 | 10.29 | 7.75 | 2.6
2005-09-19 | 104:45 | 11.74 | 0.064 | 10.3 | 7.76 | 2.6
2005-09-19 | 105:00 | 11.77 | 0.064 | 10.3 | 7.77 | 2.5
2005-09-19 | 105:15 | 11.8 | 0.064 | 10.31 | 7.78 | 2.5
2005-09-19 | 105:30 | 11.84 | 0.064 | 10.3 | 7.8 | 2.5
2005-09-19 | 105:45 | 11.88 | 0.064 | 10.31 | 7.81 | 2.3
2005-09-19 | 106:00 | 11.94 | 0.064 | 10.31 | 7.83 | 2.4
2005-09-19 | 106:15 | 12.01 | 0.064 | 10.3 | 7.85 | 2.3
2005-09-19 | 106:30 | 12.09 | 0.064 | 10.3 | 7.87 | 2.5
2005-09-19 | 106:45 | 12.17 | 0.064 | 10.28 | 7.88 | 2.4
2005-09-19 | 107:00 | 12.25 | 0.064 | 10.27 | 7.91 | 2.3
2005-09-19 | 107:15 | 12.35 | 0.064 | 10.25 | 7.93 | 2.3
2005-09-19 | 107:30 | 12.45 | 0.064 | 10.23 | 7.97 | 2.3
2005-09-19 | 107:45 | 12.54 | 0.064 | 10.21 | 7.98 | 2.2
2005-09-19 | 108:00 | 12.64 | 0.064 | 10.19 | 8.01 | 2.3
2005-09-19 | 108:15 | 12.74 | 0.064 | 10.16 | 8.04 | 2.3
2005-09-19 | 108:30 | 12.85 | 0.063 | 10.14 | 8.07 | 2.2
2005-09-19 | 108:45 | 12.95 | 0.063 | 10.12 | 8.1 | 2.1
2005-09-19 | 109:00 | 13.06 | 0.063 | 10.09 | 8.13 | 2.1
2005-09-19 | 109:15 | 13.16 | 0.063 | 10.06 | 8.16 | 2.1
2005-09-19 | 109:30 | 13.27 | 0.063 | 10.04 | 8.19 | 2.2
2005-09-19 | 109:45 | 13.37 | 0.063 | 10.01 | 8.22 | 2.1
2005-09-19 | 110:00 | 13.46 | 0.063 | 9.98 | 8.25 | 2.2
2005-09-19 | 110:15 | 13.55 | 0.063 | 9.95 | 8.28 | 2.1
2005-09-19 | 110:30 | 13.64 | 0.063 | 9.93 | 8.3 | 2.1
2005-09-19 | 110:45 | 13.72 | 0.063 | 9.9 | 8.32 | 2.1
2005-09-19 | 111:00 | 13.79 | 0.064 | 9.87 | 8.33 | 2.1
2005-09-19 | 111:15 | 13.86 | 0.064 | 9.84 | 8.35 | 2.2
2005-09-19 | 111:30 | 13.92 | 0.064 | 9.8 | 8.36 | 2.1
2005-09-19 | 111:45 | 13.98 | 0.064 | 9.78 | 8.38 | 2.2
2005-09-19 | 112:00 | 14.03 | 0.064 | 9.74 | 8.4 | 2.3
2005-09-19 | 112:15 | 14.08 | 0.064 | 9.72 | 8.41 | 2.3
2005-09-19 | 112:30 | 14.13 | 0.064 | 9.69 | 8.41 | 2.2
2005-09-19 | 112:45 | 14.16 | 0.064 | 9.66 | 8.41 | 2.6
2005-09-19 | 113:00 | 14.19 | 0.064 | 9.64 | 8.41 | 2.3
2005-09-19 | 113:15 | 14.21 | 0.064 | 9.61 | 8.41 | 2.4
2005-09-19 | 113:30 | 14.23 | 0.064 | 9.6 | 8.4 | 2.5
2005-09-19 | 113:45 | 14.23 | 0.064 | 9.59 | 8.39 | 2.3
2005-09-19 | 114:00 | 14.24 | 0.064 | 9.58 | 8.38 | 2.3
2005-09-19 | 114:15 | 14.24 | 0.064 | 9.58 | 8.36 | 2.5
2005-09-19 | 114:30 | 14.23 | 0.064 | 9.58 | 8.35 | 2.2
2005-09-19 | 114:45 | 14.22 | 0.064 | 9.58 | 8.32 | 2.3
2005-09-19 | 115:00 | 14.19 | 0.064 | 9.58 | 8.29 | 2.4
2005-09-19 | 115:15 | 14.17 | 0.064 | 9.58 | 8.25 | 2.3
2005-09-19 | 115:30 | 14.14 | 0.064 | 9.58 | 8.21 | 2.5
2005-09-19 | 115:45 | 14.1 | 0.064 | 9.58 | 8.17 | 2.5
2005-09-19 | 116:00 | 14.07 | 0.064 | 9.6 | 8.13 | 2.6
2005-09-19 | 116:15 | 14.03 | 0.064 | 9.62 | 8.08 | 2.4
2005-09-19 | 116:30 | 13.98 | 0.065 | 9.63 | 8.05 | 2.3
2005-09-19 | 116:45 | 13.93 | 0.065 | 9.65 | 8.02 | 2.3
2005-09-19 | 117:00 | 13.87 | 0.065 | 9.66 | 7.98 | 2.3
2005-09-19 | 117:15 | 13.81 | 0.065 | 9.67 | 7.95 | 2.3
2005-09-19 | 117:30 | 13.75 | 0.065 | 9.7 | 7.93 | 2.3
2005-09-19 | 117:45 | 13.69 | 0.065 | 9.71 | 7.9 | 2.4
2005-09-19 | 118:00 | 13.62 | 0.065 | 9.73 | 7.87 | 2.3
2005-09-19 | 118:15 | 13.55 | 0.065 | 9.74 | 7.85 | 2.3
2005-09-19 | 118:30 | 13.47 | 0.065 | 9.76 | 7.83 | 2.2
2005-09-19 | 118:45 | 13.4 | 0.065 | 9.78 | 7.81 | 2.5
2005-09-19 | 119:00 | 13.32 | 0.065 | 9.8 | 7.8 | 2.3
2005-09-19 | 119:15 | 13.25 | 0.065 | 9.81 | 7.79 | 2.5
2005-09-19 | 119:30 | 13.18 | 0.065 | 9.83 | 7.78 | 2.2
2005-09-19 | 119:45 | 13.12 | 0.065 | 9.85 | 7.77 | 2.3
2005-09-20 | 120:00 | 13.06 | 0.065 | 9.86 | 7.76 | 2.3
2005-09-20 | 120:15 | 13 | 0.065 | 9.88 | 7.75 | 2.3
2005-09-20 | 120:30 | 12.94 | 0.065 | 9.89 | 7.75 | 2.3
2005-09-20 | 120:45 | 12.89 | 0.065 | 9.91 | 7.74 | 2.3
2005-09-20 | 121:00 | 12.83 | 0.065 | 9.92 | 7.74 | 2.3
2005-09-20 | 121:15 | 12.78 | 0.065 | 9.94 | 7.74 | 2.3
2005-09-20 | 121:30 | 12.73 | 0.065 | 9.95 | 7.73 | 2.3
2005-09-20 | 121:45 | 12.68 | 0.065 | 9.96 | 7.73 | 2.3
2005-09-20 | 122:00 | 12.65 | 0.065 | 9.97 | 7.73 | 2.5
2005-09-20 | 122:15 | 12.6 | 0.065 | 9.98 | 7.73 | 2.3
2005-09-20 | 122:30 | 12.56 | 0.065 | 9.99 | 7.73 | 2.3
2005-09-20 | 122:45 | 12.52 | 0.065 | 9.99 | 7.72 | 2.2
2005-09-20 | 123:00 | 12.49 | 0.065 | 10 | 7.72 | 2.3
2005-09-20 | 123:15 | 12.46 | 0.065 | 10.02 | 7.72 | 2.5
2005-09-20 | 123:30 | 12.43 | 0.064 | 10.02 | 7.72 | 2.3
2005-09-20 | 123:45 | 12.4 | 0.064 | 10.02 | 7.72 | 2.3
2005-09-20 | 124:00 | 12.37 | 0.064 | 10.03 | 7.72 | 2.3
2005-09-20 | 124:15 | 12.35 | 0.064 | 10.04 | 7.71 | 2.4
2005-09-20 | 124:30 | 12.32 | 0.064 | 10.04 | 7.71 | 2.4
2005-09-20 | 124:45 | 12.3 | 0.064 | 10.05 | 7.71 | 2.4
2005-09-20 | 125:00 | 12.28 | 0.064 | 10.05 | 7.72 | 2.5
2005-09-20 | 125:15 | 12.26 | 0.064 | 10.06 | 7.71 | 2.4
2005-09-20 | 125:30 | 12.24 | 0.064 | 10.06 | 7.71 | 2.4
2005-09-20 | 125:45 | 12.22 | 0.064 | 10.07 | 7.71 | 2.5
2005-09-20 | 126:00 | 12.2 | 0.064 | 10.07 | 7.71 | 2.5
2005-09-20 | 126:15 | 12.18 | 0.064 | 10.08 | 7.71 | 2.5
2005-09-20 | 126:30 | 12.17 | 0.064 | 10.09 | 7.71 | 2.5
2005-09-20 | 126:45 | 12.15 | 0.064 | 10.09 | 7.7 | 2.5
2005-09-20 | 127:00 | 12.14 | 0.064 | 10.1 | 7.71 | 2.6
2005-09-20 | 127:15 | 12.13 | 0.064 | 10.11 | 7.71 | 2.6
2005-09-20 | 127:30 | 12.13 | 0.064 | 10.12 | 7.71 | 2.7
2005-09-20 | 127:45 | 12.13 | 0.064 | 10.13 | 7.72 | 2.6
2005-09-20 | 128:00 | 12.13 | 0.064 | 10.13 | 7.72 | 2.7
2005-09-20 | 128:15 | 12.13 | 0.064 | 10.13 | 7.72 | 2.7
2005-09-20 | 128:30 | 12.15 | 0.064 | 10.13 | 7.73 | 2.8
2005-09-20 | 128:45 | 12.18 | 0.064 | 10.14 | 7.75 | 2.5
2005-09-20 | 129:00 | 12.22 | 0.064 | 10.15 | 7.76 | 2.4
2005-09-20 | 129:15 | 12.26 | 0.064 | 10.15 | 7.78 | 2.5
2005-09-20 | 129:30 | 12.31 | 0.064 | 10.15 | 7.79 | 2.4
2005-09-20 | 129:45 | 12.36 | 0.064 | 10.15 | 7.8 | 2.5
2005-09-20 | 130:00 | 12.42 | 0.064 | 10.15 | 7.82 | 2.4
2005-09-20 | 130:15 | 12.49 | 0.063 | 10.14 | 7.84 | 2.5
2005-09-20 | 130:30 | 12.57 | 0.063 | 10.13 | 7.87 | 2.4
2005-09-20 | 130:45 | 12.65 | 0.063 | 10.12 | 7.89 | 2.4
2005-09-20 | 131:00 | 12.73 | 0.063 | 10.11 | 7.91 | 2.4
2005-09-20 | 131:15 | 12.83 | 0.063 | 10.1 | 7.93 | 2.4
2005-09-20 | 131:30 | 12.92 | 0.063 | 10.08 | 7.96 | 2.4
2005-09-20 | 131:45 | 13.02 | 0.063 | 10.06 | 7.98 | 2.5
2005-09-20 | 132:00 | 13.13 | 0.063 | 10.03 | 8.01 | 2.3
2005-09-20 | 132:15 | 13.23 | 0.063 | 10 | 8.04 | 3.4
2005-09-20 | 132:30 | 13.33 | 0.063 | 9.99 | 8.07 | 2.4
2005-09-20 | 132:45 | 13.44 | 0.063 | 9.96 | 8.11 | 2.3
2005-09-20 | 133:00 | 13.54 | 0.063 | 9.94 | 8.14 | 2.4
2005-09-20 | 133:15 | 13.64 | 0.063 | 9.91 | 8.17 | 2.4
2005-09-20 | 133:30 | 13.74 | 0.063 | 9.87 | 8.21 | 2.3
2005-09-20 | 133:45 | 13.84 | 0.063 | 9.84 | 8.24 | 2.4
2005-09-20 | 134:00 | 13.93 | 0.063 | 9.82 | 8.26 | 2.4
2005-09-20 | 134:15 | 14.01 | 0.063 | 9.79 | 8.29 | 2.3
2005-09-20 | 134:30 | 14.09 | 0.063 | 9.77 | 8.32 | 2.3
2005-09-20 | 134:45 | 14.16 | 0.063 | 9.74 | 8.34 | 2.4
2005-09-20 | 135:00 | 14.22 | 0.063 | 9.71 | 8.36 | 2.4
2005-09-20 | 135:15 | 14.27 | 0.063 | 9.68 | 8.37 | 2.3
2005-09-20 | 135:30 | 14.32 | 0.064 | 9.66 | 8.38 | 2.4
2005-09-20 | 135:45 | 14.37 | 0.064 | 9.64 | 8.4 | 2.4
2005-09-20 | 136:00 | 14.42 | 0.064 | 9.62 | 8.4 | 2.4
2005-09-20 | 136:15 | 14.46 | 0.064 | 9.59 | 8.41 | 2.5
2005-09-20 | 136:30 | 14.47 | 0.064 | 9.57 | 8.42 | 2.5
2005-09-20 | 136:45 | 14.47 | 0.064 | 9.53 | 8.41 | 2.5
2005-09-20 | 137:00 | 14.48 | 0.064 | 9.52 | 8.42 | 2.6
2005-09-20 | 137:15 | 14.47 | 0.064 | 9.52 | 8.41 | 2.5
2005-09-20 | 137:30 | 14.47 | 0.064 | 9.52 | 8.39 | 2.5
2005-09-20 | 137:45 | 14.46 | 0.064 | 9.5 | 8.39 | 2.5
2005-09-20 | 138:00 | 14.45 | 0.064 | 9.49 | 8.38 | 2.9
2005-09-20 | 138:15 | 14.43 | 0.064 | 9.48 | 8.36 | 2.8
2005-09-20 | 138:30 | 14.41 | 0.064 | 9.48 | 8.34 | 2.8
2005-09-20 | 138:45 | 14.38 | 0.064 | 9.48 | 8.31 | 2.8
2005-09-20 | 139:00 | 14.34 | 0.064 | 9.49 | 8.28 | 2.6
2005-09-20 | 139:15 | 14.31 | 0.064 | 9.49 | 8.25 | 2.6
2005-09-20 | 139:30 | 14.26 | 0.064 | 9.51 | 8.2 | 2.7
2005-09-20 | 139:45 | 14.21 | 0.065 | 9.51 | 8.16 | 2.8
2005-09-20 | 140:00 | 14.16 | 0.065 | 9.53 | 8.12 | 2.8
2005-09-20 | 140:15 | 14.11 | 0.065 | 9.55 | 8.09 | 2.8
2005-09-20 | 140:30 | 14.05 | 0.065 | 9.56 | 8.05 | 2.9
2005-09-20 | 140:45 | 13.99 | 0.065 | 9.58 | 8.01 | 2.9
2005-09-20 | 141:00 | 13.92 | 0.065 | 9.59 | 7.98 | 3.1
2005-09-20 | 141:15 | 13.85 | 0.065 | 9.61 | 7.96 | 3.1
2005-09-20 | 141:30 | 13.78 | 0.065 | 9.63 | 7.92 | 2.8
2005-09-20 | 141:45 | 13.7 | 0.065 | 9.65 | 7.9 | 2.8
2005-09-20 | 142:00 | 13.63 | 0.065 | 9.68 | 7.88 | 2.9
2005-09-20 | 142:15 | 13.55 | 0.065 | 9.7 | 7.86 | 2.9
2005-09-20 | 142:30 | 13.47 | 0.065 | 9.72 | 7.85 | 2.8
2005-09-20 | 142:45 | 13.39 | 0.065 | 9.75 | 7.83 | 2.8
2005-09-20 | 143:00 | 13.31 | 0.065 | 9.76 | 7.82 | 2.8
2005-09-20 | 143:15 | 13.23 | 0.065 | 9.78 | 7.8 | 2.7
2005-09-20 | 143:30 | 13.15 | 0.065 | 9.79 | 7.79 | 2.8
2005-09-20 | 143:45 | 13.07 | 0.065 | 9.82 | 7.79 | 2.8
2005-09-21 | 144:00 | 13.01 | 0.065 | 9.82 | 7.78 | 2.8
2005-09-21 | 144:15 | 12.94 | 0.065 | 9.85 | 7.77 | 2.7
2005-09-21 | 144:30 | 12.88 | 0.065 | 9.86 | 7.76 | 2.7
2005-09-21 | 144:45 | 12.82 | 0.065 | 9.88 | 7.76 | 2.7
2005-09-21 | 145:00 | 12.76 | 0.065 | 9.89 | 7.75 | 2.8
2005-09-21 | 145:15 | 12.7 | 0.065 | 9.91 | 7.75 | 2.8
2005-09-21 | 145:30 | 12.65 | 0.065 | 9.91 | 7.74 | 2.8
2005-09-21 | 145:45 | 12.61 | 0.065 | 9.92 | 7.74 | 2.8
2005-09-21 | 146:00 | 12.56 | 0.065 | 9.93 | 7.73 | 2.8
2005-09-21 | 146:15 | 12.52 | 0.065 | 9.96 | 7.73 | 2.8
2005-09-21 | 146:30 | 12.48 | 0.065 | 9.95 | 7.73 | 2.8
2005-09-21 | 146:45 | 12.44 | 0.065 | 9.96 | 7.73 | 2.6
2005-09-21 | 147:00 | 12.4 | 0.065 | 9.97 | 7.73 | 2.8
2005-09-21 | 147:15 | 12.37 | 0.065 | 9.98 | 7.72 | 2.8
2005-09-21 | 147:30 | 12.33 | 0.065 | 9.99 | 7.72 | 2.7
2005-09-21 | 147:45 | 12.3 | 0.065 | 10 | 7.72 | 2.8
2005-09-21 | 148:00 | 12.27 | 0.065 | 10 | 7.72 | 2.7
2005-09-21 | 148:15 | 12.24 | 0.065 | 10.01 | 7.72 | 2.7
2005-09-21 | 148:30 | 12.21 | 0.065 | 10.01 | 7.71 | 2.6
2005-09-21 | 148:45 | 12.18 | 0.065 | 10.02 | 7.71 | 2.7
2005-09-21 | 149:00 | 12.15 | 0.065 | 10.03 | 7.71 | 2.8
2005-09-21 | 149:15 | 12.13 | 0.065 | 10.04 | 7.71 | 2.6
2005-09-21 | 149:30 | 12.1 | 0.065 | 10.05 | 7.71 | 2.6
2005-09-21 | 149:45 | 12.07 | 0.065 | 10.05 | 7.7 | 2.7
2005-09-21 | 150:00 | 12.04 | 0.065 | 10.06 | 7.7 | 2.7
2005-09-21 | 150:15 | 12.01 | 0.065 | 10.06 | 7.7 | 2.8
2005-09-21 | 150:30 | 11.98 | 0.065 | 10.07 | 7.7 | 2.6
2005-09-21 | 150:45 | 11.95 | 0.065 | 10.08 | 7.7 | 2.7
2005-09-21 | 151:00 | 11.92 | 0.065 | 10.09 | 7.7 | 2.6
2005-09-21 | 151:15 | 11.89 | 0.065 | 10.1 | 7.69 | 2.7
2005-09-21 | 151:30 | 11.87 | 0.065 | 10.11 | 7.69 | 2.8
2005-09-21 | 151:45 | 11.84 | 0.065 | 10.13 | 7.7 | 2.8
2005-09-21 | 152:00 | 11.83 | 0.064 | 10.15 | 7.7 | 2.9
2005-09-21 | 152:15 | 11.83 | 0.064 | 10.16 | 7.72 | 2.7
2005-09-21 | 152:30 | 11.84 | 0.064 | 10.17 | 7.73 | 2.8
2005-09-21 | 152:45 | 11.85 | 0.064 | 10.17 | 7.74 | 2.8
2005-09-21 | 153:00 | 11.87 | 0.064 | 10.18 | 7.76 | 2.6
2005-09-21 | 153:15 | 11.9 | 0.064 | 10.19 | 7.77 | 2.7
2005-09-21 | 153:30 | 11.94 | 0.064 | 10.2 | 7.79 | 2.9
2005-09-21 | 153:45 | 11.98 | 0.064 | 10.2 | 7.8 | 2.8
2005-09-21 | 154:00 | 12.02 | 0.064 | 10.2 | 7.82 | 2.8
2005-09-21 | 154:15 | 12.07 | 0.064 | 10.21 | 7.84 | 2.8
2005-09-21 | 154:30 | 12.14 | 0.064 | 10.2 | 7.86 | 2.8
2005-09-21 | 154:45 | 12.22 | 0.064 | 10.19 | 7.88 | 2.9
2005-09-21 | 155:00 | 12.3 | 0.064 | 10.18 | 7.9 | 2.9
2005-09-21 | 155:15 | 12.38 | 0.064 | 10.17 | 7.92 | 2.8
2005-09-21 | 155:30 | 12.47 | 0.064 | 10.15 | 7.95 | 3
2005-09-21 | 155:45 | 12.57 | 0.064 | 10.14 | 7.97 | 2.9
2005-09-21 | 156:00 | 12.67 | 0.064 | 10.11 | 7.99 | 2.9
2005-09-21 | 156:15 | 12.77 | 0.064 | 10.09 | 8.02 | 2.9
2005-09-21 | 156:30 | 12.87 | 0.064 | 10.06 | 8.05 | 2.8
2005-09-21 | 156:45 | 12.98 | 0.064 | 10.04 | 8.08 | 2.8
2005-09-21 | 157:00 | 13.08 | 0.064 | 10.01 | 8.11 | 2.9
2005-09-21 | 157:15 | 13.19 | 0.064 | 9.98 | 8.15 | 2.8
2005-09-21 | 157:30 | 13.29 | 0.064 | 9.95 | 8.17 | 2.8
2005-09-21 | 157:45 | 13.39 | 0.064 | 9.93 | 8.2 | 2.8
2005-09-21 | 158:00 | 13.49 | 0.064 | 9.89 | 8.24 | 3.7
2005-09-21 | 158:15 | 13.57 | 0.064 | 9.87 | 8.26 | 2.8
2005-09-21 | 158:30 | 13.65 | 0.064 | 9.83 | 8.29 | 2.8
2005-09-21 | 158:45 | 13.72 | 0.064 | 9.81 | 8.31 | 2.8
2005-09-21 | 159:00 | 13.79 | 0.064 | 9.78 | 8.34 | 2.8
2005-09-21 | 159:15 | 13.84 | 0.064 | 9.75 | 8.35 | 2.8
2005-09-21 | 159:30 | 13.9 | 0.064 | 9.72 | 8.37 | 2.8
2005-09-21 | 159:45 | 13.95 | 0.064 | 9.7 | 8.38 | 2.6
2005-09-21 | 160:00 | 13.99 | 0.064 | 9.67 | 8.39 | 2.8
2005-09-21 | 160:15 | 14.03 | 0.064 | 9.64 | 8.39 | 2.8
2005-09-21 | 160:30 | 14.06 | 0.064 | 9.62 | 8.4 | 2.8
2005-09-21 | 160:45 | 14.09 | 0.064 | 9.59 | 8.4 | 2.8
2005-09-21 | 161:00 | 14.11 | 0.064 | 9.57 | 8.4 | 2.7
2005-09-21 | 161:15 | 14.12 | 0.064 | 9.54 | 8.4 | 2.9
2005-09-21 | 161:30 | 14.13 | 0.064 | 9.53 | 8.39 | 3
2005-09-21 | 161:45 | 14.13 | 0.064 | 9.52 | 8.38 | 2.9
2005-09-21 | 162:00 | 14.14 | 0.064 | 9.51 | 8.38 | 2.9
2005-09-21 | 162:15 | 14.14 | 0.064 | 9.51 | 8.37 | 2.9
2005-09-21 | 162:30 | 14.12 | 0.064 | 9.5 | 8.34 | 2.9
2005-09-21 | 162:45 | 14.1 | 0.064 | 9.5 | 8.32 | 2.9
2005-09-21 | 163:00 | 14.07 | 0.065 | 9.5 | 8.29 | 2.9
2005-09-21 | 163:15 | 14.03 | 0.065 | 9.5 | 8.25 | 2.9
2005-09-21 | 163:30 | 13.99 | 0.065 | 9.51 | 8.21 | 2.9
2005-09-21 | 163:45 | 13.94 | 0.065 | 9.52 | 8.17 | 2.8
2005-09-21 | 164:00 | 13.89 | 0.065 | 9.53 | 8.14 | 2.7
2005-09-21 | 164:15 | 13.84 | 0.065 | 9.55 | 8.1 | 2.8
2005-09-21 | 164:30 | 13.78 | 0.065 | 9.57 | 8.07 | 2.9
2005-09-21 | 164:45 | 13.72 | 0.065 | 9.58 | 8.03 | 2.8
2005-09-21 | 165:00 | 13.65 | 0.065 | 9.6 | 8 | 5.3
2005-09-21 | 165:15 | 13.58 | 0.065 | 9.62 | 7.97 | 2.8
2005-09-21 | 165:30 | 13.51 | 0.065 | 9.64 | 7.93 | 2.8
2005-09-21 | 165:45 | 13.43 | 0.065 | 9.66 | 7.91 | 2.6
2005-09-21 | 166:00 | 13.34 | 0.065 | 9.68 | 7.88 | 2.7
2005-09-21 | 166:15 | 13.25 | 0.065 | 9.7 | 7.86 | 2.7
2005-09-21 | 166:30 | 13.16 | 0.065 | 9.72 | 7.84 | 2.8
2005-09-21 | 166:45 | 13.06 | 0.065 | 9.75 | 7.82 | 2.8
2005-09-21 | 167:00 | 12.97 | 0.065 | 9.78 | 7.81 | 2.6
2005-09-21 | 167:15 | 12.87 | 0.065 | 9.8 | 7.8 | 2.5
2005-09-21 | 167:30 | 12.78 | 0.065 | 9.82 | 7.79 | 2.7
2005-09-21 | 167:45 | 12.7 | 0.065 | 9.84 | 7.78 | 2.3
2005-09-22 | 168:00 | 12.61 | 0.065 | 9.85 | 7.77 | 2.3
2005-09-22 | 168:15 | 12.54 | 0.065 | 9.87 | 7.76 | 2.5
2005-09-22 | 168:30 | 12.46 | 0.065 | 9.89 | 7.75 | 2.3
2005-09-22 | 168:45 | 12.39 | 0.065 | 9.91 | 7.75 | 2.9
2005-09-22 | 169:00 | 12.33 | 0.065 | 9.92 | 7.75 | 2.2
2005-09-22 | 169:15 | 12.26 | 0.065 | 9.94 | 7.74 | 2.3
2005-09-22 | 169:30 | 12.2 | 0.065 | 9.95 | 7.73 | 2.3
2005-09-22 | 169:45 | 12.14 | 0.065 | 9.97 | 7.73 | 2.3
2005-09-22 | 170:00 | 12.08 | 0.065 | 9.99 | 7.72 | 2.3
2005-09-22 | 170:15 | 12.03 | 0.065 | 10 | 7.72 | 2.3
2005-09-22 | 170:30 | 11.99 | 0.065 | 10 | 7.72 | 2.3
2005-09-22 | 170:45 | 11.94 | 0.065 | 10.01 | 7.72 | 2.6
2005-09-22 | 171:00 | 11.9 | 0.065 | 10.03 | 7.72 | 2.9
2005-09-22 | 171:15 | 11.86 | 0.065 | 10.03 | 7.71 | 2.6
2005-09-22 | 171:30 | 11.82 | 0.065 | 10.04 | 7.71 | 2.6
2005-09-22 | 171:45 | 11.79 | 0.065 | 10.05 | 7.71 | 2.6
2005-09-22 | 172:00 | 11.75 | 0.065 | 10.06 | 7.71 | 2.7
2005-09-22 | 172:15 | 11.71 | 0.065 | 10.07 | 7.71 | 2.4
2005-09-22 | 172:30 | 11.67 | 0.065 | 10.08 | 7.71 | 2.5
2005-09-22 | 172:45 | 11.64 | 0.065 | 10.08 | 7.71 | 2.5
2005-09-22 | 173:00 | 11.61 | 0.065 | 10.09 | 7.7 | 2.4
2005-09-22 | 173:15 | 11.59 | 0.065 | 10.1 | 7.71 | 2.5
2005-09-22 | 173:30 | 11.57 | 0.065 | 10.1 | 7.71 | 2.4
2005-09-22 | 173:45 | 11.55 | 0.065 | 10.1 | 7.71 | 2.5
2005-09-22 | 174:00 | 11.53 | 0.065 | 10.11 | 7.7 | 2.5
2005-09-22 | 174:15 | 11.52 | 0.065 | 10.12 | 7.71 | 2.4
2005-09-22 | 174:30 | 11.5 | 0.065 | 10.11 | 7.7 | 2.5
2005-09-22 | 174:45 | 11.49 | 0.065 | 10.12 | 7.7 | 2.5
2005-09-22 | 175:00 | 11.48 | 0.065 | 10.12 | 7.7 | 2.7
2005-09-22 | 175:15 | 11.47 | 0.065 | 10.13 | 7.7 | 2.5
2005-09-22 | 175:30 | 11.47 | 0.065 | 10.15 | 7.71 | 2.5
2005-09-22 | 175:45 | 11.47 | 0.065 | 10.17 | 7.72 | 2.6
2005-09-22 | 176:00 | 11.48 | 0.064 | 10.18 | 7.73 | 2.4
2005-09-22 | 176:15 | 11.49 | 0.064 | 10.18 | 7.75 | 2.3
2005-09-22 | 176:30 | 11.51 | 0.064 | 10.18 | 7.75 | 2.4
2005-09-22 | 176:45 | 11.54 | 0.064 | 10.18 | 7.76 | 2.5
2005-09-22 | 177:00 | 11.57 | 0.064 | 10.17 | 7.77 | 2.4
2005-09-22 | 177:15 | 11.61 | 0.064 | 10.18 | 7.78 | 2.5
2005-09-22 | 177:30 | 11.66 | 0.064 | 10.18 | 7.8 | 2.4
2005-09-22 | 177:45 | 11.72 | 0.064 | 10.18 | 7.82 | 2.8
2005-09-22 | 178:00 | 11.77 | 0.064 | 10.17 | 7.83 | 2.5
2005-09-22 | 178:15 | 11.84 | 0.064 | 10.17 | 7.85 | 2.3
2005-09-22 | 178:30 | 11.9 | 0.064 | 10.16 | 7.87 | 2.4
2005-09-22 | 178:45 | 11.99 | 0.064 | 10.15 | 7.89 | 2.3
2005-09-22 | 179:00 | 12.08 | 0.064 | 10.14 | 7.91 | 2.3