## cont multiparameter audit
Organization | City Of Bend |  | Station ID |  | DR 172.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Southern Urban Growth Boundary
Contact Name | Drexell Branes |  | Latitude
Number | 541-322-6368 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | 04F11310AD
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2006-06-06 14:45 | 16.6 | 7.43 | 9.12 | 0 | 59.4 | 0
Sonde |  | 16.68 | 7.57 | 9.41 | 96.7 | 62.5
Grade |  | A | A | A |  | A
Audit 2 | 2006-06-15 13:30 | 14.59 | 7.34 | 9.03 | 0 | 57.9 | 0
Sonde |  | 14.6 | 7.56 | 10.86 | 106.8 | 61
Grade |  | A | A | Fail |  | A
Audit 3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Sonde
Grade |  | #REF! | #REF! | #REF! |  | #REF!
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | City Of Bend |  | Station ID |  | DR 172.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Southern Urban Growth Boundary
Contact Name | Drexell Branes |  | Latitude |  | 0
Number | 541-322-6368 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-06-06 14:45 | 16.6 | 7.43 | 9.12 |  | 59.4
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | City Of Bend |  | Station ID |  | DR 172.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Southern Urban Growth Boundary
Contact Name | Drexell Branes |  | Latitude |  | 0
Number | 541-322-6368 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-06-15 13:30 | 14.59 | 7.34 | 9.03 |  | 57.9
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | DO Conc | DO% | SpCond | TDS
|  | °C |  | mg/L | % | uS/cm | ug/L
2006-06-06 | 14:45 | 16.68 | 7.57 |  |  | 62.5 | 40.625
2006-06-06 | 15:00 | 16.74 | 7.59 |  |  | 62.5 | 40.651
2006-06-06 | 15:15 | 16.71 | 7.6 |  |  | 62.5 | 40.653
2006-06-06 | 15:30 | 16.74 | 7.6 |  |  | 62.5 | 40.647
2006-06-06 | 15:45 | 16.76 | 7.6 |  |  | 62.5 | 40.637
2006-06-06 | 16:00 | 16.8 | 7.6 |  |  | 62.5 | 40.61
2006-06-06 | 16:15 | 16.84 | 7.6 |  |  | 62.5 | 40.604
2006-06-06 | 16:30 | 16.86 | 7.61 |  |  | 62.5 | 40.608
2006-06-06 | 16:45 | 16.86 | 7.6 |  |  | 62.5 | 40.596
2006-06-06 | 17:00 | 16.86 | 7.59 |  |  | 62.4 | 40.591
2006-06-06 | 17:15 | 16.85 | 7.59 |  |  | 62.4 | 40.58
2006-06-06 | 17:30 | 16.83 | 7.57 |  |  | 62.4 | 40.579
2006-06-06 | 17:45 | 16.7 | 7.55 |  |  | 62.4 | 40.584
2006-06-06 | 18:00 | 16.68 | 7.52 |  |  | 62.4 | 40.567
2006-06-06 | 18:15 | 16.66 | 7.51 |  |  | 62.4 | 40.558
2006-06-06 | 18:30 | 16.62 | 7.47 |  |  | 62.4 | 40.558
2006-06-06 | 18:45 | 16.6 | 7.46 |  |  | 62.4 | 40.538
2006-06-06 | 19:00 | 16.59 | 7.45 |  |  | 62.3 | 40.523
2006-06-06 | 19:15 | 16.58 | 7.43 |  |  | 62.3 | 40.524
2006-06-06 | 19:30 | 16.56 | 7.41 |  |  | 62.3 | 40.518
2006-06-06 | 19:45 | 16.55 | 7.39 |  |  | 62.3 | 40.515
2006-06-06 | 20:00 | 16.54 | 7.37 |  |  | 62.3 | 40.498
2006-06-06 | 20:15 | 16.54 | 7.35 |  |  | 62.3 | 40.504
2006-06-06 | 20:30 | 16.54 | 7.34 |  |  | 62.3 | 40.497
2006-06-06 | 20:45 | 16.54 | 7.33 |  |  | 62.3 | 40.49
2006-06-06 | 21:00 | 16.54 | 7.3 |  |  | 62.3 | 40.482
2006-06-06 | 21:15 | 16.54 | 7.29 |  |  | 62.3 | 40.468
2006-06-06 | 21:30 | 16.54 | 7.28 |  |  | 62.2 | 40.461
2006-06-06 | 21:45 | 16.54 | 7.27 |  |  | 62.2 | 40.455
2006-06-06 | 22:00 | 16.55 | 7.27 |  |  | 62.2 | 40.437
2006-06-06 | 22:15 | 16.55 | 7.26 |  |  | 62.2 | 40.427
2006-06-06 | 22:30 | 16.55 | 7.26 |  |  | 62.2 | 40.415
2006-06-06 | 22:45 | 16.55 | 7.25 |  |  | 62.2 | 40.412
2006-06-06 | 23:00 | 16.54 | 7.25 |  |  | 62.2 | 40.398
2006-06-06 | 23:15 | 16.54 | 7.25 |  |  | 62.1 | 40.394
2006-06-06 | 23:30 | 16.54 | 7.24 |  |  | 62.1 | 40.389
2006-06-06 | 23:45 | 16.54 | 7.23 |  |  | 62.1 | 40.386
2006-06-07 | 0:00 | 16.54 | 7.23 |  |  | 62.1 | 40.384
2006-06-07 | 0:15 | 16.54 | 7.22 |  |  | 62.1 | 40.383
2006-06-07 | 0:30 | 16.53 | 7.21 |  |  | 62.1 | 40.392
2006-06-07 | 0:45 | 16.53 | 7.22 |  |  | 62.1 | 40.388
2006-06-07 | 1:01 | 16.53 | 7.21 |  |  | 62.1 | 40.397
2006-06-07 | 1:16 | 16.52 | 7.21 |  |  | 62.1 | 40.396
2006-06-07 | 1:31 | 16.52 | 7.2 |  |  | 62.2 | 40.406
2006-06-07 | 1:46 | 16.51 | 7.21 |  |  | 62.2 | 40.409
2006-06-07 | 2:01 | 16.51 | 7.2 |  |  | 62.2 | 40.42
2006-06-07 | 2:16 | 16.49 | 7.2 |  |  | 62.2 | 40.414
2006-06-07 | 2:31 | 16.48 | 7.19 |  |  | 62.2 | 40.422
2006-06-07 | 2:46 | 16.46 | 7.2 |  |  | 62.2 | 40.43
2006-06-07 | 3:01 | 16.44 | 7.2 |  |  | 62.2 | 40.446
2006-06-07 | 3:16 | 16.42 | 7.19 |  |  | 62.2 | 40.44
2006-06-07 | 3:31 | 16.39 | 7.19 |  |  | 62.2 | 40.442
2006-06-07 | 3:46 | 16.37 | 7.19 |  |  | 62.2 | 40.443
2006-06-07 | 4:01 | 16.34 | 7.19 |  |  | 62.2 | 40.458
2006-06-07 | 4:16 | 16.32 | 7.2 |  |  | 62.2 | 40.452
2006-06-07 | 4:31 | 16.28 | 7.2 |  |  | 62.2 | 40.451
2006-06-07 | 4:46 | 16.25 | 7.2 |  |  | 62.2 | 40.461
2006-06-07 | 5:01 | 16.21 | 7.19 |  |  | 62.2 | 40.45
2006-06-07 | 5:16 | 16.17 | 7.19 |  |  | 62.2 | 40.446
2006-06-07 | 5:31 | 16.13 | 7.2 |  |  | 62.2 | 40.449
2006-06-07 | 5:46 | 16.08 | 7.2 |  |  | 62.2 | 40.459
2006-06-07 | 6:01 | 16.02 | 7.19 |  |  | 62.2 | 40.449
2006-06-07 | 6:16 | 15.97 | 7.21 |  |  | 62.2 | 40.411
2006-06-07 | 6:31 | 15.92 | 7.22 |  |  | 62.2 | 40.41
2006-06-07 | 6:46 | 15.87 | 7.23 |  |  | 62.2 | 40.409
2006-06-07 | 7:01 | 15.82 | 7.24 |  |  | 62.1 | 40.395
2006-06-07 | 7:16 | 15.76 | 7.25 |  |  | 62.1 | 40.392
2006-06-07 | 7:31 | 15.72 | 7.26 |  |  | 62.1 | 40.375
2006-06-07 | 7:46 | 15.69 | 7.28 |  |  | 62.1 | 40.372
2006-06-07 | 8:01 | 15.66 | 7.29 |  |  | 62.1 | 40.369
2006-06-07 | 8:16 | 15.64 | 7.31 |  |  | 62.1 | 40.352
2006-06-07 | 8:31 | 15.62 | 7.32 |  |  | 62.1 | 40.347
2006-06-07 | 8:46 | 15.62 | 7.33 |  |  | 62.1 | 40.352
2006-06-07 | 9:01 | 15.63 | 7.35 |  |  | 62.1 | 40.348
2006-06-07 | 9:16 | 15.65 | 7.36 |  |  | 62.1 | 40.342
2006-06-07 | 9:31 | 15.67 | 7.38 |  |  | 62.1 | 40.342
2006-06-07 | 9:46 | 15.7 | 7.4 |  |  | 62.1 | 40.339
2006-06-07 | 10:01 | 15.74 | 7.42 |  |  | 62.1 | 40.34
2006-06-07 | 10:16 | 15.78 | 7.44 |  |  | 62.1 | 40.339
2006-06-07 | 10:31 | 15.83 | 7.45 |  |  | 62.1 | 40.346
2006-06-07 | 10:46 | 15.89 | 7.47 |  |  | 62.1 | 40.346
2006-06-07 | 11:01 | 15.95 | 7.48 |  |  | 62.1 | 40.34
2006-06-07 | 11:16 | 16.01 | 7.5 |  |  | 62.1 | 40.357
2006-06-07 | 11:31 | 16.08 | 7.51 |  |  | 62.1 | 40.36
2006-06-07 | 11:46 | 16.17 | 7.52 |  |  | 62.1 | 40.359
2006-06-07 | 12:01 | 16.25 | 7.54 |  |  | 62.1 | 40.356
2006-06-07 | 12:16 | 16.32 | 7.55 |  |  | 62.1 | 40.355
2006-06-07 | 12:31 | 16.39 | 7.56 |  |  | 62.1 | 40.363
2006-06-07 | 12:46 | 16.45 | 7.58 |  |  | 62.1 | 40.364
2006-06-07 | 13:00 | 16.52 | 7.58 |  |  | 62.1 | 40.34
2006-06-07 | 13:15 | 16.59 | 7.59 |  |  | 62.1 | 40.343
2006-06-07 | 13:30 | 16.65 | 7.6 |  |  | 62.1 | 40.34
2006-06-07 | 13:45 | 16.73 | 7.6 |  |  | 62 | 40.332
2006-06-07 | 14:00 | 16.78 | 7.61 |  |  | 62 | 40.312
2006-06-07 | 14:15 | 16.8 | 7.61 |  |  | 62 | 40.318
2006-06-07 | 14:30 | 16.85 | 7.62 |  |  | 62 | 40.298
2006-06-07 | 14:45 | 16.91 | 7.62 |  |  | 62 | 40.284
2006-06-07 | 15:00 | 16.94 | 7.62 |  |  | 62 | 40.276
2006-06-07 | 15:15 | 16.98 | 7.63 |  |  | 61.9 | 40.264
2006-06-07 | 15:30 | 17.02 | 7.62 |  |  | 61.9 | 40.251
2006-06-07 | 15:45 | 17.04 | 7.62 |  |  | 61.9 | 40.25
2006-06-07 | 16:00 | 17.07 | 7.63 |  |  | 61.9 | 40.229
2006-06-07 | 16:15 | 17.01 | 7.6 |  |  | 61.9 | 40.217
2006-06-07 | 16:30 | 17.08 | 7.61 |  |  | 61.9 | 40.205
2006-06-07 | 16:45 | 16.99 | 7.58 |  |  | 61.8 | 40.201
2006-06-07 | 17:00 | 16.91 | 7.57 |  |  | 61.8 | 40.193
2006-06-07 | 17:15 | 16.8 | 7.54 |  |  | 61.8 | 40.178
2006-06-07 | 17:30 | 16.76 | 7.53 |  |  | 61.8 | 40.165
2006-06-07 | 17:45 | 16.79 | 7.54 |  |  | 61.7 | 40.137
2006-06-07 | 18:00 | 16.73 | 7.51 |  |  | 61.8 | 40.146
2006-06-07 | 18:15 | 16.7 | 7.5 |  |  | 61.7 | 40.118
2006-06-07 | 18:30 | 16.66 | 7.49 |  |  | 61.7 | 40.104
2006-06-07 | 18:45 | 16.64 | 7.46 |  |  | 61.7 | 40.11
2006-06-07 | 19:00 | 16.6 | 7.43 |  |  | 61.7 | 40.119
2006-06-07 | 19:15 | 16.57 | 7.41 |  |  | 61.7 | 40.106
2006-06-07 | 19:30 | 16.54 | 7.38 |  |  | 61.7 | 40.107
2006-06-07 | 19:45 | 16.54 | 7.37 |  |  | 61.7 | 40.077
2006-06-07 | 20:00 | 16.52 | 7.36 |  |  | 61.6 | 40.071
2006-06-07 | 20:15 | 16.5 | 7.34 |  |  | 61.6 | 40.064
2006-06-07 | 20:30 | 16.48 | 7.33 |  |  | 61.6 | 40.059
2006-06-07 | 20:45 | 16.46 | 7.31 |  |  | 61.6 | 40.058
2006-06-07 | 21:00 | 16.45 | 7.3 |  |  | 61.6 | 40.045
2006-06-07 | 21:15 | 16.43 | 7.29 |  |  | 61.6 | 40.041
2006-06-07 | 21:30 | 16.41 | 7.28 |  |  | 61.6 | 40.037
2006-06-07 | 21:45 | 16.38 | 7.27 |  |  | 61.6 | 40.022
2006-06-07 | 22:00 | 16.36 | 7.26 |  |  | 61.6 | 40.011
2006-06-07 | 22:15 | 16.34 | 7.26 |  |  | 61.5 | 40
2006-06-07 | 22:30 | 16.32 | 7.26 |  |  | 61.5 | 39.994
2006-06-07 | 22:45 | 16.31 | 7.25 |  |  | 61.5 | 39.969
2006-06-07 | 23:00 | 16.3 | 7.25 |  |  | 61.5 | 39.961
2006-06-07 | 23:15 | 16.3 | 7.24 |  |  | 61.5 | 39.947
2006-06-07 | 23:30 | 16.3 | 7.23 |  |  | 61.4 | 39.942
2006-06-07 | 23:45 | 16.29 | 7.24 |  |  | 61.4 | 39.932
2006-06-08 | 0:00 | 16.29 | 7.23 |  |  | 61.4 | 39.939
2006-06-08 | 0:15 | 16.28 | 7.23 |  |  | 61.4 | 39.941
2006-06-08 | 0:30 | 16.27 | 7.22 |  |  | 61.5 | 39.951
2006-06-08 | 0:45 | 16.26 | 7.22 |  |  | 61.5 | 39.959
2006-06-08 | 1:01 | 16.25 | 7.22 |  |  | 61.5 | 39.96
2006-06-08 | 1:16 | 16.23 | 7.21 |  |  | 61.5 | 39.956
2006-06-08 | 1:31 | 16.22 | 7.21 |  |  | 61.5 | 39.966
2006-06-08 | 1:46 | 16.2 | 7.2 |  |  | 61.5 | 39.996
2006-06-08 | 2:01 | 16.18 | 7.21 |  |  | 61.5 | 39.99
2006-06-08 | 2:16 | 16.16 | 7.21 |  |  | 61.6 | 40.01
2006-06-08 | 2:31 | 16.14 | 7.2 |  |  | 61.6 | 40.015
2006-06-08 | 2:46 | 16.12 | 7.2 |  |  | 61.6 | 40.021
2006-06-08 | 3:01 | 16.09 | 7.2 |  |  | 61.6 | 40.03
2006-06-08 | 3:16 | 16.07 | 7.2 |  |  | 61.6 | 40.017
2006-06-08 | 3:31 | 16.04 | 7.2 |  |  | 61.6 | 40.02
2006-06-08 | 3:46 | 16.02 | 7.2 |  |  | 61.6 | 40.039
2006-06-08 | 4:01 | 15.99 | 7.2 |  |  | 61.6 | 40.033
2006-06-08 | 4:16 | 15.96 | 7.2 |  |  | 61.6 | 40.031
2006-06-08 | 4:31 | 15.92 | 7.2 |  |  | 61.6 | 40.036
2006-06-08 | 4:46 | 15.89 | 7.2 |  |  | 61.6 | 40.053
2006-06-08 | 5:01 | 15.85 | 7.2 |  |  | 61.6 | 40.046
2006-06-08 | 5:16 | 15.81 | 7.2 |  |  | 61.6 | 40.066
2006-06-08 | 5:31 | 15.77 | 7.2 |  |  | 61.6 | 40.058
2006-06-08 | 5:46 | 15.73 | 7.21 |  |  | 61.7 | 40.095
2006-06-08 | 6:01 | 15.69 | 7.21 |  |  | 61.7 | 40.077
2006-06-08 | 6:16 | 15.65 | 7.21 |  |  | 61.6 | 40.045
2006-06-08 | 6:31 | 15.6 | 7.22 |  |  | 61.6 | 40.044
2006-06-08 | 6:46 | 15.56 | 7.23 |  |  | 61.6 | 40.015
2006-06-08 | 7:01 | 15.5 | 7.24 |  |  | 61.6 | 40.01
2006-06-08 | 7:16 | 15.46 | 7.25 |  |  | 61.5 | 40.007
2006-06-08 | 7:31 | 15.41 | 7.26 |  |  | 61.6 | 40.009
2006-06-08 | 7:46 | 15.36 | 7.26 |  |  | 61.5 | 39.994
2006-06-08 | 8:01 | 15.32 | 7.28 |  |  | 61.5 | 39.989
2006-06-08 | 8:16 | 15.29 | 7.3 |  |  | 61.5 | 39.979
2006-06-08 | 8:31 | 15.26 | 7.31 |  |  | 61.5 | 39.977
2006-06-08 | 8:46 | 15.24 | 7.33 |  |  | 61.5 | 39.983
2006-06-08 | 9:01 | 15.23 | 7.34 |  |  | 61.5 | 39.981
2006-06-08 | 9:16 | 15.23 | 7.36 |  |  | 61.5 | 39.965
2006-06-08 | 9:31 | 15.24 | 7.38 |  |  | 61.5 | 39.981
2006-06-08 | 9:46 | 15.26 | 7.4 |  |  | 61.5 | 39.977
2006-06-08 | 10:01 | 15.29 | 7.41 |  |  | 61.5 | 39.969
2006-06-08 | 10:16 | 15.33 | 7.43 |  |  | 61.5 | 39.966
2006-06-08 | 10:31 | 15.36 | 7.45 |  |  | 61.5 | 39.982
2006-06-08 | 10:46 | 15.4 | 7.46 |  |  | 61.5 | 39.985
2006-06-08 | 11:01 | 15.47 | 7.48 |  |  | 61.5 | 39.983
2006-06-08 | 11:16 | 15.54 | 7.5 |  |  | 61.5 | 39.991
2006-06-08 | 11:31 | 15.6 | 7.51 |  |  | 61.5 | 39.992
2006-06-08 | 11:46 | 15.67 | 7.53 |  |  | 61.5 | 40.002
2006-06-08 | 12:01 | 15.74 | 7.54 |  |  | 61.5 | 39.999
2006-06-08 | 12:16 | 15.8 | 7.55 |  |  | 61.5 | 40.003
2006-06-08 | 12:31 | 15.87 | 7.56 |  |  | 61.5 | 40.006
2006-06-08 | 12:46 | 15.94 | 7.57 |  |  | 61.6 | 40.015
2006-06-08 | 13:00 | 15.95 | 7.58 |  |  | 61.5 | 40.002
2006-06-08 | 13:15 | 15.99 | 7.58 |  |  | 61.5 | 40.001
2006-06-08 | 13:30 | 16.1 | 7.58 |  |  | 61.5 | 39.996
2006-06-08 | 13:45 | 16.18 | 7.6 |  |  | 61.5 | 39.986
2006-06-08 | 14:00 | 16.21 | 7.61 |  |  | 61.5 | 39.978
2006-06-08 | 14:15 | 16.23 | 7.61 |  |  | 61.5 | 39.97
2006-06-08 | 14:30 | 16.24 | 7.62 |  |  | 61.5 | 39.961
2006-06-08 | 14:45 | 16.3 | 7.62 |  |  | 61.5 | 39.947
2006-06-08 | 15:00 | 16.37 | 7.62 |  |  | 61.4 | 39.935
2006-06-08 | 15:15 | 16.42 | 7.63 |  |  | 61.4 | 39.941
2006-06-08 | 15:30 | 16.46 | 7.63 |  |  | 61.4 | 39.941
2006-06-08 | 15:45 | 16.48 | 7.63 |  |  | 61.4 | 39.936
2006-06-08 | 16:00 | 16.48 | 7.62 |  |  | 61.5 | 39.964
2006-06-08 | 16:15 | 16.47 | 7.62 |  |  | 61.5 | 39.97
2006-06-08 | 16:30 | 16.44 | 7.6 |  |  | 61.5 | 39.979
2006-06-08 | 16:45 | 16.41 | 7.61 |  |  | 61.5 | 39.981
2006-06-08 | 17:00 | 16.39 | 7.6 |  |  | 61.6 | 40.015
2006-06-08 | 17:15 | 16.37 | 7.59 |  |  | 61.6 | 40.02
2006-06-08 | 17:30 | 16.33 | 7.58 |  |  | 61.6 | 40.034
2006-06-08 | 17:45 | 16.3 | 7.57 |  |  | 61.6 | 40.043
2006-06-08 | 18:00 | 16.23 | 7.56 |  |  | 61.6 | 40.04
2006-06-08 | 18:15 | 16.16 | 7.55 |  |  | 61.6 | 40.04
2006-06-08 | 18:30 | 16.07 | 7.53 |  |  | 61.6 | 40.042
2006-06-08 | 18:45 | 15.97 | 7.5 |  |  | 61.6 | 40.045
2006-06-08 | 19:00 | 15.9 | 7.47 |  |  | 61.6 | 40.042
2006-06-08 | 19:15 | 15.83 | 7.45 |  |  | 61.6 | 40.038
2006-06-08 | 19:30 | 15.75 | 7.41 |  |  | 61.6 | 40.027
2006-06-08 | 19:45 | 15.71 | 7.4 |  |  | 61.6 | 40.019
2006-06-08 | 20:00 | 15.64 | 7.37 |  |  | 61.6 | 40.015
2006-06-08 | 20:15 | 15.6 | 7.35 |  |  | 61.5 | 39.996
2006-06-08 | 20:30 | 15.58 | 7.33 |  |  | 61.5 | 39.978
2006-06-08 | 20:45 | 15.57 | 7.32 |  |  | 61.5 | 39.959
2006-06-08 | 21:00 | 15.56 | 7.31 |  |  | 61.4 | 39.935
2006-06-08 | 21:15 | 15.54 | 7.29 |  |  | 61.4 | 39.912
2006-06-08 | 21:30 | 15.53 | 7.28 |  |  | 61.4 | 39.896
2006-06-08 | 21:45 | 15.52 | 7.28 |  |  | 61.4 | 39.879
2006-06-08 | 22:00 | 15.51 | 7.27 |  |  | 61.3 | 39.846
2006-06-08 | 22:15 | 15.5 | 7.27 |  |  | 61.3 | 39.825
2006-06-08 | 22:30 | 15.49 | 7.26 |  |  | 61.2 | 39.802
2006-06-08 | 22:45 | 15.49 | 7.25 |  |  | 61.2 | 39.786
2006-06-08 | 23:00 | 15.48 | 7.25 |  |  | 61.2 | 39.767
2006-06-08 | 23:15 | 15.47 | 7.25 |  |  | 61.1 | 39.745
2006-06-08 | 23:30 | 15.46 | 7.24 |  |  | 61.1 | 39.739
2006-06-08 | 23:45 | 15.45 | 7.24 |  |  | 61.1 | 39.726
2006-06-09 | 0:00 | 15.43 | 7.24 |  |  | 61.1 | 39.724
2006-06-09 | 0:15 | 15.41 | 7.23 |  |  | 61.1 | 39.715
2006-06-09 | 0:30 | 15.39 | 7.23 |  |  | 61.1 | 39.707
2006-06-09 | 0:45 | 15.37 | 7.22 |  |  | 61.1 | 39.706
2006-06-09 | 1:01 | 15.36 | 7.22 |  |  | 61.1 | 39.713
2006-06-09 | 1:16 | 15.34 | 7.21 |  |  | 61.1 | 39.713
2006-06-09 | 1:31 | 15.32 | 7.21 |  |  | 61.1 | 39.709
2006-06-09 | 1:46 | 15.31 | 7.21 |  |  | 61.1 | 39.706
2006-06-09 | 2:01 | 15.3 | 7.21 |  |  | 61.1 | 39.714
2006-06-09 | 2:16 | 15.29 | 7.2 |  |  | 61.1 | 39.723
2006-06-09 | 2:31 | 15.27 | 7.2 |  |  | 61.1 | 39.734
2006-06-09 | 2:46 | 15.26 | 7.2 |  |  | 61.1 | 39.742
2006-06-09 | 3:01 | 15.25 | 7.2 |  |  | 61.2 | 39.75
2006-06-09 | 3:16 | 15.23 | 7.2 |  |  | 61.2 | 39.759
2006-06-09 | 3:31 | 15.21 | 7.2 |  |  | 61.2 | 39.765
2006-06-09 | 3:46 | 15.19 | 7.2 |  |  | 61.2 | 39.759
2006-06-09 | 4:01 | 15.16 | 7.2 |  |  | 61.2 | 39.77
2006-06-09 | 4:16 | 15.13 | 7.19 |  |  | 61.2 | 39.778
2006-06-09 | 4:31 | 15.09 | 7.19 |  |  | 61.2 | 39.785
2006-06-09 | 4:46 | 15.05 | 7.19 |  |  | 61.2 | 39.777
2006-06-09 | 5:01 | 15 | 7.19 |  |  | 61.2 | 39.785
2006-06-09 | 5:16 | 14.95 | 7.19 |  |  | 61.2 | 39.791
2006-06-09 | 5:31 | 14.9 | 7.19 |  |  | 61.2 | 39.784
2006-06-09 | 5:46 | 14.85 | 7.19 |  |  | 61.2 | 39.777
2006-06-09 | 6:01 | 14.8 | 7.2 |  |  | 61.2 | 39.788
2006-06-09 | 6:16 | 14.74 | 7.2 |  |  | 61.2 | 39.784
2006-06-09 | 6:31 | 14.69 | 7.21 |  |  | 61.2 | 39.768
2006-06-09 | 6:46 | 14.64 | 7.21 |  |  | 61.2 | 39.764
2006-06-09 | 7:01 | 14.59 | 7.22 |  |  | 61.2 | 39.751
2006-06-09 | 7:16 | 14.55 | 7.23 |  |  | 61.2 | 39.748
2006-06-09 | 7:31 | 14.52 | 7.24 |  |  | 61.2 | 39.756
2006-06-09 | 7:46 | 14.49 | 7.26 |  |  | 61.1 | 39.743
2006-06-09 | 8:01 | 14.47 | 7.26 |  |  | 61.1 | 39.74
2006-06-09 | 8:16 | 14.45 | 7.28 |  |  | 61.2 | 39.749
2006-06-09 | 8:31 | 14.44 | 7.3 |  |  | 61.2 | 39.755
2006-06-09 | 8:46 | 14.45 | 7.31 |  |  | 61.2 | 39.755
2006-06-09 | 9:01 | 14.47 | 7.33 |  |  | 61.2 | 39.754
2006-06-09 | 9:16 | 14.49 | 7.34 |  |  | 61.2 | 39.752
2006-06-09 | 9:31 | 14.52 | 7.36 |  |  | 61.2 | 39.76
2006-06-09 | 9:46 | 14.54 | 7.38 |  |  | 61.2 | 39.755
2006-06-09 | 10:01 | 14.59 | 7.39 |  |  | 61.2 | 39.76
2006-06-09 | 10:16 | 14.64 | 7.41 |  |  | 61.2 | 39.771
2006-06-09 | 10:31 | 14.68 | 7.43 |  |  | 61.2 | 39.773
2006-06-09 | 10:46 | 14.75 | 7.45 |  |  | 61.2 | 39.777
2006-06-09 | 11:01 | 14.8 | 7.47 |  |  | 61.2 | 39.776
2006-06-09 | 11:16 | 14.87 | 7.48 |  |  | 61.2 | 39.758
2006-06-09 | 11:31 | 14.95 | 7.49 |  |  | 61.2 | 39.766
2006-06-09 | 11:46 | 15.02 | 7.51 |  |  | 61.2 | 39.761
2006-06-09 | 12:01 | 15.1 | 7.52 |  |  | 61.2 | 39.755
2006-06-09 | 12:16 | 15.17 | 7.53 |  |  | 61.1 | 39.742
2006-06-09 | 12:31 | 15.24 | 7.54 |  |  | 61.2 | 39.75
2006-06-09 | 12:46 | 15.31 | 7.56 |  |  | 61.1 | 39.737
2006-06-09 | 13:00 | 15.39 | 7.56 |  |  | 61.1 | 39.736
2006-06-09 | 13:15 | 15.44 | 7.56 |  |  | 61.1 | 39.729
2006-06-09 | 13:30 | 15.55 | 7.58 |  |  | 61.1 | 39.724
2006-06-09 | 13:45 | 15.65 | 7.59 |  |  | 61.1 | 39.716
2006-06-09 | 14:00 | 15.72 | 7.6 |  |  | 61.1 | 39.706
2006-06-09 | 14:15 | 15.77 | 7.6 |  |  | 61.1 | 39.691
2006-06-09 | 14:30 | 15.81 | 7.61 |  |  | 61 | 39.674
2006-06-09 | 14:45 | 15.87 | 7.61 |  |  | 61 | 39.659
2006-06-09 | 15:00 | 15.89 | 7.61 |  |  | 61 | 39.651
2006-06-09 | 15:15 | 15.83 | 7.59 |  |  | 61 | 39.638
2006-06-09 | 15:30 | 15.73 | 7.56 |  |  | 60.9 | 39.609
2006-06-09 | 15:45 | 15.78 | 7.58 |  |  | 60.9 | 39.589
2006-06-09 | 16:00 | 15.86 | 7.58 |  |  | 60.9 | 39.577
2006-06-09 | 16:15 | 15.87 | 7.59 |  |  | 60.8 | 39.545
2006-06-09 | 16:30 | 15.91 | 7.59 |  |  | 60.8 | 39.532
2006-06-09 | 16:45 | 15.93 | 7.59 |  |  | 60.8 | 39.525
2006-06-09 | 17:00 | 15.92 | 7.58 |  |  | 60.8 | 39.514
2006-06-09 | 17:15 | 15.89 | 7.57 |  |  | 60.8 | 39.493
2006-06-09 | 17:30 | 15.84 | 7.56 |  |  | 60.7 | 39.474
2006-06-09 | 17:45 | 15.77 | 7.52 |  |  | 60.7 | 39.455
2006-06-09 | 18:00 | 15.78 | 7.54 |  |  | 60.7 | 39.439
2006-06-09 | 18:15 | 15.78 | 7.54 |  |  | 60.7 | 39.429
2006-06-09 | 18:30 | 15.72 | 7.52 |  |  | 60.6 | 39.417
2006-06-09 | 18:45 | 15.64 | 7.5 |  |  | 60.6 | 39.403
2006-06-09 | 19:00 | 15.56 | 7.47 |  |  | 60.6 | 39.403
2006-06-09 | 19:15 | 15.5 | 7.45 |  |  | 60.6 | 39.397
2006-06-09 | 19:30 | 15.44 | 7.41 |  |  | 60.6 | 39.385
2006-06-09 | 19:45 | 15.39 | 7.39 |  |  | 60.6 | 39.376
2006-06-09 | 20:00 | 15.34 | 7.36 |  |  | 60.6 | 39.365
2006-06-09 | 20:15 | 15.3 | 7.34 |  |  | 60.5 | 39.344
2006-06-09 | 20:30 | 15.26 | 7.33 |  |  | 60.5 | 39.331
2006-06-09 | 20:45 | 15.23 | 7.31 |  |  | 60.5 | 39.325
2006-06-09 | 21:00 | 15.2 | 7.3 |  |  | 60.5 | 39.318
2006-06-09 | 21:15 | 15.16 | 7.28 |  |  | 60.5 | 39.299
2006-06-09 | 21:30 | 15.13 | 7.27 |  |  | 60.4 | 39.283
2006-06-09 | 21:45 | 15.12 | 7.26 |  |  | 60.4 | 39.274
2006-06-09 | 22:00 | 15.1 | 7.26 |  |  | 60.4 | 39.25
2006-06-09 | 22:15 | 15.09 | 7.25 |  |  | 60.4 | 39.249
2006-06-09 | 22:30 | 15.07 | 7.25 |  |  | 60.4 | 39.239
2006-06-09 | 22:45 | 15.06 | 7.24 |  |  | 60.4 | 39.238
2006-06-09 | 23:00 | 15.04 | 7.23 |  |  | 60.4 | 39.237
2006-06-09 | 23:15 | 15.02 | 7.22 |  |  | 60.4 | 39.228
2006-06-09 | 23:30 | 15 | 7.22 |  |  | 60.4 | 39.23
2006-06-09 | 23:45 | 14.98 | 7.22 |  |  | 60.4 | 39.233
2006-06-10 | 0:00 | 14.95 | 7.21 |  |  | 60.4 | 39.24
2006-06-10 | 0:15 | 14.92 | 7.21 |  |  | 60.4 | 39.236
2006-06-10 | 0:30 | 14.9 | 7.21 |  |  | 60.4 | 39.243
2006-06-10 | 0:45 | 14.88 | 7.21 |  |  | 60.4 | 39.249
2006-06-10 | 1:01 | 14.85 | 7.21 |  |  | 60.4 | 39.254
2006-06-10 | 1:16 | 14.84 | 7.2 |  |  | 60.4 | 39.261
2006-06-10 | 1:31 | 14.83 | 7.2 |  |  | 60.4 | 39.27
2006-06-10 | 1:46 | 14.82 | 7.2 |  |  | 60.5 | 39.294
2006-06-10 | 2:01 | 14.82 | 7.2 |  |  | 60.5 | 39.299
2006-06-10 | 2:16 | 14.81 | 7.19 |  |  | 60.5 | 39.316
2006-06-10 | 2:31 | 14.81 | 7.19 |  |  | 60.5 | 39.319
2006-06-10 | 2:46 | 14.8 | 7.18 |  |  | 60.5 | 39.331
2006-06-10 | 3:01 | 14.79 | 7.18 |  |  | 60.5 | 39.335
2006-06-10 | 3:16 | 14.77 | 7.18 |  |  | 60.5 | 39.353
2006-06-10 | 3:31 | 14.74 | 7.19 |  |  | 60.5 | 39.355
2006-06-10 | 3:46 | 14.71 | 7.18 |  |  | 60.6 | 39.372
2006-06-10 | 4:01 | 14.67 | 7.18 |  |  | 60.6 | 39.365
2006-06-10 | 4:16 | 14.63 | 7.18 |  |  | 60.6 | 39.388
2006-06-10 | 4:31 | 14.58 | 7.18 |  |  | 60.6 | 39.39
2006-06-10 | 4:46 | 14.53 | 7.18 |  |  | 60.6 | 39.383
2006-06-10 | 5:01 | 14.48 | 7.18 |  |  | 60.6 | 39.398
2006-06-10 | 5:16 | 14.42 | 7.17 |  |  | 60.6 | 39.398
2006-06-10 | 5:31 | 14.37 | 7.17 |  |  | 60.6 | 39.4
2006-06-10 | 5:46 | 14.32 | 7.17 |  |  | 60.6 | 39.4
2006-06-10 | 6:01 | 14.28 | 7.18 |  |  | 60.6 | 39.41
2006-06-10 | 6:16 | 14.23 | 7.17 |  |  | 60.6 | 39.404
2006-06-10 | 6:31 | 14.18 | 7.19 |  |  | 60.6 | 39.384
2006-06-10 | 6:46 | 14.14 | 7.19 |  |  | 60.6 | 39.38
2006-06-10 | 7:01 | 14.1 | 7.21 |  |  | 60.6 | 39.385
2006-06-10 | 7:16 | 14.06 | 7.21 |  |  | 60.6 | 39.36
2006-06-10 | 7:31 | 14.02 | 7.23 |  |  | 60.5 | 39.35
2006-06-10 | 7:46 | 14 | 7.24 |  |  | 60.5 | 39.34
2006-06-10 | 8:01 | 13.97 | 7.27 |  |  | 60.5 | 39.345
2006-06-10 | 8:16 | 13.94 | 7.3 |  |  | 60.5 | 39.348
2006-06-10 | 8:31 | 13.92 | 7.31 |  |  | 60.5 | 39.349
2006-06-10 | 8:46 | 13.92 | 7.33 |  |  | 60.5 | 39.344
2006-06-10 | 9:01 | 13.89 | 7.34 |  |  | 60.6 | 39.376
2006-06-10 | 9:16 | 13.86 | 7.35 |  |  | 60.6 | 39.378
2006-06-10 | 9:31 | 13.86 | 7.36 |  |  | 60.6 | 39.41
2006-06-10 | 9:46 | 13.87 | 7.37 |  |  | 60.7 | 39.468
2006-06-10 | 10:01 | 13.93 | 7.39 |  |  | 60.8 | 39.498
2006-06-10 | 10:16 | 13.97 | 7.39 |  |  | 60.9 | 39.575
2006-06-10 | 10:31 | 14.06 | 7.41 |  |  | 61 | 39.63
2006-06-10 | 10:46 | 14.15 | 7.43 |  |  | 61.1 | 39.685
2006-06-10 | 11:01 | 14.28 | 7.46 |  |  | 61.1 | 39.723
2006-06-10 | 11:16 | 14.38 | 7.48 |  |  | 61.2 | 39.751
2006-06-10 | 11:31 | 14.37 | 7.49 |  |  | 61.2 | 39.773
2006-06-10 | 11:46 | 14.38 | 7.51 |  |  | 61.2 | 39.797
2006-06-10 | 12:01 | 14.41 | 7.5 |  |  | 61.2 | 39.805
2006-06-10 | 12:16 | 14.44 | 7.5 |  |  | 61.2 | 39.796
2006-06-10 | 12:31 | 14.56 | 7.51 |  |  | 61.2 | 39.771
2006-06-10 | 12:46 | 14.62 | 7.52 |  |  | 61.2 | 39.761
2006-06-10 | 13:00 | 14.73 | 7.54 |  |  | 61.1 | 39.734
2006-06-10 | 13:15 | 14.86 | 7.53 |  |  | 61.1 | 39.695
2006-06-10 | 13:30 | 14.84 | 7.53 |  |  | 61 | 39.666
2006-06-10 | 13:45 | 14.79 | 7.53 |  |  | 61 | 39.628
2006-06-10 | 14:00 | 14.86 | 7.54 |  |  | 60.9 | 39.589
2006-06-10 | 14:15 | 14.94 | 7.54 |  |  | 60.9 | 39.558
2006-06-10 | 14:30 | 14.92 | 7.53 |  |  | 60.8 | 39.531
2006-06-10 | 14:45 | 14.85 | 7.55 |  |  | 60.8 | 39.489
2006-06-10 | 15:00 | 14.81 | 7.55 |  |  | 60.7 | 39.468
2006-06-10 | 15:15 | 14.86 | 7.56 |  |  | 60.7 | 39.442
2006-06-10 | 15:30 | 14.92 | 7.54 |  |  | 60.6 | 39.412
2006-06-10 | 15:45 | 15.04 | 7.55 |  |  | 60.6 | 39.409
2006-06-10 | 16:00 | 15.13 | 7.55 |  |  | 60.6 | 39.365
2006-06-10 | 16:15 | 15.14 | 7.55 |  |  | 60.6 | 39.367
2006-06-10 | 16:30 | 15.2 | 7.55 |  |  | 60.6 | 39.36
2006-06-10 | 16:45 | 15.24 | 7.55 |  |  | 60.5 | 39.342
2006-06-10 | 17:00 | 15.19 | 7.53 |  |  | 60.5 | 39.331
2006-06-10 | 17:15 | 15.2 | 7.52 |  |  | 60.5 | 39.305
2006-06-10 | 17:30 | 15.22 | 7.52 |  |  | 60.5 | 39.295
2006-06-10 | 17:45 | 15.25 | 7.52 |  |  | 60.4 | 39.271
2006-06-10 | 18:00 | 15.29 | 7.52 |  |  | 60.4 | 39.25
2006-06-10 | 18:15 | 15.24 | 7.5 |  |  | 60.4 | 39.266
2006-06-10 | 18:30 | 15.24 | 7.48 |  |  | 60.4 | 39.241
2006-06-10 | 18:45 | 15.21 | 7.47 |  |  | 60.3 | 39.218
2006-06-10 | 19:00 | 15.19 | 7.45 |  |  | 60.4 | 39.231
2006-06-10 | 19:15 | 15.14 | 7.43 |  |  | 60.3 | 39.225
2006-06-10 | 19:30 | 15.09 | 7.4 |  |  | 60.3 | 39.212
2006-06-10 | 19:45 | 15.05 | 7.39 |  |  | 60.3 | 39.215
2006-06-10 | 20:00 | 15.02 | 7.35 |  |  | 60.3 | 39.207
2006-06-10 | 20:15 | 15.01 | 7.33 |  |  | 60.3 | 39.2
2006-06-10 | 20:30 | 15 | 7.31 |  |  | 60.3 | 39.182
2006-06-10 | 20:45 | 15 | 7.28 |  |  | 60.3 | 39.168
2006-06-10 | 21:00 | 14.99 | 7.28 |  |  | 60.2 | 39.162
2006-06-10 | 21:15 | 14.99 | 7.27 |  |  | 60.2 | 39.143
2006-06-10 | 21:30 | 14.99 | 7.26 |  |  | 60.2 | 39.145
2006-06-10 | 21:45 | 14.99 | 7.26 |  |  | 60.2 | 39.137
2006-06-10 | 22:00 | 14.98 | 7.24 |  |  | 60.2 | 39.125
2006-06-10 | 22:15 | 14.98 | 7.23 |  |  | 60.2 | 39.125
2006-06-10 | 22:30 | 14.96 | 7.22 |  |  | 60.2 | 39.103
2006-06-10 | 22:45 | 14.95 | 7.22 |  |  | 60.2 | 39.125
2006-06-10 | 23:00 | 14.93 | 7.22 |  |  | 60.2 | 39.106
2006-06-10 | 23:15 | 14.92 | 7.21 |  |  | 60.2 | 39.117
2006-06-10 | 23:30 | 14.9 | 7.2 |  |  | 60.2 | 39.112
2006-06-10 | 23:45 | 14.89 | 7.2 |  |  | 60.2 | 39.115
2006-06-11 | 0:00 | 14.87 | 7.2 |  |  | 60.2 | 39.113
2006-06-11 | 0:15 | 14.87 | 7.2 |  |  | 60.2 | 39.127
2006-06-11 | 0:30 | 14.86 | 7.19 |  |  | 60.2 | 39.126
2006-06-11 | 0:45 | 14.85 | 7.19 |  |  | 60.2 | 39.129
2006-06-11 | 1:01 | 14.82 | 7.19 |  |  | 60.2 | 39.132
2006-06-11 | 1:16 | 14.8 | 7.19 |  |  | 60.2 | 39.147
2006-06-11 | 1:31 | 14.78 | 7.19 |  |  | 60.2 | 39.137
2006-06-11 | 1:46 | 14.76 | 7.18 |  |  | 60.2 | 39.152
2006-06-11 | 2:01 | 14.73 | 7.18 |  |  | 60.2 | 39.156
2006-06-11 | 2:16 | 14.7 | 7.17 |  |  | 60.3 | 39.169
2006-06-11 | 2:31 | 14.67 | 7.18 |  |  | 60.3 | 39.179
2006-06-11 | 2:46 | 14.65 | 7.17 |  |  | 60.3 | 39.176
2006-06-11 | 3:01 | 14.62 | 7.17 |  |  | 60.3 | 39.189
2006-06-11 | 3:16 | 14.59 | 7.17 |  |  | 60.3 | 39.204
2006-06-11 | 3:31 | 14.56 | 7.17 |  |  | 60.3 | 39.211
2006-06-11 | 3:46 | 14.53 | 7.17 |  |  | 60.3 | 39.211
2006-06-11 | 4:01 | 14.5 | 7.17 |  |  | 60.3 | 39.226
2006-06-11 | 4:16 | 14.47 | 7.17 |  |  | 60.4 | 39.228
2006-06-11 | 4:31 | 14.42 | 7.17 |  |  | 60.4 | 39.239
2006-06-11 | 4:46 | 14.38 | 7.17 |  |  | 60.4 | 39.246
2006-06-11 | 5:01 | 14.32 | 7.17 |  |  | 60.4 | 39.242
2006-06-11 | 5:16 | 14.27 | 7.17 |  |  | 60.4 | 39.246
2006-06-11 | 5:31 | 14.22 | 7.17 |  |  | 60.4 | 39.253
2006-06-11 | 5:46 | 14.17 | 7.17 |  |  | 60.4 | 39.25
2006-06-11 | 6:01 | 14.13 | 7.17 |  |  | 60.4 | 39.252
2006-06-11 | 6:16 | 14.08 | 7.17 |  |  | 60.4 | 39.247
2006-06-11 | 6:31 | 14.03 | 7.19 |  |  | 60.4 | 39.234
2006-06-11 | 6:46 | 14 | 7.19 |  |  | 60.4 | 39.236
2006-06-11 | 7:01 | 13.96 | 7.2 |  |  | 60.3 | 39.221
2006-06-11 | 7:16 | 13.93 | 7.22 |  |  | 60.3 | 39.205
2006-06-11 | 7:31 | 13.9 | 7.23 |  |  | 60.3 | 39.209
2006-06-11 | 7:46 | 13.88 | 7.25 |  |  | 60.3 | 39.207
2006-06-11 | 8:01 | 13.85 | 7.26 |  |  | 60.3 | 39.206
2006-06-11 | 8:16 | 13.85 | 7.28 |  |  | 60.3 | 39.203
2006-06-11 | 8:31 | 13.87 | 7.31 |  |  | 60.3 | 39.199
2006-06-11 | 8:46 | 13.86 | 7.33 |  |  | 60.3 | 39.199
2006-06-11 | 9:01 | 13.85 | 7.34 |  |  | 60.3 | 39.201
2006-06-11 | 9:16 | 13.86 | 7.36 |  |  | 60.3 | 39.209
2006-06-11 | 9:31 | 13.87 | 7.37 |  |  | 60.3 | 39.202
2006-06-11 | 9:46 | 13.9 | 7.39 |  |  | 60.3 | 39.205
2006-06-11 | 10:01 | 13.93 | 7.39 |  |  | 60.3 | 39.21
2006-06-11 | 10:16 | 13.99 | 7.42 |  |  | 60.3 | 39.196
2006-06-11 | 10:31 | 14.06 | 7.44 |  |  | 60.3 | 39.196
2006-06-11 | 10:46 | 14.13 | 7.46 |  |  | 60.3 | 39.201
2006-06-11 | 11:01 | 14.17 | 7.47 |  |  | 60.3 | 39.204
2006-06-11 | 11:16 | 14.2 | 7.47 |  |  | 60.3 | 39.212
2006-06-11 | 11:31 | 14.26 | 7.49 |  |  | 60.3 | 39.2
2006-06-11 | 11:46 | 14.36 | 7.5 |  |  | 60.3 | 39.211
2006-06-11 | 12:01 | 14.41 | 7.51 |  |  | 60.3 | 39.218
2006-06-11 | 12:16 | 14.39 | 7.51 |  |  | 60.3 | 39.21
2006-06-11 | 12:31 | 14.46 | 7.51 |  |  | 60.3 | 39.211
2006-06-11 | 12:46 | 14.56 | 7.53 |  |  | 60.3 | 39.21
2006-06-11 | 13:00 | 14.64 | 7.54 |  |  | 60.3 | 39.212
2006-06-11 | 13:15 | 14.66 | 7.55 |  |  | 60.3 | 39.21
2006-06-11 | 13:30 | 14.71 | 7.55 |  |  | 60.3 | 39.203
2006-06-11 | 13:45 | 14.86 | 7.57 |  |  | 60.3 | 39.204
2006-06-11 | 14:00 | 14.97 | 7.57 |  |  | 60.3 | 39.197
2006-06-11 | 14:15 | 15.04 | 7.57 |  |  | 60.3 | 39.2
2006-06-11 | 14:30 | 15.11 | 7.58 |  |  | 60.3 | 39.2
2006-06-11 | 14:45 | 15.18 | 7.59 |  |  | 60.3 | 39.184
2006-06-11 | 15:00 | 15.25 | 7.58 |  |  | 60.3 | 39.185
2006-06-11 | 15:15 | 15.32 | 7.58 |  |  | 60.3 | 39.184
2006-06-11 | 15:30 | 15.32 | 7.59 |  |  | 60.3 | 39.165
2006-06-11 | 15:45 | 15.37 | 7.58 |  |  | 60.2 | 39.155
2006-06-11 | 16:00 | 15.42 | 7.58 |  |  | 60.2 | 39.155
2006-06-11 | 16:15 | 15.47 | 7.57 |  |  | 60.2 | 39.14
2006-06-11 | 16:30 | 15.46 | 7.58 |  |  | 60.2 | 39.131
2006-06-11 | 16:45 | 15.38 | 7.57 |  |  | 60.2 | 39.112
2006-06-11 | 17:00 | 15.24 | 7.52 |  |  | 60.2 | 39.126
2006-06-11 | 17:15 | 15.13 | 7.48 |  |  | 60.2 | 39.141
2006-06-11 | 17:30 | 15.07 | 7.44 |  |  | 60.2 | 39.107
2006-06-11 | 17:45 | 15.02 | 7.4 |  |  | 60.1 | 39.075
2006-06-11 | 18:00 | 14.97 | 7.37 |  |  | 60.1 | 39.049
2006-06-11 | 18:15 | 14.95 | 7.36 |  |  | 60 | 39.001
2006-06-11 | 18:30 | 14.94 | 7.35 |  |  | 60 | 38.992
2006-06-11 | 18:45 | 14.94 | 7.34 |  |  | 60 | 38.97
2006-06-11 | 19:00 | 14.95 | 7.34 |  |  | 59.9 | 38.963
2006-06-11 | 19:15 | 14.95 | 7.34 |  |  | 59.9 | 38.965
2006-06-11 | 19:30 | 14.96 | 7.33 |  |  | 59.9 | 38.937
2006-06-11 | 19:45 | 14.98 | 7.34 |  |  | 59.9 | 38.936
2006-06-11 | 20:00 | 14.98 | 7.32 |  |  | 59.9 | 38.962
2006-06-11 | 20:15 | 14.97 | 7.3 |  |  | 59.8 | 38.9
2006-06-11 | 20:30 | 14.97 | 7.28 |  |  | 59.8 | 38.869
2006-06-11 | 20:45 | 14.96 | 7.27 |  |  | 59.8 | 38.855
2006-06-11 | 21:00 | 14.95 | 7.26 |  |  | 59.8 | 38.855
2006-06-11 | 21:15 | 14.94 | 7.25 |  |  | 59.8 | 38.851
2006-06-11 | 21:30 | 14.93 | 7.23 |  |  | 59.8 | 38.843
2006-06-11 | 21:45 | 14.93 | 7.24 |  |  | 59.8 | 38.852
2006-06-11 | 22:00 | 14.94 | 7.23 |  |  | 59.8 | 38.851
2006-06-11 | 22:15 | 14.95 | 7.22 |  |  | 59.8 | 38.848
2006-06-11 | 22:30 | 14.96 | 7.22 |  |  | 59.8 | 38.868
2006-06-11 | 22:45 | 14.97 | 7.21 |  |  | 59.8 | 38.877
2006-06-11 | 23:00 | 14.98 | 7.21 |  |  | 59.8 | 38.887
2006-06-11 | 23:15 | 14.98 | 7.2 |  |  | 59.8 | 38.89
2006-06-11 | 23:30 | 14.99 | 7.19 |  |  | 59.8 | 38.881
2006-06-11 | 23:45 | 15 | 7.19 |  |  | 59.8 | 38.886
2006-06-12 | 0:00 | 15.01 | 7.19 |  |  | 59.8 | 38.885
2006-06-12 | 0:15 | 15.04 | 7.19 |  |  | 59.8 | 38.887
2006-06-12 | 0:30 | 15.05 | 7.19 |  |  | 59.8 | 38.881
2006-06-12 | 0:45 | 15.07 | 7.19 |  |  | 59.9 | 38.907
2006-06-12 | 1:01 | 15.08 | 7.18 |  |  | 59.8 | 38.895
2006-06-12 | 1:16 | 15.09 | 7.17 |  |  | 59.9 | 38.914
2006-06-12 | 1:31 | 15.09 | 7.18 |  |  | 59.9 | 38.926
2006-06-12 | 1:46 | 15.09 | 7.17 |  |  | 59.9 | 38.954
2006-06-12 | 2:01 | 15.1 | 7.18 |  |  | 60 | 38.969
2006-06-12 | 2:16 | 15.1 | 7.17 |  |  | 60 | 38.971
2006-06-12 | 2:31 | 15.09 | 7.17 |  |  | 60 | 38.997
2006-06-12 | 2:46 | 15.09 | 7.17 |  |  | 60 | 39.011
2006-06-12 | 3:01 | 15.08 | 7.16 |  |  | 60 | 39.032
2006-06-12 | 3:16 | 15.05 | 7.17 |  |  | 60.1 | 39.047
2006-06-12 | 3:31 | 15.03 | 7.16 |  |  | 60.1 | 39.053
2006-06-12 | 3:46 | 15 | 7.16 |  |  | 60.1 | 39.058
2006-06-12 | 4:01 | 14.96 | 7.16 |  |  | 60.1 | 39.058
2006-06-12 | 4:16 | 14.92 | 7.16 |  |  | 60.1 | 39.073
2006-06-12 | 4:31 | 14.88 | 7.16 |  |  | 60.1 | 39.077
2006-06-12 | 4:46 | 14.83 | 7.16 |  |  | 60.1 | 39.09
2006-06-12 | 5:01 | 14.77 | 7.15 |  |  | 60.2 | 39.104
2006-06-12 | 5:16 | 14.72 | 7.15 |  |  | 60.2 | 39.103
2006-06-12 | 5:31 | 14.68 | 7.13 |  |  | 60.1 | 39.096
2006-06-12 | 5:46 | 14.63 | 7.15 |  |  | 60.2 | 39.108
2006-06-12 | 6:01 | 14.59 | 7.15 |  |  | 60.2 | 39.106
2006-06-12 | 6:16 | 14.55 | 7.15 |  |  | 60.1 | 39.097
2006-06-12 | 6:31 | 14.52 | 7.17 |  |  | 60.1 | 39.077
2006-06-12 | 6:46 | 14.49 | 7.18 |  |  | 60.1 | 39.063
2006-06-12 | 7:01 | 14.47 | 7.2 |  |  | 60.1 | 39.049
2006-06-12 | 7:16 | 14.46 | 7.21 |  |  | 60.1 | 39.041
2006-06-12 | 7:31 | 14.44 | 7.22 |  |  | 60.1 | 39.038
2006-06-12 | 7:46 | 14.41 | 7.24 |  |  | 60.1 | 39.043
2006-06-12 | 8:01 | 14.4 | 7.25 |  |  | 60.1 | 39.04
2006-06-12 | 8:16 | 14.39 | 7.27 |  |  | 60.1 | 39.044
2006-06-12 | 8:31 | 14.39 | 7.28 |  |  | 60 | 39.027
2006-06-12 | 8:46 | 14.38 | 7.29 |  |  | 60 | 39.03
2006-06-12 | 9:01 | 14.39 | 7.31 |  |  | 60 | 39.024
2006-06-12 | 9:16 | 14.39 | 7.33 |  |  | 60.1 | 39.035
2006-06-12 | 9:31 | 14.41 | 7.35 |  |  | 60 | 39.031
2006-06-12 | 9:46 | 14.43 | 7.36 |  |  | 60.1 | 39.041
2006-06-12 | 10:01 | 14.45 | 7.38 |  |  | 60.1 | 39.042
2006-06-12 | 10:16 | 14.53 | 7.4 |  |  | 60.1 | 39.042
2006-06-12 | 10:31 | 14.61 | 7.42 |  |  | 60.1 | 39.045
2006-06-12 | 10:46 | 14.74 | 7.44 |  |  | 60.1 | 39.054
2006-06-12 | 11:01 | 14.74 | 7.44 |  |  | 60.1 | 39.057
2006-06-12 | 11:16 | 14.78 | 7.46 |  |  | 60.1 | 39.06
2006-06-12 | 11:31 | 14.81 | 7.48 |  |  | 60.1 | 39.071
2006-06-12 | 11:46 | 14.92 | 7.49 |  |  | 60.1 | 39.078
2006-06-12 | 12:01 | 15.02 | 7.5 |  |  | 60.1 | 39.092
2006-06-12 | 12:16 | 15.1 | 7.51 |  |  | 60.1 | 39.088
2006-06-12 | 12:31 | 15.1 | 7.51 |  |  | 60.1 | 39.076
2006-06-12 | 12:46 | 15.21 | 7.53 |  |  | 60.1 | 39.086
2006-06-12 | 13:00 | 15.32 | 7.55 |  |  | 60.1 | 39.084
2006-06-12 | 13:15 | 15.41 | 7.55 |  |  | 60.1 | 39.091
2006-06-12 | 13:30 | 15.47 | 7.55 |  |  | 60.1 | 39.089
2006-06-12 | 13:45 | 15.39 | 7.53 |  |  | 60.2 | 39.099
2006-06-12 | 14:00 | 15.28 | 7.49 |  |  | 60.1 | 39.096
2006-06-12 | 14:15 | 15.18 | 7.45 |  |  | 60.2 | 39.124
2006-06-12 | 14:30 | 14.6 | 7.38 |  |  | 59.3 | 38.575
2006-06-12 | 14:45 | 14.49 | 7.35 |  |  | 59.1 | 38.415
2006-06-12 | 15:00 | 14.62 | 7.33 |  |  | 59.2 | 38.491
2006-06-12 | 15:15 | 14.79 | 7.39 |  |  | 59.4 | 38.604
2006-06-12 | 15:30 | 14.79 | 7.41 |  |  | 59.3 | 38.524
2006-06-12 | 15:45 | 14.87 | 7.44 |  |  | 59.2 | 38.496
2006-06-12 | 16:00 | 15.01 | 7.47 |  |  | 59.3 | 38.54
2006-06-12 | 16:15 | 15.02 | 7.46 |  |  | 59.3 | 38.522
2006-06-12 | 16:30 | 14.99 | 7.45 |  |  | 59.3 | 38.527
2006-06-12 | 16:45 | 15.11 | 7.47 |  |  | 59.4 | 38.579
2006-06-12 | 17:00 | 15.16 | 7.46 |  |  | 59.4 | 38.634
2006-06-12 | 17:15 | 15.14 | 7.44 |  |  | 59.5 | 38.651
2006-06-12 | 17:30 | 15.05 | 7.4 |  |  | 59.5 | 38.68
2006-06-12 | 17:45 | 14.97 | 7.38 |  |  | 59.3 | 38.544
2006-06-12 | 18:00 | 14.96 | 7.39 |  |  | 59.3 | 38.522
2006-06-12 | 18:15 | 14.98 | 7.39 |  |  | 59.3 | 38.549
2006-06-12 | 18:30 | 14.99 | 7.41 |  |  | 59.3 | 38.575
2006-06-12 | 18:45 | 14.97 | 7.38 |  |  | 59.4 | 38.606
2006-06-12 | 19:00 | 14.97 | 7.38 |  |  | 59.5 | 38.643
2006-06-12 | 19:15 | 14.94 | 7.36 |  |  | 59.5 | 38.646
2006-06-12 | 19:30 | 14.88 | 7.33 |  |  | 59.5 | 38.653
2006-06-12 | 19:45 | 14.8 | 7.33 |  |  | 59.4 | 38.633
2006-06-12 | 20:00 | 14.72 | 7.3 |  |  | 59.4 | 38.623
2006-06-12 | 20:15 | 14.66 | 7.28 |  |  | 59.4 | 38.609
2006-06-12 | 20:30 | 14.63 | 7.26 |  |  | 59.4 | 38.613
2006-06-12 | 20:45 | 14.66 | 7.23 |  |  | 59.5 | 38.66
2006-06-12 | 21:00 | 14.72 | 7.25 |  |  | 59.5 | 38.694
2006-06-12 | 21:15 | 14.77 | 7.23 |  |  | 59.6 | 38.733
2006-06-12 | 21:30 | 14.8 | 7.23 |  |  | 59.6 | 38.748
2006-06-12 | 21:45 | 14.82 | 7.21 |  |  | 59.6 | 38.76
2006-06-12 | 22:00 | 14.84 | 7.21 |  |  | 59.7 | 38.774
2006-06-12 | 22:15 | 14.84 | 7.2 |  |  | 59.6 | 38.763
2006-06-12 | 22:30 | 14.83 | 7.2 |  |  | 59.6 | 38.756
2006-06-12 | 22:45 | 14.82 | 7.19 |  |  | 59.8 | 38.849
2006-06-12 | 23:00 | 14.8 | 7.19 |  |  | 59.6 | 38.753
2006-06-12 | 23:15 | 14.8 | 7.18 |  |  | 59.6 | 38.763
2006-06-12 | 23:30 | 14.81 | 7.18 |  |  | 59.6 | 38.761
2006-06-12 | 23:45 | 14.82 | 7.18 |  |  | 59.6 | 38.772
2006-06-13 | 0:00 | 14.84 | 7.18 |  |  | 59.7 | 38.78
2006-06-13 | 0:15 | 14.85 | 7.17 |  |  | 59.6 | 38.772
2006-06-13 | 0:30 | 14.86 | 7.17 |  |  | 59.7 | 38.8
2006-06-13 | 0:45 | 14.85 | 7.17 |  |  | 59.7 | 38.806
2006-06-13 | 1:01 | 14.84 | 7.17 |  |  | 59.8 | 38.841
2006-06-13 | 1:16 | 14.84 | 7.17 |  |  | 59.8 | 38.848
2006-06-13 | 1:31 | 14.84 | 7.16 |  |  | 59.8 | 38.869
2006-06-13 | 1:46 | 14.84 | 7.16 |  |  | 59.8 | 38.883
2006-06-13 | 2:01 | 14.83 | 7.16 |  |  | 59.9 | 38.913
2006-06-13 | 2:16 | 14.81 | 7.16 |  |  | 59.9 | 38.922
2006-06-13 | 2:31 | 14.78 | 7.16 |  |  | 59.9 | 38.936
2006-06-13 | 2:46 | 14.74 | 7.15 |  |  | 59.9 | 38.933
2006-06-13 | 3:01 | 14.7 | 7.15 |  |  | 59.9 | 38.955
2006-06-13 | 3:16 | 14.67 | 7.15 |  |  | 59.9 | 38.96
2006-06-13 | 3:31 | 14.63 | 7.15 |  |  | 59.9 | 38.965
2006-06-13 | 3:46 | 14.6 | 7.14 |  |  | 60 | 38.988
2006-06-13 | 4:01 | 14.58 | 7.14 |  |  | 60 | 39.004
2006-06-13 | 4:16 | 14.56 | 7.14 |  |  | 60 | 39.01
2006-06-13 | 4:31 | 14.54 | 7.14 |  |  | 60 | 39.023
2006-06-13 | 4:46 | 14.52 | 7.14 |  |  | 60.1 | 39.049
2006-06-13 | 5:01 | 14.49 | 7.15 |  |  | 60.1 | 39.065
2006-06-13 | 5:16 | 14.46 | 7.14 |  |  | 60.1 | 39.083
2006-06-13 | 5:31 | 14.43 | 7.15 |  |  | 60.1 | 39.093
2006-06-13 | 5:46 | 14.41 | 7.15 |  |  | 60.1 | 39.095
2006-06-13 | 6:01 | 14.39 | 7.15 |  |  | 60.2 | 39.102
2006-06-13 | 6:16 | 14.36 | 7.15 |  |  | 60.2 | 39.13
2006-06-13 | 6:31 | 14.33 | 7.16 |  |  | 60.2 | 39.114
2006-06-13 | 6:46 | 14.29 | 7.16 |  |  | 60.2 | 39.115
2006-06-13 | 7:01 | 14.27 | 7.17 |  |  | 60.2 | 39.104
2006-06-13 | 7:16 | 14.22 | 7.19 |  |  | 60.2 | 39.105
2006-06-13 | 7:31 | 14.2 | 7.2 |  |  | 60.2 | 39.115
2006-06-13 | 7:46 | 14.2 | 7.23 |  |  | 60.1 | 39.092
2006-06-13 | 8:01 | 14.21 | 7.26 |  |  | 60.2 | 39.107
2006-06-13 | 8:16 | 14.26 | 7.29 |  |  | 60.2 | 39.109
2006-06-13 | 8:31 | 14.26 | 7.32 |  |  | 60.2 | 39.123
2006-06-13 | 8:46 | 14.26 | 7.33 |  |  | 60.2 | 39.131
2006-06-13 | 9:01 | 14.26 | 7.34 |  |  | 60.2 | 39.139
2006-06-13 | 9:16 | 14.23 | 7.35 |  |  | 60.3 | 39.17
2006-06-13 | 9:31 | 14.22 | 7.37 |  |  | 60.3 | 39.174
2006-06-13 | 9:46 | 14.25 | 7.38 |  |  | 60.3 | 39.189
2006-06-13 | 10:01 | 14.3 | 7.4 |  |  | 60.3 | 39.207
2006-06-13 | 10:16 | 14.25 | 7.39 |  |  | 60.4 | 39.237
2006-06-13 | 10:31 | 14.23 | 7.39 |  |  | 60.4 | 39.243
2006-06-13 | 10:46 | 14.29 | 7.4 |  |  | 60.4 | 39.25
2006-06-13 | 11:01 | 14.3 | 7.41 |  |  | 60.4 | 39.261
2006-06-13 | 11:16 | 14.3 | 7.42 |  |  | 60.4 | 39.274
2006-06-13 | 11:31 | 14.34 | 7.44 |  |  | 60.4 | 39.277
2006-06-13 | 11:46 | 14.36 | 7.43 |  |  | 60.4 | 39.288
2006-06-13 | 12:01 | 14.34 | 7.41 |  |  | 60.4 | 39.239
2006-06-13 | 12:16 | 14.38 | 7.42 |  |  | 60.4 | 39.228
2006-06-13 | 12:31 | 14.35 | 7.4 |  |  | 60.5 | 39.298
2006-06-13 | 12:46 | 14.47 | 7.39 |  |  | 60.4 | 39.284
2006-06-13 | 13:00 | 14.68 | 7.44 |  |  | 60.4 | 39.272
2006-06-13 | 13:15 | 14.67 | 7.45 |  |  | 60.4 | 39.291
2006-06-13 | 13:30 | 14.54 | 7.45 |  |  | 60.5 | 39.304
2006-06-13 | 13:45 | 14.47 | 7.43 |  |  | 60.5 | 39.335
2006-06-13 | 14:00 | 14.43 | 7.38 |  |  | 60.5 | 39.323
2006-06-13 | 14:15 | 14.46 | 7.42 |  |  | 60.4 | 39.29
2006-06-13 | 14:30 | 14.53 | 7.44 |  |  | 60.4 | 39.283
2006-06-13 | 14:45 | 14.58 | 7.46 |  |  | 60.4 | 39.283
2006-06-13 | 15:00 | 14.57 | 7.46 |  |  | 60.4 | 39.291
2006-06-13 | 15:15 | 14.52 | 7.43 |  |  | 60.4 | 39.273
2006-06-13 | 15:30 | 14.51 | 7.41 |  |  | 60.4 | 39.26
2006-06-13 | 15:45 | 14.56 | 7.43 |  |  | 60.4 | 39.244
2006-06-13 | 16:00 | 14.58 | 7.43 |  |  | 60.3 | 39.224
2006-06-13 | 16:15 | 14.61 | 7.46 |  |  | 60.3 | 39.205
2006-06-13 | 16:30 | 14.61 | 7.46 |  |  | 60.3 | 39.189
2006-06-13 | 16:45 | 14.56 | 7.43 |  |  | 60.3 | 39.217
2006-06-13 | 17:00 | 14.55 | 7.41 |  |  | 60.3 | 39.213
2006-06-13 | 17:15 | 14.57 | 7.42 |  |  | 60.3 | 39.18
2006-06-13 | 17:30 | 14.56 | 7.42 |  |  | 60.3 | 39.171
2006-06-13 | 17:45 | 14.53 | 7.42 |  |  | 60.2 | 39.144
2006-06-13 | 18:00 | 14.51 | 7.39 |  |  | 60.1 | 39.081
2006-06-13 | 18:15 | 14.48 | 7.38 |  |  | 60.1 | 39.087
2006-06-13 | 18:30 | 14.45 | 7.36 |  |  | 60.2 | 39.108
2006-06-13 | 18:45 | 14.42 | 7.34 |  |  | 60.1 | 39.067
2006-06-13 | 19:00 | 14.39 | 7.32 |  |  | 60 | 38.987
2006-06-13 | 19:15 | 14.37 | 7.32 |  |  | 59.8 | 38.883
2006-06-13 | 19:30 | 14.38 | 7.33 |  |  | 59.9 | 38.916
2006-06-13 | 19:45 | 14.41 | 7.34 |  |  | 59.9 | 38.935
2006-06-13 | 20:00 | 14.42 | 7.33 |  |  | 60 | 38.971
2006-06-13 | 20:15 | 14.42 | 7.33 |  |  | 60 | 39.013
2006-06-13 | 20:30 | 14.42 | 7.31 |  |  | 60.1 | 39.033
2006-06-13 | 20:45 | 14.41 | 7.29 |  |  | 60.1 | 39.074
2006-06-13 | 21:00 | 14.39 | 7.27 |  |  | 60.2 | 39.111
2006-06-13 | 21:15 | 14.39 | 7.26 |  |  | 60.2 | 39.112
2006-06-13 | 21:30 | 14.39 | 7.24 |  |  | 60.2 | 39.12
2006-06-13 | 21:45 | 14.38 | 7.22 |  |  | 60.2 | 39.147
2006-06-13 | 22:00 | 14.37 | 7.22 |  |  | 60.2 | 39.139
2006-06-13 | 22:15 | 14.36 | 7.21 |  |  | 60.2 | 39.143
2006-06-13 | 22:30 | 14.34 | 7.2 |  |  | 60.2 | 39.148
2006-06-13 | 22:45 | 14.31 | 7.2 |  |  | 60.2 | 39.154
2006-06-13 | 23:00 | 14.29 | 7.21 |  |  | 60.2 | 39.153
2006-06-13 | 23:15 | 14.26 | 7.2 |  |  | 60.2 | 39.141
2006-06-13 | 23:30 | 14.24 | 7.2 |  |  | 60.2 | 39.146
2006-06-13 | 23:45 | 14.21 | 7.2 |  |  | 60.2 | 39.158
2006-06-14 | 0:00 | 14.19 | 7.2 |  |  | 60.2 | 39.156
2006-06-14 | 0:15 | 14.16 | 7.19 |  |  | 60.2 | 39.154
2006-06-14 | 0:30 | 14.13 | 7.19 |  |  | 60.3 | 39.165
2006-06-14 | 0:45 | 14.1 | 7.18 |  |  | 60.3 | 39.17
2006-06-14 | 1:01 | 14.07 | 7.18 |  |  | 60.3 | 39.183
2006-06-14 | 1:16 | 14.03 | 7.19 |  |  | 60.3 | 39.186
2006-06-14 | 1:31 | 13.99 | 7.19 |  |  | 60.3 | 39.187
2006-06-14 | 1:46 | 13.95 | 7.19 |  |  | 60.3 | 39.194
2006-06-14 | 2:01 | 13.92 | 7.18 |  |  | 60.3 | 39.199
2006-06-14 | 2:16 | 13.88 | 7.18 |  |  | 60.3 | 39.215
2006-06-14 | 2:31 | 13.85 | 7.18 |  |  | 60.3 | 39.214
2006-06-14 | 2:46 | 13.81 | 7.18 |  |  | 60.4 | 39.236
2006-06-14 | 3:01 | 13.77 | 7.17 |  |  | 60.4 | 39.246
2006-06-14 | 3:16 | 13.74 | 7.17 |  |  | 60.4 | 39.257
2006-06-14 | 3:31 | 13.71 | 7.18 |  |  | 60.4 | 39.26
2006-06-14 | 3:46 | 13.68 | 7.17 |  |  | 60.4 | 39.27
2006-06-14 | 4:01 | 13.64 | 7.17 |  |  | 60.4 | 39.272
2006-06-14 | 4:16 | 13.62 | 7.16 |  |  | 60.4 | 39.272
2006-06-14 | 4:31 | 13.59 | 7.17 |  |  | 60.4 | 39.274
2006-06-14 | 4:46 | 13.56 | 7.17 |  |  | 60.4 | 39.274
2006-06-14 | 5:01 | 13.53 | 7.16 |  |  | 60.4 | 39.284
2006-06-14 | 5:16 | 13.49 | 7.17 |  |  | 60.4 | 39.288
2006-06-14 | 5:31 | 13.44 | 7.17 |  |  | 60.4 | 39.29
2006-06-14 | 5:46 | 13.39 | 7.17 |  |  | 60.4 | 39.284
2006-06-14 | 6:01 | 13.34 | 7.18 |  |  | 60.4 | 39.276
2006-06-14 | 6:16 | 13.29 | 7.17 |  |  | 60.4 | 39.292
2006-06-14 | 6:31 | 13.24 | 7.19 |  |  | 60.4 | 39.291
2006-06-14 | 6:46 | 13.19 | 7.19 |  |  | 60.4 | 39.291
2006-06-14 | 7:01 | 13.13 | 7.2 |  |  | 60.5 | 39.303
2006-06-14 | 7:16 | 13.1 | 7.21 |  |  | 60.4 | 39.286
2006-06-14 | 7:31 | 13.05 | 7.23 |  |  | 60.5 | 39.293
2006-06-14 | 7:46 | 13.02 | 7.24 |  |  | 60.4 | 39.283
2006-06-14 | 8:01 | 13 | 7.25 |  |  | 60.5 | 39.296
2006-06-14 | 8:16 | 12.99 | 7.27 |  |  | 60.4 | 39.282
2006-06-14 | 8:31 | 12.99 | 7.29 |  |  | 60.5 | 39.295
2006-06-14 | 8:46 | 12.99 | 7.3 |  |  | 60.5 | 39.3
2006-06-14 | 9:01 | 13 | 7.31 |  |  | 60.5 | 39.323
2006-06-14 | 9:16 | 13.03 | 7.34 |  |  | 60.5 | 39.325
2006-06-14 | 9:31 | 13.11 | 7.37 |  |  | 60.5 | 39.317
2006-06-14 | 9:46 | 13.14 | 7.39 |  |  | 60.5 | 39.351
2006-06-14 | 10:01 | 13.19 | 7.41 |  |  | 60.6 | 39.368
2006-06-14 | 10:16 | 13.2 | 7.42 |  |  | 60.6 | 39.394
2006-06-14 | 10:31 | 13.27 | 7.44 |  |  | 60.6 | 39.414
2006-06-14 | 10:46 | 13.37 | 7.46 |  |  | 60.7 | 39.423
2006-06-14 | 11:01 | 13.43 | 7.46 |  |  | 60.7 | 39.423
2006-06-14 | 11:16 | 13.48 | 7.48 |  |  | 60.7 | 39.447
2006-06-14 | 11:31 | 13.56 | 7.49 |  |  | 60.7 | 39.45
2006-06-14 | 11:46 | 13.63 | 7.5 |  |  | 60.7 | 39.456
2006-06-14 | 12:01 | 13.61 | 7.5 |  |  | 60.7 | 39.461
2006-06-14 | 12:16 | 13.61 | 7.51 |  |  | 60.7 | 39.457
2006-06-14 | 12:31 | 13.64 | 7.51 |  |  | 60.7 | 39.452
2006-06-14 | 12:46 | 13.67 | 7.52 |  |  | 60.7 | 39.447
2006-06-14 | 13:00 | 13.75 | 7.53 |  |  | 60.7 | 39.435
2006-06-14 | 13:15 | 13.73 | 7.53 |  |  | 60.6 | 39.415
2006-06-14 | 13:30 | 13.81 | 7.54 |  |  | 60.6 | 39.406
2006-06-14 | 13:45 | 13.96 | 7.55 |  |  | 60.6 | 39.403
2006-06-14 | 14:00 | 14.05 | 7.57 |  |  | 60.6 | 39.404
2006-06-14 | 14:15 | 14.09 | 7.56 |  |  | 60.6 | 39.387
2006-06-14 | 14:30 | 14.22 | 7.58 |  |  | 60.6 | 39.394
2006-06-14 | 14:45 | 14.28 | 7.58 |  |  | 60.6 | 39.382
2006-06-14 | 15:00 | 14.22 | 7.57 |  |  | 60.6 | 39.391
2006-06-14 | 15:15 | 14.27 | 7.57 |  |  | 60.6 | 39.377
2006-06-14 | 15:30 | 14.26 | 7.56 |  |  | 60.6 | 39.364
2006-06-14 | 15:45 | 14.37 | 7.57 |  |  | 60.6 | 39.36
2006-06-14 | 16:00 | 14.42 | 7.57 |  |  | 60.5 | 39.35
2006-06-14 | 16:15 | 14.47 | 7.58 |  |  | 60.5 | 39.351
2006-06-14 | 16:30 | 14.49 | 7.57 |  |  | 60.5 | 39.347
2006-06-14 | 16:45 | 14.46 | 7.57 |  |  | 60.5 | 39.34
2006-06-14 | 17:00 | 14.46 | 7.55 |  |  | 60.5 | 39.322
2006-06-14 | 17:15 | 14.44 | 7.54 |  |  | 60.5 | 39.31
2006-06-14 | 17:30 | 14.48 | 7.55 |  |  | 60.5 | 39.303
2006-06-14 | 17:45 | 14.43 | 7.53 |  |  | 60.4 | 39.291
2006-06-14 | 18:00 | 14.4 | 7.52 |  |  | 60.4 | 39.274
2006-06-14 | 18:15 | 14.43 | 7.51 |  |  | 60.4 | 39.254
2006-06-14 | 18:30 | 14.4 | 7.5 |  |  | 60.4 | 39.247
2006-06-14 | 18:45 | 14.33 | 7.48 |  |  | 60.4 | 39.249
2006-06-14 | 19:00 | 14.26 | 7.46 |  |  | 60.4 | 39.242
2006-06-14 | 19:15 | 14.22 | 7.43 |  |  | 60.4 | 39.245
2006-06-14 | 19:30 | 14.17 | 7.41 |  |  | 60.4 | 39.244
2006-06-14 | 19:45 | 14.12 | 7.39 |  |  | 60.4 | 39.244
2006-06-14 | 20:00 | 14.08 | 7.34 |  |  | 60.4 | 39.234
2006-06-14 | 20:15 | 14.02 | 7.35 |  |  | 60.4 | 39.252
2006-06-14 | 20:30 | 13.98 | 7.3 |  |  | 60.4 | 39.246
2006-06-14 | 20:45 | 13.94 | 7.31 |  |  | 60.4 | 39.259
2006-06-14 | 21:00 | 13.91 | 7.28 |  |  | 60.4 | 39.247
2006-06-14 | 21:15 | 13.88 | 7.27 |  |  | 60.4 | 39.253
2006-06-14 | 21:30 | 13.85 | 7.27 |  |  | 60.4 | 39.241
2006-06-14 | 21:45 | 13.84 | 7.26 |  |  | 60.4 | 39.24
2006-06-14 | 22:00 | 13.82 | 7.24 |  |  | 60.4 | 39.246
2006-06-14 | 22:15 | 13.8 | 7.25 |  |  | 60.4 | 39.244
2006-06-14 | 22:30 | 13.79 | 7.23 |  |  | 60.4 | 39.239
2006-06-14 | 22:45 | 13.77 | 7.23 |  |  | 60.4 | 39.246
2006-06-14 | 23:00 | 13.75 | 7.22 |  |  | 60.4 | 39.243
2006-06-14 | 23:15 | 13.74 | 7.2 |  |  | 60.4 | 39.238
2006-06-14 | 23:30 | 13.73 | 7.21 |  |  | 60.4 | 39.246
2006-06-14 | 23:45 | 13.73 | 7.19 |  |  | 60.4 | 39.252
2006-06-15 | 0:00 | 13.73 | 7.21 |  |  | 60.4 | 39.252
2006-06-15 | 0:15 | 13.73 | 7.2 |  |  | 60.4 | 39.259
2006-06-15 | 0:30 | 13.74 | 7.2 |  |  | 60.4 | 39.26
2006-06-15 | 0:45 | 13.75 | 7.2 |  |  | 60.4 | 39.268
2006-06-15 | 1:01 | 13.77 | 7.2 |  |  | 60.4 | 39.281
2006-06-15 | 1:16 | 13.78 | 7.2 |  |  | 60.4 | 39.284
2006-06-15 | 1:31 | 13.79 | 7.2 |  |  | 60.5 | 39.298
2006-06-15 | 1:46 | 13.79 | 7.21 |  |  | 60.5 | 39.299
2006-06-15 | 2:01 | 13.79 | 7.2 |  |  | 60.5 | 39.321
2006-06-15 | 2:16 | 13.79 | 7.19 |  |  | 60.5 | 39.335
2006-06-15 | 2:31 | 13.78 | 7.21 |  |  | 60.5 | 39.357
2006-06-15 | 2:46 | 13.77 | 7.2 |  |  | 60.6 | 39.368
2006-06-15 | 3:01 | 13.75 | 7.19 |  |  | 60.6 | 39.377
2006-06-15 | 3:16 | 13.73 | 7.18 |  |  | 60.6 | 39.409
2006-06-15 | 3:31 | 13.71 | 7.19 |  |  | 60.6 | 39.411
2006-06-15 | 3:46 | 13.69 | 7.19 |  |  | 60.7 | 39.424
2006-06-15 | 4:01 | 13.67 | 7.18 |  |  | 60.7 | 39.438
2006-06-15 | 4:16 | 13.64 | 7.18 |  |  | 60.7 | 39.446
2006-06-15 | 4:31 | 13.62 | 7.18 |  |  | 60.7 | 39.47
2006-06-15 | 4:46 | 13.61 | 7.18 |  |  | 60.7 | 39.486
2006-06-15 | 5:01 | 13.58 | 7.17 |  |  | 60.8 | 39.492
2006-06-15 | 5:16 | 13.55 | 7.17 |  |  | 60.8 | 39.503
2006-06-15 | 5:31 | 13.51 | 7.17 |  |  | 60.8 | 39.511
2006-06-15 | 5:46 | 13.48 | 7.17 |  |  | 60.8 | 39.507
2006-06-15 | 6:01 | 13.44 | 7.18 |  |  | 60.8 | 39.513
2006-06-15 | 6:16 | 13.4 | 7.18 |  |  | 60.8 | 39.517
2006-06-15 | 6:31 | 13.36 | 7.18 |  |  | 60.8 | 39.519
2006-06-15 | 6:46 | 13.32 | 7.19 |  |  | 60.8 | 39.521
2006-06-15 | 7:01 | 13.28 | 7.19 |  |  | 60.8 | 39.509
2006-06-15 | 7:16 | 13.25 | 7.2 |  |  | 60.8 | 39.511
2006-06-15 | 7:31 | 13.22 | 7.22 |  |  | 60.8 | 39.513
2006-06-15 | 7:46 | 13.22 | 7.24 |  |  | 60.8 | 39.495
2006-06-15 | 8:01 | 13.19 | 7.25 |  |  | 60.8 | 39.526
2006-06-15 | 8:16 | 13.18 | 7.26 |  |  | 60.8 | 39.52
2006-06-15 | 8:31 | 13.2 | 7.28 |  |  | 60.8 | 39.503
2006-06-15 | 8:46 | 13.26 | 7.32 |  |  | 60.8 | 39.49
2006-06-15 | 9:01 | 13.26 | 7.34 |  |  | 60.8 | 39.513
2006-06-15 | 9:16 | 13.27 | 7.36 |  |  | 60.8 | 39.517
2006-06-15 | 9:31 | 13.27 | 7.38 |  |  | 60.8 | 39.538
2006-06-15 | 9:46 | 13.3 | 7.39 |  |  | 60.8 | 39.527
2006-06-15 | 10:01 | 13.36 | 7.4 |  |  | 60.8 | 39.547
2006-06-15 | 10:16 | 13.5 | 7.43 |  |  | 60.9 | 39.56
2006-06-15 | 10:31 | 13.39 | 7.43 |  |  | 60.9 | 39.579
2006-06-15 | 10:46 | 13.4 | 7.43 |  |  | 60.9 | 39.586
2006-06-15 | 11:01 | 13.43 | 7.44 |  |  | 60.9 | 39.581
2006-06-15 | 11:16 | 13.43 | 7.45 |  |  | 60.9 | 39.594
2006-06-15 | 11:31 | 13.55 | 7.48 |  |  | 60.9 | 39.593
2006-06-15 | 11:46 | 13.6 | 7.48 |  |  | 60.9 | 39.611
2006-06-15 | 12:01 | 13.57 | 7.48 |  |  | 61 | 39.626
2006-06-15 | 12:16 | 13.7 | 7.48 |  |  | 61 | 39.618
2006-06-15 | 12:31 | 13.92 | 7.51 |  |  | 61 | 39.633
2006-06-15 | 12:46 | 14.16 | 7.53 |  |  | 61 | 39.633
2006-06-15 | 13:00 | 14.34 | 7.54 |  |  | 61 | 39.651
2006-06-15 | 13:15 | 14.48 | 7.57 |  |  | 61 | 39.67
2006-06-15 | 13:30 | 14.6 | 7.56 |  |  | 61 | 39.664