## cont multiparameter audit
Organization | City Of Bend |  | Station ID |  | DR 160.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Tumalo Boulder Filed
Contact Name | Drexell Barnes |  | Latitude
Number | 541-322-6368 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | 04F11310AC
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2006-10-19 10:30 | 7.6 | 6.52 | 11.53 | 0 | 67.1 | 0
Sonde |  | 7.56 | 6.65 | 12.38 | 103.4 | 70.2
Grade |  | A | A | B |  | A
Audit 2 | 2006-10-25 10:45 | 5.48 | 6.74 | 12.03 | 0 | 73.3 | 0
Sonde |  | 5.42 | 6.65 | 14.11 | 111.7 | 77
Grade |  | A | A | 0 |  | A
Audit 3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Sonde
Grade |  | #REF! | #REF! | #REF! |  | #REF!
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | City Of Bend |  | Station ID |  | DR 160.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Tumalo Boulder Filed
Contact Name | Drexell Barnes |  | Latitude |  | 0
Number | 541-322-6368 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-10-19 10:30 | 7.6 | 6.52 | 11.53 |  | 67.1
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | City Of Bend |  | Station ID |  | DR 160.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Tumalo Boulder Filed
Contact Name | Drexell Barnes |  | Latitude |  | 0
Number | 541-322-6368 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-10-25 10:45 | 5.48 | 6.74 | 12.03 |  | 73.3
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | DO Conc | DO% | SpCond | TDS
|  | °C |  | mg/L | % | uS/cm | mg/L
2006-10-19 | 10:31 | 7.56 | 6.65 |  |  | 70.2 | 45.606
2006-10-19 | 10:46 | 7.6 | 6.67 |  |  | 70.2 | 45.613
2006-10-19 | 11:01 | 7.65 | 6.69 |  |  | 70.2 | 45.616
2006-10-19 | 11:16 | 7.71 | 6.7 |  |  | 70.2 | 45.642
2006-10-19 | 11:31 | 7.74 | 6.71 |  |  | 70.2 | 45.655
2006-10-19 | 11:46 | 7.75 | 6.72 |  |  | 70.3 | 45.677
2006-10-19 | 12:01 | 7.76 | 6.72 |  |  | 70.3 | 45.686
2006-10-19 | 12:16 | 7.77 | 6.72 |  |  | 70.3 | 45.673
2006-10-19 | 12:31 | 7.78 | 6.72 |  |  | 70.1 | 45.588
2006-10-19 | 12:46 | 7.78 | 6.72 |  |  | 70 | 45.48
2006-10-19 | 13:00 | 7.78 | 6.71 |  |  | 69.9 | 45.406
2006-10-19 | 13:15 | 7.82 | 6.73 |  |  | 69.8 | 45.355
2006-10-19 | 13:30 | 7.88 | 6.74 |  |  | 69.7 | 45.325
2006-10-19 | 13:45 | 7.92 | 6.74 |  |  | 69.7 | 45.308
2006-10-19 | 14:00 | 7.97 | 6.74 |  |  | 69.7 | 45.313
2006-10-19 | 14:15 | 8.02 | 6.75 |  |  | 69.7 | 45.283
2006-10-19 | 14:30 | 8.05 | 6.75 |  |  | 69.6 | 45.265
2006-10-19 | 14:45 | 8.09 | 6.76 |  |  | 69.6 | 45.236
2006-10-19 | 15:00 | 8.12 | 6.76 |  |  | 69.5 | 45.2
2006-10-19 | 15:15 | 8.14 | 6.76 |  |  | 69.5 | 45.186
2006-10-19 | 15:30 | 8.19 | 6.77 |  |  | 69.5 | 45.169
2006-10-19 | 15:45 | 8.21 | 6.76 |  |  | 69.5 | 45.154
2006-10-19 | 16:00 | 8.21 | 6.75 |  |  | 69.5 | 45.17
2006-10-19 | 16:15 | 8.2 | 6.74 |  |  | 69.5 | 45.172
2006-10-19 | 16:30 | 8.2 | 6.73 |  |  | 69.5 | 45.167
2006-10-19 | 16:45 | 8.21 | 6.72 |  |  | 69.5 | 45.192
2006-10-19 | 17:00 | 8.22 | 6.72 |  |  | 69.7 | 45.293
2006-10-19 | 17:15 | 8.24 | 6.7 |  |  | 69.8 | 45.369
2006-10-19 | 17:30 | 8.24 | 6.7 |  |  | 69.9 | 45.438
2006-10-19 | 17:45 | 8.24 | 6.69 |  |  | 70 | 45.483
2006-10-19 | 18:00 | 8.26 | 6.68 |  |  | 70 | 45.478
2006-10-19 | 18:15 | 8.28 | 6.67 |  |  | 70 | 45.482
2006-10-19 | 18:30 | 8.29 | 6.67 |  |  | 70 | 45.491
2006-10-19 | 18:45 | 8.29 | 6.65 |  |  | 70 | 45.474
2006-10-19 | 19:00 | 8.28 | 6.65 |  |  | 70 | 45.468
2006-10-19 | 19:15 | 8.27 | 6.64 |  |  | 70 | 45.471
2006-10-19 | 19:30 | 8.26 | 6.63 |  |  | 69.9 | 45.457
2006-10-19 | 19:45 | 8.25 | 6.63 |  |  | 70 | 45.521
2006-10-19 | 20:00 | 8.24 | 6.62 |  |  | 70.1 | 45.571
2006-10-19 | 20:15 | 8.21 | 6.62 |  |  | 70.2 | 45.604
2006-10-19 | 20:30 | 8.19 | 6.61 |  |  | 70.2 | 45.617
2006-10-19 | 20:45 | 8.17 | 6.61 |  |  | 70.1 | 45.592
2006-10-19 | 21:00 | 8.14 | 6.61 |  |  | 70 | 45.531
2006-10-19 | 21:15 | 8.12 | 6.61 |  |  | 70 | 45.505
2006-10-19 | 21:30 | 8.1 | 6.6 |  |  | 70 | 45.484
2006-10-19 | 21:45 | 8.08 | 6.6 |  |  | 70 | 45.478
2006-10-19 | 22:00 | 8.07 | 6.59 |  |  | 69.9 | 45.447
2006-10-19 | 22:15 | 8.05 | 6.59 |  |  | 69.9 | 45.419
2006-10-19 | 22:30 | 8.03 | 6.6 |  |  | 69.8 | 45.367
2006-10-19 | 22:45 | 8.02 | 6.59 |  |  | 69.8 | 45.374
2006-10-19 | 23:00 | 8.01 | 6.59 |  |  | 69.7 | 45.332
2006-10-19 | 23:15 | 8 | 6.59 |  |  | 69.7 | 45.312
2006-10-19 | 23:30 | 7.99 | 6.59 |  |  | 69.6 | 45.23
2006-10-19 | 23:45 | 7.99 | 6.59 |  |  | 69.4 | 45.134
2006-10-20 | 0:00 | 7.98 | 6.6 |  |  | 69.3 | 45.028
2006-10-20 | 0:15 | 7.97 | 6.59 |  |  | 69.1 | 44.934
2006-10-20 | 0:30 | 7.97 | 6.59 |  |  | 69 | 44.874
2006-10-20 | 0:45 | 7.97 | 6.59 |  |  | 68.9 | 44.811
2006-10-20 | 1:01 | 7.96 | 6.58 |  |  | 68.9 | 44.77
2006-10-20 | 1:16 | 7.95 | 6.59 |  |  | 68.8 | 44.726
2006-10-20 | 1:31 | 7.95 | 6.59 |  |  | 68.8 | 44.691
2006-10-20 | 1:46 | 7.94 | 6.59 |  |  | 68.7 | 44.679
2006-10-20 | 2:01 | 7.93 | 6.6 |  |  | 68.8 | 44.72
2006-10-20 | 2:16 | 7.93 | 6.59 |  |  | 68.9 | 44.794
2006-10-20 | 2:31 | 7.93 | 6.59 |  |  | 69 | 44.879
2006-10-20 | 2:46 | 7.93 | 6.6 |  |  | 69.2 | 44.952
2006-10-20 | 3:01 | 7.93 | 6.6 |  |  | 69.2 | 45.011
2006-10-20 | 3:16 | 7.92 | 6.6 |  |  | 69.3 | 45.044
2006-10-20 | 3:31 | 7.92 | 6.6 |  |  | 69.4 | 45.105
2006-10-20 | 3:46 | 7.91 | 6.6 |  |  | 69.4 | 45.141
2006-10-20 | 4:01 | 7.9 | 6.59 |  |  | 69.5 | 45.17
2006-10-20 | 4:16 | 7.89 | 6.59 |  |  | 69.6 | 45.219
2006-10-20 | 4:31 | 7.88 | 6.59 |  |  | 69.6 | 45.259
2006-10-20 | 4:46 | 7.87 | 6.6 |  |  | 69.7 | 45.286
2006-10-20 | 5:01 | 7.86 | 6.59 |  |  | 69.7 | 45.305
2006-10-20 | 5:16 | 7.83 | 6.59 |  |  | 69.7 | 45.329
2006-10-20 | 5:31 | 7.8 | 6.59 |  |  | 69.8 | 45.381
2006-10-20 | 5:46 | 7.77 | 6.58 |  |  | 69.9 | 45.441
2006-10-20 | 6:01 | 7.73 | 6.59 |  |  | 69.9 | 45.465
2006-10-20 | 6:16 | 7.69 | 6.59 |  |  | 70 | 45.522
2006-10-20 | 6:31 | 7.65 | 6.59 |  |  | 70.1 | 45.541
2006-10-20 | 6:46 | 7.6 | 6.59 |  |  | 70.1 | 45.563
2006-10-20 | 7:01 | 7.56 | 6.59 |  |  | 70.1 | 45.581
2006-10-20 | 7:16 | 7.51 | 6.59 |  |  | 70.2 | 45.626
2006-10-20 | 7:31 | 7.46 | 6.58 |  |  | 70.2 | 45.652
2006-10-20 | 7:46 | 7.42 | 6.59 |  |  | 70.2 | 45.645
2006-10-20 | 8:01 | 7.37 | 6.59 |  |  | 70.2 | 45.634
2006-10-20 | 8:16 | 7.32 | 6.59 |  |  | 70.2 | 45.64
2006-10-20 | 8:31 | 7.28 | 6.59 |  |  | 70.2 | 45.652
2006-10-20 | 8:46 | 7.24 | 6.59 |  |  | 70.2 | 45.627
2006-10-20 | 9:01 | 7.21 | 6.61 |  |  | 70.2 | 45.636
2006-10-20 | 9:16 | 7.17 | 6.61 |  |  | 70.1 | 45.534
2006-10-20 | 9:31 | 7.11 | 6.61 |  |  | 69.7 | 45.324
2006-10-20 | 9:46 | 7.02 | 6.62 |  |  | 69.1 | 44.89
2006-10-20 | 10:01 | 6.93 | 6.61 |  |  | 68.3 | 44.37
2006-10-20 | 10:16 | 6.84 | 6.62 |  |  | 67.4 | 43.83
2006-10-20 | 10:31 | 6.76 | 6.63 |  |  | 66.6 | 43.312
2006-10-20 | 10:46 | 6.81 | 6.62 |  |  | 67.1 | 43.606
2006-10-20 | 11:01 | 7.06 | 6.65 |  |  | 69.2 | 44.952
2006-10-20 | 11:16 | 7.11 | 6.66 |  |  | 69.3 | 45.049
2006-10-20 | 11:31 | 7.11 | 6.68 |  |  | 68.8 | 44.717
2006-10-20 | 11:46 | 7.11 | 6.68 |  |  | 68.2 | 44.351
2006-10-20 | 12:01 | 7.11 | 6.7 |  |  | 67.7 | 43.973
2006-10-20 | 12:16 | 7.12 | 6.7 |  |  | 67.1 | 43.589
2006-10-20 | 12:31 | 7.13 | 6.7 |  |  | 66.5 | 43.227
2006-10-20 | 12:46 | 7.17 | 6.72 |  |  | 66.1 | 42.936
2006-10-20 | 13:00 | 7.22 | 6.75 |  |  | 65.7 | 42.692
2006-10-20 | 13:15 | 7.26 | 6.76 |  |  | 65.4 | 42.485
2006-10-20 | 13:30 | 7.3 | 6.75 |  |  | 64.9 | 42.2
2006-10-20 | 13:45 | 7.33 | 6.76 |  |  | 64.7 | 42.025
2006-10-20 | 14:00 | 7.39 | 6.77 |  |  | 64.4 | 41.858
2006-10-20 | 14:15 | 7.43 | 6.78 |  |  | 64.1 | 41.656
2006-10-20 | 14:30 | 7.46 | 6.77 |  |  | 63.9 | 41.546
2006-10-20 | 14:45 | 7.51 | 6.77 |  |  | 63.8 | 41.446
2006-10-20 | 15:00 | 7.54 | 6.77 |  |  | 63.7 | 41.426
2006-10-20 | 15:15 | 7.56 | 6.76 |  |  | 63.7 | 41.411
2006-10-20 | 15:30 | 7.58 | 6.76 |  |  | 63.9 | 41.518
2006-10-20 | 15:45 | 8.26 | 6.83 |  |  | 74 | 48.123
2006-10-20 | 16:00 | 8.29 | 6.78 |  |  | 75.4 | 49.006
2006-10-20 | 16:15 | 8.36 | 6.76 |  |  | 75.5 | 49.106
2006-10-20 | 16:30 | 8.42 | 6.73 |  |  | 75.7 | 49.205
2006-10-20 | 16:45 | 8.35 | 6.71 |  |  | 75.7 | 49.227
2006-10-20 | 17:00 | 8.22 | 6.68 |  |  | 75.9 | 49.304
2006-10-20 | 17:15 | 8.13 | 6.68 |  |  | 76 | 49.378
2006-10-20 | 17:30 | 8.08 | 6.67 |  |  | 75.9 | 49.314
2006-10-20 | 17:45 | 8.03 | 6.66 |  |  | 75.8 | 49.271
2006-10-20 | 18:00 | 8 | 6.66 |  |  | 75.7 | 49.231
2006-10-20 | 18:15 | 7.97 | 6.66 |  |  | 75.6 | 49.17
2006-10-20 | 18:30 | 7.95 | 6.65 |  |  | 75.6 | 49.152
2006-10-20 | 18:45 | 7.94 | 6.65 |  |  | 75.6 | 49.127
2006-10-20 | 19:00 | 7.94 | 6.65 |  |  | 75.6 | 49.124
2006-10-20 | 19:15 | 7.94 | 6.65 |  |  | 75.6 | 49.146
2006-10-20 | 19:30 | 7.94 | 6.64 |  |  | 75.5 | 49.101
2006-10-20 | 19:45 | 7.95 | 6.64 |  |  | 75.6 | 49.127
2006-10-20 | 20:00 | 7.96 | 6.64 |  |  | 75.6 | 49.122
2006-10-20 | 20:15 | 7.97 | 6.64 |  |  | 75.6 | 49.117
2006-10-20 | 20:30 | 7.97 | 6.63 |  |  | 75.6 | 49.137
2006-10-20 | 20:45 | 7.98 | 6.63 |  |  | 75.6 | 49.121
2006-10-20 | 21:00 | 7.98 | 6.63 |  |  | 75.5 | 49.095
2006-10-20 | 21:15 | 7.99 | 6.63 |  |  | 75.5 | 49.088
2006-10-20 | 21:30 | 7.99 | 6.63 |  |  | 75.5 | 49.073
2006-10-20 | 21:45 | 7.99 | 6.63 |  |  | 75.4 | 49.036
2006-10-20 | 22:00 | 7.97 | 6.63 |  |  | 75.4 | 49.016
2006-10-20 | 22:15 | 7.97 | 6.63 |  |  | 75.4 | 49.037
2006-10-20 | 22:30 | 7.97 | 6.63 |  |  | 75.4 | 49.019
2006-10-20 | 22:45 | 7.96 | 6.63 |  |  | 75.4 | 49.004
2006-10-20 | 23:00 | 7.96 | 6.63 |  |  | 75.4 | 49.004
2006-10-20 | 23:15 | 7.96 | 6.63 |  |  | 75.3 | 48.975
2006-10-20 | 23:30 | 7.96 | 6.63 |  |  | 75.4 | 48.988
2006-10-20 | 23:45 | 7.95 | 6.63 |  |  | 75.3 | 48.959
2006-10-21 | 0:00 | 7.95 | 6.63 |  |  | 75.3 | 48.973
2006-10-21 | 0:15 | 7.94 | 6.63 |  |  | 75.4 | 48.984
2006-10-21 | 0:30 | 7.93 | 6.63 |  |  | 75.4 | 48.985
2006-10-21 | 0:45 | 7.93 | 6.62 |  |  | 75.4 | 49.015
2006-10-21 | 1:01 | 7.91 | 6.62 |  |  | 75.4 | 49.001
2006-10-21 | 1:16 | 7.9 | 6.63 |  |  | 75.4 | 49.021
2006-10-21 | 1:31 | 7.89 | 6.63 |  |  | 75.4 | 49.013
2006-10-21 | 1:46 | 7.88 | 6.62 |  |  | 75.4 | 49.005
2006-10-21 | 2:01 | 7.86 | 6.62 |  |  | 75.4 | 49.014
2006-10-21 | 2:16 | 7.85 | 6.62 |  |  | 75.4 | 49.039
2006-10-21 | 2:31 | 7.83 | 6.62 |  |  | 75.5 | 49.06
2006-10-21 | 2:46 | 7.81 | 6.62 |  |  | 75.5 | 49.088
2006-10-21 | 3:01 | 7.78 | 6.62 |  |  | 75.5 | 49.074
2006-10-21 | 3:16 | 7.75 | 6.62 |  |  | 75.6 | 49.109
2006-10-21 | 3:31 | 7.73 | 6.62 |  |  | 75.6 | 49.142
2006-10-21 | 3:46 | 7.69 | 6.62 |  |  | 75.6 | 49.146
2006-10-21 | 4:01 | 7.66 | 6.62 |  |  | 75.6 | 49.143
2006-10-21 | 4:16 | 7.62 | 6.62 |  |  | 75.6 | 49.139
2006-10-21 | 4:31 | 7.59 | 6.63 |  |  | 75.6 | 49.157
2006-10-21 | 4:46 | 7.56 | 6.63 |  |  | 75.7 | 49.188
2006-10-21 | 5:01 | 7.52 | 6.62 |  |  | 75.7 | 49.201
2006-10-21 | 5:16 | 7.48 | 6.62 |  |  | 75.7 | 49.202
2006-10-21 | 5:31 | 7.45 | 6.62 |  |  | 75.7 | 49.232
2006-10-21 | 5:46 | 7.41 | 6.62 |  |  | 75.7 | 49.229
2006-10-21 | 6:01 | 7.37 | 6.61 |  |  | 75.7 | 49.237
2006-10-21 | 6:16 | 7.32 | 6.61 |  |  | 75.8 | 49.266
2006-10-21 | 6:31 | 7.27 | 6.62 |  |  | 75.8 | 49.293
2006-10-21 | 6:46 | 7.22 | 6.61 |  |  | 75.9 | 49.334
2006-10-21 | 7:01 | 7.17 | 6.61 |  |  | 75.9 | 49.352
2006-10-21 | 7:16 | 7.11 | 6.61 |  |  | 75.9 | 49.367
2006-10-21 | 7:31 | 7.05 | 6.62 |  |  | 76 | 49.425
2006-10-21 | 7:46 | 6.99 | 6.62 |  |  | 76.1 | 49.444
2006-10-21 | 8:01 | 6.93 | 6.62 |  |  | 76.1 | 49.461
2006-10-21 | 8:16 | 6.86 | 6.62 |  |  | 76.1 | 49.497
2006-10-21 | 8:31 | 6.8 | 6.61 |  |  | 76.2 | 49.542
2006-10-21 | 8:46 | 6.74 | 6.61 |  |  | 76.3 | 49.567
2006-10-21 | 9:01 | 6.69 | 6.62 |  |  | 76.3 | 49.598
2006-10-21 | 9:16 | 6.65 | 6.64 |  |  | 76.3 | 49.595
2006-10-21 | 9:31 | 6.6 | 6.64 |  |  | 76.3 | 49.627
2006-10-21 | 9:46 | 6.54 | 6.64 |  |  | 76.2 | 49.555
2006-10-21 | 10:01 | 6.48 | 6.64 |  |  | 76 | 49.42
2006-10-21 | 10:16 | 6.44 | 6.65 |  |  | 75.9 | 49.307
2006-10-21 | 10:31 | 6.41 | 6.66 |  |  | 75.6 | 49.161
2006-10-21 | 10:46 | 6.41 | 6.65 |  |  | 75.6 | 49.127
2006-10-21 | 11:01 | 6.42 | 6.66 |  |  | 75.6 | 49.117
2006-10-21 | 11:16 | 6.46 | 6.67 |  |  | 75.7 | 49.179
2006-10-21 | 11:31 | 6.5 | 6.68 |  |  | 75.9 | 49.316
2006-10-21 | 11:46 | 6.55 | 6.69 |  |  | 76.1 | 49.444
2006-10-21 | 12:01 | 6.6 | 6.7 |  |  | 76.3 | 49.587
2006-10-21 | 12:16 | 6.66 | 6.71 |  |  | 76.5 | 49.697
2006-10-21 | 12:31 | 6.71 | 6.72 |  |  | 76.5 | 49.754
2006-10-21 | 12:46 | 6.76 | 6.73 |  |  | 76.7 | 49.857
2006-10-21 | 13:00 | 6.81 | 6.72 |  |  | 76.8 | 49.896
2006-10-21 | 13:15 | 6.86 | 6.74 |  |  | 76.8 | 49.938
2006-10-21 | 13:30 | 6.9 | 6.73 |  |  | 76.8 | 49.934
2006-10-21 | 13:45 | 6.94 | 6.74 |  |  | 76.8 | 49.935
2006-10-21 | 14:00 | 6.97 | 6.74 |  |  | 76.8 | 49.944
2006-10-21 | 14:15 | 7 | 6.74 |  |  | 76.9 | 49.965
2006-10-21 | 14:30 | 7.01 | 6.74 |  |  | 76.9 | 49.96
2006-10-21 | 14:45 | 7.03 | 6.73 |  |  | 76.9 | 49.967
2006-10-21 | 15:00 | 7.03 | 6.73 |  |  | 76.9 | 49.962
2006-10-21 | 15:15 | 7.04 | 6.73 |  |  | 76.8 | 49.947
2006-10-21 | 15:30 | 7.03 | 6.73 |  |  | 76.8 | 49.938
2006-10-21 | 15:45 | 7.03 | 6.72 |  |  | 76.8 | 49.926
2006-10-21 | 16:00 | 7.02 | 6.71 |  |  | 76.8 | 49.945
2006-10-21 | 16:15 | 7.02 | 6.7 |  |  | 76.9 | 49.957
2006-10-21 | 16:30 | 7.02 | 6.69 |  |  | 76.8 | 49.952
2006-10-21 | 16:45 | 7.02 | 6.69 |  |  | 76.9 | 49.972
2006-10-21 | 17:00 | 7.02 | 6.68 |  |  | 76.9 | 50.014
2006-10-21 | 17:15 | 7.01 | 6.68 |  |  | 77 | 50.058
2006-10-21 | 17:30 | 7.01 | 6.67 |  |  | 77 | 50.073
2006-10-21 | 17:45 | 7.01 | 6.67 |  |  | 77 | 50.056
2006-10-21 | 18:00 | 7.01 | 6.67 |  |  | 76.9 | 50.014
2006-10-21 | 18:15 | 7.01 | 6.65 |  |  | 77 | 50.025
2006-10-21 | 18:30 | 7.02 | 6.66 |  |  | 76.9 | 49.969
2006-10-21 | 18:45 | 7.02 | 6.64 |  |  | 76.8 | 49.933
2006-10-21 | 19:00 | 7.03 | 6.64 |  |  | 76.8 | 49.889
2006-10-21 | 19:15 | 7.05 | 6.64 |  |  | 76.7 | 49.836
2006-10-21 | 19:30 | 7.06 | 6.64 |  |  | 76.6 | 49.792
2006-10-21 | 19:45 | 7.08 | 6.63 |  |  | 76.5 | 49.744
2006-10-21 | 20:00 | 7.1 | 6.63 |  |  | 76.5 | 49.696
2006-10-21 | 20:15 | 7.12 | 6.63 |  |  | 76.4 | 49.675
2006-10-21 | 20:30 | 7.14 | 6.64 |  |  | 76.4 | 49.653
2006-10-21 | 20:45 | 7.16 | 6.63 |  |  | 76.3 | 49.627
2006-10-21 | 21:00 | 7.18 | 6.62 |  |  | 76.3 | 49.626
2006-10-21 | 21:15 | 7.2 | 6.63 |  |  | 76.3 | 49.612
2006-10-21 | 21:30 | 7.22 | 6.62 |  |  | 76.3 | 49.596
2006-10-21 | 21:45 | 7.23 | 6.63 |  |  | 76.3 | 49.585
2006-10-21 | 22:00 | 7.24 | 6.63 |  |  | 76.2 | 49.554
2006-10-21 | 22:15 | 7.25 | 6.63 |  |  | 76.2 | 49.516
2006-10-21 | 22:30 | 7.26 | 6.63 |  |  | 76.2 | 49.504
2006-10-21 | 22:45 | 7.26 | 6.62 |  |  | 76.1 | 49.475
2006-10-21 | 23:00 | 7.27 | 6.62 |  |  | 76.1 | 49.459
2006-10-21 | 23:15 | 7.27 | 6.62 |  |  | 76.1 | 49.436
2006-10-21 | 23:30 | 7.28 | 6.63 |  |  | 76.1 | 49.436
2006-10-21 | 23:45 | 7.29 | 6.62 |  |  | 76 | 49.425
2006-10-22 | 0:00 | 7.29 | 6.62 |  |  | 76 | 49.419
2006-10-22 | 0:15 | 7.3 | 6.62 |  |  | 76 | 49.429
2006-10-22 | 0:30 | 7.3 | 6.62 |  |  | 76 | 49.424
2006-10-22 | 0:45 | 7.3 | 6.62 |  |  | 76 | 49.405
2006-10-22 | 1:01 | 7.3 | 6.62 |  |  | 76 | 49.423
2006-10-22 | 1:16 | 7.3 | 6.62 |  |  | 76 | 49.412
2006-10-22 | 1:31 | 7.3 | 6.62 |  |  | 76 | 49.415
2006-10-22 | 1:46 | 7.3 | 6.62 |  |  | 76.1 | 49.445
2006-10-22 | 2:01 | 7.3 | 6.61 |  |  | 76.1 | 49.44
2006-10-22 | 2:16 | 7.29 | 6.61 |  |  | 76.1 | 49.434
2006-10-22 | 2:31 | 7.28 | 6.61 |  |  | 76.1 | 49.461
2006-10-22 | 2:46 | 7.27 | 6.61 |  |  | 76.1 | 49.473
2006-10-22 | 3:01 | 7.25 | 6.61 |  |  | 76.1 | 49.483
2006-10-22 | 3:16 | 7.24 | 6.61 |  |  | 76.2 | 49.516
2006-10-22 | 3:31 | 7.22 | 6.61 |  |  | 76.2 | 49.526
2006-10-22 | 3:46 | 7.19 | 6.61 |  |  | 76.2 | 49.522
2006-10-22 | 4:01 | 7.17 | 6.61 |  |  | 76.2 | 49.536
2006-10-22 | 4:16 | 7.14 | 6.61 |  |  | 76.2 | 49.562
2006-10-22 | 4:31 | 7.12 | 6.6 |  |  | 76.3 | 49.566
2006-10-22 | 4:46 | 7.08 | 6.6 |  |  | 76.3 | 49.564
2006-10-22 | 5:01 | 7.06 | 6.6 |  |  | 76.3 | 49.568
2006-10-22 | 5:16 | 7.03 | 6.6 |  |  | 76.3 | 49.6
2006-10-22 | 5:31 | 7 | 6.61 |  |  | 76.3 | 49.606
2006-10-22 | 5:46 | 6.97 | 6.62 |  |  | 76.3 | 49.617
2006-10-22 | 6:01 | 6.94 | 6.6 |  |  | 76.3 | 49.627
2006-10-22 | 6:16 | 6.9 | 6.6 |  |  | 76.4 | 49.649
2006-10-22 | 6:31 | 6.86 | 6.6 |  |  | 76.4 | 49.67
2006-10-22 | 6:46 | 6.82 | 6.6 |  |  | 76.4 | 49.689
2006-10-22 | 7:01 | 6.78 | 6.6 |  |  | 76.5 | 49.695
2006-10-22 | 7:16 | 6.73 | 6.6 |  |  | 76.5 | 49.701
2006-10-22 | 7:31 | 6.68 | 6.59 |  |  | 76.5 | 49.745
2006-10-22 | 7:46 | 6.63 | 6.6 |  |  | 76.5 | 49.754
2006-10-22 | 8:01 | 6.57 | 6.6 |  |  | 76.6 | 49.778
2006-10-22 | 8:16 | 6.52 | 6.6 |  |  | 76.6 | 49.811
2006-10-22 | 8:31 | 6.46 | 6.6 |  |  | 76.6 | 49.818
2006-10-22 | 8:46 | 6.41 | 6.6 |  |  | 76.7 | 49.848
2006-10-22 | 9:01 | 6.37 | 6.6 |  |  | 76.8 | 49.888
2006-10-22 | 9:16 | 6.32 | 6.61 |  |  | 76.8 | 49.892
2006-10-22 | 9:31 | 6.28 | 6.62 |  |  | 76.8 | 49.927
2006-10-22 | 9:46 | 6.24 | 6.64 |  |  | 76.8 | 49.948
2006-10-22 | 10:01 | 6.22 | 6.63 |  |  | 76.9 | 49.968
2006-10-22 | 10:16 | 6.2 | 6.64 |  |  | 76.9 | 49.974
2006-10-22 | 10:31 | 6.19 | 6.63 |  |  | 76.9 | 49.963
2006-10-22 | 10:46 | 6.2 | 6.65 |  |  | 76.9 | 49.971
2006-10-22 | 11:01 | 6.21 | 6.65 |  |  | 76.9 | 50.007
2006-10-22 | 11:16 | 6.22 | 6.66 |  |  | 77 | 50.023
2006-10-22 | 11:31 | 6.25 | 6.67 |  |  | 76.9 | 50.015
2006-10-22 | 11:46 | 6.28 | 6.68 |  |  | 77 | 50.038
2006-10-22 | 12:01 | 6.31 | 6.68 |  |  | 77 | 50.033
2006-10-22 | 12:16 | 6.36 | 6.69 |  |  | 77 | 50.054
2006-10-22 | 12:31 | 6.41 | 6.7 |  |  | 77 | 50.079
2006-10-22 | 12:46 | 6.46 | 6.7 |  |  | 77 | 50.064
2006-10-22 | 13:00 | 6.51 | 6.71 |  |  | 77 | 50.044
2006-10-22 | 13:15 | 6.56 | 6.71 |  |  | 77 | 50.047
2006-10-22 | 13:30 | 6.6 | 6.72 |  |  | 77 | 50.019
2006-10-22 | 13:45 | 6.64 | 6.72 |  |  | 76.9 | 49.992
2006-10-22 | 14:00 | 6.67 | 6.72 |  |  | 76.9 | 49.992
2006-10-22 | 14:15 | 6.7 | 6.73 |  |  | 76.8 | 49.949
2006-10-22 | 14:30 | 6.72 | 6.72 |  |  | 76.9 | 49.953
2006-10-22 | 14:45 | 6.73 | 6.73 |  |  | 76.8 | 49.927
2006-10-22 | 15:00 | 6.74 | 6.73 |  |  | 76.8 | 49.917
2006-10-22 | 15:15 | 6.75 | 6.72 |  |  | 76.7 | 49.885
2006-10-22 | 15:30 | 6.76 | 6.72 |  |  | 76.7 | 49.876
2006-10-22 | 15:45 | 6.76 | 6.71 |  |  | 76.7 | 49.859
2006-10-22 | 16:00 | 6.77 | 6.7 |  |  | 76.7 | 49.835
2006-10-22 | 16:15 | 6.77 | 6.7 |  |  | 76.7 | 49.827
2006-10-22 | 16:30 | 6.78 | 6.69 |  |  | 76.7 | 49.844
2006-10-22 | 16:45 | 6.79 | 6.69 |  |  | 76.7 | 49.837
2006-10-22 | 17:00 | 6.81 | 6.68 |  |  | 76.7 | 49.842
2006-10-22 | 17:15 | 6.81 | 6.67 |  |  | 76.6 | 49.803
2006-10-22 | 17:30 | 6.82 | 6.67 |  |  | 76.6 | 49.786
2006-10-22 | 17:45 | 6.83 | 6.67 |  |  | 76.6 | 49.763
2006-10-22 | 18:00 | 6.84 | 6.67 |  |  | 76.5 | 49.733
2006-10-22 | 18:15 | 6.85 | 6.65 |  |  | 76.5 | 49.696
2006-10-22 | 18:30 | 6.86 | 6.64 |  |  | 76.4 | 49.689
2006-10-22 | 18:45 | 6.88 | 6.65 |  |  | 76.4 | 49.662
2006-10-22 | 19:00 | 6.9 | 6.64 |  |  | 76.4 | 49.629
2006-10-22 | 19:15 | 6.92 | 6.63 |  |  | 76.3 | 49.609
2006-10-22 | 19:30 | 6.95 | 6.63 |  |  | 76.3 | 49.579
2006-10-22 | 19:45 | 6.97 | 6.63 |  |  | 76.2 | 49.547
2006-10-22 | 20:00 | 7 | 6.62 |  |  | 76.2 | 49.547
2006-10-22 | 20:15 | 7.03 | 6.62 |  |  | 76.2 | 49.514
2006-10-22 | 20:30 | 7.05 | 6.62 |  |  | 76.1 | 49.495
2006-10-22 | 20:45 | 7.08 | 6.62 |  |  | 76.1 | 49.47
2006-10-22 | 21:00 | 7.1 | 6.62 |  |  | 76.1 | 49.454
2006-10-22 | 21:15 | 7.12 | 6.62 |  |  | 76 | 49.429
2006-10-22 | 21:30 | 7.14 | 6.62 |  |  | 76 | 49.416
2006-10-22 | 21:45 | 7.16 | 6.62 |  |  | 76 | 49.391
2006-10-22 | 22:00 | 7.18 | 6.62 |  |  | 76 | 49.388
2006-10-22 | 22:15 | 7.19 | 6.61 |  |  | 75.9 | 49.352
2006-10-22 | 22:30 | 7.2 | 6.62 |  |  | 75.9 | 49.34
2006-10-22 | 22:45 | 7.21 | 6.61 |  |  | 75.9 | 49.341
2006-10-22 | 23:00 | 7.22 | 6.61 |  |  | 75.9 | 49.309
2006-10-22 | 23:15 | 7.23 | 6.62 |  |  | 75.9 | 49.306
2006-10-22 | 23:30 | 7.23 | 6.62 |  |  | 75.8 | 49.281
2006-10-22 | 23:45 | 7.24 | 6.62 |  |  | 75.8 | 49.287
2006-10-23 | 0:00 | 7.25 | 6.61 |  |  | 75.8 | 49.282
2006-10-23 | 0:15 | 7.26 | 6.61 |  |  | 75.8 | 49.293
2006-10-23 | 0:30 | 7.26 | 6.61 |  |  | 75.8 | 49.28
2006-10-23 | 0:45 | 7.27 | 6.61 |  |  | 75.8 | 49.295
2006-10-23 | 1:01 | 7.28 | 6.61 |  |  | 75.8 | 49.296
2006-10-23 | 1:16 | 7.28 | 6.61 |  |  | 75.8 | 49.266
2006-10-23 | 1:31 | 7.28 | 6.61 |  |  | 75.8 | 49.284
2006-10-23 | 1:46 | 7.29 | 6.61 |  |  | 75.8 | 49.301
2006-10-23 | 2:01 | 7.29 | 6.61 |  |  | 75.8 | 49.295
2006-10-23 | 2:16 | 7.29 | 6.61 |  |  | 75.9 | 49.304
2006-10-23 | 2:31 | 7.29 | 6.6 |  |  | 75.8 | 49.295
2006-10-23 | 2:46 | 7.29 | 6.61 |  |  | 75.8 | 49.299
2006-10-23 | 3:01 | 7.28 | 6.61 |  |  | 75.9 | 49.307
2006-10-23 | 3:16 | 7.27 | 6.61 |  |  | 75.9 | 49.319
2006-10-23 | 3:31 | 7.26 | 6.61 |  |  | 75.9 | 49.326
2006-10-23 | 3:46 | 7.25 | 6.6 |  |  | 75.9 | 49.341
2006-10-23 | 4:01 | 7.24 | 6.6 |  |  | 75.9 | 49.336
2006-10-23 | 4:16 | 7.23 | 6.61 |  |  | 75.9 | 49.364
2006-10-23 | 4:31 | 7.22 | 6.6 |  |  | 75.9 | 49.355
2006-10-23 | 4:46 | 7.2 | 6.61 |  |  | 76 | 49.382
2006-10-23 | 5:01 | 7.19 | 6.6 |  |  | 76 | 49.386
2006-10-23 | 5:16 | 7.18 | 6.6 |  |  | 76 | 49.399
2006-10-23 | 5:31 | 7.16 | 6.6 |  |  | 76 | 49.408
2006-10-23 | 5:46 | 7.14 | 6.6 |  |  | 76 | 49.4
2006-10-23 | 6:01 | 7.11 | 6.59 |  |  | 76.1 | 49.433
2006-10-23 | 6:16 | 7.08 | 6.6 |  |  | 76 | 49.429
2006-10-23 | 6:31 | 7.05 | 6.6 |  |  | 76.1 | 49.47
2006-10-23 | 6:46 | 7.02 | 6.6 |  |  | 76.1 | 49.479
2006-10-23 | 7:01 | 6.98 | 6.59 |  |  | 76.1 | 49.496
2006-10-23 | 7:16 | 6.94 | 6.6 |  |  | 76.2 | 49.524
2006-10-23 | 7:31 | 6.89 | 6.6 |  |  | 76.2 | 49.554
2006-10-23 | 7:46 | 6.84 | 6.6 |  |  | 76.2 | 49.56
2006-10-23 | 8:01 | 6.8 | 6.6 |  |  | 76.3 | 49.598
2006-10-23 | 8:16 | 6.75 | 6.6 |  |  | 76.3 | 49.613
2006-10-23 | 8:31 | 6.71 | 6.61 |  |  | 76.3 | 49.625
2006-10-23 | 8:46 | 6.66 | 6.61 |  |  | 76.4 | 49.654
2006-10-23 | 9:01 | 6.61 | 6.61 |  |  | 76.5 | 49.696
2006-10-23 | 9:16 | 6.57 | 6.61 |  |  | 76.5 | 49.735
2006-10-23 | 9:31 | 6.52 | 6.62 |  |  | 76.6 | 49.764
2006-10-23 | 9:46 | 6.48 | 6.63 |  |  | 76.6 | 49.785
2006-10-23 | 10:01 | 6.45 | 6.63 |  |  | 76.6 | 49.809
2006-10-23 | 10:16 | 6.42 | 6.63 |  |  | 76.6 | 49.818
2006-10-23 | 10:31 | 6.41 | 6.64 |  |  | 76.7 | 49.836
2006-10-23 | 10:46 | 6.41 | 6.65 |  |  | 76.6 | 49.809
2006-10-23 | 11:01 | 6.42 | 6.66 |  |  | 76.6 | 49.819
2006-10-23 | 11:16 | 6.44 | 6.67 |  |  | 76.6 | 49.803
2006-10-23 | 11:31 | 6.46 | 6.67 |  |  | 76.6 | 49.805
2006-10-23 | 11:46 | 6.49 | 6.69 |  |  | 76.6 | 49.808
2006-10-23 | 12:01 | 6.52 | 6.7 |  |  | 76.6 | 49.821
2006-10-23 | 12:16 | 6.56 | 6.7 |  |  | 76.6 | 49.79
2006-10-23 | 12:31 | 6.6 | 6.71 |  |  | 76.6 | 49.791
2006-10-23 | 12:46 | 6.64 | 6.72 |  |  | 76.6 | 49.794
2006-10-23 | 13:00 | 6.68 | 6.73 |  |  | 76.6 | 49.789
2006-10-23 | 13:15 | 6.73 | 6.73 |  |  | 76.6 | 49.791
2006-10-23 | 13:30 | 6.78 | 6.73 |  |  | 76.6 | 49.782
2006-10-23 | 13:45 | 6.82 | 6.74 |  |  | 76.5 | 49.748
2006-10-23 | 14:00 | 6.86 | 6.74 |  |  | 76.6 | 49.762
2006-10-23 | 14:15 | 6.89 | 6.74 |  |  | 76.6 | 49.759
2006-10-23 | 14:30 | 6.92 | 6.74 |  |  | 76.5 | 49.745
2006-10-23 | 14:45 | 6.93 | 6.74 |  |  | 76.5 | 49.738
2006-10-23 | 15:00 | 6.94 | 6.74 |  |  | 76.5 | 49.734
2006-10-23 | 15:15 | 6.95 | 6.73 |  |  | 76.5 | 49.733
2006-10-23 | 15:30 | 6.95 | 6.73 |  |  | 76.5 | 49.727
2006-10-23 | 15:45 | 6.94 | 6.72 |  |  | 76.5 | 49.709
2006-10-23 | 16:00 | 6.93 | 6.72 |  |  | 76.5 | 49.716
2006-10-23 | 16:15 | 6.93 | 6.71 |  |  | 76.5 | 49.704
2006-10-23 | 16:30 | 6.93 | 6.71 |  |  | 76.4 | 49.682
2006-10-23 | 16:45 | 6.93 | 6.7 |  |  | 76.5 | 49.696
2006-10-23 | 17:00 | 6.94 | 6.7 |  |  | 76.4 | 49.654
2006-10-23 | 17:15 | 6.94 | 6.69 |  |  | 76.4 | 49.643
2006-10-23 | 17:30 | 6.93 | 6.69 |  |  | 76.3 | 49.621
2006-10-23 | 17:45 | 6.93 | 6.68 |  |  | 76.3 | 49.618
2006-10-23 | 18:00 | 6.93 | 6.67 |  |  | 76.3 | 49.611
2006-10-23 | 18:15 | 6.93 | 6.67 |  |  | 76.3 | 49.589
2006-10-23 | 18:30 | 6.92 | 6.66 |  |  | 76.3 | 49.576
2006-10-23 | 18:45 | 6.93 | 6.65 |  |  | 76.2 | 49.556
2006-10-23 | 19:00 | 6.93 | 6.65 |  |  | 76.2 | 49.551
2006-10-23 | 19:15 | 6.93 | 6.65 |  |  | 76.2 | 49.529
2006-10-23 | 19:30 | 6.94 | 6.64 |  |  | 76.2 | 49.512
2006-10-23 | 19:45 | 6.95 | 6.64 |  |  | 76.1 | 49.483
2006-10-23 | 20:00 | 6.97 | 6.63 |  |  | 76.1 | 49.463
2006-10-23 | 20:15 | 6.98 | 6.64 |  |  | 76.1 | 49.449
2006-10-23 | 20:30 | 7 | 6.63 |  |  | 76 | 49.419
2006-10-23 | 20:45 | 7.02 | 6.63 |  |  | 76 | 49.386
2006-10-23 | 21:00 | 7.04 | 6.63 |  |  | 76 | 49.389
2006-10-23 | 21:15 | 7.06 | 6.63 |  |  | 76 | 49.371
2006-10-23 | 21:30 | 7.08 | 6.63 |  |  | 75.9 | 49.332
2006-10-23 | 21:45 | 7.09 | 6.62 |  |  | 75.9 | 49.318
2006-10-23 | 22:00 | 7.1 | 6.62 |  |  | 75.8 | 49.302
2006-10-23 | 22:15 | 7.11 | 6.62 |  |  | 75.8 | 49.286
2006-10-23 | 22:30 | 7.12 | 6.62 |  |  | 75.8 | 49.266
2006-10-23 | 22:45 | 7.13 | 6.62 |  |  | 75.8 | 49.254
2006-10-23 | 23:00 | 7.14 | 6.62 |  |  | 75.8 | 49.242
2006-10-23 | 23:15 | 7.15 | 6.62 |  |  | 75.7 | 49.217
2006-10-23 | 23:30 | 7.17 | 6.62 |  |  | 75.7 | 49.211
2006-10-23 | 23:45 | 7.19 | 6.62 |  |  | 75.7 | 49.224
2006-10-24 | 0:00 | 7.2 | 6.62 |  |  | 75.7 | 49.224
2006-10-24 | 0:15 | 7.21 | 6.62 |  |  | 75.7 | 49.205
2006-10-24 | 0:30 | 7.21 | 6.62 |  |  | 75.7 | 49.208
2006-10-24 | 0:45 | 7.22 | 6.62 |  |  | 75.7 | 49.198
2006-10-24 | 1:01 | 7.22 | 6.62 |  |  | 75.7 | 49.194
2006-10-24 | 1:16 | 7.22 | 6.62 |  |  | 75.7 | 49.19
2006-10-24 | 1:31 | 7.23 | 6.62 |  |  | 75.7 | 49.191
2006-10-24 | 1:46 | 7.23 | 6.61 |  |  | 75.7 | 49.205
2006-10-24 | 2:01 | 7.23 | 6.62 |  |  | 75.7 | 49.195
2006-10-24 | 2:16 | 7.23 | 6.61 |  |  | 75.7 | 49.236
2006-10-24 | 2:31 | 7.23 | 6.61 |  |  | 75.7 | 49.235
2006-10-24 | 2:46 | 7.22 | 6.61 |  |  | 75.7 | 49.219
2006-10-24 | 3:01 | 7.22 | 6.61 |  |  | 75.8 | 49.247
2006-10-24 | 3:16 | 7.22 | 6.61 |  |  | 75.8 | 49.273
2006-10-24 | 3:31 | 7.21 | 6.61 |  |  | 75.8 | 49.265
2006-10-24 | 3:46 | 7.21 | 6.62 |  |  | 75.8 | 49.283
2006-10-24 | 4:01 | 7.2 | 6.61 |  |  | 75.8 | 49.291
2006-10-24 | 4:16 | 7.19 | 6.61 |  |  | 75.9 | 49.304
2006-10-24 | 4:31 | 7.17 | 6.61 |  |  | 75.9 | 49.309
2006-10-24 | 4:46 | 7.16 | 6.61 |  |  | 75.9 | 49.317
2006-10-24 | 5:01 | 7.14 | 6.6 |  |  | 75.9 | 49.329
2006-10-24 | 5:16 | 7.12 | 6.61 |  |  | 75.9 | 49.342
2006-10-24 | 5:31 | 7.11 | 6.61 |  |  | 75.9 | 49.364
2006-10-24 | 5:46 | 7.09 | 6.61 |  |  | 76 | 49.383
2006-10-24 | 6:01 | 7.07 | 6.61 |  |  | 76 | 49.401
2006-10-24 | 6:16 | 7.05 | 6.61 |  |  | 76 | 49.426
2006-10-24 | 6:31 | 7.03 | 6.61 |  |  | 76.1 | 49.44
2006-10-24 | 6:46 | 7.01 | 6.62 |  |  | 76.1 | 49.447
2006-10-24 | 7:01 | 6.98 | 6.61 |  |  | 76.1 | 49.479
2006-10-24 | 7:16 | 6.94 | 6.62 |  |  | 76.1 | 49.493
2006-10-24 | 7:31 | 6.9 | 6.62 |  |  | 76.2 | 49.539
2006-10-24 | 7:46 | 6.86 | 6.62 |  |  | 76.3 | 49.563
2006-10-24 | 8:01 | 6.81 | 6.62 |  |  | 76.3 | 49.586
2006-10-24 | 8:16 | 6.77 | 6.62 |  |  | 76.3 | 49.625
2006-10-24 | 8:31 | 6.72 | 6.62 |  |  | 76.4 | 49.653
2006-10-24 | 8:46 | 6.68 | 6.63 |  |  | 76.4 | 49.691
2006-10-24 | 9:01 | 6.63 | 6.63 |  |  | 76.5 | 49.708
2006-10-24 | 9:16 | 6.59 | 6.63 |  |  | 76.5 | 49.749
2006-10-24 | 9:31 | 6.55 | 6.64 |  |  | 76.5 | 49.752
2006-10-24 | 9:46 | 6.52 | 6.64 |  |  | 76.6 | 49.791
2006-10-24 | 10:01 | 6.5 | 6.65 |  |  | 76.6 | 49.805
2006-10-24 | 10:16 | 6.48 | 6.65 |  |  | 76.7 | 49.843
2006-10-24 | 10:31 | 6.46 | 6.66 |  |  | 76.6 | 49.818
2006-10-24 | 10:46 | 6.45 | 6.67 |  |  | 76.6 | 49.817
2006-10-24 | 11:01 | 6.45 | 6.68 |  |  | 76.6 | 49.815
2006-10-24 | 11:16 | 6.46 | 6.69 |  |  | 76.6 | 49.818
2006-10-24 | 11:31 | 6.48 | 6.7 |  |  | 76.6 | 49.814
2006-10-24 | 11:46 | 6.5 | 6.7 |  |  | 76.6 | 49.805
2006-10-24 | 12:01 | 6.53 | 6.72 |  |  | 76.6 | 49.78
2006-10-24 | 12:16 | 6.56 | 6.72 |  |  | 76.6 | 49.796
2006-10-24 | 12:31 | 6.6 | 6.73 |  |  | 76.6 | 49.796
2006-10-24 | 12:46 | 6.64 | 6.73 |  |  | 76.6 | 49.796
2006-10-24 | 13:00 | 6.68 | 6.74 |  |  | 76.6 | 49.801
2006-10-24 | 13:15 | 6.69 | 6.74 |  |  | 76.6 | 49.799
2006-10-24 | 13:30 | 6.69 | 6.74 |  |  | 76.6 | 49.812
2006-10-24 | 13:45 | 6.71 | 6.74 |  |  | 76.6 | 49.79
2006-10-24 | 14:00 | 6.69 | 6.74 |  |  | 76.6 | 49.81
2006-10-24 | 14:15 | 6.69 | 6.74 |  |  | 76.6 | 49.815
2006-10-24 | 14:30 | 6.73 | 6.74 |  |  | 76.6 | 49.785
2006-10-24 | 14:45 | 6.73 | 6.75 |  |  | 76.6 | 49.795
2006-10-24 | 15:00 | 6.71 | 6.75 |  |  | 76.6 | 49.767
2006-10-24 | 15:15 | 6.7 | 6.74 |  |  | 76.6 | 49.768
2006-10-24 | 15:30 | 6.68 | 6.74 |  |  | 76.6 | 49.774
2006-10-24 | 15:45 | 6.65 | 6.73 |  |  | 76.6 | 49.759
2006-10-24 | 16:00 | 6.63 | 6.73 |  |  | 76.6 | 49.776
2006-10-24 | 16:15 | 6.61 | 6.72 |  |  | 76.6 | 49.782
2006-10-24 | 16:30 | 6.56 | 6.71 |  |  | 76.6 | 49.778
2006-10-24 | 16:45 | 6.53 | 6.71 |  |  | 76.5 | 49.747
2006-10-24 | 17:00 | 6.51 | 6.7 |  |  | 76.5 | 49.697
2006-10-24 | 17:15 | 6.49 | 6.7 |  |  | 76.5 | 49.698
2006-10-24 | 17:30 | 6.48 | 6.69 |  |  | 76.4 | 49.676
2006-10-24 | 17:45 | 6.47 | 6.69 |  |  | 76.4 | 49.691
2006-10-24 | 18:00 | 6.46 | 6.68 |  |  | 76.4 | 49.676
2006-10-24 | 18:15 | 6.46 | 6.68 |  |  | 76.4 | 49.67
2006-10-24 | 18:30 | 6.47 | 6.67 |  |  | 76.4 | 49.641
2006-10-24 | 18:45 | 6.48 | 6.67 |  |  | 76.3 | 49.603
2006-10-24 | 19:00 | 6.49 | 6.67 |  |  | 76.3 | 49.609
2006-10-24 | 19:15 | 6.5 | 6.66 |  |  | 76.3 | 49.627
2006-10-24 | 19:30 | 6.51 | 6.66 |  |  | 76.4 | 49.652
2006-10-24 | 19:45 | 6.53 | 6.65 |  |  | 76.4 | 49.659
2006-10-24 | 20:00 | 6.55 | 6.65 |  |  | 76.4 | 49.645
2006-10-24 | 20:15 | 6.57 | 6.65 |  |  | 76.3 | 49.622
2006-10-24 | 20:30 | 6.59 | 6.65 |  |  | 76.3 | 49.59
2006-10-24 | 20:45 | 6.6 | 6.65 |  |  | 76.3 | 49.564
2006-10-24 | 21:00 | 6.62 | 6.64 |  |  | 76.3 | 49.573
2006-10-24 | 21:15 | 6.64 | 6.64 |  |  | 76.3 | 49.586
2006-10-24 | 21:30 | 6.66 | 6.65 |  |  | 76.2 | 49.56
2006-10-24 | 21:45 | 6.67 | 6.64 |  |  | 76.2 | 49.53
2006-10-24 | 22:00 | 6.68 | 6.64 |  |  | 76.2 | 49.512
2006-10-24 | 22:15 | 6.69 | 6.64 |  |  | 76.2 | 49.512
2006-10-24 | 22:30 | 6.7 | 6.64 |  |  | 76.1 | 49.492
2006-10-24 | 22:45 | 6.7 | 6.64 |  |  | 76.1 | 49.481
2006-10-24 | 23:00 | 6.7 | 6.63 |  |  | 76.1 | 49.485
2006-10-24 | 23:15 | 6.7 | 6.64 |  |  | 76.1 | 49.468
2006-10-24 | 23:30 | 6.69 | 6.63 |  |  | 76.1 | 49.438
2006-10-24 | 23:45 | 6.68 | 6.64 |  |  | 76 | 49.426
2006-10-25 | 0:00 | 6.67 | 6.63 |  |  | 76 | 49.399
2006-10-25 | 0:15 | 6.66 | 6.63 |  |  | 76 | 49.381
2006-10-25 | 0:30 | 6.65 | 6.63 |  |  | 76 | 49.389
2006-10-25 | 0:45 | 6.63 | 6.63 |  |  | 76 | 49.392
2006-10-25 | 1:01 | 6.61 | 6.64 |  |  | 76 | 49.401
2006-10-25 | 1:16 | 6.59 | 6.63 |  |  | 76 | 49.373
2006-10-25 | 1:31 | 6.57 | 6.63 |  |  | 76 | 49.386
2006-10-25 | 1:46 | 6.55 | 6.63 |  |  | 76 | 49.385
2006-10-25 | 2:01 | 6.53 | 6.63 |  |  | 76 | 49.374
2006-10-25 | 2:16 | 6.51 | 6.63 |  |  | 76 | 49.376
2006-10-25 | 2:31 | 6.48 | 6.62 |  |  | 76 | 49.373
2006-10-25 | 2:46 | 6.46 | 6.63 |  |  | 76 | 49.379
2006-10-25 | 3:01 | 6.43 | 6.62 |  |  | 76 | 49.372
2006-10-25 | 3:16 | 6.41 | 6.62 |  |  | 76 | 49.375
2006-10-25 | 3:31 | 6.38 | 6.63 |  |  | 75.9 | 49.363
2006-10-25 | 3:46 | 6.35 | 6.62 |  |  | 75.9 | 49.362
2006-10-25 | 4:01 | 6.33 | 6.62 |  |  | 75.9 | 49.347
2006-10-25 | 4:16 | 6.3 | 6.62 |  |  | 75.9 | 49.356
2006-10-25 | 4:31 | 6.28 | 6.62 |  |  | 75.9 | 49.364
2006-10-25 | 4:46 | 6.25 | 6.62 |  |  | 75.9 | 49.354
2006-10-25 | 5:01 | 6.23 | 6.62 |  |  | 76 | 49.381
2006-10-25 | 5:16 | 6.2 | 6.62 |  |  | 75.9 | 49.36
2006-10-25 | 5:31 | 6.17 | 6.62 |  |  | 76 | 49.409
2006-10-25 | 5:46 | 6.14 | 6.62 |  |  | 76.1 | 49.441
2006-10-25 | 6:01 | 6.11 | 6.62 |  |  | 76.1 | 49.497
2006-10-25 | 6:16 | 6.08 | 6.62 |  |  | 76.2 | 49.542
2006-10-25 | 6:31 | 6.05 | 6.62 |  |  | 76.3 | 49.607
2006-10-25 | 6:46 | 6.01 | 6.61 |  |  | 76.4 | 49.656
2006-10-25 | 7:01 | 5.97 | 6.61 |  |  | 76.4 | 49.687
2006-10-25 | 7:16 | 5.92 | 6.61 |  |  | 76.5 | 49.73
2006-10-25 | 7:31 | 5.88 | 6.61 |  |  | 76.5 | 49.753
2006-10-25 | 7:46 | 5.83 | 6.61 |  |  | 76.6 | 49.779
2006-10-25 | 8:01 | 5.78 | 6.61 |  |  | 76.7 | 49.835
2006-10-25 | 8:16 | 5.73 | 6.61 |  |  | 76.7 | 49.853
2006-10-25 | 8:31 | 5.68 | 6.62 |  |  | 76.7 | 49.867
2006-10-25 | 8:46 | 5.64 | 6.62 |  |  | 76.8 | 49.927
2006-10-25 | 9:01 | 5.59 | 6.62 |  |  | 76.9 | 49.969
2006-10-25 | 9:16 | 5.55 | 6.63 |  |  | 76.9 | 49.98
2006-10-25 | 9:31 | 5.51 | 6.63 |  |  | 77 | 50.031
2006-10-25 | 9:46 | 5.48 | 6.63 |  |  | 77 | 50.056
2006-10-25 | 10:01 | 5.46 | 6.63 |  |  | 77 | 50.069
2006-10-25 | 10:16 | 5.44 | 6.64 |  |  | 77.1 | 50.083
2006-10-25 | 10:31 | 5.42 | 6.65 |  |  | 77 | 50.08
2006-10-25 | 10:46 | 5.42 | 6.65 |  |  | 77 | 50.036