## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, OWRD Gauge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 6920 sonde #06A2389AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-07-27 |  | 2006-07-27 |  | A
Time | Computer set to atomic time |  | 22:20 |  | 22:20 |  | A
Temp | VWR NIST traceable surface water |  | 18.1 |  | 18.1 |  | A
DO | Winkler titration of surface water |  | 8.8 |  | 8.5 |  | A
SpC | 148 µS/mL |  | 59 |  | 57 |  | A
Ph | pH 7 standard solution |  | 7.3 |  | 7.7 |  | B
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2006-07-28 15:23 | 12.6 | 6.8 | 10.6 | 99.8 | 21.5 | 14
Sonde |  | 12.53 |  | 10.03 |  | 21.3
Grade |  | A | Fail | B |  | A
Audit 2 | 2006-07-29 11:30 | 8 | 6.6 | 12.2 | 102.8 | 21.2 | 14
Sonde |  | 7.98 |  | 11.3 |  | 21.6
Grade |  | A | Fail | B |  | A
Audit 3 | 2006-08-03 14:49 | 10.8 | 6.9 | 11 | 99 | 24 | 15
Sonde |  | 10.39 |  | 10.81 |  | 24.6
Grade |  | A | Fail | A |  | A
Post-sampling Field Equipment Check
Field equipment ID | YSI 6920 sonde #06A2389AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-07 |  | 2006-08-07 |  | A
Time | Computer set to atomic time |  | 14:20 |  | 14:20 |  | A
Temp | VWR NIST traceable surface water |  | 17.8 |  | 17.9 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 10.3 |  | A
SpC | 148 µS/mL |  | 59 |  | 61 |  | A
Ph | pH 7 standard solution |  | 8.5 |  | 8.7 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Sonde reading for Audit 4 is in Comments in Audit 4 sheet |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, OWRD Gauge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS 05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-07-27 |  | 2006-07-27 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 22.2 |  | 22 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 8.6 |  | A
SpC | 148 µS/mL |  | 159 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.2 |  | 7.4 |  | A
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 477 |  | B
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-07-28 15:23 | 12.6 | 6.8 | 10.6 | 99.8 | 21.5 | 14
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-02 |  | 2006-08-02 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 21.4 |  | 21.2 |  | A
DO | Winkler titration of surface water |  | 9 |  | 8.1 |  | B
SpC | 148 µS/mL |  | 159 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.1 |  | 7.5 |  | B
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 478 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, OWRD Gauge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS 05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-07-27 |  | 2006-07-27 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 22.2 |  | 22 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 8.6 |  | A
SpC | 148 µS/mL |  | 159 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.2 |  | 7.4 |  | A
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 477 |  | B
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-07-29 11:30 | 8 | 6.6 | 12.2 | 102.8 | 21.2 | 14
6
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-02 |  | 2006-08-02 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 21.4 |  | 21.2 |  | A
DO | Winkler titration of surface water |  | 9 |  | 8.1 |  | B
SpC | 148 µS/mL |  | 159 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.1 |  | 7.5 |  | B
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 478 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, OWRD Gauge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS 05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-02 |  | 2006-08-02 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 21.4 |  | 21.2 |  | A
DO | Winkler titration of surface water |  | 9 |  | 8.1 |  | B
SpC | 148 µS/mL |  | 159 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.1 |  | 7.5 |  | B
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 478 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-08-03 14:49 | 10.8 | 6.9 | 11 | 99 | 24 | 15
4
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-03 |  | 2006-08-03 |  | A
Time | Computer set to atomic time |  | 18:00 |  | 18:00 |  | A
Temp | VWR NIST traceable surface water |  | 20.5 |  | 20.2 |  | A
DO | Winkler titration of surface water |  | 9 |  | 9.1 |  | A
SpC | 148 µS/mL |  | 150 |  | 141 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.4 |  | B
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 478 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, OWRD Gauge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS 05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-03 |  | 2006-08-03 |  | A
Time | Computer set to atomic time |  | 18:00 |  | 18:00 |  | A
Temp | VWR NIST traceable surface water |  | 20.5 |  | 20.2 |  | A
DO | Winkler titration of surface water |  | 9 |  | 9.1 |  | A
SpC | 148 µS/mL |  | 150 |  | 141 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.4 |  | B
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 478 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-08-06 14:44 | 11.4 | 7.1 | 10.8 | 99 | 23 | 15
6
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-07 |  | 2006-08-07 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | VWR NIST traceable surface water |  | 21.6 |  | 21.1 |  | A
DO | Winkler titration of surface water |  | 9 |  | 8.2 |  | B
SpC | 148 µS/mL |  | 152 |  | 143 |  | A
Ph | pH 7 standard solution |  | 7.1 |  | 7.4 |  | A
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 51 |  | B
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 478 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Sonde Reading for 8/6/06 14:44 Temp: 11.35; pH 7.45; DO: 10.59; SpC: 24.6 |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | DO Conc | DO% | SpCond | TDS
|  | °C |  | mg/L | % | uS/cm | mg/L
2006-07-28 | 15:30 | 12.53 | 7.47 | 10.03 | 94.2 | 21.3 | 13.869
2006-07-28 | 15:45 | 12.76 | 7.47 | 9.96 | 94 | 21.4 | 13.905
2006-07-28 | 16:00 | 13 | 7.48 | 9.9 | 94 | 21.4 | 13.914
2006-07-28 | 16:15 | 13.27 | 7.48 | 9.83 | 93.9 | 21.5 | 13.948
2006-07-28 | 16:30 | 13.53 | 7.49 | 9.8 | 94.1 | 21.5 | 13.985
2006-07-28 | 16:45 | 13.75 | 7.49 | 9.83 | 94.8 | 21.6 | 14.016
2006-07-28 | 17:00 | 13.92 | 7.49 | 9.83 | 95.2 | 21.6 | 14.045
2006-07-28 | 17:15 | 14.06 | 7.48 | 9.77 | 94.9 | 21.7 | 14.075
2006-07-28 | 17:30 | 14.15 | 7.48 | 9.78 | 95.2 | 21.7 | 14.099
2006-07-28 | 17:45 | 14.15 | 7.48 | 9.77 | 95.1 | 21.7 | 14.103
2006-07-28 | 18:00 | 14.12 | 7.47 | 9.73 | 94.6 | 21.7 | 14.112
2006-07-28 | 18:15 | 14.06 | 7.47 | 9.71 | 94.3 | 21.7 | 14.122
2006-07-28 | 18:30 | 14.01 | 7.46 | 9.67 | 93.9 | 21.7 | 14.121
2006-07-28 | 18:45 | 13.96 | 7.47 | 9.65 | 93.6 | 21.7 | 14.107
2006-07-28 | 19:00 | 13.9 | 7.47 | 9.63 | 93.3 | 21.7 | 14.108
2006-07-28 | 19:15 | 13.81 | 7.46 | 9.62 | 93 | 21.7 | 14.102
2006-07-28 | 19:30 | 13.68 | 7.45 | 9.67 | 93.1 | 21.7 | 14.089
2006-07-28 | 19:45 | 13.54 | 7.45 | 9.69 | 93 | 21.7 | 14.073
2006-07-28 | 20:00 | 13.39 | 7.44 | 9.72 | 93.1 | 21.6 | 14.072
2006-07-28 | 20:15 | 13.25 | 7.43 | 9.74 | 92.9 | 21.6 | 14.048
2006-07-28 | 20:30 | 13.12 | 7.43 | 9.76 | 92.9 | 21.6 | 14.032
2006-07-28 | 20:45 | 13 | 7.42 | 9.77 | 92.7 | 21.6 | 14.017
2006-07-28 | 21:00 | 12.86 | 7.4 | 9.81 | 92.9 | 21.5 | 13.997
2006-07-28 | 21:15 | 12.74 | 7.4 | 9.83 | 92.8 | 21.5 | 13.986
2006-07-28 | 21:30 | 12.6 | 7.39 | 9.86 | 92.8 | 21.5 | 13.963
2006-07-28 | 21:45 | 12.46 | 7.39 | 9.9 | 92.8 | 21.5 | 13.966
2006-07-28 | 22:00 | 12.31 | 7.38 | 9.92 | 92.7 | 21.4 | 13.935
2006-07-28 | 22:15 | 12.15 | 7.38 | 9.96 | 92.7 | 21.4 | 13.918
2006-07-28 | 22:30 | 11.99 | 7.37 | 9.99 | 92.7 | 21.4 | 13.908
2006-07-28 | 22:45 | 11.83 | 7.37 | 10.03 | 92.7 | 21.4 | 13.884
2006-07-28 | 23:00 | 11.67 | 7.36 | 10.06 | 92.7 | 21.3 | 13.861
2006-07-28 | 23:15 | 11.5 | 7.36 | 10.13 | 92.9 | 21.3 | 13.837
2006-07-28 | 23:30 | 11.33 | 7.36 | 10.16 | 92.9 | 21.3 | 13.835
2006-07-28 | 23:45 | 11.15 | 7.35 | 10.2 | 92.8 | 21.2 | 13.806
2006-07-29 | 0:00 | 10.98 | 7.35 | 10.25 | 92.9 | 21.2 | 13.793
2006-07-29 | 0:15 | 10.8 | 7.33 | 10.28 | 92.8 | 21.2 | 13.775
2006-07-29 | 0:30 | 10.62 | 7.35 | 10.34 | 93 | 21.2 | 13.772
2006-07-29 | 0:45 | 10.43 | 7.33 | 10.39 | 93 | 21.2 | 13.774
2006-07-29 | 1:01 | 10.25 | 7.32 | 10.43 | 93 | 21.2 | 13.769
2006-07-29 | 1:16 | 10.08 | 7.33 | 10.49 | 93.1 | 21.1 | 13.747
2006-07-29 | 1:31 | 9.9 | 7.32 | 10.53 | 93.1 | 21.2 | 13.752
2006-07-29 | 1:46 | 9.73 | 7.33 | 10.59 | 93.2 | 21.1 | 13.745
2006-07-29 | 2:01 | 9.56 | 7.32 | 10.63 | 93.2 | 21.1 | 13.747
2006-07-29 | 2:16 | 9.39 | 7.32 | 10.67 | 93.2 | 21.2 | 13.755
2006-07-29 | 2:31 | 9.22 | 7.31 | 10.72 | 93.2 | 21.2 | 13.783
2006-07-29 | 2:46 | 9.06 | 7.31 | 10.77 | 93.3 | 21.2 | 13.783
2006-07-29 | 3:01 | 8.9 | 7.3 | 10.81 | 93.3 | 21.2 | 13.802
2006-07-29 | 3:16 | 8.74 | 7.31 | 10.86 | 93.4 | 21.2 | 13.795
2006-07-29 | 3:31 | 8.59 | 7.3 | 10.91 | 93.4 | 21.2 | 13.808
2006-07-29 | 3:46 | 8.45 | 7.3 | 10.95 | 93.5 | 21.3 | 13.818
2006-07-29 | 4:01 | 8.31 | 7.31 | 10.99 | 93.5 | 21.3 | 13.827
2006-07-29 | 4:16 | 8.19 | 7.29 | 11.02 | 93.5 | 21.3 | 13.837
2006-07-29 | 4:31 | 8.07 | 7.3 | 11.07 | 93.6 | 21.3 | 13.838
2006-07-29 | 4:46 | 7.95 | 7.29 | 11.09 | 93.5 | 21.3 | 13.846
2006-07-29 | 5:01 | 7.83 | 7.3 | 11.14 | 93.6 | 21.3 | 13.871
2006-07-29 | 5:16 | 7.72 | 7.29 | 11.17 | 93.7 | 21.3 | 13.871
2006-07-29 | 5:31 | 7.62 | 7.29 | 11.2 | 93.7 | 21.3 | 13.877
2006-07-29 | 5:46 | 7.52 | 7.3 | 11.24 | 93.8 | 21.4 | 13.903
2006-07-29 | 6:01 | 7.41 | 7.29 | 11.28 | 93.8 | 21.4 | 13.914
2006-07-29 | 6:16 | 7.31 | 7.29 | 11.32 | 93.9 | 21.4 | 13.94
2006-07-29 | 6:31 | 7.21 | 7.3 | 11.34 | 93.9 | 21.5 | 13.945
2006-07-29 | 6:46 | 7.13 | 7.3 | 11.37 | 94 | 21.5 | 13.954
2006-07-29 | 7:01 | 7.06 | 7.3 | 11.4 | 94.1 | 21.5 | 13.956
2006-07-29 | 7:16 | 7 | 7.3 | 11.43 | 94.2 | 21.5 | 13.973
2006-07-29 | 7:31 | 6.96 | 7.3 | 11.45 | 94.2 | 21.5 | 13.991
2006-07-29 | 7:46 | 6.93 | 7.32 | 11.47 | 94.3 | 21.5 | 13.981
2006-07-29 | 8:01 | 6.9 | 7.33 | 11.49 | 94.4 | 21.5 | 14.001
2006-07-29 | 8:16 | 6.87 | 7.33 | 11.52 | 94.6 | 21.6 | 14.025
2006-07-29 | 8:31 | 6.85 | 7.34 | 11.52 | 94.5 | 21.6 | 14.031
2006-07-29 | 8:46 | 6.84 | 7.34 | 11.54 | 94.7 | 21.6 | 14.033
2006-07-29 | 9:01 | 6.86 | 7.35 | 11.53 | 94.6 | 21.6 | 14.06
2006-07-29 | 9:16 | 6.9 | 7.35 | 11.52 | 94.7 | 21.7 | 14.074
2006-07-29 | 9:31 | 6.96 | 7.35 | 11.51 | 94.7 | 21.7 | 14.084
2006-07-29 | 9:46 | 7.03 | 7.35 | 11.49 | 94.7 | 21.7 | 14.11
2006-07-29 | 10:01 | 7.1 | 7.35 | 11.49 | 94.9 | 21.7 | 14.111
2006-07-29 | 10:16 | 7.18 | 7.36 | 11.45 | 94.7 | 21.7 | 14.117
2006-07-29 | 10:31 | 7.28 | 7.36 | 11.44 | 94.9 | 21.7 | 14.126
2006-07-29 | 10:46 | 7.4 | 7.37 | 11.42 | 95 | 21.7 | 14.133
2006-07-29 | 11:01 | 7.52 | 7.37 | 11.37 | 94.9 | 21.8 | 14.154
2006-07-29 | 11:16 | 7.67 | 7.36 | 11.35 | 95.1 | 21.8 | 14.164
2006-07-29 | 11:31 | 7.98 | 7.38 | 11.3 | 95.4 | 21.6 | 14.053
2006-07-29 | 11:46 | 8.17 | 7.4 | 11.26 | 95.5 | 21.5 | 14.003
2006-07-29 | 12:01 | 8.37 | 7.39 | 11.22 | 95.6 | 21.6 | 14.01
2006-07-29 | 12:16 | 8.59 | 7.39 | 11.17 | 95.7 | 21.7 | 14.085
2006-07-29 | 12:31 | 8.85 | 7.4 | 11.11 | 95.8 | 21.5 | 14.006
2006-07-29 | 12:46 | 9.14 | 7.4 | 11.06 | 96 | 21.3 | 13.853
2006-07-29 | 13:00 | 9.43 | 7.4 | 10.97 | 95.8 | 21.4 | 13.912
2006-07-29 | 13:15 | 9.71 | 7.41 | 10.88 | 95.8 | 21.6 | 14.032
2006-07-29 | 13:30 | 10.01 | 7.4 | 10.8 | 95.7 | 21.5 | 13.987
2006-07-29 | 13:45 | 10.29 | 7.42 | 10.76 | 95.9 | 21.5 | 13.976
2006-07-29 | 14:00 | 10.55 | 7.41 | 10.68 | 95.8 | 21.5 | 13.965
2006-07-29 | 14:15 | 10.82 | 7.41 | 10.6 | 95.7 | 21.4 | 13.919
2006-07-29 | 14:30 | 11.11 | 7.42 | 10.51 | 95.5 | 21.6 | 14.02
2006-07-29 | 14:45 | 11.36 | 7.43 | 10.43 | 95.4 | 21.5 | 14.005
2006-07-29 | 15:00 | 11.59 | 7.42 | 10.38 | 95.4 | 21.6 | 14.025
2006-07-29 | 15:15 | 11.81 | 7.43 | 10.31 | 95.2 | 21.6 | 14.04
2006-07-29 | 15:30 | 11.92 | 7.43 | 10.27 | 95.1 | 21.5 | 13.971
2006-07-29 | 15:45 | 12.06 | 7.43 | 10.24 | 95.2 | 21.6 | 14.057
2006-07-29 | 16:00 | 12.24 | 7.44 | 10.21 | 95.3 | 21.5 | 13.999
2006-07-29 | 16:15 | 12.41 | 7.44 | 10.19 | 95.4 | 21.6 | 14.014
2006-07-29 | 16:30 | 12.58 | 7.43 | 10.21 | 95.9 | 21.5 | 13.964
2006-07-29 | 16:45 | 12.69 | 7.45 | 10.18 | 95.9 | 21.5 | 13.95
2006-07-29 | 17:00 | 12.75 | 7.43 | 10.16 | 95.9 | 21.5 | 13.987
2006-07-29 | 17:15 | 12.83 | 7.43 | 10.13 | 95.7 | 21.9 | 14.247
2006-07-29 | 17:30 | 12.86 | 7.43 | 10.08 | 95.4 | 21.9 | 14.217
2006-07-29 | 17:45 | 12.88 | 7.43 | 10.06 | 95.2 | 21.9 | 14.205
2006-07-29 | 18:00 | 12.87 | 7.44 | 10.07 | 95.3 | 21.9 | 14.228
2006-07-29 | 18:15 | 12.84 | 7.42 | 10.06 | 95.2 | 21.9 | 14.244
2006-07-29 | 18:30 | 12.81 | 7.42 | 10.04 | 94.9 | 21.9 | 14.22
2006-07-29 | 18:45 | 12.74 | 7.43 | 10.05 | 94.8 | 21.9 | 14.218
2006-07-29 | 19:00 | 12.65 | 7.42 | 10.07 | 94.8 | 21.8 | 14.186
2006-07-29 | 19:15 | 12.55 | 7.42 | 10.07 | 94.6 | 21.8 | 14.179
2006-07-29 | 19:30 | 12.45 | 7.42 | 10.07 | 94.4 | 21.8 | 14.179
2006-07-29 | 19:45 | 12.34 | 7.41 | 10.09 | 94.4 | 21.8 | 14.144
2006-07-29 | 20:00 | 12.23 | 7.41 | 10.11 | 94.3 | 21.8 | 14.146
2006-07-29 | 20:15 | 12.12 | 7.4 | 10.13 | 94.2 | 21.7 | 14.126
2006-07-29 | 20:30 | 12.01 | 7.39 | 10.13 | 94 | 21.7 | 14.113
2006-07-29 | 20:45 | 11.9 | 7.39 | 10.16 | 94.1 | 21.7 | 14.097
2006-07-29 | 21:00 | 11.76 | 7.38 | 10.19 | 94.1 | 21.7 | 14.084
2006-07-29 | 21:15 | 11.62 | 7.38 | 10.22 | 94 | 21.6 | 14.065
2006-07-29 | 21:30 | 11.45 | 7.38 | 10.26 | 94 | 21.6 | 14.062
2006-07-29 | 21:45 | 11.27 | 7.38 | 10.31 | 94 | 21.7 | 14.081
2006-07-29 | 22:00 | 11.08 | 7.38 | 10.35 | 94 | 21.6 | 14.049
2006-07-29 | 22:15 | 10.89 | 7.39 | 10.4 | 94.1 | 21.6 | 14.016
2006-07-29 | 22:30 | 10.7 | 7.39 | 10.44 | 94 | 21.4 | 13.894
2006-07-29 | 22:45 | 10.52 | 7.39 | 10.49 | 94.1 | 21.4 | 13.888
2006-07-29 | 23:00 | 10.33 | 7.38 | 10.55 | 94.2 | 21.3 | 13.833
2006-07-29 | 23:15 | 10.14 | 7.37 | 10.61 | 94.3 | 21.2 | 13.804
2006-07-29 | 23:30 | 9.94 | 7.37 | 10.65 | 94.2 | 21.4 | 13.911
2006-07-29 | 23:45 | 9.77 | 7.35 | 10.73 | 94.5 | 21.3 | 13.858
2006-07-30 | 0:00 | 9.6 | 7.36 | 10.75 | 94.4 | 21.5 | 13.987
2006-07-30 | 0:15 | 9.45 | 7.36 | 10.8 | 94.4 | 21.5 | 13.96
2006-07-30 | 0:30 | 9.29 | 7.36 | 10.85 | 94.5 | 21.5 | 13.99
2006-07-30 | 0:45 | 9.13 | 7.34 | 10.9 | 94.5 | 21.8 | 14.179
2006-07-30 | 1:01 | 8.97 | 7.34 | 10.94 | 94.5 | 21.9 | 14.213
2006-07-30 | 1:16 | 8.8 | 7.33 | 10.99 | 94.6 | 21.9 | 14.257
2006-07-30 | 1:31 | 8.64 | 7.34 | 11.04 | 94.7 | 22 | 14.288
2006-07-30 | 1:46 | 8.49 | 7.33 | 11.08 | 94.6 | 22 | 14.289
2006-07-30 | 2:01 | 8.34 | 7.33 | 11.13 | 94.7 | 22 | 14.302
2006-07-30 | 2:16 | 8.2 | 7.34 | 11.17 | 94.8 | 22.1 | 14.336
2006-07-30 | 2:31 | 8.06 | 7.32 | 11.22 | 94.8 | 22.1 | 14.342
2006-07-30 | 2:46 | 7.92 | 7.33 | 11.27 | 95 | 22.1 | 14.347
2006-07-30 | 3:01 | 7.79 | 7.32 | 11.3 | 94.9 | 22.1 | 14.364
2006-07-30 | 3:16 | 7.67 | 7.32 | 11.33 | 94.9 | 22.1 | 14.382
2006-07-30 | 3:31 | 7.55 | 7.32 | 11.37 | 94.9 | 22.1 | 14.389
2006-07-30 | 3:46 | 7.43 | 7.31 | 11.42 | 95.1 | 22.2 | 14.402
2006-07-30 | 4:01 | 7.31 | 7.32 | 11.45 | 95.1 | 22.1 | 14.392
2006-07-30 | 4:16 | 7.19 | 7.31 | 11.49 | 95.1 | 22.2 | 14.407
2006-07-30 | 4:31 | 7.08 | 7.31 | 11.52 | 95.1 | 22.2 | 14.415
2006-07-30 | 4:46 | 6.97 | 7.32 | 11.57 | 95.3 | 22.2 | 14.436
2006-07-30 | 5:01 | 6.86 | 7.33 | 11.62 | 95.4 | 22.2 | 14.437
2006-07-30 | 5:16 | 6.75 | 7.34 | 11.63 | 95.2 | 22.2 | 14.445
2006-07-30 | 5:31 | 6.64 | 7.35 | 11.69 | 95.4 | 22.2 | 14.447
2006-07-30 | 5:46 | 6.53 | 7.32 | 11.72 | 95.4 | 22.3 | 14.467
2006-07-30 | 6:01 | 6.43 | 7.3 | 11.77 | 95.6 | 22.2 | 14.443
2006-07-30 | 6:16 | 6.32 | 7.31 | 11.79 | 95.5 | 22.2 | 14.444
2006-07-30 | 6:31 | 6.23 | 7.31 | 11.84 | 95.7 | 22.2 | 14.441
2006-07-30 | 6:46 | 6.14 | 7.31 | 11.88 | 95.8 | 22.2 | 14.44
2006-07-30 | 7:01 | 6.06 | 7.31 | 11.92 | 95.9 | 22.2 | 14.445
2006-07-30 | 7:16 | 6 | 7.31 | 11.94 | 95.9 | 22.2 | 14.451
2006-07-30 | 7:31 | 5.95 | 7.33 | 11.99 | 96.2 | 22.2 | 14.459
2006-07-30 | 7:46 | 5.92 | 7.33 | 12.01 | 96.3 | 22.2 | 14.455
2006-07-30 | 8:01 | 5.9 | 7.33 | 12.03 | 96.4 | 22.2 | 14.434
2006-07-30 | 8:16 | 5.89 | 7.34 | 12.04 | 96.5 | 22.3 | 14.466
2006-07-30 | 8:31 | 5.9 | 7.35 | 12.06 | 96.6 | 22.3 | 14.473
2006-07-30 | 8:46 | 5.92 | 7.36 | 12.06 | 96.7 | 22.3 | 14.464
2006-07-30 | 9:01 | 5.93 | 7.37 | 12.09 | 97 | 22.2 | 14.442
2006-07-30 | 9:16 | 5.98 | 7.37 | 12.07 | 96.9 | 22.2 | 14.459
2006-07-30 | 9:31 | 6.01 | 7.37 | 12.06 | 96.9 | 22.2 | 14.42
2006-07-30 | 9:46 | 6.05 | 7.36 | 12.04 | 96.9 | 22.2 | 14.428
2006-07-30 | 10:01 | 6.1 | 7.37 | 12.02 | 96.8 | 22.2 | 14.4
2006-07-30 | 10:16 | 6.24 | 7.37 | 11.99 | 96.9 | 22.1 | 14.383
2006-07-30 | 10:31 | 6.3 | 7.38 | 11.96 | 96.8 | 22.3 | 14.466
2006-07-30 | 10:46 | 6.45 | 7.37 | 11.9 | 96.7 | 22.3 | 14.472
2006-07-30 | 11:01 | 6.67 | 7.37 | 11.83 | 96.7 | 22.2 | 14.459
2006-07-30 | 11:16 | 6.75 | 7.37 | 11.8 | 96.6 | 22.2 | 14.449
2006-07-30 | 11:31 | 6.76 | 7.38 | 11.79 | 96.6 | 22.3 | 14.478
2006-07-30 | 11:46 | 6.79 | 7.38 | 11.8 | 96.7 | 22.3 | 14.495
2006-07-30 | 12:01 | 6.98 | 7.39 | 11.73 | 96.6 | 22.3 | 14.495
2006-07-30 | 12:16 | 7.16 | 7.39 | 11.68 | 96.6 | 22.3 | 14.509
2006-07-30 | 12:31 | 7.34 | 7.38 | 11.59 | 96.3 | 22.2 | 14.461
2006-07-30 | 12:46 | 7.53 | 7.39 | 11.55 | 96.4 | 22.3 | 14.503
2006-07-30 | 13:00 | 7.69 | 7.4 | 11.51 | 96.4 | 22.2 | 14.449
2006-07-30 | 13:15 | 7.88 | 7.41 | 11.45 | 96.4 | 22.2 | 14.426
2006-07-30 | 13:30 | 8.12 | 7.4 | 11.36 | 96.2 | 22.2 | 14.453
2006-07-30 | 13:45 | 8.31 | 7.4 | 11.31 | 96.2 | 22.2 | 14.438
2006-07-30 | 14:00 | 8.48 | 7.41 | 11.25 | 96.1 | 22.2 | 14.45
2006-07-30 | 14:15 | 8.68 | 7.41 | 11.19 | 96 | 22.4 | 14.562
2006-07-30 | 14:30 | 8.83 | 7.42 | 11.15 | 96 | 22.4 | 14.579
2006-07-30 | 14:45 | 9 | 7.42 | 11.11 | 96.1 | 22.7 | 14.723
2006-07-30 | 15:00 | 9.21 | 7.42 | 11.02 | 95.9 | 22.7 | 14.742
2006-07-30 | 15:15 | 9.42 | 7.41 | 10.97 | 95.9 | 22.6 | 14.699
2006-07-30 | 15:30 | 9.63 | 7.41 | 10.91 | 95.8 | 22.5 | 14.623
2006-07-30 | 15:45 | 9.87 | 7.42 | 10.83 | 95.7 | 22.4 | 14.592
2006-07-30 | 16:00 | 10.05 | 7.42 | 10.77 | 95.5 | 22.4 | 14.567
2006-07-30 | 16:15 | 10.17 | 7.43 | 10.74 | 95.5 | 22.4 | 14.567
2006-07-30 | 16:30 | 10.28 | 7.43 | 10.7 | 95.4 | 22.4 | 14.571
2006-07-30 | 16:45 | 10.4 | 7.43 | 10.65 | 95.2 | 22.4 | 14.545
2006-07-30 | 17:00 | 10.53 | 7.43 | 10.62 | 95.3 | 22.4 | 14.585
2006-07-30 | 17:15 | 10.64 | 7.43 | 10.59 | 95.3 | 22.4 | 14.588
2006-07-30 | 17:30 | 10.72 | 7.42 | 10.57 | 95.2 | 22.4 | 14.587
2006-07-30 | 17:45 | 10.76 | 7.42 | 10.56 | 95.2 | 22.5 | 14.641
2006-07-30 | 18:00 | 10.74 | 7.43 | 10.55 | 95.1 | 22.5 | 14.646
2006-07-30 | 18:15 | 10.69 | 7.42 | 10.55 | 95 | 22.6 | 14.669
2006-07-30 | 18:30 | 10.63 | 7.42 | 10.58 | 95.1 | 22.6 | 14.665
2006-07-30 | 18:45 | 10.53 | 7.42 | 10.6 | 95.1 | 22.6 | 14.659
2006-07-30 | 19:00 | 10.44 | 7.41 | 10.62 | 95 | 22.6 | 14.672
2006-07-30 | 19:15 | 10.34 | 7.41 | 10.64 | 95 | 22.6 | 14.664
2006-07-30 | 19:30 | 10.27 | 7.41 | 10.65 | 95 | 22.6 | 14.667
2006-07-30 | 19:45 | 10.22 | 7.41 | 10.67 | 95 | 22.6 | 14.669
2006-07-30 | 20:00 | 10.18 | 7.41 | 10.68 | 95 | 22.6 | 14.692
2006-07-30 | 20:15 | 10.12 | 7.4 | 10.7 | 95 | 22.6 | 14.687
2006-07-30 | 20:30 | 10.05 | 7.4 | 10.72 | 95.1 | 22.6 | 14.69
2006-07-30 | 20:45 | 9.97 | 7.39 | 10.73 | 95 | 22.6 | 14.698
2006-07-30 | 21:00 | 9.89 | 7.38 | 10.76 | 95 | 22.6 | 14.683
2006-07-30 | 21:15 | 9.8 | 7.38 | 10.79 | 95.1 | 22.6 | 14.68
2006-07-30 | 21:30 | 9.71 | 7.38 | 10.81 | 95.1 | 22.6 | 14.704
2006-07-30 | 21:45 | 9.61 | 7.38 | 10.83 | 95 | 22.6 | 14.701
2006-07-30 | 22:00 | 9.51 | 7.37 | 10.86 | 95.1 | 22.7 | 14.784
2006-07-30 | 22:15 | 9.4 | 7.37 | 10.89 | 95.1 | 22.7 | 14.771
2006-07-30 | 22:30 | 9.27 | 7.36 | 10.94 | 95.3 | 22.7 | 14.737
2006-07-30 | 22:45 | 9.13 | 7.36 | 10.97 | 95.2 | 22.7 | 14.766
2006-07-30 | 23:00 | 8.98 | 7.36 | 11.02 | 95.3 | 22.8 | 14.796
2006-07-30 | 23:15 | 8.83 | 7.35 | 11.06 | 95.3 | 22.7 | 14.765
2006-07-30 | 23:30 | 8.68 | 7.35 | 11.12 | 95.4 | 22.7 | 14.765
2006-07-30 | 23:45 | 8.52 | 7.34 | 11.16 | 95.5 | 22.7 | 14.736
2006-07-31 | 0:00 | 8.36 | 7.35 | 11.21 | 95.5 | 22.7 | 14.755
2006-07-31 | 0:15 | 8.21 | 7.34 | 11.25 | 95.5 | 22.7 | 14.734
2006-07-31 | 0:30 | 8.06 | 7.34 | 11.3 | 95.6 | 22.8 | 14.788
2006-07-31 | 0:45 | 7.9 | 7.34 | 11.34 | 95.6 | 22.8 | 14.801
2006-07-31 | 1:01 | 7.75 | 7.33 | 11.39 | 95.6 | 22.8 | 14.807
2006-07-31 | 1:16 | 7.59 | 7.33 | 11.45 | 95.7 | 22.8 | 14.818
2006-07-31 | 1:31 | 7.44 | 7.33 | 11.49 | 95.7 | 22.8 | 14.815
2006-07-31 | 1:46 | 7.29 | 7.33 | 11.55 | 95.8 | 22.8 | 14.819
2006-07-31 | 2:01 | 7.14 | 7.33 | 11.6 | 95.9 | 22.8 | 14.834
2006-07-31 | 2:16 | 6.99 | 7.33 | 11.64 | 95.8 | 22.8 | 14.817
2006-07-31 | 2:31 | 6.85 | 7.32 | 11.69 | 95.9 | 22.8 | 14.847
2006-07-31 | 2:46 | 6.71 | 7.32 | 11.73 | 95.9 | 22.9 | 14.868
2006-07-31 | 3:01 | 6.58 | 7.32 | 11.78 | 96 | 22.9 | 14.905
2006-07-31 | 3:16 | 6.44 | 7.32 | 11.82 | 96 | 22.9 | 14.905
2006-07-31 | 3:31 | 6.3 | 7.31 | 11.89 | 96.3 | 23 | 14.921
2006-07-31 | 3:46 | 6.17 | 7.32 | 11.91 | 96.1 | 23 | 14.923
2006-07-31 | 4:01 | 6.05 | 7.32 | 11.95 | 96.1 | 23 | 14.931
2006-07-31 | 4:16 | 5.92 | 7.31 | 11.99 | 96.1 | 23 | 14.944
2006-07-31 | 4:31 | 5.8 | 7.31 | 12.04 | 96.3 | 23.1 | 14.99
2006-07-31 | 4:46 | 5.68 | 7.31 | 12.08 | 96.3 | 23.1 | 14.984
2006-07-31 | 5:01 | 5.56 | 7.31 | 12.12 | 96.3 | 23.1 | 15.017
2006-07-31 | 5:16 | 5.45 | 7.32 | 12.18 | 96.5 | 23.1 | 15.013
2006-07-31 | 5:31 | 5.35 | 7.31 | 12.21 | 96.5 | 23.1 | 15.019
2006-07-31 | 5:46 | 5.25 | 7.31 | 12.23 | 96.4 | 23.1 | 15.036
2006-07-31 | 6:01 | 5.16 | 7.31 | 12.27 | 96.5 | 23.2 | 15.056
2006-07-31 | 6:16 | 5.08 | 7.31 | 12.29 | 96.4 | 23.2 | 15.066
2006-07-31 | 6:31 | 5.01 | 7.31 | 12.32 | 96.5 | 23.2 | 15.072
2006-07-31 | 6:46 | 4.95 | 7.31 | 12.36 | 96.7 | 23.2 | 15.092
2006-07-31 | 7:01 | 4.9 | 7.33 | 12.38 | 96.7 | 23.2 | 15.098
2006-07-31 | 7:16 | 4.87 | 7.33 | 12.38 | 96.6 | 23.2 | 15.106
2006-07-31 | 7:31 | 4.85 | 7.34 | 12.41 | 96.8 | 23.2 | 15.109
2006-07-31 | 7:46 | 4.84 | 7.34 | 12.41 | 96.8 | 23.2 | 15.111
2006-07-31 | 8:01 | 4.84 | 7.35 | 12.42 | 96.8 | 23.3 | 15.141
2006-07-31 | 8:16 | 4.85 | 7.35 | 12.39 | 96.7 | 23.3 | 15.118
2006-07-31 | 8:31 | 4.87 | 7.36 | 12.4 | 96.8 | 23.2 | 15.083
2006-07-31 | 8:46 | 4.89 | 7.35 | 12.4 | 96.8 | 23.2 | 15.074
2006-07-31 | 9:01 | 4.93 | 7.36 | 12.4 | 96.9 | 23.2 | 15.095
2006-07-31 | 9:16 | 4.99 | 7.36 | 12.37 | 96.9 | 23.3 | 15.118
2006-07-31 | 9:31 | 5.05 | 7.36 | 12.34 | 96.8 | 23.3 | 15.126
2006-07-31 | 9:46 | 5.13 | 7.37 | 12.33 | 96.9 | 23.3 | 15.132
2006-07-31 | 10:01 | 5.22 | 7.37 | 12.3 | 96.9 | 23.3 | 15.151
2006-07-31 | 10:16 | 5.32 | 7.37 | 12.28 | 96.9 | 23.3 | 15.141
2006-07-31 | 10:31 | 5.44 | 7.37 | 12.23 | 96.8 | 23.3 | 15.167
2006-07-31 | 10:46 | 5.59 | 7.38 | 12.19 | 96.9 | 23.3 | 15.166
2006-07-31 | 11:01 | 5.78 | 7.38 | 12.1 | 96.7 | 23.3 | 15.159
2006-07-31 | 11:16 | 5.95 | 7.39 | 12.07 | 96.8 | 23.4 | 15.188
2006-07-31 | 11:31 | 6.14 | 7.4 | 12 | 96.7 | 23.4 | 15.205
2006-07-31 | 11:46 | 6.34 | 7.39 | 11.94 | 96.7 | 23.4 | 15.201
2006-07-31 | 12:01 | 6.54 | 7.4 | 11.88 | 96.7 | 23.4 | 15.204
2006-07-31 | 12:16 | 6.68 | 7.4 | 11.81 | 96.5 | 23.4 | 15.192
2006-07-31 | 12:31 | 6.83 | 7.39 | 11.76 | 96.5 | 23.4 | 15.196
2006-07-31 | 12:46 | 6.99 | 7.4 | 11.71 | 96.5 | 23.5 | 15.251
2006-07-31 | 13:00 | 7.18 | 7.4 | 11.65 | 96.4 | 23.4 | 15.238
2006-07-31 | 13:15 | 7.4 | 7.4 | 11.59 | 96.4 | 23.5 | 15.258
2006-07-31 | 13:30 | 7.63 | 7.41 | 11.51 | 96.3 | 23.5 | 15.257
2006-07-31 | 13:45 | 7.87 | 7.42 | 11.44 | 96.3 | 23.5 | 15.243
2006-07-31 | 14:00 | 8.12 | 7.41 | 11.36 | 96.2 | 23.5 | 15.257
2006-07-31 | 14:15 | 8.31 | 7.41 | 11.27 | 95.9 | 23.4 | 15.24
2006-07-31 | 14:30 | 8.47 | 7.42 | 11.23 | 95.9 | 23.5 | 15.249
2006-07-31 | 14:45 | 8.62 | 7.42 | 11.2 | 96 | 23.5 | 15.258
2006-07-31 | 15:00 | 8.78 | 7.43 | 11.16 | 96.1 | 23.5 | 15.261
2006-07-31 | 15:15 | 8.95 | 7.42 | 11.1 | 95.9 | 23.5 | 15.265
2006-07-31 | 15:30 | 9.09 | 7.42 | 11.03 | 95.6 | 23.5 | 15.27
2006-07-31 | 15:45 | 9.26 | 7.42 | 10.98 | 95.6 | 23.5 | 15.274
2006-07-31 | 16:00 | 9.41 | 7.42 | 10.93 | 95.5 | 23.5 | 15.278
2006-07-31 | 16:15 | 9.58 | 7.42 | 10.86 | 95.3 | 23.5 | 15.275
2006-07-31 | 16:30 | 9.72 | 7.43 | 10.85 | 95.5 | 23.5 | 15.278
2006-07-31 | 16:45 | 9.87 | 7.43 | 10.79 | 95.3 | 23.5 | 15.289
2006-07-31 | 17:00 | 10.02 | 7.42 | 10.75 | 95.3 | 23.5 | 15.305
2006-07-31 | 17:15 | 10.13 | 7.43 | 10.72 | 95.3 | 23.6 | 15.324
2006-07-31 | 17:30 | 10.2 | 7.41 | 10.69 | 95.2 | 23.6 | 15.312
2006-07-31 | 17:45 | 10.22 | 7.42 | 10.69 | 95.2 | 23.5 | 15.3
2006-07-31 | 18:00 | 10.3 | 7.42 | 10.66 | 95.1 | 23.5 | 15.298
2006-07-31 | 18:15 | 10.42 | 7.43 | 10.62 | 95 | 23.8 | 15.465
2006-07-31 | 18:30 | 10.55 | 7.42 | 10.6 | 95.1 | 23.8 | 15.449
2006-07-31 | 18:45 | 10.64 | 7.42 | 10.56 | 95 | 23.8 | 15.465
2006-07-31 | 19:00 | 10.7 | 7.41 | 10.54 | 94.9 | 23.8 | 15.469
2006-07-31 | 19:15 | 10.67 | 7.42 | 10.51 | 94.6 | 23.8 | 15.48
2006-07-31 | 19:30 | 10.58 | 7.41 | 10.55 | 94.8 | 23.8 | 15.471
2006-07-31 | 19:45 | 10.46 | 7.41 | 10.6 | 94.9 | 23.8 | 15.488
2006-07-31 | 20:00 | 10.34 | 7.4 | 10.63 | 95 | 23.9 | 15.508
2006-07-31 | 20:15 | 10.24 | 7.4 | 10.66 | 94.9 | 23.8 | 15.498
2006-07-31 | 20:30 | 10.14 | 7.4 | 10.69 | 95 | 23.9 | 15.512
2006-07-31 | 20:45 | 9.99 | 7.39 | 10.73 | 95.1 | 23.9 | 15.526
2006-07-31 | 21:00 | 9.84 | 7.38 | 10.78 | 95.1 | 23.9 | 15.534
2006-07-31 | 21:15 | 9.69 | 7.38 | 10.83 | 95.2 | 23.9 | 15.546
2006-07-31 | 21:30 | 9.54 | 7.37 | 10.86 | 95.2 | 23.9 | 15.551
2006-07-31 | 21:45 | 9.39 | 7.37 | 10.91 | 95.3 | 23.9 | 15.564
2006-07-31 | 22:00 | 9.25 | 7.37 | 10.95 | 95.3 | 23.9 | 15.563
2006-07-31 | 22:15 | 9.12 | 7.36 | 11.01 | 95.5 | 24 | 15.571
2006-07-31 | 22:30 | 9 | 7.36 | 11.05 | 95.6 | 24 | 15.581
2006-07-31 | 22:45 | 8.91 | 7.35 | 11.06 | 95.4 | 24 | 15.591
2006-07-31 | 23:00 | 8.84 | 7.36 | 11.08 | 95.5 | 24 | 15.586
2006-07-31 | 23:15 | 8.78 | 7.36 | 11.1 | 95.5 | 24 | 15.591
2006-07-31 | 23:30 | 8.73 | 7.36 | 11.12 | 95.5 | 24 | 15.603
2006-07-31 | 23:45 | 8.67 | 7.36 | 11.13 | 95.5 | 24 | 15.627
2006-08-01 | 0:00 | 8.59 | 7.36 | 11.15 | 95.5 | 24 | 15.631
2006-08-01 | 0:15 | 8.5 | 7.35 | 11.18 | 95.6 | 24.1 | 15.645
2006-08-01 | 0:30 | 8.4 | 7.34 | 11.23 | 95.7 | 24.1 | 15.663
2006-08-01 | 0:45 | 8.28 | 7.36 | 11.26 | 95.7 | 24.1 | 15.683
2006-08-01 | 1:01 | 8.15 | 7.34 | 11.27 | 95.5 | 24.1 | 15.69
2006-08-01 | 1:16 | 8.01 | 7.35 | 11.34 | 95.8 | 24.2 | 15.718
2006-08-01 | 1:31 | 7.87 | 7.35 | 11.4 | 95.9 | 24.2 | 15.727
2006-08-01 | 1:46 | 7.73 | 7.34 | 11.43 | 95.9 | 24.2 | 15.737
2006-08-01 | 2:01 | 7.6 | 7.34 | 11.47 | 95.9 | 24.2 | 15.752
2006-08-01 | 2:16 | 7.47 | 7.34 | 11.52 | 96 | 24.2 | 15.759
2006-08-01 | 2:31 | 7.34 | 7.33 | 11.56 | 96 | 24.3 | 15.773
2006-08-01 | 2:46 | 7.21 | 7.33 | 11.6 | 96 | 24.3 | 15.775
2006-08-01 | 3:01 | 7.09 | 7.34 | 11.64 | 96.1 | 24.3 | 15.804
2006-08-01 | 3:16 | 6.98 | 7.33 | 11.68 | 96.2 | 24.3 | 15.795
2006-08-01 | 3:31 | 6.86 | 7.34 | 11.73 | 96.3 | 24.3 | 15.804
2006-08-01 | 3:46 | 6.75 | 7.34 | 11.76 | 96.3 | 24.3 | 15.808
2006-08-01 | 4:01 | 6.63 | 7.33 | 11.8 | 96.3 | 24.3 | 15.815
2006-08-01 | 4:16 | 6.51 | 7.33 | 11.84 | 96.4 | 24.4 | 15.843
2006-08-01 | 4:31 | 6.4 | 7.33 | 11.88 | 96.4 | 24.4 | 15.84
2006-08-01 | 4:46 | 6.31 | 7.34 | 11.91 | 96.4 | 24.4 | 15.85
2006-08-01 | 5:01 | 6.2 | 7.34 | 11.95 | 96.5 | 24.4 | 15.86
2006-08-01 | 5:16 | 6.1 | 7.33 | 11.98 | 96.5 | 24.4 | 15.855
2006-08-01 | 5:31 | 6 | 7.33 | 12.03 | 96.6 | 24.4 | 15.864
2006-08-01 | 5:46 | 5.9 | 7.33 | 12.06 | 96.6 | 24.4 | 15.862
2006-08-01 | 6:01 | 5.81 | 7.33 | 12.1 | 96.8 | 24.4 | 15.891
2006-08-01 | 6:16 | 5.72 | 7.33 | 12.13 | 96.8 | 24.4 | 15.881
2006-08-01 | 6:31 | 5.62 | 7.33 | 12.17 | 96.8 | 24.4 | 15.889
2006-08-01 | 6:46 | 5.52 | 7.33 | 12.21 | 96.9 | 24.5 | 15.893
2006-08-01 | 7:01 | 5.43 | 7.34 | 12.26 | 97 | 24.5 | 15.911
2006-08-01 | 7:16 | 5.36 | 7.33 | 12.27 | 97 | 24.5 | 15.914
2006-08-01 | 7:31 | 5.3 | 7.35 | 12.31 | 97.2 | 24.5 | 15.905
2006-08-01 | 7:46 | 5.26 | 7.35 | 12.32 | 97.1 | 24.5 | 15.908
2006-08-01 | 8:01 | 5.22 | 7.36 | 12.34 | 97.2 | 24.5 | 15.91
2006-08-01 | 8:16 | 5.2 | 7.37 | 12.35 | 97.2 | 24.5 | 15.926
2006-08-01 | 8:31 | 5.2 | 7.36 | 12.36 | 97.3 | 24.5 | 15.938
2006-08-01 | 8:46 | 5.21 | 7.37 | 12.37 | 97.4 | 24.5 | 15.923
2006-08-01 | 9:01 | 5.24 | 7.37 | 12.37 | 97.4 | 24.5 | 15.928
2006-08-01 | 9:16 | 5.28 | 7.38 | 12.34 | 97.3 | 24.5 | 15.944
2006-08-01 | 9:31 | 5.33 | 7.38 | 12.33 | 97.4 | 24.6 | 16.01
2006-08-01 | 9:46 | 5.39 | 7.38 | 12.33 | 97.5 | 24.5 | 15.932
2006-08-01 | 10:01 | 5.47 | 7.39 | 12.3 | 97.5 | 24.5 | 15.944
2006-08-01 | 10:16 | 5.55 | 7.39 | 12.27 | 97.4 | 24.5 | 15.909
2006-08-01 | 10:31 | 5.66 | 7.39 | 12.22 | 97.3 | 24.5 | 15.911
2006-08-01 | 10:46 | 5.8 | 7.39 | 12.19 | 97.4 | 24.5 | 15.916
2006-08-01 | 11:01 | 5.95 | 7.39 | 12.14 | 97.4 | 24.5 | 15.926
2006-08-01 | 11:16 | 6.11 | 7.4 | 12.08 | 97.3 | 24.5 | 15.92
2006-08-01 | 11:31 | 6.27 | 7.4 | 12.02 | 97.2 | 24.5 | 15.917
2006-08-01 | 11:46 | 6.45 | 7.4 | 11.96 | 97.1 | 24.5 | 15.934
2006-08-01 | 12:01 | 6.64 | 7.4 | 11.91 | 97.3 | 24.4 | 15.883
2006-08-01 | 12:16 | 6.82 | 7.4 | 11.85 | 97.2 | 24.5 | 15.921
2006-08-01 | 12:31 | 7.04 | 7.41 | 11.78 | 97.1 | 24.5 | 15.92
2006-08-01 | 12:46 | 7.3 | 7.42 | 11.7 | 97.1 | 24.5 | 15.928
2006-08-01 | 13:00 | 7.57 | 7.41 | 11.62 | 97 | 24.5 | 15.924
2006-08-01 | 13:15 | 7.82 | 7.42 | 11.54 | 97 | 24.5 | 15.9
2006-08-01 | 13:30 | 8.09 | 7.43 | 11.43 | 96.7 | 24.5 | 15.945
2006-08-01 | 13:45 | 8.33 | 7.42 | 11.38 | 96.8 | 24.5 | 15.952
2006-08-01 | 14:00 | 8.57 | 7.43 | 11.29 | 96.7 | 24.5 | 15.952
2006-08-01 | 14:15 | 8.82 | 7.42 | 11.22 | 96.6 | 24.5 | 15.932
2006-08-01 | 14:30 | 9.05 | 7.43 | 11.14 | 96.5 | 24.5 | 15.933
2006-08-01 | 14:45 | 9.3 | 7.43 | 11.07 | 96.5 | 24.5 | 15.936
2006-08-01 | 15:00 | 9.57 | 7.44 | 10.99 | 96.4 | 24.5 | 15.933
2006-08-01 | 15:15 | 9.93 | 7.44 | 10.87 | 96.2 | 24.5 | 15.955
2006-08-01 | 15:30 | 10.27 | 7.45 | 10.78 | 96.1 | 24.5 | 15.928
2006-08-01 | 15:45 | 10.59 | 7.45 | 10.68 | 96 | 24.5 | 15.924
2006-08-01 | 16:00 | 10.83 | 7.45 | 10.62 | 95.9 | 24.5 | 15.934
2006-08-01 | 16:15 | 11 | 7.45 | 10.56 | 95.8 | 24.5 | 15.935
2006-08-01 | 16:30 | 11.28 | 7.45 | 10.48 | 95.7 | 24.5 | 15.936
2006-08-01 | 16:45 | 11.53 | 7.46 | 10.42 | 95.6 | 24.5 | 15.92
2006-08-01 | 17:00 | 11.69 | 7.46 | 10.37 | 95.5 | 24.5 | 15.935
2006-08-01 | 17:15 | 11.83 | 7.45 | 10.33 | 95.5 | 24.5 | 15.928
2006-08-01 | 17:30 | 11.91 | 7.45 | 10.32 | 95.6 | 24.5 | 15.936
2006-08-01 | 17:45 | 11.94 | 7.46 | 10.29 | 95.4 | 24.5 | 15.926
2006-08-01 | 18:00 | 11.94 | 7.45 | 10.32 | 95.6 | 24.5 | 15.93
2006-08-01 | 18:15 | 11.9 | 7.45 | 10.3 | 95.4 | 24.5 | 15.905
2006-08-01 | 18:30 | 11.87 | 7.45 | 10.3 | 95.3 | 24.5 | 15.903
2006-08-01 | 18:45 | 11.8 | 7.44 | 10.32 | 95.4 | 24.5 | 15.909
2006-08-01 | 19:00 | 11.65 | 7.44 | 10.37 | 95.4 | 24.5 | 15.909
2006-08-01 | 19:15 | 11.43 | 7.43 | 10.42 | 95.5 | 24.5 | 15.912
2006-08-01 | 19:30 | 11.23 | 7.43 | 10.47 | 95.5 | 24.5 | 15.897
2006-08-01 | 19:45 | 11.07 | 7.43 | 10.51 | 95.4 | 24.5 | 15.893
2006-08-01 | 20:00 | 10.95 | 7.42 | 10.56 | 95.6 | 24.4 | 15.878
2006-08-01 | 20:15 | 10.85 | 7.42 | 10.57 | 95.5 | 24.4 | 15.877
2006-08-01 | 20:30 | 10.75 | 7.41 | 10.61 | 95.6 | 24.4 | 15.874
2006-08-01 | 20:45 | 10.67 | 7.4 | 10.62 | 95.6 | 24.4 | 15.863
2006-08-01 | 21:00 | 10.58 | 7.4 | 10.67 | 95.9 | 24.4 | 15.873
2006-08-01 | 21:15 | 10.49 | 7.39 | 10.67 | 95.6 | 24.4 | 15.869
2006-08-01 | 21:30 | 10.4 | 7.39 | 10.71 | 95.7 | 24.4 | 15.864
2006-08-01 | 21:45 | 10.3 | 7.38 | 10.69 | 95.4 | 24.4 | 15.872
2006-08-01 | 22:00 | 10.19 | 7.38 | 10.77 | 95.9 | 24.4 | 15.858
2006-08-01 | 22:15 | 10.08 | 7.38 | 10.8 | 95.9 | 24.4 | 15.855
2006-08-01 | 22:30 | 9.98 | 7.38 | 10.84 | 95.9 | 24.4 | 15.869
2006-08-01 | 22:45 | 9.87 | 7.38 | 10.86 | 95.9 | 24.4 | 15.852
2006-08-01 | 23:00 | 9.76 | 7.37 | 10.87 | 95.8 | 24.4 | 15.853
2006-08-01 | 23:15 | 9.63 | 7.37 | 10.94 | 96 | 24.4 | 15.865
2006-08-01 | 23:30 | 9.49 | 7.37 | 10.97 | 96.1 | 24.4 | 15.87
2006-08-01 | 23:45 | 9.35 | 7.36 | 11.02 | 96.1 | 24.4 | 15.884
2006-08-02 | 0:00 | 9.19 | 7.36 | 11.08 | 96.3 | 24.4 | 15.872
2006-08-02 | 0:15 | 9.04 | 7.36 | 11.11 | 96.2 | 24.4 | 15.874
2006-08-02 | 0:30 | 8.88 | 7.36 | 11.16 | 96.2 | 24.4 | 15.882
2006-08-02 | 0:45 | 8.72 | 7.35 | 11.21 | 96.3 | 24.4 | 15.888
2006-08-02 | 1:01 | 8.55 | 7.35 | 11.26 | 96.4 | 24.5 | 15.909
2006-08-02 | 1:16 | 8.38 | 7.35 | 11.32 | 96.4 | 24.4 | 15.884
2006-08-02 | 1:31 | 8.2 | 7.35 | 11.39 | 96.6 | 24.4 | 15.878
2006-08-02 | 1:46 | 8.02 | 7.35 | 11.43 | 96.5 | 24.5 | 15.901
2006-08-02 | 2:01 | 7.85 | 7.35 | 11.49 | 96.7 | 24.4 | 15.886
2006-08-02 | 2:16 | 7.67 | 7.35 | 11.55 | 96.8 | 24.4 | 15.884
2006-08-02 | 2:31 | 7.5 | 7.34 | 11.6 | 96.8 | 24.5 | 15.908
2006-08-02 | 2:46 | 7.33 | 7.34 | 11.62 | 96.5 | 24.5 | 15.912
2006-08-02 | 3:01 | 7.18 | 7.34 | 11.7 | 96.8 | 24.5 | 15.92
2006-08-02 | 3:16 | 7.03 | 7.34 | 11.75 | 96.9 | 24.5 | 15.92
2006-08-02 | 3:31 | 6.88 | 7.34 | 11.8 | 96.9 | 24.5 | 15.923
2006-08-02 | 3:46 | 6.73 | 7.33 | 11.86 | 97.1 | 24.5 | 15.91
2006-08-02 | 4:01 | 6.58 | 7.33 | 11.9 | 97 | 24.5 | 15.92
2006-08-02 | 4:16 | 6.44 | 7.33 | 11.95 | 97.1 | 24.5 | 15.923
2006-08-02 | 4:31 | 6.3 | 7.33 | 12.01 | 97.2 | 24.5 | 15.929
2006-08-02 | 4:46 | 6.16 | 7.33 | 12.06 | 97.3 | 24.5 | 15.939
2006-08-02 | 5:01 | 6.04 | 7.33 | 12.11 | 97.3 | 24.5 | 15.947
2006-08-02 | 5:16 | 5.91 | 7.33 | 12.15 | 97.4 | 24.6 | 15.975
2006-08-02 | 5:31 | 5.8 | 7.32 | 12.19 | 97.4 | 24.6 | 15.974
2006-08-02 | 5:46 | 5.69 | 7.33 | 12.23 | 97.5 | 24.6 | 15.963
2006-08-02 | 6:01 | 5.59 | 7.33 | 12.26 | 97.5 | 24.6 | 15.985
2006-08-02 | 6:16 | 5.5 | 7.32 | 12.29 | 97.5 | 24.6 | 15.99
2006-08-02 | 6:31 | 5.42 | 7.32 | 12.33 | 97.6 | 24.6 | 15.983
2006-08-02 | 6:46 | 5.34 | 7.33 | 12.36 | 97.6 | 24.6 | 15.993
2006-08-02 | 7:01 | 5.28 | 7.33 | 12.4 | 97.8 | 24.6 | 15.992
2006-08-02 | 7:16 | 5.23 | 7.34 | 12.37 | 97.5 | 24.6 | 15.985
2006-08-02 | 7:31 | 5.19 | 7.34 | 12.43 | 97.8 | 24.6 | 15.996
2006-08-02 | 7:46 | 5.16 | 7.34 | 12.44 | 97.8 | 24.6 | 15.985
2006-08-02 | 8:01 | 5.13 | 7.35 | 12.45 | 97.8 | 24.6 | 15.982
2006-08-02 | 8:16 | 5.12 | 7.36 | 12.48 | 98 | 24.6 | 15.998
2006-08-02 | 8:31 | 5.12 | 7.36 | 12.47 | 98 | 24.6 | 15.98
2006-08-02 | 8:46 | 5.13 | 7.37 | 12.48 | 98 | 24.6 | 15.991
2006-08-02 | 9:01 | 5.17 | 7.37 | 12.47 | 98 | 24.6 | 15.976
2006-08-02 | 9:16 | 5.23 | 7.37 | 12.44 | 98 | 24.6 | 15.975
2006-08-02 | 9:31 | 5.3 | 7.37 | 12.43 | 98.1 | 24.6 | 15.98
2006-08-02 | 9:46 | 5.39 | 7.38 | 12.39 | 98 | 24.6 | 15.971
2006-08-02 | 10:01 | 5.49 | 7.38 | 12.37 | 98.1 | 24.6 | 15.99
2006-08-02 | 10:16 | 5.61 | 7.39 | 12.33 | 98.1 | 24.6 | 15.985
2006-08-02 | 10:31 | 5.74 | 7.39 | 12.29 | 98.1 | 24.6 | 15.969
2006-08-02 | 10:46 | 5.87 | 7.39 | 12.24 | 98 | 24.6 | 15.978
2006-08-02 | 11:01 | 6.03 | 7.4 | 12.19 | 98 | 24.6 | 15.965
2006-08-02 | 11:16 | 6.21 | 7.39 | 12.12 | 97.8 | 24.5 | 15.942
2006-08-02 | 11:31 | 6.41 | 7.39 | 12.06 | 97.9 | 24.6 | 15.977
2006-08-02 | 11:46 | 6.63 | 7.4 | 11.98 | 97.7 | 24.5 | 15.94
2006-08-02 | 12:01 | 6.85 | 7.4 | 11.91 | 97.7 | 24.6 | 15.969
2006-08-02 | 12:16 | 7.08 | 7.41 | 11.83 | 97.7 | 24.6 | 15.963
2006-08-02 | 12:31 | 7.31 | 7.41 | 11.76 | 97.6 | 24.6 | 15.961
2006-08-02 | 12:46 | 7.57 | 7.41 | 11.67 | 97.5 | 24.5 | 15.899
2006-08-02 | 13:00 | 7.84 | 7.42 | 11.59 | 97.5 | 24.6 | 15.965
2006-08-02 | 13:15 | 8.12 | 7.42 | 11.5 | 97.4 | 24.6 | 15.959
2006-08-02 | 13:30 | 8.39 | 7.42 | 11.41 | 97.3 | 24.6 | 15.959
2006-08-02 | 13:45 | 8.68 | 7.42 | 11.33 | 97.2 | 24.6 | 15.969
2006-08-02 | 14:00 | 8.95 | 7.43 | 11.24 | 97.2 | 24.5 | 15.956
2006-08-02 | 14:15 | 9.22 | 7.43 | 11.16 | 97 | 24.5 | 15.956
2006-08-02 | 14:30 | 9.5 | 7.43 | 11.08 | 97 | 24.5 | 15.948
2006-08-02 | 14:45 | 9.76 | 7.44 | 10.99 | 96.8 | 24.5 | 15.937
2006-08-02 | 15:00 | 10.01 | 7.44 | 10.91 | 96.7 | 24.5 | 15.947
2006-08-02 | 15:15 | 10.26 | 7.44 | 10.84 | 96.6 | 24.5 | 15.934
2006-08-02 | 15:30 | 10.49 | 7.45 | 10.78 | 96.6 | 24.5 | 15.923
2006-08-02 | 15:45 | 10.7 | 7.44 | 10.71 | 96.5 | 24.5 | 15.944
2006-08-02 | 16:00 | 10.89 | 7.44 | 10.65 | 96.3 | 24.6 | 15.97
2006-08-02 | 16:15 | 11.06 | 7.45 | 10.6 | 96.2 | 24.6 | 15.965
2006-08-02 | 16:30 | 11.22 | 7.45 | 10.55 | 96.2 | 24.5 | 15.928
2006-08-02 | 16:45 | 11.36 | 7.45 | 10.5 | 96.1 | 24.5 | 15.942
2006-08-02 | 17:00 | 11.48 | 7.45 | 10.48 | 96.1 | 24.5 | 15.942
2006-08-02 | 17:15 | 11.57 | 7.45 | 10.45 | 96.1 | 24.5 | 15.946
2006-08-02 | 17:30 | 11.64 | 7.45 | 10.43 | 96 | 24.6 | 15.966
2006-08-02 | 17:45 | 11.7 | 7.44 | 10.41 | 96 | 24.6 | 15.974
2006-08-02 | 18:00 | 11.74 | 7.45 | 10.4 | 95.9 | 24.6 | 15.972
2006-08-02 | 18:15 | 11.75 | 7.44 | 10.39 | 95.9 | 24.6 | 15.97
2006-08-02 | 18:30 | 11.74 | 7.43 | 10.38 | 95.8 | 24.6 | 15.965
2006-08-02 | 18:45 | 11.72 | 7.43 | 10.38 | 95.7 | 24.6 | 15.961
2006-08-02 | 19:00 | 11.68 | 7.42 | 10.38 | 95.6 | 24.5 | 15.955
2006-08-02 | 19:15 | 11.63 | 7.42 | 10.41 | 95.8 | 24.6 | 15.974
2006-08-02 | 19:30 | 11.56 | 7.41 | 10.42 | 95.7 | 24.6 | 15.98
2006-08-02 | 19:45 | 11.48 | 7.4 | 10.44 | 95.7 | 24.6 | 15.975
2006-08-02 | 20:00 | 11.38 | 7.39 | 10.44 | 95.5 | 24.5 | 15.955
2006-08-02 | 20:15 | 11.28 | 7.39 | 10.49 | 95.7 | 24.6 | 15.98
2006-08-02 | 20:30 | 11.18 | 7.38 | 10.51 | 95.7 | 24.6 | 15.978
2006-08-02 | 20:45 | 11.08 | 7.37 | 10.54 | 95.7 | 24.6 | 15.991
2006-08-02 | 21:00 | 10.98 | 7.37 | 10.56 | 95.7 | 24.6 | 15.976
2006-08-02 | 21:15 | 10.87 | 7.37 | 10.59 | 95.7 | 24.5 | 15.954
2006-08-02 | 21:30 | 10.77 | 7.37 | 10.63 | 95.8 | 24.6 | 15.994
2006-08-02 | 21:45 | 10.66 | 7.37 | 10.66 | 95.9 | 24.6 | 15.981
2006-08-02 | 22:00 | 10.56 | 7.37 | 10.68 | 95.8 | 24.6 | 15.98
2006-08-02 | 22:15 | 10.45 | 7.37 | 10.72 | 96 | 24.6 | 15.994
2006-08-02 | 22:30 | 10.34 | 7.36 | 10.75 | 96 | 24.6 | 15.969
2006-08-02 | 22:45 | 10.22 | 7.36 | 10.78 | 96 | 24.6 | 15.975
2006-08-02 | 23:00 | 10.11 | 7.36 | 10.81 | 96 | 24.6 | 15.975
2006-08-02 | 23:15 | 10 | 7.36 | 10.84 | 96 | 24.5 | 15.953
2006-08-02 | 23:30 | 9.89 | 7.35 | 10.88 | 96.1 | 24.6 | 15.96
2006-08-02 | 23:45 | 9.77 | 7.35 | 10.91 | 96.2 | 24.6 | 15.963
2006-08-03 | 0:00 | 9.65 | 7.34 | 10.94 | 96.1 | 24.5 | 15.954
2006-08-03 | 0:15 | 9.52 | 7.35 | 10.98 | 96.2 | 24.5 | 15.951
2006-08-03 | 0:30 | 9.4 | 7.35 | 11.02 | 96.2 | 24.5 | 15.943
2006-08-03 | 0:45 | 9.27 | 7.35 | 11.06 | 96.3 | 24.5 | 15.939
2006-08-03 | 1:01 | 9.14 | 7.35 | 11.09 | 96.3 | 24.5 | 15.937
2006-08-03 | 1:16 | 9 | 7.34 | 11.13 | 96.3 | 24.5 | 15.946
2006-08-03 | 1:31 | 8.87 | 7.35 | 11.18 | 96.4 | 24.5 | 15.948
2006-08-03 | 1:46 | 8.73 | 7.34 | 11.21 | 96.4 | 24.5 | 15.934
2006-08-03 | 2:01 | 8.6 | 7.35 | 11.25 | 96.4 | 24.5 | 15.934
2006-08-03 | 2:16 | 8.46 | 7.34 | 11.3 | 96.5 | 24.5 | 15.938
2006-08-03 | 2:31 | 8.33 | 7.34 | 11.34 | 96.5 | 24.5 | 15.953
2006-08-03 | 2:46 | 8.2 | 7.35 | 11.39 | 96.6 | 24.5 | 15.938
2006-08-03 | 3:01 | 8.07 | 7.34 | 11.42 | 96.6 | 24.6 | 15.988
2006-08-03 | 3:16 | 7.94 | 7.34 | 11.45 | 96.5 | 24.6 | 15.972
2006-08-03 | 3:31 | 7.82 | 7.34 | 11.51 | 96.8 | 24.6 | 15.963
2006-08-03 | 3:46 | 7.7 | 7.35 | 11.54 | 96.7 | 24.6 | 15.996
2006-08-03 | 4:01 | 7.57 | 7.34 | 11.59 | 96.8 | 24.6 | 15.983
2006-08-03 | 4:16 | 7.46 | 7.34 | 11.61 | 96.7 | 24.6 | 15.982
2006-08-03 | 4:31 | 7.34 | 7.34 | 11.65 | 96.8 | 24.6 | 15.982
2006-08-03 | 4:46 | 7.23 | 7.34 | 11.69 | 96.9 | 24.6 | 15.984
2006-08-03 | 5:01 | 7.12 | 7.34 | 11.72 | 96.9 | 24.6 | 15.976
2006-08-03 | 5:16 | 7.02 | 7.34 | 11.76 | 96.9 | 24.6 | 15.966
2006-08-03 | 5:31 | 6.92 | 7.34 | 11.79 | 97 | 24.6 | 15.987
2006-08-03 | 5:46 | 6.82 | 7.34 | 11.83 | 97 | 24.6 | 16.002
2006-08-03 | 6:01 | 6.72 | 7.34 | 11.87 | 97.1 | 24.6 | 16.002
2006-08-03 | 6:16 | 6.64 | 7.34 | 11.9 | 97.2 | 24.6 | 15.979
2006-08-03 | 6:31 | 6.55 | 7.34 | 11.91 | 97 | 24.6 | 15.973
2006-08-03 | 6:46 | 6.48 | 7.34 | 11.97 | 97.3 | 24.6 | 15.98
2006-08-03 | 7:01 | 6.42 | 7.35 | 11.99 | 97.3 | 24.6 | 15.98
2006-08-03 | 7:16 | 6.37 | 7.35 | 12.01 | 97.4 | 24.6 | 15.965
2006-08-03 | 7:31 | 6.33 | 7.36 | 12.03 | 97.5 | 24.6 | 15.989
2006-08-03 | 7:46 | 6.29 | 7.36 | 12.04 | 97.4 | 24.6 | 15.996
2006-08-03 | 8:01 | 6.26 | 7.36 | 12.06 | 97.5 | 24.6 | 15.979
2006-08-03 | 8:16 | 6.24 | 7.37 | 12.06 | 97.5 | 24.6 | 15.963
2006-08-03 | 8:31 | 6.23 | 7.38 | 12.08 | 97.6 | 24.6 | 15.974
2006-08-03 | 8:46 | 6.23 | 7.38 | 12.08 | 97.6 | 24.5 | 15.955
2006-08-03 | 9:01 | 6.25 | 7.38 | 12.07 | 97.6 | 24.6 | 15.987
2006-08-03 | 9:16 | 6.28 | 7.39 | 12.06 | 97.6 | 24.6 | 15.972
2006-08-03 | 9:31 | 6.32 | 7.39 | 12.06 | 97.7 | 24.6 | 15.984
2006-08-03 | 9:46 | 6.39 | 7.39 | 12.03 | 97.6 | 24.6 | 15.965
2006-08-03 | 10:01 | 6.46 | 7.4 | 12.01 | 97.6 | 24.6 | 15.971
2006-08-03 | 10:16 | 6.55 | 7.4 | 11.99 | 97.6 | 24.6 | 15.984
2006-08-03 | 10:31 | 6.65 | 7.4 | 11.96 | 97.6 | 24.6 | 15.977
2006-08-03 | 10:46 | 6.77 | 7.41 | 11.93 | 97.7 | 24.6 | 15.981
2006-08-03 | 11:01 | 6.91 | 7.41 | 11.87 | 97.6 | 24.6 | 16
2006-08-03 | 11:16 | 7.07 | 7.41 | 11.83 | 97.6 | 24.6 | 15.992
2006-08-03 | 11:31 | 7.24 | 7.41 | 11.77 | 97.6 | 24.6 | 16.001
2006-08-03 | 11:46 | 7.42 | 7.41 | 11.71 | 97.5 | 24.6 | 15.991
2006-08-03 | 12:01 | 7.62 | 7.42 | 11.64 | 97.4 | 24.6 | 15.981
2006-08-03 | 12:16 | 7.84 | 7.43 | 11.58 | 97.4 | 24.6 | 15.979
2006-08-03 | 12:31 | 8.07 | 7.42 | 11.51 | 97.3 | 24.6 | 15.99
2006-08-03 | 12:46 | 8.33 | 7.43 | 11.43 | 97.3 | 24.6 | 15.988
2006-08-03 | 13:00 | 8.6 | 7.43 | 11.34 | 97.2 | 24.6 | 15.996
2006-08-03 | 13:15 | 8.89 | 7.44 | 11.25 | 97.1 | 24.6 | 15.97
2006-08-03 | 13:30 | 9.19 | 7.44 | 11.16 | 97 | 24.6 | 16.008
2006-08-03 | 13:45 | 9.49 | 7.44 | 11.06 | 96.8 | 24.6 | 15.995
2006-08-03 | 14:00 | 9.79 | 7.45 | 11 | 96.9 | 24.6 | 15.98
2006-08-03 | 14:15 | 10.09 | 7.44 | 10.89 | 96.7 | 24.6 | 15.986
2006-08-03 | 14:30 | 10.39 | 7.43 | 10.81 | 96.6 | 24.6 | 15.979
2006-08-03 | 14:45 | 10.67 | 7.45 | 10.71 | 96.4 | 24.6 | 15.965
2006-08-03 | 15:00 | 10.94 | 7.44 | 10.63 | 96.3 | 24.6 | 15.981
2006-08-03 | 15:15 | 11.19 | 7.46 | 10.57 | 96.2 | 24.5 | 15.956
2006-08-03 | 15:30 | 11.43 | 7.45 | 10.49 | 96.1 | 24.6 | 15.989
2006-08-03 | 15:45 | 11.66 | 7.44 | 10.43 | 96 | 24.6 | 15.985
2006-08-03 | 16:00 | 11.86 | 7.46 | 10.38 | 96 | 24.6 | 15.983
2006-08-03 | 16:15 | 12.04 | 7.46 | 10.3 | 95.7 | 24.6 | 15.986
2006-08-03 | 16:30 | 12.2 | 7.45 | 10.25 | 95.5 | 24.6 | 15.988
2006-08-03 | 16:45 | 12.32 | 7.46 | 10.24 | 95.7 | 24.6 | 16.013
2006-08-03 | 17:00 | 12.42 | 7.46 | 10.22 | 95.7 | 24.6 | 15.994
2006-08-03 | 17:15 | 12.5 | 7.46 | 10.19 | 95.6 | 24.6 | 15.974
2006-08-03 | 17:30 | 12.56 | 7.46 | 10.16 | 95.5 | 24.6 | 15.993
2006-08-03 | 17:45 | 12.58 | 7.46 | 10.16 | 95.5 | 24.6 | 16.005
2006-08-03 | 18:00 | 12.59 | 7.45 | 10.15 | 95.4 | 24.6 | 16.014
2006-08-03 | 18:15 | 12.58 | 7.45 | 10.15 | 95.4 | 24.6 | 16.008
2006-08-03 | 18:30 | 12.54 | 7.45 | 10.16 | 95.4 | 24.6 | 16.005
2006-08-03 | 18:45 | 12.48 | 7.43 | 10.17 | 95.4 | 24.6 | 16.022
2006-08-03 | 19:00 | 12.39 | 7.43 | 10.18 | 95.3 | 24.6 | 16.001
2006-08-03 | 19:15 | 12.3 | 7.43 | 10.21 | 95.4 | 24.6 | 16.005
2006-08-03 | 19:30 | 12.19 | 7.41 | 10.24 | 95.4 | 24.6 | 16.015
2006-08-03 | 19:45 | 12.04 | 7.41 | 10.27 | 95.4 | 24.6 | 15.997
2006-08-03 | 20:00 | 11.89 | 7.4 | 10.31 | 95.4 | 24.6 | 15.995
2006-08-03 | 20:15 | 11.75 | 7.39 | 10.34 | 95.4 | 24.6 | 15.989
2006-08-03 | 20:30 | 11.62 | 7.39 | 10.38 | 95.5 | 24.6 | 15.985
2006-08-03 | 20:45 | 11.51 | 7.38 | 10.41 | 95.5 | 24.6 | 15.974
2006-08-03 | 21:00 | 11.4 | 7.38 | 10.44 | 95.6 | 24.6 | 16.015
2006-08-03 | 21:15 | 11.3 | 7.38 | 10.48 | 95.7 | 24.5 | 15.952
2006-08-03 | 21:30 | 11.19 | 7.38 | 10.51 | 95.7 | 24.5 | 15.955
2006-08-03 | 21:45 | 11.08 | 7.38 | 10.53 | 95.7 | 24.6 | 15.969
2006-08-03 | 22:00 | 10.96 | 7.37 | 10.57 | 95.8 | 24.5 | 15.95
2006-08-03 | 22:15 | 10.85 | 7.37 | 10.61 | 95.8 | 24.5 | 15.953
2006-08-03 | 22:30 | 10.73 | 7.37 | 10.64 | 95.9 | 24.5 | 15.956
2006-08-03 | 22:45 | 10.62 | 7.37 | 10.67 | 95.9 | 24.5 | 15.954
2006-08-03 | 23:00 | 10.5 | 7.36 | 10.7 | 96 | 24.5 | 15.953
2006-08-03 | 23:15 | 10.38 | 7.36 | 10.74 | 96 | 24.5 | 15.954
2006-08-03 | 23:30 | 10.25 | 7.36 | 10.78 | 96.1 | 24.5 | 15.954
2006-08-03 | 23:45 | 10.11 | 7.35 | 10.82 | 96.1 | 24.5 | 15.95
2006-08-04 | 0:00 | 9.97 | 7.36 | 10.86 | 96.2 | 24.5 | 15.955
2006-08-04 | 0:15 | 9.83 | 7.36 | 10.89 | 96.1 | 24.5 | 15.952
2006-08-04 | 0:30 | 9.69 | 7.35 | 10.95 | 96.3 | 24.5 | 15.949
2006-08-04 | 0:45 | 9.55 | 7.36 | 11 | 96.4 | 24.5 | 15.941
2006-08-04 | 1:01 | 9.41 | 7.36 | 11.03 | 96.3 | 24.5 | 15.942
2006-08-04 | 1:16 | 9.29 | 7.35 | 11.06 | 96.3 | 24.5 | 15.939
2006-08-04 | 1:31 | 9.16 | 7.35 | 11.1 | 96.4 | 24.5 | 15.939
2006-08-04 | 1:46 | 9.04 | 7.35 | 11.14 | 96.4 | 24.5 | 15.939
2006-08-04 | 2:01 | 8.92 | 7.35 | 11.17 | 96.4 | 24.5 | 15.943
2006-08-04 | 2:16 | 8.8 | 7.35 | 11.21 | 96.5 | 24.5 | 15.943
2006-08-04 | 2:31 | 8.68 | 7.34 | 11.24 | 96.5 | 24.5 | 15.95
2006-08-04 | 2:46 | 8.55 | 7.34 | 11.28 | 96.5 | 24.6 | 15.958
2006-08-04 | 3:01 | 8.43 | 7.34 | 11.33 | 96.7 | 24.6 | 15.964
2006-08-04 | 3:16 | 8.31 | 7.33 | 11.37 | 96.7 | 24.6 | 15.963
2006-08-04 | 3:31 | 8.19 | 7.34 | 11.39 | 96.6 | 24.6 | 15.966
2006-08-04 | 3:46 | 8.08 | 7.34 | 11.44 | 96.7 | 24.6 | 15.972
2006-08-04 | 4:01 | 7.97 | 7.34 | 11.46 | 96.7 | 24.6 | 15.973
2006-08-04 | 4:16 | 7.87 | 7.34 | 11.5 | 96.8 | 24.6 | 15.979
2006-08-04 | 4:31 | 7.76 | 7.33 | 11.54 | 96.8 | 24.6 | 15.984
2006-08-04 | 4:46 | 7.66 | 7.34 | 11.56 | 96.8 | 24.6 | 15.995
2006-08-04 | 5:01 | 7.56 | 7.33 | 11.6 | 96.9 | 24.6 | 15.975
2006-08-04 | 5:16 | 7.46 | 7.33 | 11.61 | 96.8 | 24.6 | 15.982
2006-08-04 | 5:31 | 7.36 | 7.33 | 11.67 | 97 | 24.6 | 15.984
2006-08-04 | 5:46 | 7.27 | 7.33 | 11.7 | 97 | 24.6 | 15.983
2006-08-04 | 6:01 | 7.18 | 7.33 | 11.73 | 97.1 | 24.6 | 16.002
2006-08-04 | 6:16 | 7.1 | 7.33 | 11.76 | 97.1 | 24.6 | 15.989
2006-08-04 | 6:31 | 7.02 | 7.34 | 11.79 | 97.1 | 24.6 | 15.981
2006-08-04 | 6:46 | 6.95 | 7.34 | 11.82 | 97.2 | 24.6 | 15.996
2006-08-04 | 7:01 | 6.89 | 7.35 | 11.85 | 97.3 | 24.6 | 15.977
2006-08-04 | 7:16 | 6.85 | 7.36 | 11.87 | 97.4 | 24.6 | 15.976
2006-08-04 | 7:31 | 6.81 | 7.36 | 11.88 | 97.4 | 24.6 | 15.977
2006-08-04 | 7:46 | 6.78 | 7.37 | 11.89 | 97.4 | 24.6 | 15.984
2006-08-04 | 8:01 | 6.76 | 7.37 | 11.91 | 97.5 | 24.6 | 15.969
2006-08-04 | 8:16 | 6.74 | 7.38 | 11.9 | 97.4 | 24.6 | 15.971
2006-08-04 | 8:31 | 6.74 | 7.39 | 11.92 | 97.5 | 24.6 | 15.97
2006-08-04 | 8:46 | 6.73 | 7.39 | 11.93 | 97.6 | 24.6 | 15.983
2006-08-04 | 9:01 | 6.75 | 7.39 | 11.92 | 97.6 | 24.6 | 15.962
2006-08-04 | 9:16 | 6.78 | 7.4 | 11.93 | 97.7 | 24.6 | 15.97
2006-08-04 | 9:31 | 6.83 | 7.4 | 11.9 | 97.6 | 24.6 | 15.984
2006-08-04 | 9:46 | 6.89 | 7.4 | 11.88 | 97.6 | 24.6 | 15.969
2006-08-04 | 10:01 | 6.96 | 7.41 | 11.86 | 97.6 | 24.6 | 15.978
2006-08-04 | 10:16 | 7.04 | 7.41 | 11.83 | 97.5 | 24.6 | 15.971
2006-08-04 | 10:31 | 7.14 | 7.41 | 11.81 | 97.6 | 24.6 | 15.985
2006-08-04 | 10:46 | 7.26 | 7.41 | 11.76 | 97.5 | 24.6 | 15.989
2006-08-04 | 11:01 | 7.38 | 7.42 | 11.73 | 97.5 | 24.6 | 15.974
2006-08-04 | 11:16 | 7.52 | 7.42 | 11.69 | 97.5 | 24.6 | 15.977
2006-08-04 | 11:31 | 7.67 | 7.42 | 11.64 | 97.5 | 24.6 | 15.973
2006-08-04 | 11:46 | 7.85 | 7.42 | 11.58 | 97.4 | 24.6 | 15.979
2006-08-04 | 12:01 | 8.04 | 7.43 | 11.52 | 97.4 | 24.6 | 15.977
2006-08-04 | 12:16 | 8.25 | 7.43 | 11.46 | 97.4 | 24.6 | 15.978
2006-08-04 | 12:31 | 8.48 | 7.43 | 11.39 | 97.3 | 24.6 | 15.982
2006-08-04 | 12:46 | 8.72 | 7.43 | 11.32 | 97.2 | 24.6 | 15.978
2006-08-04 | 13:00 | 8.97 | 7.44 | 11.24 | 97.2 | 24.6 | 15.98
2006-08-04 | 13:15 | 9.24 | 7.44 | 11.15 | 97 | 24.6 | 15.994
2006-08-04 | 13:30 | 9.52 | 7.44 | 11.07 | 97 | 24.6 | 15.995
2006-08-04 | 13:45 | 9.8 | 7.44 | 10.99 | 96.9 | 24.6 | 15.993
2006-08-04 | 14:00 | 10.08 | 7.46 | 10.91 | 96.8 | 24.6 | 15.983
2006-08-04 | 14:15 | 10.34 | 7.45 | 10.83 | 96.7 | 24.6 | 15.995
2006-08-04 | 14:30 | 10.59 | 7.46 | 10.76 | 96.7 | 24.6 | 15.999
2006-08-04 | 14:45 | 10.83 | 7.46 | 10.69 | 96.5 | 24.6 | 16.009
2006-08-04 | 15:00 | 11.06 | 7.47 | 10.62 | 96.4 | 24.6 | 16.02
2006-08-04 | 15:15 | 11.27 | 7.46 | 10.56 | 96.3 | 24.6 | 16.012
2006-08-04 | 15:30 | 11.47 | 7.47 | 10.49 | 96.2 | 24.6 | 16.014
2006-08-04 | 15:45 | 11.66 | 7.47 | 10.43 | 96.1 | 24.6 | 16.022
2006-08-04 | 16:00 | 11.84 | 7.47 | 10.39 | 96 | 24.6 | 16.02
2006-08-04 | 16:15 | 12.01 | 7.47 | 10.34 | 96 | 24.6 | 16
2006-08-04 | 16:30 | 12.18 | 7.46 | 10.3 | 95.9 | 24.6 | 16
2006-08-04 | 16:45 | 12.32 | 7.46 | 10.25 | 95.8 | 24.6 | 16.006
2006-08-04 | 17:00 | 12.44 | 7.47 | 10.22 | 95.7 | 24.6 | 16.001
2006-08-04 | 17:15 | 12.54 | 7.46 | 10.19 | 95.7 | 24.6 | 15.995
2006-08-04 | 17:30 | 12.63 | 7.46 | 10.16 | 95.6 | 24.6 | 16.003
2006-08-04 | 17:45 | 12.69 | 7.46 | 10.15 | 95.7 | 24.6 | 16.018
2006-08-04 | 18:00 | 12.73 | 7.46 | 10.14 | 95.6 | 24.6 | 16.017
2006-08-04 | 18:15 | 12.74 | 7.46 | 10.12 | 95.5 | 24.6 | 15.997
2006-08-04 | 18:30 | 12.71 | 7.46 | 10.14 | 95.6 | 24.6 | 16.004
2006-08-04 | 18:45 | 12.66 | 7.45 | 10.13 | 95.4 | 24.6 | 16.012
2006-08-04 | 19:00 | 12.59 | 7.44 | 10.17 | 95.6 | 24.6 | 16.021
2006-08-04 | 19:15 | 12.49 | 7.45 | 10.19 | 95.6 | 24.6 | 15.993
2006-08-04 | 19:30 | 12.38 | 7.44 | 10.22 | 95.7 | 24.6 | 15.989
2006-08-04 | 19:45 | 12.26 | 7.43 | 10.24 | 95.5 | 24.6 | 15.979
2006-08-04 | 20:00 | 12.15 | 7.43 | 10.28 | 95.7 | 24.6 | 15.989
2006-08-04 | 20:15 | 12.03 | 7.42 | 10.3 | 95.7 | 24.6 | 15.968
2006-08-04 | 20:30 | 11.93 | 7.42 | 10.33 | 95.7 | 24.6 | 15.967
2006-08-04 | 20:45 | 11.82 | 7.41 | 10.36 | 95.8 | 24.6 | 15.961
2006-08-04 | 21:00 | 11.72 | 7.41 | 10.4 | 95.8 | 24.5 | 15.954
2006-08-04 | 21:15 | 11.61 | 7.4 | 10.42 | 95.9 | 24.5 | 15.947
2006-08-04 | 21:30 | 11.49 | 7.4 | 10.46 | 95.9 | 24.5 | 15.941
2006-08-04 | 21:45 | 11.37 | 7.4 | 10.49 | 96 | 24.5 | 15.935
2006-08-04 | 22:00 | 11.24 | 7.4 | 10.55 | 96.2 | 24.5 | 15.93
2006-08-04 | 22:15 | 11.09 | 7.39 | 10.57 | 96.1 | 24.5 | 15.925
2006-08-04 | 22:30 | 10.95 | 7.39 | 10.6 | 96 | 24.5 | 15.919
2006-08-04 | 22:45 | 10.8 | 7.39 | 10.65 | 96.1 | 24.5 | 15.93
2006-08-04 | 23:00 | 10.65 | 7.38 | 10.7 | 96.3 | 24.5 | 15.906
2006-08-04 | 23:15 | 10.49 | 7.37 | 10.74 | 96.2 | 24.5 | 15.912
2006-08-04 | 23:30 | 10.32 | 7.38 | 10.79 | 96.3 | 24.4 | 15.892
2006-08-04 | 23:45 | 10.15 | 7.37 | 10.84 | 96.4 | 24.4 | 15.883
2006-08-05 | 0:00 | 9.96 | 7.37 | 10.89 | 96.4 | 24.4 | 15.876
2006-08-05 | 0:15 | 9.76 | 7.36 | 10.96 | 96.5 | 24.4 | 15.875
2006-08-05 | 0:30 | 9.56 | 7.36 | 11.02 | 96.6 | 24.4 | 15.872
2006-08-05 | 0:45 | 9.37 | 7.36 | 11.09 | 96.8 | 24.4 | 15.875
2006-08-05 | 1:01 | 9.17 | 7.35 | 11.14 | 96.8 | 24.4 | 15.856
2006-08-05 | 1:16 | 8.97 | 7.35 | 11.2 | 96.8 | 24.4 | 15.849
2006-08-05 | 1:31 | 8.79 | 7.36 | 11.25 | 96.8 | 24.4 | 15.846
2006-08-05 | 1:46 | 8.6 | 7.35 | 11.31 | 96.9 | 24.4 | 15.847
2006-08-05 | 2:01 | 8.42 | 7.35 | 11.37 | 96.9 | 24.4 | 15.843
2006-08-05 | 2:16 | 8.24 | 7.35 | 11.42 | 97 | 24.4 | 15.856
2006-08-05 | 2:31 | 8.07 | 7.35 | 11.48 | 97.1 | 24.4 | 15.85
2006-08-05 | 2:46 | 7.9 | 7.34 | 11.53 | 97.1 | 24.4 | 15.857
2006-08-05 | 3:01 | 7.74 | 7.34 | 11.58 | 97.2 | 24.4 | 15.858
2006-08-05 | 3:16 | 7.58 | 7.35 | 11.64 | 97.3 | 24.4 | 15.877
2006-08-05 | 3:31 | 7.43 | 7.35 | 11.69 | 97.3 | 24.4 | 15.862
2006-08-05 | 3:46 | 7.29 | 7.34 | 11.74 | 97.4 | 24.4 | 15.878
2006-08-05 | 4:01 | 7.14 | 7.34 | 11.8 | 97.5 | 24.4 | 15.858
2006-08-05 | 4:16 | 6.99 | 7.34 | 11.84 | 97.5 | 24.4 | 15.865
2006-08-05 | 4:31 | 6.85 | 7.33 | 11.89 | 97.6 | 24.4 | 15.862
2006-08-05 | 4:46 | 6.71 | 7.34 | 11.9 | 97.3 | 24.4 | 15.879
2006-08-05 | 5:01 | 6.58 | 7.33 | 11.98 | 97.6 | 24.4 | 15.866
2006-08-05 | 5:16 | 6.46 | 7.33 | 12.03 | 97.8 | 24.4 | 15.863
2006-08-05 | 5:31 | 6.33 | 7.33 | 12.07 | 97.8 | 24.4 | 15.861
2006-08-05 | 5:46 | 6.21 | 7.33 | 12.13 | 97.9 | 24.4 | 15.859
2006-08-05 | 6:01 | 6.1 | 7.32 | 12.15 | 97.8 | 24.4 | 15.875
2006-08-05 | 6:16 | 5.99 | 7.33 | 12.19 | 97.9 | 24.4 | 15.864
2006-08-05 | 6:31 | 5.9 | 7.33 | 12.23 | 98 | 24.4 | 15.88
2006-08-05 | 6:46 | 5.81 | 7.33 | 12.27 | 98.1 | 24.4 | 15.857
2006-08-05 | 7:01 | 5.74 | 7.34 | 12.3 | 98.1 | 24.4 | 15.852
2006-08-05 | 7:16 | 5.69 | 7.34 | 12.33 | 98.2 | 24.4 | 15.849
2006-08-05 | 7:31 | 5.64 | 7.35 | 12.34 | 98.2 | 24.4 | 15.844
2006-08-05 | 7:46 | 5.61 | 7.36 | 12.36 | 98.3 | 24.4 | 15.84
2006-08-05 | 8:01 | 5.59 | 7.35 | 12.38 | 98.4 | 24.4 | 15.832
2006-08-05 | 8:16 | 5.57 | 7.37 | 12.38 | 98.4 | 24.4 | 15.839
2006-08-05 | 8:31 | 5.57 | 7.36 | 12.38 | 98.4 | 24.4 | 15.828
2006-08-05 | 8:46 | 5.58 | 7.37 | 12.39 | 98.4 | 24.4 | 15.841
2006-08-05 | 9:01 | 5.61 | 7.38 | 12.38 | 98.5 | 24.4 | 15.838
2006-08-05 | 9:16 | 5.66 | 7.38 | 12.36 | 98.4 | 24.4 | 15.833
2006-08-05 | 9:31 | 5.72 | 7.39 | 12.34 | 98.5 | 24.3 | 15.814
2006-08-05 | 9:46 | 5.8 | 7.38 | 12.3 | 98.3 | 24.3 | 15.809
2006-08-05 | 10:01 | 5.89 | 7.39 | 12.29 | 98.4 | 24.3 | 15.809
2006-08-05 | 10:16 | 6 | 7.39 | 12.25 | 98.4 | 24.3 | 15.811
2006-08-05 | 10:31 | 6.12 | 7.39 | 12.21 | 98.4 | 24.3 | 15.799
2006-08-05 | 10:46 | 6.26 | 7.4 | 12.16 | 98.3 | 24.3 | 15.819
2006-08-05 | 11:01 | 6.42 | 7.4 | 12.11 | 98.3 | 24.3 | 15.801
2006-08-05 | 11:16 | 6.59 | 7.4 | 12.03 | 98.1 | 24.3 | 15.818
2006-08-05 | 11:31 | 6.78 | 7.41 | 11.99 | 98.2 | 24.3 | 15.803
2006-08-05 | 11:46 | 6.98 | 7.4 | 11.92 | 98.1 | 24.3 | 15.812
2006-08-05 | 12:01 | 7.21 | 7.41 | 11.84 | 98.1 | 24.3 | 15.814
2006-08-05 | 12:16 | 7.45 | 7.42 | 11.77 | 98 | 24.4 | 15.83
2006-08-05 | 12:31 | 7.72 | 7.42 | 11.68 | 98 | 24.3 | 15.821
2006-08-05 | 12:46 | 7.99 | 7.42 | 11.59 | 97.8 | 24.3 | 15.823
2006-08-05 | 13:00 | 8.29 | 7.43 | 11.5 | 97.8 | 24.3 | 15.825
2006-08-05 | 13:15 | 8.59 | 7.43 | 11.41 | 97.7 | 24.4 | 15.838
2006-08-05 | 13:30 | 8.9 | 7.43 | 11.31 | 97.6 | 24.4 | 15.832
2006-08-05 | 13:45 | 9.23 | 7.44 | 11.2 | 97.5 | 24.3 | 15.827
2006-08-05 | 14:00 | 9.55 | 7.45 | 11.1 | 97.3 | 24.4 | 15.838
2006-08-05 | 14:15 | 9.87 | 7.45 | 11.02 | 97.3 | 24.4 | 15.832
2006-08-05 | 14:30 | 10.18 | 7.45 | 10.91 | 97.1 | 24.4 | 15.832
2006-08-05 | 14:45 | 10.48 | 7.45 | 10.82 | 97 | 24.4 | 15.835
2006-08-05 | 15:00 | 10.77 | 7.45 | 10.73 | 96.8 | 24.4 | 15.843
2006-08-05 | 15:15 | 11.04 | 7.46 | 10.66 | 96.8 | 24.4 | 15.845
2006-08-05 | 15:30 | 11.3 | 7.46 | 10.58 | 96.6 | 24.4 | 15.853
2006-08-05 | 15:45 | 11.54 | 7.46 | 10.52 | 96.6 | 24.4 | 15.873
2006-08-05 | 16:00 | 11.76 | 7.46 | 10.44 | 96.4 | 24.4 | 15.857
2006-08-05 | 16:15 | 11.97 | 7.46 | 10.38 | 96.2 | 24.4 | 15.871
2006-08-05 | 16:30 | 12.16 | 7.46 | 10.33 | 96.2 | 24.4 | 15.885
2006-08-05 | 16:45 | 12.32 | 7.46 | 10.29 | 96.2 | 24.4 | 15.877
2006-08-05 | 17:00 | 12.45 | 7.47 | 10.25 | 96.1 | 24.5 | 15.9
2006-08-05 | 17:15 | 12.56 | 7.46 | 10.23 | 96.1 | 24.4 | 15.89
2006-08-05 | 17:30 | 12.65 | 7.46 | 10.19 | 96 | 24.5 | 15.91
2006-08-05 | 17:45 | 12.71 | 7.46 | 10.2 | 96.1 | 24.5 | 15.919
2006-08-05 | 18:00 | 12.76 | 7.45 | 10.15 | 95.8 | 24.5 | 15.902
2006-08-05 | 18:15 | 12.78 | 7.46 | 10.15 | 95.9 | 24.5 | 15.906
2006-08-05 | 18:30 | 12.79 | 7.45 | 10.14 | 95.8 | 24.5 | 15.909
2006-08-05 | 18:45 | 12.77 | 7.44 | 10.15 | 95.8 | 24.5 | 15.936
2006-08-05 | 19:00 | 12.72 | 7.44 | 10.16 | 95.8 | 24.5 | 15.914
2006-08-05 | 19:15 | 12.64 | 7.43 | 10.17 | 95.8 | 24.5 | 15.916
2006-08-05 | 19:30 | 12.53 | 7.41 | 10.19 | 95.7 | 24.5 | 15.92
2006-08-05 | 19:45 | 12.42 | 7.41 | 10.23 | 95.8 | 24.5 | 15.937
2006-08-05 | 20:00 | 12.3 | 7.4 | 10.25 | 95.8 | 24.5 | 15.933
2006-08-05 | 20:15 | 12.19 | 7.39 | 10.29 | 95.9 | 24.5 | 15.939
2006-08-05 | 20:30 | 12.07 | 7.39 | 10.32 | 95.9 | 24.5 | 15.942
2006-08-05 | 20:45 | 11.96 | 7.38 | 10.35 | 95.9 | 24.5 | 15.919
2006-08-05 | 21:00 | 11.85 | 7.38 | 10.38 | 96 | 24.5 | 15.914
2006-08-05 | 21:15 | 11.75 | 7.38 | 10.42 | 96.1 | 24.5 | 15.906
2006-08-05 | 21:30 | 11.63 | 7.38 | 10.44 | 96.1 | 24.5 | 15.918
2006-08-05 | 21:45 | 11.52 | 7.37 | 10.47 | 96.1 | 24.5 | 15.9
2006-08-05 | 22:00 | 11.4 | 7.37 | 10.47 | 95.8 | 24.5 | 15.901
2006-08-05 | 22:15 | 11.27 | 7.37 | 10.54 | 96.2 | 24.5 | 15.899
2006-08-05 | 22:30 | 11.14 | 7.37 | 10.57 | 96.1 | 24.5 | 15.91
2006-08-05 | 22:45 | 11 | 7.37 | 10.62 | 96.3 | 24.5 | 15.893
2006-08-05 | 23:00 | 10.87 | 7.36 | 10.65 | 96.3 | 24.4 | 15.892
2006-08-05 | 23:15 | 10.74 | 7.36 | 10.69 | 96.4 | 24.5 | 15.903
2006-08-05 | 23:30 | 10.61 | 7.36 | 10.73 | 96.4 | 24.4 | 15.89
2006-08-05 | 23:45 | 10.48 | 7.35 | 10.76 | 96.4 | 24.5 | 15.903
2006-08-06 | 0:00 | 10.35 | 7.35 | 10.8 | 96.5 | 24.4 | 15.885
2006-08-06 | 0:15 | 10.21 | 7.35 | 10.84 | 96.5 | 24.4 | 15.89
2006-08-06 | 0:30 | 10.08 | 7.35 | 10.88 | 96.6 | 24.4 | 15.881
2006-08-06 | 0:45 | 9.94 | 7.34 | 10.91 | 96.6 | 24.4 | 15.884
2006-08-06 | 1:01 | 9.8 | 7.34 | 10.96 | 96.6 | 24.4 | 15.879
2006-08-06 | 1:16 | 9.66 | 7.34 | 11 | 96.7 | 24.4 | 15.877
2006-08-06 | 1:31 | 9.53 | 7.34 | 11.04 | 96.7 | 24.4 | 15.877
2006-08-06 | 1:46 | 9.39 | 7.34 | 11.08 | 96.8 | 24.4 | 15.877
2006-08-06 | 2:01 | 9.25 | 7.34 | 11.13 | 96.8 | 24.4 | 15.88
2006-08-06 | 2:16 | 9.12 | 7.34 | 11.16 | 96.8 | 24.5 | 15.903
2006-08-06 | 2:31 | 9 | 7.34 | 11.2 | 96.9 | 24.4 | 15.891
2006-08-06 | 2:46 | 8.88 | 7.33 | 11.24 | 96.9 | 24.5 | 15.915
2006-08-06 | 3:01 | 8.76 | 7.33 | 11.27 | 96.9 | 24.5 | 15.907
2006-08-06 | 3:16 | 8.64 | 7.33 | 11.31 | 97 | 24.5 | 15.907
2006-08-06 | 3:31 | 8.52 | 7.33 | 11.34 | 97 | 24.5 | 15.913
2006-08-06 | 3:46 | 8.41 | 7.33 | 11.38 | 97.1 | 24.5 | 15.91
2006-08-06 | 4:01 | 8.3 | 7.33 | 11.42 | 97.2 | 24.5 | 15.927
2006-08-06 | 4:16 | 8.19 | 7.33 | 11.45 | 97.1 | 24.5 | 15.913
2006-08-06 | 4:31 | 8.09 | 7.33 | 11.48 | 97.2 | 24.5 | 15.929
2006-08-06 | 4:46 | 8 | 7.33 | 11.51 | 97.2 | 24.5 | 15.914
2006-08-06 | 5:01 | 7.9 | 7.33 | 11.54 | 97.2 | 24.5 | 15.916
2006-08-06 | 5:16 | 7.81 | 7.33 | 11.58 | 97.3 | 24.5 | 15.912
2006-08-06 | 5:31 | 7.73 | 7.33 | 11.57 | 97.1 | 24.5 | 15.918
2006-08-06 | 5:46 | 7.65 | 7.33 | 11.62 | 97.3 | 24.5 | 15.919
2006-08-06 | 6:01 | 7.57 | 7.33 | 11.66 | 97.4 | 24.5 | 15.931
2006-08-06 | 6:16 | 7.5 | 7.33 | 11.68 | 97.4 | 24.5 | 15.923
2006-08-06 | 6:31 | 7.43 | 7.33 | 11.71 | 97.5 | 24.5 | 15.94
2006-08-06 | 6:46 | 7.37 | 7.33 | 11.74 | 97.6 | 24.5 | 15.934
2006-08-06 | 7:01 | 7.32 | 7.34 | 11.75 | 97.6 | 24.5 | 15.915
2006-08-06 | 7:16 | 7.27 | 7.35 | 11.77 | 97.6 | 24.5 | 15.919
2006-08-06 | 7:31 | 7.24 | 7.35 | 11.78 | 97.6 | 24.5 | 15.915
2006-08-06 | 7:46 | 7.22 | 7.36 | 11.8 | 97.7 | 24.5 | 15.912
2006-08-06 | 8:01 | 7.2 | 7.37 | 11.81 | 97.7 | 24.5 | 15.925
2006-08-06 | 8:16 | 7.19 | 7.37 | 11.81 | 97.8 | 24.5 | 15.915
2006-08-06 | 8:31 | 7.19 | 7.38 | 11.81 | 97.8 | 24.5 | 15.906
2006-08-06 | 8:46 | 7.19 | 7.39 | 11.83 | 97.9 | 24.5 | 15.905
2006-08-06 | 9:01 | 7.22 | 7.4 | 11.83 | 98 | 24.5 | 15.897
2006-08-06 | 9:16 | 7.27 | 7.4 | 11.81 | 97.9 | 24.5 | 15.908
2006-08-06 | 9:31 | 7.31 | 7.41 | 11.8 | 97.9 | 24.5 | 15.897
2006-08-06 | 9:46 | 7.35 | 7.42 | 11.79 | 97.9 | 24.4 | 15.892
2006-08-06 | 10:01 | 7.4 | 7.41 | 11.77 | 97.9 | 24.5 | 15.899
2006-08-06 | 10:16 | 7.51 | 7.42 | 11.72 | 97.8 | 24.5 | 15.902
2006-08-06 | 10:31 | 7.65 | 7.42 | 11.69 | 97.9 | 24.5 | 15.908
2006-08-06 | 10:46 | 7.79 | 7.43 | 11.65 | 97.8 | 24.5 | 15.907
2006-08-06 | 11:01 | 7.93 | 7.43 | 11.6 | 97.8 | 24.5 | 15.909
2006-08-06 | 11:16 | 8.09 | 7.43 | 11.57 | 97.9 | 24.5 | 15.917
2006-08-06 | 11:31 | 8.24 | 7.44 | 11.5 | 97.7 | 24.5 | 15.919
2006-08-06 | 11:46 | 8.39 | 7.44 | 11.45 | 97.6 | 24.5 | 15.921
2006-08-06 | 12:01 | 8.56 | 7.43 | 11.37 | 97.3 | 24.5 | 15.936
2006-08-06 | 12:16 | 8.78 | 7.44 | 11.33 | 97.5 | 24.5 | 15.925
2006-08-06 | 12:31 | 9.02 | 7.44 | 11.26 | 97.5 | 24.5 | 15.941
2006-08-06 | 12:46 | 9.32 | 7.44 | 11.17 | 97.4 | 24.5 | 15.946
2006-08-06 | 13:00 | 9.59 | 7.45 | 11.09 | 97.3 | 24.5 | 15.936
2006-08-06 | 13:15 | 9.82 | 7.45 | 11.01 | 97.1 | 24.5 | 15.942
2006-08-06 | 13:30 | 9.99 | 7.45 | 10.97 | 97.2 | 24.5 | 15.954
2006-08-06 | 13:45 | 10.22 | 7.45 | 10.91 | 97.1 | 24.5 | 15.945
2006-08-06 | 14:00 | 10.46 | 7.45 | 10.85 | 97.2 | 24.5 | 15.943
2006-08-06 | 14:15 | 10.75 | 7.46 | 10.77 | 97.1 | 24.5 | 15.929
2006-08-06 | 14:30 | 11.08 | 7.45 | 10.68 | 97 | 24.5 | 15.954
2006-08-06 | 14:45 | 11.35 | 7.45 | 10.59 | 96.9 | 24.6 | 15.967