## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River lower bridge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 6920 # 06A2389AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable #51063402 surface water | [5 empty] | A
Ph | pH 7 standard solution | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148.0 µS/mL | [5 empty] | #DIV/0!
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2006-08-29 14:19 | 21.49 | 8.48 | 9.77 | 110.6 | 69.8 | 45.348
Field Equipment | 2006-08-29 14:16 | 21.47 | 8.8 | 10.06 | 113.9 | 71.9 | 46.759
Grade |  | A | B | A |  | A
Audit 2 | 2006-09-09 14:41 | 19.68 |  | 10.34 | 113 | 68.9 | 44.782
Field Equipment | 2006-09-09 14:46 | 19.69 |  | 9.59 | 104.8 | 71.5 | 46.464
Grade |  | A | A | B |  | A
Audit 3 | 2006-09-15 12:37 | 15.31 | 8 | 10.74 | 107.2 | 69 | 44.832
Field Equipment | 2006-09-15 12:31 | 15.15 | 8.5 | 9.77 | 97.2 | 71.5 | 46.495
Grade |  | A | B | B |  | A
Post-sampling Field Equipment Check
Field equipment ID | YSI 6920 # 06A2389AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-15 |  | 2006-09-15 |  | A
Time | Computer set to atomic time |  | 15:07 |  | 16:05 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 16.9 |  | 16.4 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.1 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 9.2 |  | B
SpC | 148.0 µS/mL |  | 148 |  | 158 |  | A
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
|  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River lower bridge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-29 |  | 2006-08-29 |  | A
Time | Computer set to atomic time |  | 9:23 |  | 9:23 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 17.6 |  | 17.4 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7 |  | A
DO | Winkler titration of surface water |  | 8.7 |  | 9.2 |  | B
SpC | 148.0 µS/mL |  | 148 |  | 141 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
Field equipment ID | Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
YSI 556 MPS #05G1083AM | 2006-08-29 14:19 | 21.49 | 8.48 | 9.77 | 110.6 | 69.8 | 45.348
Hach 2100P #50828
Post-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-30 |  | 2006-08-30 |  | A
Time | Computer set to atomic time |  | 10:00 |  | 10:00 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 16.6 |  | 16.3 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7 |  | A
DO | Winkler titration of surface water |  | 9.5 |  | 9.6 |  | A
SpC | 148.0 µS/mL |  | 148 |  | 142 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
|  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River lower bridge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-08 |  | 2006-09-08 |  | A
Time | Computer set to atomic time |  | 9:25 |  | 9:24 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 17.1 |  | 16.4 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 9.6 |  | B
SpC | 148.0 µS/mL |  | 148 |  | 147 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
````````````````````````````````````````````````````````````````````````````````````````
Field Equipment Grab Sampling Results
Field equipment ID | Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
YSI 556 MPS #05G1083AM | 2006-09-09 14:41 | 19.68 | 8.28 | 10.34 | 113 | 68.9 | 44.782
Hach 2100P #50828
Post-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-12 |  | 2006-09-12 |  | A
Time | Computer set to atomic time |  | 12:05 |  | 12:04 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 18.9 |  | 18.4 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7 |  | A
DO | Winkler titration of surface water |  | 9.7 |  | 9.7 |  | A
SpC | 148.0 µS/mL |  | 148 |  | 143 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
|  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River lower bridge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-15 |  | 2006-09-15 |  | A
Time | Computer set to atomic time |  | 9:16 |  | 9:15 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 14.1 |  | 13.7 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.1 |  | A
DO | Winkler titration of surface water |  | 9.2 |  | 9.9 |  | B
SpC | 148.0 µS/mL |  | 148 |  | 143 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
Field equipment ID | Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
YSI 556 MPS #05G1083AM | 2006-09-15 12:37 | 15.31 | 7.97 | 10.74 | 107.2 | 69 | 44.832
Hach 2100P #50828
Post-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-15 |  | 2006-09-15 |  | A
Time | Computer set to atomic time |  | 15:52 |  | 15:51 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 16.9 |  | 16.1 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.1 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 9.6 |  | B
SpC | 148.0 µS/mL |  | 148 |  | 142 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
|  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1

## data
Date | Time | Temp | pH | DO Conc | DO% | SpCond | TDS
|  | °C |  | mg/L | % | uS/cm | mg/L
2006-08-29 | 14:15 | 21.47 | 8.8 | 10.06 | 113.9 | 71.9 | 46.759
2006-08-29 | 14:30 | 21.55 | 8.85 | 9.94 | 112.7 | 72 | 46.777
2006-08-29 | 14:45 | 21.65 | 8.88 | 9.84 | 111.8 | 72.1 | 46.856
2006-08-29 | 15:00 | 21.72 | 8.89 | 9.75 | 110.9 | 72.1 | 46.853
2006-08-29 | 15:15 | 21.76 | 8.91 | 9.69 | 110.3 | 72.2 | 46.899
2006-08-29 | 15:30 | 21.77 | 8.92 | 9.61 | 109.5 | 72.2 | 46.916
2006-08-29 | 15:45 | 21.79 | 8.94 | 9.58 | 109.1 | 72.1 | 46.895
2006-08-29 | 16:00 | 21.79 | 8.95 | 9.52 | 108.5 | 72.2 | 46.934
2006-08-29 | 16:15 | 21.78 | 8.96 | 9.48 | 107.9 | 72.2 | 46.916
2006-08-29 | 16:30 | 21.77 | 8.97 | 9.42 | 107.3 | 72.1 | 46.895
2006-08-29 | 16:45 | 21.74 | 8.95 | 9.35 | 106.4 | 72.2 | 46.905
2006-08-29 | 17:00 | 21.69 | 8.96 | 9.28 | 105.5 | 72.2 | 46.935
2006-08-29 | 17:15 | 21.65 | 8.96 | 9.23 | 104.8 | 72.2 | 46.948
2006-08-29 | 17:30 | 21.59 | 8.95 | 9.16 | 103.9 | 72.2 | 46.911
2006-08-29 | 17:45 | 21.54 | 8.94 | 9.09 | 103 | 72.2 | 46.949
2006-08-29 | 18:00 | 21.44 | 8.95 | 9.06 | 102.5 | 72.2 | 46.943
2006-08-29 | 18:15 | 21.31 | 8.94 | 9 | 101.6 | 72.2 | 46.954
2006-08-29 | 18:30 | 21.15 | 8.92 | 8.95 | 100.7 | 72.3 | 46.996
2006-08-29 | 18:45 | 20.98 | 8.9 | 8.86 | 99.4 | 72.3 | 47.015
2006-08-29 | 19:00 | 20.8 | 8.86 | 8.74 | 97.7 | 72.3 | 46.994
2006-08-29 | 19:15 | 20.6 | 8.83 | 8.62 | 96 | 72.4 | 47.044
2006-08-29 | 19:30 | 20.47 | 8.78 | 8.49 | 94.3 | 72.4 | 47.07
2006-08-29 | 19:45 | 20.33 | 8.72 | 8.37 | 92.7 | 72.4 | 47.074
2006-08-29 | 20:00 | 20.12 | 8.65 | 8.25 | 90.9 | 72.4 | 47.088
2006-08-29 | 20:15 | 19.97 | 8.59 | 8.16 | 89.7 | 72.5 | 47.12
2006-08-29 | 20:30 | 19.84 | 8.5 | 8.08 | 88.6 | 72.6 | 47.166
2006-08-29 | 20:45 | 19.7 | 8.4 | 7.99 | 87.3 | 72.7 | 47.27
2006-08-29 | 21:00 | 19.54 | 8.31 | 7.94 | 86.5 | 72.7 | 47.269
2006-08-29 | 21:15 | 19.38 | 8.21 | 7.88 | 85.7 | 72.8 | 47.34
2006-08-29 | 21:30 | 19.24 | 8.11 | 7.85 | 85 | 72.9 | 47.383
2006-08-29 | 21:45 | 19.12 | 8.02 | 7.8 | 84.3 | 73.1 | 47.494
2006-08-29 | 22:00 | 18.99 | 7.94 | 7.78 | 83.9 | 73.2 | 47.563
2006-08-29 | 22:15 | 18.81 | 7.87 | 7.75 | 83.3 | 73.3 | 47.647
2006-08-29 | 22:30 | 18.67 | 7.81 | 7.75 | 83 | 73.4 | 47.687
2006-08-29 | 22:45 | 18.53 | 7.76 | 7.73 | 82.6 | 73.5 | 47.754
2006-08-29 | 23:00 | 18.41 | 7.72 | 7.71 | 82.2 | 73.6 | 47.845
2006-08-29 | 23:15 | 18.3 | 7.68 | 7.72 | 82 | 73.6 | 47.859
2006-08-29 | 23:30 | 18.17 | 7.65 | 7.72 | 81.8 | 73.6 | 47.866
2006-08-29 | 23:45 | 18.05 | 7.63 | 7.72 | 81.7 | 73.6 | 47.843
2006-08-30 | 0:00 | 17.95 | 7.61 | 7.72 | 81.5 | 73.7 | 47.883
2006-08-30 | 0:15 | 17.86 | 7.59 | 7.74 | 81.5 | 73.5 | 47.794
2006-08-30 | 0:30 | 17.76 | 7.57 | 7.73 | 81.3 | 73.5 | 47.761
2006-08-30 | 0:45 | 17.68 | 7.56 | 7.73 | 81.2 | 73.4 | 47.712
2006-08-30 | 1:01 | 17.59 | 7.55 | 7.75 | 81.2 | 73.3 | 47.636
2006-08-30 | 1:16 | 17.49 | 7.53 | 7.76 | 81.2 | 73.2 | 47.598
2006-08-30 | 1:31 | 17.37 | 7.52 | 7.78 | 81.1 | 73.1 | 47.536
2006-08-30 | 1:46 | 17.26 | 7.52 | 7.8 | 81.2 | 73.1 | 47.515
2006-08-30 | 2:01 | 17.17 | 7.51 | 7.82 | 81.2 | 73 | 47.458
2006-08-30 | 2:16 | 17.06 | 7.51 | 7.83 | 81.2 | 73 | 47.435
2006-08-30 | 2:31 | 16.99 | 7.5 | 7.85 | 81.2 | 72.9 | 47.397
2006-08-30 | 2:46 | 16.93 | 7.49 | 7.86 | 81.3 | 72.9 | 47.357
2006-08-30 | 3:01 | 16.89 | 7.49 | 7.88 | 81.4 | 72.8 | 47.333
2006-08-30 | 3:16 | 16.84 | 7.49 | 7.89 | 81.4 | 72.8 | 47.349
2006-08-30 | 3:31 | 16.74 | 7.49 | 7.91 | 81.4 | 72.9 | 47.378
2006-08-30 | 3:46 | 16.65 | 7.48 | 7.92 | 81.3 | 72.8 | 47.319
2006-08-30 | 4:01 | 16.62 | 7.48 | 7.93 | 81.4 | 72.8 | 47.303
2006-08-30 | 4:16 | 16.56 | 7.48 | 7.94 | 81.4 | 72.7 | 47.264
2006-08-30 | 4:31 | 16.52 | 7.48 | 7.95 | 81.4 | 72.7 | 47.247
2006-08-30 | 4:46 | 16.49 | 7.47 | 7.96 | 81.5 | 72.6 | 47.204
2006-08-30 | 5:01 | 16.47 | 7.47 | 7.97 | 81.6 | 72.6 | 47.17
2006-08-30 | 5:16 | 16.43 | 7.47 | 7.98 | 81.5 | 72.6 | 47.162
2006-08-30 | 5:31 | 16.4 | 7.47 | 7.97 | 81.5 | 72.5 | 47.095
2006-08-30 | 5:46 | 16.36 | 7.47 | 7.99 | 81.6 | 72.5 | 47.095
2006-08-30 | 6:01 | 16.33 | 7.46 | 8 | 81.6 | 72.4 | 47.052
2006-08-30 | 6:16 | 16.32 | 7.46 | 8 | 81.6 | 72.3 | 47.019
2006-08-30 | 6:31 | 16.32 | 7.46 | 8.01 | 81.8 | 72.3 | 46.997
2006-08-30 | 6:46 | 16.32 | 7.47 | 8.04 | 82 | 72.2 | 46.944
2006-08-30 | 7:01 | 16.31 | 7.48 | 8.07 | 82.3 | 72.2 | 46.921
2006-08-30 | 7:16 | 16.3 | 7.49 | 8.13 | 82.9 | 72.1 | 46.884
2006-08-30 | 7:31 | 16.3 | 7.51 | 8.23 | 84 | 72 | 46.815
2006-08-30 | 7:46 | 16.33 | 7.54 | 8.35 | 85.2 | 72 | 46.793
2006-08-30 | 8:01 | 16.36 | 7.57 | 8.47 | 86.5 | 71.9 | 46.746
2006-08-30 | 8:16 | 16.41 | 7.6 | 8.57 | 87.6 | 71.9 | 46.716
2006-08-30 | 8:31 | 16.47 | 7.63 | 8.68 | 88.8 | 71.9 | 46.71
2006-08-30 | 8:46 | 16.54 | 7.67 | 8.8 | 90.2 | 71.7 | 46.637
2006-08-30 | 9:01 | 16.6 | 7.72 | 8.93 | 91.6 | 71.7 | 46.587
2006-08-30 | 9:16 | 16.71 | 7.77 | 9.05 | 93.1 | 71.6 | 46.526
2006-08-30 | 9:31 | 16.83 | 7.83 | 9.17 | 94.6 | 71.5 | 46.464
2006-08-30 | 9:46 | 16.94 | 7.9 | 9.3 | 96.1 | 71.4 | 46.434
2006-08-30 | 10:01 | 17.03 | 7.95 | 9.38 | 97.1 | 71.4 | 46.384
2006-08-30 | 10:16 | 17.16 | 8.02 | 9.48 | 98.4 | 71.3 | 46.331
2006-08-30 | 10:31 | 17.3 | 8.1 | 9.57 | 99.7 | 71.2 | 46.298
2006-08-30 | 10:46 | 17.44 | 8.16 | 9.64 | 100.7 | 71.2 | 46.272
2006-08-30 | 11:01 | 17.59 | 8.25 | 9.72 | 101.9 | 71.2 | 46.251
2006-08-30 | 11:16 | 17.74 | 8.31 | 9.78 | 102.8 | 71.2 | 46.25
2006-08-30 | 11:31 | 17.87 | 8.37 | 9.82 | 103.5 | 71.1 | 46.211
2006-08-30 | 11:46 | 18 | 8.43 | 9.88 | 104.3 | 71.1 | 46.235
2006-08-30 | 12:01 | 18.07 | 8.48 | 9.89 | 104.6 | 71.1 | 46.231
2006-08-30 | 12:16 | 18.3 | 8.54 | 9.94 | 105.6 | 71.2 | 46.264
2006-08-30 | 12:31 | 18.38 | 8.57 | 9.93 | 105.7 | 71.2 | 46.25
2006-08-30 | 12:46 | 18.6 | 8.63 | 9.97 | 106.6 | 71.2 | 46.278
2006-08-30 | 13:00 | 18.75 | 8.66 | 9.99 | 107.2 | 71.2 | 46.308
2006-08-30 | 13:15 | 18.91 | 8.7 | 9.98 | 107.4 | 71.2 | 46.306
2006-08-30 | 13:30 | 19 | 8.74 | 9.96 | 107.4 | 71.3 | 46.318
2006-08-30 | 13:45 | 19.11 | 8.76 | 9.96 | 107.6 | 71.3 | 46.347
2006-08-30 | 14:00 | 19.25 | 8.79 | 9.95 | 107.9 | 71.3 | 46.362
2006-08-30 | 14:15 | 19.39 | 8.83 | 9.96 | 108.2 | 71.4 | 46.418
2006-08-30 | 14:30 | 19.51 | 8.85 | 9.95 | 108.4 | 71.4 | 46.428
2006-08-30 | 14:45 | 19.62 | 8.87 | 9.92 | 108.3 | 71.5 | 46.475
2006-08-30 | 15:00 | 19.71 | 8.9 | 9.91 | 108.3 | 71.6 | 46.534
2006-08-30 | 15:15 | 19.78 | 8.91 | 9.88 | 108.2 | 71.6 | 46.51
2006-08-30 | 15:30 | 19.83 | 8.93 | 9.86 | 108.1 | 71.6 | 46.531
2006-08-30 | 15:45 | 19.88 | 8.94 | 9.81 | 107.7 | 71.6 | 46.563
2006-08-30 | 16:00 | 19.94 | 8.95 | 9.76 | 107.2 | 71.7 | 46.594
2006-08-30 | 16:15 | 19.95 | 8.97 | 9.73 | 106.9 | 71.8 | 46.652
2006-08-30 | 16:30 | 19.99 | 8.98 | 9.69 | 106.6 | 71.8 | 46.681
2006-08-30 | 16:45 | 19.98 | 8.97 | 9.62 | 105.8 | 71.9 | 46.722
2006-08-30 | 17:00 | 19.95 | 8.98 | 9.56 | 105.1 | 72 | 46.781
2006-08-30 | 17:15 | 19.91 | 8.98 | 9.5 | 104.3 | 72.1 | 46.836
2006-08-30 | 17:30 | 19.87 | 8.98 | 9.45 | 103.7 | 72.1 | 46.861
2006-08-30 | 17:45 | 19.8 | 8.97 | 9.38 | 102.8 | 72.2 | 46.916
2006-08-30 | 18:00 | 19.72 | 8.96 | 9.31 | 101.9 | 72.2 | 46.958
2006-08-30 | 18:15 | 19.62 | 8.95 | 9.23 | 100.8 | 72.3 | 47.024
2006-08-30 | 18:30 | 19.52 | 8.92 | 9.14 | 99.6 | 72.4 | 47.068
2006-08-30 | 18:45 | 19.39 | 8.9 | 9.04 | 98.3 | 72.4 | 47.088
2006-08-30 | 19:00 | 19.24 | 8.86 | 8.91 | 96.6 | 72.6 | 47.178
2006-08-30 | 19:15 | 19.1 | 8.81 | 8.8 | 95 | 72.6 | 47.205
2006-08-30 | 19:30 | 18.96 | 8.77 | 8.7 | 93.7 | 72.7 | 47.255
2006-08-30 | 19:45 | 18.81 | 8.71 | 8.59 | 92.3 | 72.8 | 47.3
2006-08-30 | 20:00 | 18.67 | 8.64 | 8.49 | 90.9 | 72.9 | 47.399
2006-08-30 | 20:15 | 18.53 | 8.56 | 8.38 | 89.5 | 73 | 47.445
2006-08-30 | 20:30 | 18.39 | 8.47 | 8.32 | 88.6 | 73.1 | 47.522
2006-08-30 | 20:45 | 18.25 | 8.39 | 8.27 | 87.8 | 73.3 | 47.645
2006-08-30 | 21:00 | 18.11 | 8.29 | 8.21 | 87 | 73.4 | 47.726
2006-08-30 | 21:15 | 17.96 | 8.2 | 8.2 | 86.5 | 73.6 | 47.825
2006-08-30 | 21:30 | 17.82 | 8.1 | 8.14 | 85.7 | 73.7 | 47.91
2006-08-30 | 21:45 | 17.66 | 8.01 | 8.12 | 85.2 | 73.9 | 48.008
2006-08-30 | 22:00 | 17.52 | 7.93 | 8.11 | 84.8 | 74.1 | 48.14
2006-08-30 | 22:15 | 17.38 | 7.86 | 8.09 | 84.4 | 74.1 | 48.192
2006-08-30 | 22:30 | 17.23 | 7.81 | 8.07 | 84 | 74.2 | 48.257
2006-08-30 | 22:45 | 17.07 | 7.77 | 8.07 | 83.7 | 74.3 | 48.323
2006-08-30 | 23:00 | 16.92 | 7.73 | 8.07 | 83.4 | 74.5 | 48.426
2006-08-30 | 23:15 | 16.77 | 7.69 | 8.08 | 83.2 | 74.5 | 48.447
2006-08-30 | 23:30 | 16.62 | 7.67 | 8.1 | 83.1 | 74.6 | 48.464
2006-08-30 | 23:45 | 16.48 | 7.64 | 8.1 | 82.9 | 74.6 | 48.497
2006-08-31 | 0:00 | 16.34 | 7.62 | 8.12 | 82.9 | 74.5 | 48.445
2006-08-31 | 0:15 | 16.21 | 7.6 | 8.13 | 82.8 | 74.5 | 48.438
2006-08-31 | 0:30 | 16.09 | 7.59 | 8.16 | 82.8 | 74.4 | 48.386
2006-08-31 | 0:45 | 15.97 | 7.58 | 8.17 | 82.7 | 74.4 | 48.355
2006-08-31 | 1:01 | 15.86 | 7.57 | 8.2 | 82.8 | 74.3 | 48.299
2006-08-31 | 1:16 | 15.74 | 7.56 | 8.22 | 82.9 | 74.2 | 48.242
2006-08-31 | 1:31 | 15.64 | 7.55 | 8.25 | 82.9 | 74.1 | 48.17
2006-08-31 | 1:46 | 15.53 | 7.54 | 8.27 | 83 | 74 | 48.123
2006-08-31 | 2:01 | 15.42 | 7.54 | 8.3 | 83 | 74 | 48.084
2006-08-31 | 2:16 | 15.32 | 7.54 | 8.34 | 83.2 | 73.9 | 48.043
2006-08-31 | 2:31 | 15.21 | 7.53 | 8.34 | 83.1 | 73.9 | 48.046
2006-08-31 | 2:46 | 15.12 | 7.52 | 8.37 | 83.3 | 73.8 | 47.976
2006-08-31 | 3:01 | 15.03 | 7.52 | 8.39 | 83.3 | 73.7 | 47.911
2006-08-31 | 3:16 | 14.94 | 7.52 | 8.41 | 83.3 | 73.7 | 47.896
2006-08-31 | 3:31 | 14.86 | 7.51 | 8.43 | 83.4 | 73.6 | 47.843
2006-08-31 | 3:46 | 14.8 | 7.52 | 8.46 | 83.5 | 73.5 | 47.783
2006-08-31 | 4:01 | 14.74 | 7.51 | 8.46 | 83.4 | 73.4 | 47.726
2006-08-31 | 4:16 | 14.68 | 7.51 | 8.48 | 83.5 | 73.4 | 47.697
2006-08-31 | 4:31 | 14.64 | 7.51 | 8.49 | 83.5 | 73.3 | 47.644
2006-08-31 | 4:46 | 14.61 | 7.5 | 8.49 | 83.5 | 73.2 | 47.603
2006-08-31 | 5:01 | 14.56 | 7.5 | 8.5 | 83.5 | 73.2 | 47.555
2006-08-31 | 5:16 | 14.52 | 7.5 | 8.5 | 83.5 | 73.1 | 47.517
2006-08-31 | 5:31 | 14.5 | 7.5 | 8.51 | 83.5 | 73 | 47.48
2006-08-31 | 5:46 | 14.47 | 7.49 | 8.52 | 83.5 | 73 | 47.437
2006-08-31 | 6:01 | 14.44 | 7.5 | 8.52 | 83.5 | 72.9 | 47.387
2006-08-31 | 6:16 | 14.42 | 7.49 | 8.52 | 83.5 | 72.8 | 47.342
2006-08-31 | 6:31 | 14.4 | 7.5 | 8.54 | 83.6 | 72.8 | 47.295
2006-08-31 | 6:46 | 14.39 | 7.5 | 8.57 | 83.8 | 72.7 | 47.256
2006-08-31 | 7:01 | 14.37 | 7.51 | 8.61 | 84.2 | 72.6 | 47.219
2006-08-31 | 7:16 | 14.36 | 7.52 | 8.67 | 84.8 | 72.6 | 47.182
2006-08-31 | 7:31 | 14.4 | 7.54 | 8.76 | 85.8 | 72.5 | 47.105
2006-08-31 | 7:46 | 14.44 | 7.56 | 8.86 | 86.9 | 72.4 | 47.052
2006-08-31 | 8:01 | 14.49 | 7.59 | 8.97 | 88 | 72.3 | 46.997
2006-08-31 | 8:16 | 14.56 | 7.63 | 9.09 | 89.3 | 72.2 | 46.926
2006-08-31 | 8:31 | 14.63 | 7.66 | 9.2 | 90.5 | 72.1 | 46.874
2006-08-31 | 8:46 | 14.71 | 7.7 | 9.31 | 91.8 | 72.1 | 46.84
2006-08-31 | 9:01 | 14.82 | 7.74 | 9.42 | 93.1 | 71.9 | 46.736
2006-08-31 | 9:16 | 14.94 | 7.8 | 9.53 | 94.4 | 71.8 | 46.7
2006-08-31 | 9:31 | 15.07 | 7.85 | 9.63 | 95.6 | 71.7 | 46.622
2006-08-31 | 9:46 | 15.23 | 7.91 | 9.74 | 97.1 | 71.6 | 46.529
2006-08-31 | 10:01 | 15.39 | 7.98 | 9.83 | 98.3 | 71.6 | 46.517
2006-08-31 | 10:16 | 15.58 | 8.04 | 9.9 | 99.4 | 71.4 | 46.417
2006-08-31 | 10:31 | 15.77 | 8.11 | 9.97 | 100.5 | 71.3 | 46.369
2006-08-31 | 10:46 | 15.96 | 8.19 | 10.03 | 101.6 | 71.3 | 46.365
2006-08-31 | 11:01 | 16.17 | 8.26 | 10.09 | 102.6 | 71.3 | 46.343
2006-08-31 | 11:16 | 16.36 | 8.32 | 10.12 | 103.4 | 71.3 | 46.321
2006-08-31 | 11:31 | 16.59 | 8.39 | 10.16 | 104.2 | 71.3 | 46.343
2006-08-31 | 11:46 | 16.81 | 8.46 | 10.2 | 105.1 | 71.3 | 46.331
2006-08-31 | 12:01 | 17.03 | 8.51 | 10.22 | 105.8 | 71.3 | 46.376
2006-08-31 | 12:16 | 17.25 | 8.56 | 10.22 | 106.3 | 71.3 | 46.375
2006-08-31 | 12:31 | 17.45 | 8.61 | 10.23 | 106.9 | 71.4 | 46.43
2006-08-31 | 12:46 | 17.65 | 8.66 | 10.23 | 107.3 | 71.5 | 46.467
2006-08-31 | 13:00 | 17.85 | 8.7 | 10.24 | 107.8 | 71.6 | 46.509
2006-08-31 | 13:15 | 18.03 | 8.73 | 10.23 | 108.1 | 71.6 | 46.555
2006-08-31 | 13:30 | 18.21 | 8.76 | 10.21 | 108.3 | 71.7 | 46.626
2006-08-31 | 13:45 | 18.35 | 8.79 | 10.19 | 108.4 | 71.8 | 46.68
2006-08-31 | 14:00 | 18.5 | 8.83 | 10.18 | 108.7 | 71.9 | 46.721
2006-08-31 | 14:15 | 18.61 | 8.86 | 10.17 | 108.8 | 72 | 46.784
2006-08-31 | 14:30 | 18.68 | 8.87 | 10.14 | 108.6 | 72 | 46.808
2006-08-31 | 14:45 | 18.77 | 8.89 | 10.12 | 108.6 | 72 | 46.805
2006-08-31 | 15:00 | 18.83 | 8.91 | 10.1 | 108.5 | 72 | 46.832
2006-08-31 | 15:15 | 18.9 | 8.92 | 10.07 | 108.4 | 72.1 | 46.87
2006-08-31 | 15:30 | 18.97 | 8.94 | 10.04 | 108.2 | 72.1 | 46.87
2006-08-31 | 15:45 | 19.01 | 8.95 | 10 | 107.8 | 72.2 | 46.908
2006-08-31 | 16:00 | 19.04 | 8.96 | 9.95 | 107.3 | 72.1 | 46.854
2006-08-31 | 16:15 | 19.05 | 8.97 | 9.92 | 107.1 | 72.1 | 46.87
2006-08-31 | 16:30 | 19.06 | 8.98 | 9.87 | 106.5 | 72.1 | 46.866
2006-08-31 | 16:45 | 19.05 | 8.99 | 9.83 | 106.1 | 72.1 | 46.864
2006-08-31 | 17:00 | 19.03 | 8.99 | 9.77 | 105.4 | 72.1 | 46.895
2006-08-31 | 17:15 | 18.99 | 8.99 | 9.72 | 104.7 | 72.1 | 46.863
2006-08-31 | 17:30 | 18.92 | 8.98 | 9.65 | 103.9 | 72.2 | 46.918
2006-08-31 | 17:45 | 18.85 | 8.98 | 9.59 | 103 | 72.1 | 46.893
2006-08-31 | 18:00 | 18.76 | 8.97 | 9.52 | 102.1 | 72.2 | 46.904
2006-08-31 | 18:15 | 18.66 | 8.95 | 9.44 | 101.1 | 72.2 | 46.909
2006-08-31 | 18:30 | 18.56 | 8.93 | 9.36 | 100 | 72.2 | 46.928
2006-08-31 | 18:45 | 18.44 | 8.91 | 9.26 | 98.7 | 72.3 | 46.963
2006-08-31 | 19:00 | 18.31 | 8.88 | 9.18 | 97.6 | 72.3 | 47.002
2006-08-31 | 19:15 | 18.19 | 8.84 | 9.07 | 96.2 | 72.4 | 47.028
2006-08-31 | 19:30 | 18.06 | 8.8 | 8.97 | 94.9 | 72.5 | 47.119
2006-08-31 | 19:45 | 17.94 | 8.74 | 8.86 | 93.4 | 72.6 | 47.184
2006-08-31 | 20:00 | 17.81 | 8.67 | 8.75 | 92.1 | 72.7 | 47.253
2006-08-31 | 20:15 | 17.68 | 8.59 | 8.65 | 90.8 | 72.7 | 47.267
2006-08-31 | 20:30 | 17.56 | 8.52 | 8.58 | 89.8 | 72.8 | 47.352
2006-08-31 | 20:45 | 17.42 | 8.44 | 8.51 | 88.9 | 73 | 47.448
2006-08-31 | 21:00 | 17.3 | 8.34 | 8.46 | 88.1 | 73.2 | 47.567
2006-08-31 | 21:15 | 17.16 | 8.24 | 8.41 | 87.3 | 73.4 | 47.701
2006-08-31 | 21:30 | 17.05 | 8.15 | 8.38 | 86.8 | 73.5 | 47.787
2006-08-31 | 21:45 | 16.93 | 8.05 | 8.35 | 86.2 | 73.6 | 47.864
2006-08-31 | 22:00 | 16.8 | 7.98 | 8.33 | 85.8 | 73.8 | 47.989
2006-08-31 | 22:15 | 16.66 | 7.91 | 8.31 | 85.4 | 74 | 48.117
2006-08-31 | 22:30 | 16.53 | 7.85 | 8.3 | 85 | 74.1 | 48.195
2006-08-31 | 22:45 | 16.4 | 7.8 | 8.3 | 84.8 | 74.3 | 48.3
2006-08-31 | 23:00 | 16.29 | 7.76 | 8.28 | 84.4 | 74.4 | 48.361
2006-08-31 | 23:15 | 16.16 | 7.73 | 8.29 | 84.3 | 74.5 | 48.43
2006-08-31 | 23:30 | 16.03 | 7.69 | 8.29 | 84 | 74.6 | 48.465
2006-08-31 | 23:45 | 15.91 | 7.67 | 8.3 | 83.9 | 74.6 | 48.463
2006-09-01 | 0:00 | 15.78 | 7.65 | 8.31 | 83.8 | 74.5 | 48.454
2006-09-01 | 0:15 | 15.66 | 7.63 | 8.32 | 83.7 | 74.5 | 48.442
2006-09-01 | 0:30 | 15.54 | 7.61 | 8.34 | 83.7 | 74.5 | 48.403
2006-09-01 | 0:45 | 15.43 | 7.6 | 8.35 | 83.6 | 74.4 | 48.377
2006-09-01 | 1:01 | 15.31 | 7.58 | 8.38 | 83.7 | 74.4 | 48.341
2006-09-01 | 1:16 | 15.21 | 7.58 | 8.41 | 83.7 | 74.3 | 48.294
2006-09-01 | 1:31 | 15.1 | 7.57 | 8.43 | 83.8 | 74.2 | 48.249
2006-09-01 | 1:46 | 14.99 | 7.56 | 8.45 | 83.8 | 74.2 | 48.215
2006-09-01 | 2:01 | 14.9 | 7.55 | 8.47 | 83.8 | 74.1 | 48.163
2006-09-01 | 2:16 | 14.81 | 7.55 | 8.5 | 83.9 | 74 | 48.132
2006-09-01 | 2:31 | 14.72 | 7.54 | 8.52 | 84 | 74 | 48.084
2006-09-01 | 2:46 | 14.65 | 7.54 | 8.54 | 84.1 | 73.9 | 48.043
2006-09-01 | 3:01 | 14.58 | 7.53 | 8.56 | 84.1 | 73.8 | 48.001
2006-09-01 | 3:16 | 14.51 | 7.53 | 8.57 | 84.1 | 73.8 | 47.971
2006-09-01 | 3:31 | 14.46 | 7.53 | 8.58 | 84.1 | 73.8 | 47.964
2006-09-01 | 3:46 | 14.4 | 7.52 | 8.58 | 84 | 73.7 | 47.88
2006-09-01 | 4:01 | 14.36 | 7.52 | 8.6 | 84.1 | 73.6 | 47.827
2006-09-01 | 4:16 | 14.32 | 7.51 | 8.61 | 84.1 | 73.5 | 47.792
2006-09-01 | 4:31 | 14.29 | 7.52 | 8.61 | 84.1 | 73.4 | 47.721
2006-09-01 | 4:46 | 14.26 | 7.51 | 8.62 | 84.1 | 73.4 | 47.68
2006-09-01 | 5:01 | 14.23 | 7.51 | 8.63 | 84.2 | 73.3 | 47.619
2006-09-01 | 5:16 | 14.21 | 7.51 | 8.62 | 84.1 | 73.2 | 47.588
2006-09-01 | 5:31 | 14.19 | 7.5 | 8.62 | 84 | 73.1 | 47.54
2006-09-01 | 5:46 | 14.19 | 7.5 | 8.63 | 84.1 | 73 | 47.435
2006-09-01 | 6:01 | 14.16 | 7.51 | 8.62 | 83.9 | 72.9 | 47.371
2006-09-01 | 6:16 | 14.14 | 7.5 | 8.63 | 84 | 72.8 | 47.347
2006-09-01 | 6:31 | 14.14 | 7.5 | 8.64 | 84.1 | 72.7 | 47.256
2006-09-01 | 6:46 | 14.13 | 7.5 | 8.66 | 84.3 | 72.6 | 47.197
2006-09-01 | 7:01 | 14.11 | 7.51 | 8.68 | 84.4 | 72.6 | 47.164
2006-09-01 | 7:16 | 14.12 | 7.52 | 8.75 | 85.1 | 72.4 | 47.09
2006-09-01 | 7:31 | 14.17 | 7.54 | 8.82 | 85.9 | 72.3 | 46.988
2006-09-01 | 7:46 | 14.21 | 7.56 | 8.92 | 87 | 72.2 | 46.928
2006-09-01 | 8:01 | 14.26 | 7.59 | 9.02 | 88 | 72.1 | 46.885
2006-09-01 | 8:16 | 14.33 | 7.62 | 9.12 | 89.2 | 72.1 | 46.845
2006-09-01 | 8:31 | 14.42 | 7.66 | 9.22 | 90.3 | 71.9 | 46.751
2006-09-01 | 8:46 | 14.52 | 7.69 | 9.32 | 91.5 | 71.9 | 46.708
2006-09-01 | 9:01 | 14.63 | 7.73 | 9.43 | 92.8 | 71.8 | 46.658
2006-09-01 | 9:16 | 14.76 | 7.79 | 9.54 | 94.1 | 71.7 | 46.584
2006-09-01 | 9:31 | 14.89 | 7.84 | 9.63 | 95.3 | 71.6 | 46.563
2006-09-01 | 9:46 | 15.04 | 7.89 | 9.73 | 96.6 | 71.5 | 46.478
2006-09-01 | 10:01 | 15.2 | 7.95 | 9.82 | 97.8 | 71.5 | 46.444
2006-09-01 | 10:16 | 15.38 | 8.03 | 9.9 | 98.9 | 71.3 | 46.368
2006-09-01 | 10:31 | 15.55 | 8.09 | 9.97 | 100.1 | 71.3 | 46.338
2006-09-01 | 10:46 | 15.73 | 8.16 | 10.03 | 101 | 71.3 | 46.346
2006-09-01 | 11:01 | 15.92 | 8.23 | 10.08 | 101.9 | 71.2 | 46.311
2006-09-01 | 11:16 | 16.11 | 8.3 | 10.12 | 102.8 | 71.3 | 46.342
2006-09-01 | 11:31 | 16.3 | 8.36 | 10.14 | 103.4 | 71.3 | 46.347
2006-09-01 | 11:46 | 16.49 | 8.41 | 10.18 | 104.2 | 71.3 | 46.356
2006-09-01 | 12:01 | 16.68 | 8.47 | 10.2 | 104.8 | 71.3 | 46.374
2006-09-01 | 12:16 | 16.88 | 8.52 | 10.21 | 105.4 | 71.4 | 46.44
2006-09-01 | 12:31 | 17.05 | 8.56 | 10.21 | 105.7 | 71.5 | 46.453
2006-09-01 | 12:46 | 17.24 | 8.6 | 10.21 | 106.2 | 71.5 | 46.463
2006-09-01 | 13:00 | 17.42 | 8.64 | 10.2 | 106.5 | 71.5 | 46.494
2006-09-01 | 13:15 | 17.59 | 8.68 | 10.2 | 106.8 | 71.5 | 46.493
2006-09-01 | 13:30 | 17.74 | 8.71 | 10.19 | 107 | 71.5 | 46.452
2006-09-01 | 13:45 | 17.9 | 8.74 | 10.15 | 107 | 71.4 | 46.415
2006-09-01 | 14:00 | 18.03 | 8.77 | 10.15 | 107.3 | 71.4 | 46.407
2006-09-01 | 14:15 | 18.17 | 8.79 | 10.12 | 107.2 | 71.4 | 46.4
2006-09-01 | 14:30 | 18.3 | 8.81 | 10.09 | 107.3 | 71.3 | 46.327
2006-09-01 | 14:45 | 18.4 | 8.83 | 10.07 | 107.2 | 71.3 | 46.317
2006-09-01 | 15:00 | 18.5 | 8.85 | 10.03 | 107.1 | 71.2 | 46.298
2006-09-01 | 15:15 | 18.6 | 8.86 | 10 | 107 | 71.1 | 46.234
2006-09-01 | 15:30 | 18.67 | 8.88 | 9.97 | 106.7 | 71.2 | 46.267
2006-09-01 | 15:45 | 18.75 | 8.89 | 9.91 | 106.3 | 71.1 | 46.225
2006-09-01 | 16:00 | 18.81 | 8.91 | 9.86 | 105.9 | 71.1 | 46.188
2006-09-01 | 16:15 | 18.86 | 8.92 | 9.84 | 105.8 | 71.1 | 46.198
2006-09-01 | 16:30 | 18.89 | 8.92 | 9.76 | 105 | 71.1 | 46.238
2006-09-01 | 16:45 | 18.9 | 8.92 | 9.73 | 104.7 | 71.1 | 46.232
2006-09-01 | 17:00 | 18.9 | 8.92 | 9.65 | 103.9 | 71.1 | 46.244
2006-09-01 | 17:15 | 18.87 | 8.92 | 9.6 | 103.3 | 71.2 | 46.26
2006-09-01 | 17:30 | 18.83 | 8.91 | 9.55 | 102.6 | 71.2 | 46.302
2006-09-01 | 17:45 | 18.78 | 8.9 | 9.48 | 101.8 | 71.3 | 46.317
2006-09-01 | 18:00 | 18.72 | 8.89 | 9.41 | 100.9 | 71.3 | 46.35
2006-09-01 | 18:15 | 18.62 | 8.87 | 9.32 | 99.7 | 71.4 | 46.42
2006-09-01 | 18:30 | 18.51 | 8.85 | 9.24 | 98.6 | 71.4 | 46.408
2006-09-01 | 18:45 | 18.38 | 8.81 | 9.13 | 97.2 | 71.5 | 46.46
2006-09-01 | 19:00 | 18.26 | 8.77 | 9.03 | 95.9 | 71.5 | 46.497
2006-09-01 | 19:15 | 18.13 | 8.71 | 8.92 | 94.5 | 71.6 | 46.536
2006-09-01 | 19:30 | 18.01 | 8.66 | 8.84 | 93.4 | 71.7 | 46.611
2006-09-01 | 19:45 | 17.88 | 8.59 | 8.72 | 91.9 | 71.9 | 46.71
2006-09-01 | 20:00 | 17.74 | 8.51 | 8.64 | 90.8 | 71.9 | 46.737
2006-09-01 | 20:15 | 17.6 | 8.42 | 8.56 | 89.7 | 72.1 | 46.833
2006-09-01 | 20:30 | 17.46 | 8.32 | 8.51 | 88.9 | 72.1 | 46.881
2006-09-01 | 20:45 | 17.32 | 8.23 | 8.47 | 88.2 | 72.3 | 46.991
2006-09-01 | 21:00 | 17.17 | 8.13 | 8.41 | 87.4 | 72.4 | 47.032
2006-09-01 | 21:15 | 17.03 | 8.04 | 8.37 | 86.7 | 72.5 | 47.152
2006-09-01 | 21:30 | 16.88 | 7.97 | 8.35 | 86.2 | 72.6 | 47.185
2006-09-01 | 21:45 | 16.73 | 7.9 | 8.34 | 85.8 | 72.7 | 47.271
2006-09-01 | 22:00 | 16.57 | 7.84 | 8.32 | 85.3 | 72.8 | 47.307
2006-09-01 | 22:15 | 16.43 | 7.79 | 8.31 | 85 | 72.9 | 47.356
2006-09-01 | 22:30 | 16.29 | 7.75 | 8.31 | 84.7 | 72.9 | 47.4
2006-09-01 | 22:45 | 16.16 | 7.71 | 8.31 | 84.5 | 73 | 47.465
2006-09-01 | 23:00 | 16.03 | 7.69 | 8.3 | 84.2 | 73 | 47.46
2006-09-01 | 23:15 | 15.9 | 7.66 | 8.3 | 83.9 | 73 | 47.461
2006-09-01 | 23:30 | 15.78 | 7.64 | 8.3 | 83.7 | 73 | 47.435
2006-09-01 | 23:45 | 15.67 | 7.62 | 8.31 | 83.7 | 72.9 | 47.385
2006-09-02 | 0:00 | 15.57 | 7.61 | 8.33 | 83.6 | 72.9 | 47.364
2006-09-02 | 0:15 | 15.47 | 7.59 | 8.33 | 83.4 | 72.8 | 47.34
2006-09-02 | 0:30 | 15.37 | 7.58 | 8.35 | 83.5 | 72.7 | 47.268
2006-09-02 | 0:45 | 15.28 | 7.57 | 8.35 | 83.3 | 72.7 | 47.234
2006-09-02 | 1:01 | 15.2 | 7.56 | 8.36 | 83.3 | 72.6 | 47.189
2006-09-02 | 1:16 | 15.12 | 7.55 | 8.37 | 83.2 | 72.6 | 47.175
2006-09-02 | 1:31 | 15.04 | 7.55 | 8.38 | 83.2 | 72.6 | 47.168
2006-09-02 | 1:46 | 14.97 | 7.54 | 8.4 | 83.3 | 72.5 | 47.138
2006-09-02 | 2:01 | 14.9 | 7.54 | 8.4 | 83.2 | 72.5 | 47.097
2006-09-02 | 2:16 | 14.84 | 7.53 | 8.41 | 83.1 | 72.4 | 47.075
2006-09-02 | 2:31 | 14.79 | 7.53 | 8.43 | 83.2 | 72.4 | 47.088
2006-09-02 | 2:46 | 14.74 | 7.53 | 8.44 | 83.2 | 72.5 | 47.104
2006-09-02 | 3:01 | 14.7 | 7.52 | 8.45 | 83.2 | 72.4 | 47.062
2006-09-02 | 3:16 | 14.66 | 7.52 | 8.44 | 83.1 | 72.4 | 47.058
2006-09-02 | 3:31 | 14.63 | 7.52 | 8.46 | 83.3 | 72.4 | 47.036
2006-09-02 | 3:46 | 14.6 | 7.51 | 8.45 | 83.1 | 72.4 | 47.037
2006-09-02 | 4:01 | 14.58 | 7.51 | 8.44 | 83 | 72.4 | 47.037
2006-09-02 | 4:16 | 14.55 | 7.51 | 8.46 | 83.1 | 72.3 | 47.021
2006-09-02 | 4:31 | 14.54 | 7.51 | 8.46 | 83.1 | 72.4 | 47.05
2006-09-02 | 4:46 | 14.53 | 7.5 | 8.46 | 83.1 | 72.3 | 47
2006-09-02 | 5:01 | 14.5 | 7.5 | 8.46 | 83 | 72.3 | 47.009
2006-09-02 | 5:16 | 14.5 | 7.5 | 8.48 | 83.2 | 72.3 | 46.968
2006-09-02 | 5:31 | 14.48 | 7.5 | 8.46 | 82.9 | 72.3 | 46.964
2006-09-02 | 5:46 | 14.49 | 7.49 | 8.46 | 83 | 72.3 | 46.965
2006-09-02 | 6:01 | 14.46 | 7.49 | 8.44 | 82.7 | 72.2 | 46.952
2006-09-02 | 6:16 | 14.45 | 7.49 | 8.46 | 82.9 | 72.1 | 46.86
2006-09-02 | 6:31 | 14.45 | 7.49 | 8.46 | 82.9 | 72.1 | 46.839
2006-09-02 | 6:46 | 14.43 | 7.5 | 8.48 | 83.1 | 72 | 46.799
2006-09-02 | 7:01 | 14.43 | 7.5 | 8.53 | 83.5 | 71.9 | 46.737
2006-09-02 | 7:16 | 14.43 | 7.52 | 8.58 | 84.1 | 71.9 | 46.707
2006-09-02 | 7:31 | 14.45 | 7.53 | 8.67 | 84.9 | 71.8 | 46.676
2006-09-02 | 7:46 | 14.49 | 7.55 | 8.75 | 85.9 | 71.7 | 46.594
2006-09-02 | 8:01 | 14.54 | 7.58 | 8.85 | 86.9 | 71.6 | 46.543
2006-09-02 | 8:16 | 14.61 | 7.61 | 8.98 | 88.3 | 71.6 | 46.509
2006-09-02 | 8:31 | 14.67 | 7.65 | 9.09 | 89.5 | 71.4 | 46.431
2006-09-02 | 8:46 | 14.76 | 7.69 | 9.21 | 90.8 | 71.4 | 46.424
2006-09-02 | 9:01 | 14.86 | 7.74 | 9.33 | 92.3 | 71.4 | 46.387
2006-09-02 | 9:16 | 14.98 | 7.78 | 9.45 | 93.6 | 71.3 | 46.32
2006-09-02 | 9:31 | 15.12 | 7.84 | 9.55 | 94.9 | 71.2 | 46.264
2006-09-02 | 9:46 | 15.27 | 7.9 | 9.64 | 96.2 | 71.2 | 46.248
2006-09-02 | 10:01 | 15.44 | 7.97 | 9.75 | 97.6 | 71.1 | 46.233
2006-09-02 | 10:16 | 15.62 | 8.04 | 9.84 | 98.9 | 71 | 46.159
2006-09-02 | 10:31 | 15.79 | 8.11 | 9.9 | 99.9 | 71 | 46.151
2006-09-02 | 10:46 | 15.99 | 8.19 | 9.96 | 100.9 | 71 | 46.118
2006-09-02 | 11:01 | 16.18 | 8.27 | 10.02 | 101.9 | 70.9 | 46.103
2006-09-02 | 11:16 | 16.38 | 8.35 | 10.07 | 102.9 | 70.9 | 46.099
2006-09-02 | 11:31 | 16.57 | 8.41 | 10.11 | 103.7 | 70.9 | 46.088
2006-09-02 | 11:46 | 16.76 | 8.47 | 10.15 | 104.5 | 71 | 46.143
2006-09-02 | 12:01 | 16.97 | 8.53 | 10.17 | 105.2 | 71 | 46.141
2006-09-02 | 12:16 | 17.16 | 8.59 | 10.18 | 105.7 | 71 | 46.155
2006-09-02 | 12:31 | 17.37 | 8.63 | 10.2 | 106.3 | 70.9 | 46.083
2006-09-02 | 12:46 | 17.57 | 8.67 | 10.19 | 106.7 | 71 | 46.126
2006-09-02 | 13:00 | 17.77 | 8.72 | 10.18 | 107 | 70.9 | 46.108
2006-09-02 | 13:15 | 17.93 | 8.76 | 10.19 | 107.5 | 71.1 | 46.185
2006-09-02 | 13:30 | 18.1 | 8.79 | 10.16 | 107.6 | 71.1 | 46.205
2006-09-02 | 13:45 | 18.27 | 8.82 | 10.14 | 107.7 | 71.1 | 46.184
2006-09-02 | 14:00 | 18.42 | 8.84 | 10.12 | 107.8 | 71.1 | 46.189
2006-09-02 | 14:15 | 18.55 | 8.87 | 10.08 | 107.7 | 71.1 | 46.213
2006-09-02 | 14:30 | 18.68 | 8.89 | 10.07 | 107.8 | 71.1 | 46.211
2006-09-02 | 14:45 | 18.78 | 8.91 | 10.04 | 107.7 | 71.1 | 46.209
2006-09-02 | 15:00 | 18.87 | 8.93 | 9.99 | 107.4 | 71.1 | 46.203
2006-09-02 | 15:15 | 18.95 | 8.94 | 9.98 | 107.4 | 71.1 | 46.197
2006-09-02 | 15:30 | 19 | 8.95 | 9.94 | 107.2 | 71 | 46.175
2006-09-02 | 15:45 | 19.07 | 8.97 | 9.88 | 106.7 | 71 | 46.168
2006-09-02 | 16:00 | 19.1 | 8.98 | 9.84 | 106.3 | 71 | 46.153
2006-09-02 | 16:15 | 19.13 | 8.99 | 9.82 | 106.2 | 71 | 46.142
2006-09-02 | 16:30 | 19.15 | 9 | 9.76 | 105.5 | 71 | 46.126
2006-09-02 | 16:45 | 19.15 | 9 | 9.7 | 104.9 | 70.9 | 46.09
2006-09-02 | 17:00 | 19.14 | 8.99 | 9.63 | 104.1 | 70.9 | 46.101
2006-09-02 | 17:15 | 19.11 | 8.99 | 9.58 | 103.5 | 70.9 | 46.095
2006-09-02 | 17:30 | 19.07 | 8.99 | 9.5 | 102.5 | 70.9 | 46.091
2006-09-02 | 17:45 | 18.98 | 8.97 | 9.4 | 101.3 | 70.9 | 46.087
2006-09-02 | 18:00 | 18.92 | 8.95 | 9.32 | 100.3 | 70.9 | 46.112
2006-09-02 | 18:15 | 18.86 | 8.94 | 9.27 | 99.6 | 70.9 | 46.101
2006-09-02 | 18:30 | 18.79 | 8.93 | 9.21 | 98.9 | 70.9 | 46.094
2006-09-02 | 18:45 | 18.67 | 8.9 | 9.09 | 97.4 | 71 | 46.132
2006-09-02 | 19:00 | 18.56 | 8.86 | 8.99 | 96.1 | 71 | 46.157
2006-09-02 | 19:15 | 18.43 | 8.82 | 8.88 | 94.7 | 71.1 | 46.21
2006-09-02 | 19:30 | 18.33 | 8.77 | 8.76 | 93.2 | 71.1 | 46.239
2006-09-02 | 19:45 | 18.22 | 8.7 | 8.65 | 91.8 | 71.2 | 46.272
2006-09-02 | 20:00 | 18.12 | 8.63 | 8.54 | 90.4 | 71.4 | 46.391
2006-09-02 | 20:15 | 17.99 | 8.55 | 8.45 | 89.2 | 71.4 | 46.438
2006-09-02 | 20:30 | 17.88 | 8.46 | 8.38 | 88.3 | 71.6 | 46.542
2006-09-02 | 20:45 | 17.75 | 8.36 | 8.32 | 87.5 | 71.7 | 46.635
2006-09-02 | 21:00 | 17.63 | 8.26 | 8.27 | 86.7 | 71.9 | 46.764
2006-09-02 | 21:15 | 17.49 | 8.17 | 8.22 | 86 | 72.1 | 46.866
2006-09-02 | 21:30 | 17.36 | 8.07 | 8.19 | 85.4 | 72.2 | 46.953
2006-09-02 | 21:45 | 17.22 | 7.98 | 8.16 | 84.9 | 72.4 | 47.03
2006-09-02 | 22:00 | 17.1 | 7.91 | 8.14 | 84.4 | 72.5 | 47.138
2006-09-02 | 22:15 | 16.98 | 7.85 | 8.12 | 84 | 72.7 | 47.271
2006-09-02 | 22:30 | 16.85 | 7.8 | 8.11 | 83.7 | 72.8 | 47.331
2006-09-02 | 22:45 | 16.72 | 7.76 | 8.1 | 83.3 | 72.9 | 47.365
2006-09-02 | 23:00 | 16.59 | 7.72 | 8.1 | 83.1 | 72.9 | 47.387
2006-09-02 | 23:15 | 16.47 | 7.69 | 8.1 | 82.9 | 72.9 | 47.399
2006-09-02 | 23:30 | 16.36 | 7.66 | 8.1 | 82.7 | 72.9 | 47.379
2006-09-02 | 23:45 | 16.25 | 7.64 | 8.1 | 82.5 | 72.9 | 47.361
2006-09-03 | 0:00 | 16.15 | 7.62 | 8.12 | 82.5 | 72.8 | 47.313
2006-09-03 | 0:15 | 16.05 | 7.61 | 8.12 | 82.4 | 72.8 | 47.314
2006-09-03 | 0:30 | 15.95 | 7.59 | 8.14 | 82.4 | 72.8 | 47.298
2006-09-03 | 0:45 | 15.86 | 7.58 | 8.15 | 82.3 | 72.6 | 47.208
2006-09-03 | 1:01 | 15.78 | 7.57 | 8.17 | 82.3 | 72.5 | 47.143
2006-09-03 | 1:16 | 15.7 | 7.56 | 8.18 | 82.3 | 72.5 | 47.117
2006-09-03 | 1:31 | 15.64 | 7.56 | 8.2 | 82.4 | 72.4 | 47.058
2006-09-03 | 1:46 | 15.57 | 7.55 | 8.21 | 82.4 | 72.3 | 47.007
2006-09-03 | 2:01 | 15.51 | 7.54 | 8.21 | 82.3 | 72.3 | 46.964
2006-09-03 | 2:16 | 15.45 | 7.54 | 8.22 | 82.3 | 72.2 | 46.957
2006-09-03 | 2:31 | 15.41 | 7.54 | 8.23 | 82.3 | 72.2 | 46.898
2006-09-03 | 2:46 | 15.38 | 7.53 | 8.22 | 82.2 | 72.1 | 46.871
2006-09-03 | 3:01 | 15.35 | 7.53 | 8.23 | 82.2 | 72.1 | 46.897
2006-09-03 | 3:16 | 15.33 | 7.52 | 8.23 | 82.2 | 72 | 46.805
2006-09-03 | 3:31 | 15.3 | 7.52 | 8.22 | 82.1 | 72 | 46.788
2006-09-03 | 3:46 | 15.27 | 7.51 | 8.22 | 82 | 71.9 | 46.73
2006-09-03 | 4:01 | 15.27 | 7.51 | 8.22 | 82 | 71.8 | 46.649
2006-09-03 | 4:16 | 15.24 | 7.5 | 8.22 | 81.9 | 71.8 | 46.641
2006-09-03 | 4:31 | 15.23 | 7.5 | 8.22 | 81.9 | 71.7 | 46.609
2006-09-03 | 4:46 | 15.22 | 7.5 | 8.22 | 81.9 | 71.6 | 46.544
2006-09-03 | 5:01 | 15.2 | 7.49 | 8.22 | 81.8 | 71.6 | 46.528
2006-09-03 | 5:16 | 15.19 | 7.5 | 8.21 | 81.8 | 71.5 | 46.482
2006-09-03 | 5:31 | 15.18 | 7.49 | 8.2 | 81.6 | 71.4 | 46.436
2006-09-03 | 5:46 | 15.17 | 7.49 | 8.2 | 81.6 | 71.5 | 46.471
2006-09-03 | 6:01 | 15.17 | 7.5 | 8.21 | 81.8 | 71.4 | 46.409
2006-09-03 | 6:16 | 15.17 | 7.49 | 8.2 | 81.6 | 71.4 | 46.403
2006-09-03 | 6:31 | 15.17 | 7.49 | 8.21 | 81.7 | 71.4 | 46.442
2006-09-03 | 6:46 | 15.16 | 7.49 | 8.23 | 81.9 | 71.4 | 46.389
2006-09-03 | 7:01 | 15.15 | 7.49 | 8.26 | 82.2 | 71.4 | 46.398
2006-09-03 | 7:16 | 15.16 | 7.5 | 8.32 | 82.8 | 71.3 | 46.364
2006-09-03 | 7:31 | 15.2 | 7.53 | 8.4 | 83.6 | 71.3 | 46.365
2006-09-03 | 7:46 | 15.24 | 7.54 | 8.49 | 84.7 | 71.3 | 46.362
2006-09-03 | 8:01 | 15.28 | 7.57 | 8.58 | 85.6 | 71.3 | 46.361
2006-09-03 | 8:16 | 15.35 | 7.6 | 8.7 | 86.9 | 71.4 | 46.378
2006-09-03 | 8:31 | 15.43 | 7.63 | 8.82 | 88.3 | 71.3 | 46.349
2006-09-03 | 8:46 | 15.52 | 7.68 | 8.95 | 89.7 | 71.3 | 46.317
2006-09-03 | 9:01 | 15.64 | 7.73 | 9.07 | 91.2 | 71.3 | 46.317
2006-09-03 | 9:16 | 15.7 | 7.77 | 9.15 | 92.1 | 71.2 | 46.308
2006-09-03 | 9:31 | 15.84 | 7.82 | 9.27 | 93.6 | 71.2 | 46.276
2006-09-03 | 9:46 | 16 | 7.89 | 9.36 | 94.9 | 71.2 | 46.248
2006-09-03 | 10:01 | 16.16 | 7.97 | 9.51 | 96.7 | 71.2 | 46.284
2006-09-03 | 10:16 | 16.23 | 8.02 | 9.56 | 97.4 | 71.1 | 46.205
2006-09-03 | 10:31 | 16.42 | 8.1 | 9.65 | 98.7 | 71 | 46.178
2006-09-03 | 10:46 | 16.6 | 8.18 | 9.74 | 99.9 | 71.1 | 46.183
2006-09-03 | 11:01 | 16.81 | 8.27 | 9.81 | 101.1 | 70.9 | 46.077
2006-09-03 | 11:16 | 16.9 | 8.34 | 9.86 | 101.8 | 70.9 | 46.117
2006-09-03 | 11:31 | 16.96 | 8.37 | 9.85 | 101.9 | 70.9 | 46.111
2006-09-03 | 11:46 | 17.07 | 8.43 | 9.88 | 102.4 | 70.7 | 45.968
2006-09-03 | 12:01 | 17.09 | 8.48 | 9.87 | 102.3 | 70.8 | 46.005
2006-09-03 | 12:16 | 17.13 | 8.5 | 9.84 | 102.1 | 70.8 | 46.018
2006-09-03 | 12:31 | 17.14 | 8.51 | 9.8 | 101.7 | 70.8 | 46.043
2006-09-03 | 12:46 | 17.23 | 8.55 | 9.79 | 101.8 | 71 | 46.127
2006-09-03 | 13:00 | 17.3 | 8.57 | 9.79 | 101.9 | 70.8 | 46.035
2006-09-03 | 13:15 | 17.55 | 8.63 | 9.85 | 103.1 | 71 | 46.141
2006-09-03 | 13:30 | 17.58 | 8.66 | 9.83 | 103 | 71 | 46.181
2006-09-03 | 13:45 | 17.63 | 8.68 | 9.84 | 103.2 | 71.1 | 46.207
2006-09-03 | 14:00 | 17.85 | 8.71 | 9.86 | 103.8 | 71.2 | 46.274
2006-09-03 | 14:15 | 18.07 | 8.78 | 9.94 | 105.1 | 71.1 | 46.242
2006-09-03 | 14:30 | 18.22 | 8.8 | 9.91 | 105.2 | 71.2 | 46.29
2006-09-03 | 14:45 | 18.34 | 8.83 | 9.91 | 105.5 | 71.3 | 46.372
2006-09-03 | 15:00 | 18.45 | 8.85 | 9.94 | 105.9 | 71.4 | 46.407
2006-09-03 | 15:15 | 18.51 | 8.87 | 9.92 | 105.9 | 71.5 | 46.462
2006-09-03 | 15:30 | 18.62 | 8.89 | 9.89 | 105.8 | 71.5 | 46.45
2006-09-03 | 15:45 | 18.72 | 8.91 | 9.87 | 105.8 | 71.5 | 46.502
2006-09-03 | 16:00 | 18.79 | 8.93 | 9.87 | 105.9 | 71.6 | 46.549
2006-09-03 | 16:15 | 18.84 | 8.94 | 9.82 | 105.6 | 71.6 | 46.543
2006-09-03 | 16:30 | 18.89 | 8.95 | 9.78 | 105.2 | 71.7 | 46.582
2006-09-03 | 16:45 | 18.92 | 8.96 | 9.72 | 104.7 | 71.7 | 46.632
2006-09-03 | 17:00 | 18.95 | 8.97 | 9.67 | 104.2 | 71.8 | 46.653
2006-09-03 | 17:15 | 18.96 | 8.97 | 9.61 | 103.5 | 71.8 | 46.661
2006-09-03 | 17:30 | 18.96 | 8.97 | 9.54 | 102.8 | 71.8 | 46.697
2006-09-03 | 17:45 | 18.94 | 8.97 | 9.46 | 101.9 | 71.9 | 46.723
2006-09-03 | 18:00 | 18.91 | 8.96 | 9.38 | 101 | 71.9 | 46.741
2006-09-03 | 18:15 | 18.87 | 8.94 | 9.28 | 99.8 | 72 | 46.799
2006-09-03 | 18:30 | 18.79 | 8.92 | 9.19 | 98.7 | 72 | 46.829
2006-09-03 | 18:45 | 18.72 | 8.89 | 9.04 | 96.9 | 72.2 | 46.902
2006-09-03 | 19:00 | 18.66 | 8.84 | 8.91 | 95.4 | 72.1 | 46.859
2006-09-03 | 19:15 | 18.59 | 8.79 | 8.78 | 93.9 | 72.3 | 46.967
2006-09-03 | 19:30 | 18.52 | 8.73 | 8.65 | 92.4 | 72.3 | 47.007
2006-09-03 | 19:45 | 18.4 | 8.66 | 8.55 | 91 | 72.5 | 47.112
2006-09-03 | 20:00 | 18.3 | 8.58 | 8.44 | 89.7 | 72.5 | 47.146
2006-09-03 | 20:15 | 18.19 | 8.49 | 8.33 | 88.3 | 72.7 | 47.227
2006-09-03 | 20:30 | 18.06 | 8.39 | 8.27 | 87.5 | 72.8 | 47.332
2006-09-03 | 20:45 | 17.93 | 8.29 | 8.23 | 86.8 | 72.9 | 47.362
2006-09-03 | 21:00 | 17.82 | 8.19 | 8.16 | 85.9 | 73 | 47.48
2006-09-03 | 21:15 | 17.69 | 8.09 | 8.11 | 85.1 | 73.1 | 47.524
2006-09-03 | 21:30 | 17.59 | 8.01 | 8.08 | 84.6 | 73.3 | 47.622
2006-09-03 | 21:45 | 17.47 | 7.94 | 8.06 | 84.3 | 73.4 | 47.73
2006-09-03 | 22:00 | 17.37 | 7.87 | 8.02 | 83.6 | 73.6 | 47.834
2006-09-03 | 22:15 | 17.27 | 7.81 | 7.98 | 83.1 | 73.7 | 47.918
2006-09-03 | 22:30 | 17.17 | 7.77 | 7.99 | 83 | 73.8 | 47.969
2006-09-03 | 22:45 | 17.06 | 7.73 | 7.97 | 82.6 | 73.9 | 48.023
2006-09-03 | 23:00 | 16.95 | 7.69 | 7.96 | 82.3 | 73.9 | 48.046
2006-09-03 | 23:15 | 16.87 | 7.67 | 7.96 | 82.1 | 73.9 | 48.062
2006-09-03 | 23:30 | 16.78 | 7.65 | 7.95 | 81.9 | 73.9 | 48.044
2006-09-03 | 23:45 | 16.68 | 7.63 | 7.96 | 81.8 | 73.8 | 48.002
2006-09-04 | 0:00 | 16.58 | 7.61 | 7.96 | 81.7 | 73.9 | 48.051
2006-09-04 | 0:15 | 16.47 | 7.6 | 7.98 | 81.7 | 73.8 | 47.975
2006-09-04 | 0:30 | 16.38 | 7.58 | 8 | 81.7 | 73.7 | 47.897
2006-09-04 | 0:45 | 16.3 | 7.57 | 8.01 | 81.6 | 73.6 | 47.837
2006-09-04 | 1:01 | 16.23 | 7.57 | 8.01 | 81.5 | 73.6 | 47.836
2006-09-04 | 1:16 | 16.15 | 7.56 | 8.02 | 81.5 | 73.5 | 47.744
2006-09-04 | 1:31 | 16.08 | 7.55 | 8.04 | 81.6 | 73.4 | 47.69
2006-09-04 | 1:46 | 16.02 | 7.54 | 8.04 | 81.4 | 73.3 | 47.664
2006-09-04 | 2:01 | 15.96 | 7.54 | 8.05 | 81.5 | 73.3 | 47.634
2006-09-04 | 2:16 | 15.9 | 7.53 | 8.07 | 81.6 | 73.2 | 47.586
2006-09-04 | 2:31 | 15.85 | 7.53 | 8.07 | 81.5 | 73.2 | 47.578
2006-09-04 | 2:46 | 15.81 | 7.52 | 8.07 | 81.5 | 73.1 | 47.529
2006-09-04 | 3:01 | 15.77 | 7.52 | 8.08 | 81.5 | 73 | 47.479
2006-09-04 | 3:16 | 15.73 | 7.51 | 8.09 | 81.5 | 73.1 | 47.486
2006-09-04 | 3:31 | 15.68 | 7.51 | 8.1 | 81.5 | 73 | 47.436
2006-09-04 | 3:46 | 15.64 | 7.51 | 8.1 | 81.4 | 72.9 | 47.395
2006-09-04 | 4:01 | 15.59 | 7.5 | 8.11 | 81.5 | 73 | 47.419
2006-09-04 | 4:16 | 15.56 | 7.51 | 8.12 | 81.5 | 72.8 | 47.345
2006-09-04 | 4:31 | 15.5 | 7.5 | 8.13 | 81.5 | 72.8 | 47.329
2006-09-04 | 4:46 | 15.46 | 7.5 | 8.14 | 81.5 | 72.8 | 47.311
2006-09-04 | 5:01 | 15.43 | 7.5 | 8.14 | 81.5 | 72.7 | 47.254
2006-09-04 | 5:16 | 15.39 | 7.5 | 8.16 | 81.6 | 72.7 | 47.252
2006-09-04 | 5:31 | 15.36 | 7.5 | 8.16 | 81.6 | 72.6 | 47.215
2006-09-04 | 5:46 | 15.33 | 7.49 | 8.17 | 81.6 | 72.6 | 47.196
2006-09-04 | 6:01 | 15.31 | 7.5 | 8.16 | 81.5 | 72.6 | 47.175
2006-09-04 | 6:16 | 15.29 | 7.49 | 8.17 | 81.5 | 72.6 | 47.17
2006-09-04 | 6:31 | 15.26 | 7.49 | 8.17 | 81.5 | 72.5 | 47.138
2006-09-04 | 6:46 | 15.26 | 7.5 | 8.2 | 81.8 | 72.5 | 47.12
2006-09-04 | 7:01 | 15.26 | 7.5 | 8.23 | 82 | 72.5 | 47.11
2006-09-04 | 7:16 | 15.27 | 7.51 | 8.29 | 82.7 | 72.4 | 47.036
2006-09-04 | 7:31 | 15.29 | 7.53 | 8.37 | 83.5 | 72.3 | 47.004
2006-09-04 | 7:46 | 15.33 | 7.55 | 8.46 | 84.5 | 72.3 | 46.98
2006-09-04 | 8:01 | 15.39 | 7.58 | 8.58 | 85.8 | 72.1 | 46.877
2006-09-04 | 8:16 | 15.46 | 7.61 | 8.69 | 87 | 72.1 | 46.846
2006-09-04 | 8:31 | 15.55 | 7.65 | 8.81 | 88.4 | 72 | 46.775
2006-09-04 | 8:46 | 15.64 | 7.69 | 8.93 | 89.8 | 71.9 | 46.724
2006-09-04 | 9:01 | 15.77 | 7.74 | 9.05 | 91.2 | 71.8 | 46.657
2006-09-04 | 9:16 | 15.89 | 7.8 | 9.16 | 92.6 | 71.7 | 46.604
2006-09-04 | 9:31 | 16.04 | 7.85 | 9.28 | 94.1 | 71.6 | 46.542
2006-09-04 | 9:46 | 16.2 | 7.91 | 9.38 | 95.5 | 71.5 | 46.484
2006-09-04 | 10:01 | 16.37 | 7.98 | 9.49 | 96.9 | 71.4 | 46.438
2006-09-04 | 10:16 | 16.55 | 8.06 | 9.58 | 98.2 | 71.3 | 46.354
2006-09-04 | 10:31 | 16.75 | 8.15 | 9.67 | 99.5 | 71.2 | 46.295
2006-09-04 | 10:46 | 16.93 | 8.23 | 9.74 | 100.6 | 71.2 | 46.251
2006-09-04 | 11:01 | 17.13 | 8.31 | 9.79 | 101.6 | 71.1 | 46.192
2006-09-04 | 11:16 | 17.33 | 8.38 | 9.84 | 102.5 | 71 | 46.177
2006-09-04 | 11:31 | 17.54 | 8.46 | 9.9 | 103.6 | 71 | 46.131
2006-09-04 | 11:46 | 17.75 | 8.53 | 9.92 | 104.3 | 70.9 | 46.107
2006-09-04 | 12:01 | 17.97 | 8.59 | 9.96 | 105.2 | 70.9 | 46.081
2006-09-04 | 12:16 | 18.17 | 8.66 | 9.98 | 105.8 | 70.9 | 46.1
2006-09-04 | 12:31 | 18.39 | 8.71 | 9.99 | 106.3 | 70.9 | 46.094
2006-09-04 | 12:46 | 18.59 | 8.76 | 9.99 | 106.8 | 70.9 | 46.096
2006-09-04 | 13:00 | 18.77 | 8.8 | 9.99 | 107.2 | 70.9 | 46.098
2006-09-04 | 13:15 | 18.96 | 8.83 | 9.98 | 107.5 | 71 | 46.15
2006-09-04 | 13:30 | 19.14 | 8.86 | 9.96 | 107.7 | 71 | 46.14
2006-09-04 | 13:45 | 19.32 | 8.9 | 9.94 | 107.9 | 71 | 46.158
2006-09-04 | 14:00 | 19.47 | 8.93 | 9.92 | 108 | 71.1 | 46.204
2006-09-04 | 14:15 | 19.63 | 8.95 | 9.89 | 108 | 71.1 | 46.233
2006-09-04 | 14:30 | 19.77 | 8.97 | 9.86 | 108 | 71.2 | 46.279
2006-09-04 | 14:45 | 19.88 | 8.99 | 9.84 | 108 | 71.3 | 46.339
2006-09-04 | 15:00 | 20 | 9.01 | 9.79 | 107.7 | 71.3 | 46.36
2006-09-04 | 15:15 | 20.09 | 9.02 | 9.75 | 107.4 | 71.4 | 46.395
2006-09-04 | 15:30 | 20.17 | 9.03 | 9.71 | 107.1 | 71.4 | 46.43
2006-09-04 | 15:45 | 20.23 | 9.04 | 9.67 | 106.9 | 71.5 | 46.443
2006-09-04 | 16:00 | 20.29 | 9.05 | 9.62 | 106.4 | 71.5 | 46.444
2006-09-04 | 16:15 | 20.33 | 9.05 | 9.57 | 105.9 | 71.6 | 46.508
2006-09-04 | 16:30 | 20.36 | 9.07 | 9.49 | 105.1 | 71.5 | 46.473
2006-09-04 | 16:45 | 20.37 | 9.06 | 9.44 | 104.6 | 71.5 | 46.5
2006-09-04 | 17:00 | 20.37 | 9.06 | 9.39 | 104 | 71.6 | 46.517
2006-09-04 | 17:15 | 20.35 | 9.07 | 9.32 | 103.2 | 71.3 | 46.335
2006-09-04 | 17:30 | 20.31 | 9.06 | 9.24 | 102.3 | 71.4 | 46.381
2006-09-04 | 17:45 | 20.25 | 9.05 | 9.16 | 101.2 | 71.4 | 46.415
2006-09-04 | 18:00 | 20.18 | 9.03 | 9.07 | 100.2 | 71.4 | 46.429
2006-09-04 | 18:15 | 20.1 | 9.01 | 8.97 | 98.9 | 71.5 | 46.499
2006-09-04 | 18:30 | 20.01 | 8.99 | 8.88 | 97.7 | 71.5 | 46.488
2006-09-04 | 18:45 | 19.92 | 8.96 | 8.77 | 96.3 | 71.6 | 46.542
2006-09-04 | 19:00 | 19.82 | 8.92 | 8.65 | 94.8 | 71.6 | 46.568
2006-09-04 | 19:15 | 19.69 | 8.87 | 8.52 | 93.2 | 71.7 | 46.633
2006-09-04 | 19:30 | 19.58 | 8.81 | 8.39 | 91.6 | 71.8 | 46.702
2006-09-04 | 19:45 | 19.48 | 8.74 | 8.28 | 90.1 | 72 | 46.768
2006-09-04 | 20:00 | 19.36 | 8.67 | 8.19 | 89 | 72.1 | 46.839
2006-09-04 | 20:15 | 19.24 | 8.58 | 8.11 | 87.8 | 72.2 | 46.946
2006-09-04 | 20:30 | 19.11 | 8.49 | 8.01 | 86.5 | 72.4 | 47.05
2006-09-04 | 20:45 | 18.98 | 8.37 | 7.94 | 85.6 | 72.6 | 47.185
2006-09-04 | 21:00 | 18.85 | 8.27 | 7.89 | 84.8 | 72.8 | 47.34
2006-09-04 | 21:15 | 18.72 | 8.15 | 7.83 | 83.9 | 73 | 47.448
2006-09-04 | 21:30 | 18.59 | 8.06 | 7.81 | 83.5 | 73.2 | 47.559
2006-09-04 | 21:45 | 18.45 | 7.97 | 7.79 | 83.1 | 73.3 | 47.655
2006-09-04 | 22:00 | 18.32 | 7.9 | 7.77 | 82.6 | 73.6 | 47.815
2006-09-04 | 22:15 | 18.18 | 7.84 | 7.76 | 82.3 | 73.6 | 47.866
2006-09-04 | 22:30 | 18.04 | 7.79 | 7.73 | 81.7 | 73.7 | 47.928
2006-09-04 | 22:45 | 17.91 | 7.74 | 7.72 | 81.4 | 73.8 | 47.953
2006-09-04 | 23:00 | 17.78 | 7.71 | 7.72 | 81.2 | 73.8 | 47.97
2006-09-04 | 23:15 | 17.67 | 7.68 | 7.73 | 81.1 | 73.8 | 47.94
2006-09-04 | 23:30 | 17.55 | 7.65 | 7.72 | 80.8 | 73.7 | 47.904
2006-09-04 | 23:45 | 17.43 | 7.64 | 7.73 | 80.7 | 73.6 | 47.869
2006-09-05 | 0:00 | 17.33 | 7.61 | 7.75 | 80.7 | 73.6 | 47.834
2006-09-05 | 0:15 | 17.23 | 7.6 | 7.76 | 80.7 | 73.4 | 47.728
2006-09-05 | 0:30 | 17.13 | 7.59 | 7.76 | 80.5 | 73.4 | 47.708
2006-09-05 | 0:45 | 17.03 | 7.58 | 7.78 | 80.5 | 73.3 | 47.642
2006-09-05 | 1:01 | 16.94 | 7.56 | 7.8 | 80.6 | 73.2 | 47.581
2006-09-05 | 1:16 | 16.86 | 7.56 | 7.82 | 80.7 | 73.1 | 47.53
2006-09-05 | 1:31 | 16.78 | 7.55 | 7.84 | 80.7 | 73 | 47.431
2006-09-05 | 1:46 | 16.71 | 7.54 | 7.84 | 80.7 | 72.9 | 47.399
2006-09-05 | 2:01 | 16.64 | 7.54 | 7.86 | 80.7 | 72.9 | 47.379
2006-09-05 | 2:16 | 16.58 | 7.54 | 7.87 | 80.8 | 72.8 | 47.303
2006-09-05 | 2:31 | 16.53 | 7.53 | 7.88 | 80.8 | 72.7 | 47.257
2006-09-05 | 2:46 | 16.49 | 7.53 | 7.89 | 80.8 | 72.7 | 47.258
2006-09-05 | 3:01 | 16.45 | 7.53 | 7.9 | 80.8 | 72.6 | 47.19
2006-09-05 | 3:16 | 16.42 | 7.52 | 7.9 | 80.8 | 72.6 | 47.164
2006-09-05 | 3:31 | 16.39 | 7.52 | 7.91 | 80.8 | 72.5 | 47.122
2006-09-05 | 3:46 | 16.38 | 7.52 | 7.92 | 80.9 | 72.5 | 47.106
2006-09-05 | 4:01 | 16.36 | 7.51 | 7.91 | 80.8 | 72.4 | 47.057
2006-09-05 | 4:16 | 16.35 | 7.51 | 7.92 | 80.8 | 72.3 | 47.022
2006-09-05 | 4:31 | 16.35 | 7.51 | 7.91 | 80.8 | 72.2 | 46.929
2006-09-05 | 4:46 | 16.34 | 7.51 | 7.91 | 80.8 | 72.1 | 46.868
2006-09-05 | 5:01 | 16.32 | 7.5 | 7.9 | 80.6 | 72 | 46.828
2006-09-05 | 5:16 | 16.32 | 7.5 | 7.9 | 80.6 | 71.9 | 46.743
2006-09-05 | 5:31 | 16.31 | 7.5 | 7.9 | 80.5 | 71.8 | 46.689
2006-09-05 | 5:46 | 16.31 | 7.5 | 7.89 | 80.5 | 71.7 | 46.603
2006-09-05 | 6:01 | 16.31 | 7.5 | 7.9 | 80.6 | 71.6 | 46.51
2006-09-05 | 6:16 | 16.32 | 7.49 | 7.89 | 80.5 | 71.5 | 46.447
2006-09-05 | 6:31 | 16.31 | 7.49 | 7.89 | 80.4 | 71.4 | 46.399
2006-09-05 | 6:46 | 16.3 | 7.49 | 7.89 | 80.5 | 71.3 | 46.343
2006-09-05 | 7:01 | 16.32 | 7.49 | 7.94 | 80.9 | 71.2 | 46.266
2006-09-05 | 7:16 | 16.32 | 7.5 | 7.98 | 81.4 | 71.1 | 46.238
2006-09-05 | 7:31 | 16.35 | 7.52 | 8.05 | 82.2 | 71 | 46.175
2006-09-05 | 7:46 | 16.38 | 7.54 | 8.15 | 83.2 | 71 | 46.162
2006-09-05 | 8:01 | 16.44 | 7.56 | 8.26 | 84.4 | 71 | 46.129
2006-09-05 | 8:16 | 16.5 | 7.59 | 8.37 | 85.7 | 71 | 46.152
2006-09-05 | 8:31 | 16.6 | 7.63 | 8.5 | 87.2 | 70.9 | 46.085
2006-09-05 | 8:46 | 16.67 | 7.67 | 8.61 | 88.5 | 70.9 | 46.084
2006-09-05 | 9:01 | 16.78 | 7.71 | 8.75 | 90.1 | 70.9 | 46.086
2006-09-05 | 9:16 | 16.89 | 7.77 | 8.87 | 91.6 | 70.9 | 46.066
2006-09-05 | 9:31 | 17.02 | 7.83 | 8.98 | 93 | 70.9 | 46.074
2006-09-05 | 9:46 | 17.16 | 7.88 | 9.1 | 94.5 | 70.8 | 46.031
2006-09-05 | 10:01 | 17.33 | 7.95 | 9.22 | 96.1 | 70.8 | 45.991
2006-09-05 | 10:16 | 17.49 | 8.02 | 9.31 | 97.4 | 70.8 | 46
2006-09-05 | 10:31 | 17.67 | 8.1 | 9.41 | 98.7 | 70.7 | 45.965
2006-09-05 | 10:46 | 17.86 | 8.18 | 9.5 | 100.1 | 70.7 | 45.971
2006-09-05 | 11:01 | 18.06 | 8.27 | 9.57 | 101.2 | 70.7 | 45.93
2006-09-05 | 11:16 | 18.25 | 8.35 | 9.64 | 102.3 | 70.7 | 45.927
2006-09-05 | 11:31 | 18.45 | 8.42 | 9.68 | 103.2 | 70.6 | 45.92
2006-09-05 | 11:46 | 18.65 | 8.5 | 9.72 | 104.1 | 70.6 | 45.916
2006-09-05 | 12:01 | 18.84 | 8.56 | 9.75 | 104.7 | 70.7 | 45.947
2006-09-05 | 12:16 | 19.04 | 8.63 | 9.77 | 105.4 | 70.7 | 45.978
2006-09-05 | 12:31 | 19.24 | 8.68 | 9.78 | 106 | 70.8 | 45.995
2006-09-05 | 12:46 | 19.42 | 8.73 | 9.79 | 106.4 | 70.8 | 46.046
2006-09-05 | 13:00 | 19.61 | 8.77 | 9.8 | 106.9 | 70.9 | 46.089
2006-09-05 | 13:15 | 19.79 | 8.81 | 9.79 | 107.2 | 70.9 | 46.113
2006-09-05 | 13:30 | 19.96 | 8.86 | 9.79 | 107.6 | 71.1 | 46.183
2006-09-05 | 13:45 | 20.13 | 8.88 | 9.79 | 107.9 | 71.1 | 46.219
2006-09-05 | 14:00 | 20.28 | 8.92 | 9.76 | 107.9 | 71.1 | 46.237
2006-09-05 | 14:15 | 20.42 | 8.95 | 9.75 | 108.1 | 71.2 | 46.307
2006-09-05 | 14:30 | 20.55 | 8.97 | 9.71 | 108 | 71.3 | 46.338
2006-09-05 | 14:45 | 20.66 | 8.98 | 9.67 | 107.8 | 71.3 | 46.37
2006-09-05 | 15:00 | 20.76 | 9.01 | 9.65 | 107.8 | 71.4 | 46.427
2006-09-05 | 15:15 | 20.86 | 9.02 | 9.6 | 107.4 | 71.4 | 46.417
2006-09-05 | 15:30 | 20.93 | 9.03 | 9.55 | 106.9 | 71.5 | 46.481
2006-09-05 | 15:45 | 21 | 9.05 | 9.5 | 106.6 | 71.5 | 46.48
2006-09-05 | 16:00 | 21.05 | 9.05 | 9.47 | 106.3 | 71.5 | 46.488
2006-09-05 | 16:15 | 21.08 | 9.06 | 9.41 | 105.8 | 71.6 | 46.515
2006-09-05 | 16:30 | 21.12 | 9.07 | 9.34 | 105 | 71.7 | 46.576
2006-09-05 | 16:45 | 21.14 | 9.07 | 9.29 | 104.5 | 71.6 | 46.563
2006-09-05 | 17:00 | 21.15 | 9.07 | 9.23 | 103.8 | 71.8 | 46.648
2006-09-05 | 17:15 | 21.12 | 9.07 | 9.14 | 102.8 | 71.8 | 46.648
2006-09-05 | 17:30 | 21.08 | 9.06 | 9.06 | 101.9 | 71.9 | 46.708
2006-09-05 | 17:45 | 21.04 | 9.05 | 8.98 | 100.8 | 71.9 | 46.742
2006-09-05 | 18:00 | 20.97 | 9.04 | 8.89 | 99.6 | 72 | 46.815
2006-09-05 | 18:15 | 20.89 | 9.02 | 8.77 | 98.2 | 72.1 | 46.834
2006-09-05 | 18:30 | 20.8 | 9 | 8.66 | 96.7 | 72.2 | 46.909
2006-09-05 | 18:45 | 20.7 | 8.96 | 8.54 | 95.2 | 72.2 | 46.934
2006-09-05 | 19:00 | 20.6 | 8.91 | 8.42 | 93.7 | 72.3 | 46.998
2006-09-05 | 19:15 | 20.48 | 8.87 | 8.3 | 92.2 | 72.4 | 47.074
2006-09-05 | 19:30 | 20.36 | 8.8 | 8.18 | 90.6 | 72.5 | 47.146
2006-09-05 | 19:45 | 20.24 | 8.74 | 8.05 | 89 | 72.7 | 47.251
2006-09-05 | 20:00 | 20.11 | 8.65 | 7.93 | 87.4 | 72.9 | 47.364
2006-09-05 | 20:15 | 19.99 | 8.57 | 7.84 | 86.2 | 73 | 47.462
2006-09-05 | 20:30 | 19.86 | 8.47 | 7.78 | 85.4 | 73.3 | 47.622
2006-09-05 | 20:45 | 19.75 | 8.36 | 7.7 | 84.3 | 73.5 | 47.75
2006-09-05 | 21:00 | 19.61 | 8.25 | 7.66 | 83.6 | 73.6 | 47.846
2006-09-05 | 21:15 | 19.48 | 8.15 | 7.61 | 82.8 | 73.9 | 48.011
2006-09-05 | 21:30 | 19.35 | 8.05 | 7.56 | 82.1 | 74 | 48.128
2006-09-05 | 21:45 | 19.22 | 7.96 | 7.53 | 81.6 | 74.3 | 48.306
2006-09-05 | 22:00 | 19.08 | 7.88 | 7.52 | 81.2 | 74.4 | 48.37
2006-09-05 | 22:15 | 18.95 | 7.82 | 7.5 | 80.8 | 74.6 | 48.499
2006-09-05 | 22:30 | 18.8 | 7.77 | 7.49 | 80.4 | 74.7 | 48.545
2006-09-05 | 22:45 | 18.67 | 7.73 | 7.47 | 80 | 74.8 | 48.642
2006-09-05 | 23:00 | 18.54 | 7.7 | 7.49 | 80 | 74.9 | 48.675
2006-09-05 | 23:15 | 18.41 | 7.66 | 7.49 | 79.8 | 74.9 | 48.657
2006-09-05 | 23:30 | 18.29 | 7.64 | 7.5 | 79.7 | 74.7 | 48.583
2006-09-05 | 23:45 | 18.18 | 7.62 | 7.49 | 79.4 | 74.7 | 48.582
2006-09-06 | 0:00 | 18.07 | 7.6 | 7.5 | 79.3 | 74.6 | 48.514
2006-09-06 | 0:15 | 17.97 | 7.59 | 7.53 | 79.5 | 74.5 | 48.447
2006-09-06 | 0:30 | 17.86 | 7.58 | 7.54 | 79.5 | 74.4 | 48.39
2006-09-06 | 0:45 | 17.77 | 7.57 | 7.55 | 79.4 | 74.3 | 48.314
2006-09-06 | 1:01 | 17.68 | 7.56 | 7.56 | 79.4 | 74.2 | 48.254
2006-09-06 | 1:16 | 17.6 | 7.55 | 7.57 | 79.3 | 74.2 | 48.216
2006-09-06 | 1:31 | 17.51 | 7.55 | 7.59 | 79.4 | 74.1 | 48.154
2006-09-06 | 1:46 | 17.44 | 7.54 | 7.6 | 79.4 | 73.9 | 48.059
2006-09-06 | 2:01 | 17.37 | 7.53 | 7.61 | 79.4 | 73.9 | 48.042
2006-09-06 | 2:16 | 17.3 | 7.53 | 7.61 | 79.3 | 73.9 | 48.012
2006-09-06 | 2:31 | 17.25 | 7.52 | 7.65 | 79.5 | 73.6 | 47.859
2006-09-06 | 2:46 | 17.19 | 7.53 | 7.65 | 79.5 | 73.6 | 47.816
2006-09-06 | 3:01 | 17.16 | 7.52 | 7.67 | 79.7 | 73.4 | 47.741
2006-09-06 | 3:16 | 17.12 | 7.51 | 7.68 | 79.6 | 73.3 | 47.673
2006-09-06 | 3:31 | 17.09 | 7.51 | 7.68 | 79.7 | 73.3 | 47.663
2006-09-06 | 3:46 | 17.05 | 7.51 | 7.69 | 79.6 | 73.3 | 47.613
2006-09-06 | 4:01 | 17.04 | 7.51 | 7.7 | 79.7 | 73.2 | 47.549
2006-09-06 | 4:16 | 17.01 | 7.51 | 7.69 | 79.6 | 73.1 | 47.533
2006-09-06 | 4:31 | 17 | 7.51 | 7.69 | 79.6 | 73.1 | 47.498
2006-09-06 | 4:46 | 17 | 7.5 | 7.69 | 79.5 | 73 | 47.427
2006-09-06 | 5:01 | 16.97 | 7.5 | 7.69 | 79.6 | 72.9 | 47.405
2006-09-06 | 5:16 | 16.97 | 7.5 | 7.69 | 79.6 | 72.8 | 47.315
2006-09-06 | 5:31 | 16.96 | 7.5 | 7.69 | 79.5 | 72.7 | 47.26
2006-09-06 | 5:46 | 16.95 | 7.49 | 7.7 | 79.6 | 72.6 | 47.195
2006-09-06 | 6:01 | 16.95 | 7.49 | 7.68 | 79.4 | 72.5 | 47.138
2006-09-06 | 6:16 | 16.94 | 7.49 | 7.68 | 79.4 | 72.4 | 47.046
2006-09-06 | 6:31 | 16.94 | 7.48 | 7.67 | 79.3 | 72.3 | 46.983
2006-09-06 | 6:46 | 16.94 | 7.49 | 7.71 | 79.7 | 72.2 | 46.912
2006-09-06 | 7:01 | 16.93 | 7.49 | 7.73 | 79.9 | 72.1 | 46.851
2006-09-06 | 7:16 | 16.95 | 7.5 | 7.77 | 80.3 | 72 | 46.796
2006-09-06 | 7:31 | 16.96 | 7.52 | 7.85 | 81.2 | 71.9 | 46.719
2006-09-06 | 7:46 | 17 | 7.53 | 7.94 | 82.2 | 71.8 | 46.678
2006-09-06 | 8:01 | 17.05 | 7.55 | 8.06 | 83.5 | 71.8 | 46.638
2006-09-06 | 8:16 | 17.12 | 7.59 | 8.17 | 84.8 | 71.7 | 46.602
2006-09-06 | 8:31 | 17.19 | 7.62 | 8.3 | 86.3 | 71.6 | 46.538
2006-09-06 | 8:46 | 17.28 | 7.67 | 8.44 | 87.8 | 71.6 | 46.544
2006-09-06 | 9:01 | 17.37 | 7.71 | 8.55 | 89.2 | 71.6 | 46.51
2006-09-06 | 9:16 | 17.49 | 7.76 | 8.69 | 90.8 | 71.5 | 46.486
2006-09-06 | 9:31 | 17.62 | 7.81 | 8.78 | 92 | 71.5 | 46.462
2006-09-06 | 9:46 | 17.76 | 7.88 | 8.94 | 93.9 | 71.4 | 46.42
2006-09-06 | 10:01 | 17.91 | 7.95 | 9.05 | 95.5 | 71.4 | 46.414
2006-09-06 | 10:16 | 18.07 | 8.02 | 9.16 | 96.9 | 71.4 | 46.396
2006-09-06 | 10:31 | 18.24 | 8.1 | 9.25 | 98.2 | 71.4 | 46.383
2006-09-06 | 10:46 | 18.43 | 8.18 | 9.34 | 99.5 | 71.3 | 46.356
2006-09-06 | 11:01 | 18.61 | 8.27 | 9.42 | 100.8 | 71.3 | 46.33
2006-09-06 | 11:16 | 18.79 | 8.35 | 9.48 | 101.8 | 71.3 | 46.329
2006-09-06 | 11:31 | 18.99 | 8.43 | 9.53 | 102.8 | 71.3 | 46.33
2006-09-06 | 11:46 | 19.18 | 8.5 | 9.57 | 103.6 | 71.3 | 46.347
2006-09-06 | 12:01 | 19.36 | 8.57 | 9.62 | 104.5 | 71.3 | 46.355
2006-09-06 | 12:16 | 19.56 | 8.64 | 9.65 | 105.2 | 71.4 | 46.399
2006-09-06 | 12:31 | 19.74 | 8.69 | 9.66 | 105.7 | 71.4 | 46.417
2006-09-06 | 12:46 | 19.93 | 8.75 | 9.68 | 106.3 | 71.5 | 46.455
2006-09-06 | 13:00 | 20.11 | 8.79 | 9.69 | 106.8 | 71.6 | 46.558
2006-09-06 | 13:15 | 20.28 | 8.84 | 9.68 | 107.1 | 71.6 | 46.517
2006-09-06 | 13:30 | 20.44 | 8.88 | 9.69 | 107.5 | 71.7 | 46.575
2006-09-06 | 13:45 | 20.6 | 8.91 | 9.66 | 107.5 | 71.7 | 46.584
2006-09-06 | 14:00 | 20.73 | 8.94 | 9.66 | 107.8 | 71.7 | 46.635
2006-09-06 | 14:15 | 20.85 | 8.97 | 9.65 | 108 | 71.8 | 46.7
2006-09-06 | 14:30 | 20.96 | 8.99 | 9.61 | 107.7 | 71.9 | 46.703
2006-09-06 | 14:45 | 21.05 | 9.01 | 9.58 | 107.6 | 71.9 | 46.748
2006-09-06 | 15:00 | 21.15 | 9.04 | 9.54 | 107.3 | 72 | 46.769
2006-09-06 | 15:15 | 21.21 | 9.05 | 9.52 | 107.3 | 72 | 46.809
2006-09-06 | 15:30 | 21.27 | 9.07 | 9.47 | 106.8 | 72.1 | 46.846
2006-09-06 | 15:45 | 21.32 | 9.09 | 9.43 | 106.5 | 72.1 | 46.871
2006-09-06 | 16:00 | 21.34 | 9.09 | 9.39 | 106.1 | 72.2 | 46.937
2006-09-06 | 16:15 | 21.35 | 9.1 | 9.37 | 105.8 | 72.2 | 46.947
2006-09-06 | 16:30 | 21.37 | 9.11 | 9.29 | 105 | 72.3 | 47.011
2006-09-06 | 16:45 | 21.35 | 9.12 | 9.25 | 104.5 | 72.4 | 47.062
2006-09-06 | 17:00 | 21.33 | 9.11 | 9.18 | 103.6 | 72.5 | 47.122
2006-09-06 | 17:15 | 21.29 | 9.11 | 9.09 | 102.6 | 72.5 | 47.147
2006-09-06 | 17:30 | 21.23 | 9.11 | 9.03 | 101.7 | 72.6 | 47.211
2006-09-06 | 17:45 | 21.16 | 9.1 | 8.95 | 100.8 | 72.7 | 47.272
2006-09-06 | 18:00 | 21.07 | 9.08 | 8.86 | 99.6 | 72.9 | 47.363
2006-09-06 | 18:15 | 20.98 | 9.07 | 8.76 | 98.2 | 72.9 | 47.411
2006-09-06 | 18:30 | 20.88 | 9.04 | 8.64 | 96.7 | 73.1 | 47.529
2006-09-06 | 18:45 | 20.77 | 9.01 | 8.53 | 95.2 | 73.3 | 47.625
2006-09-06 | 19:00 | 20.64 | 8.98 | 8.4 | 93.6 | 73.4 | 47.683
2006-09-06 | 19:15 | 20.54 | 8.92 | 8.27 | 91.9 | 73.5 | 47.784
2006-09-06 | 19:30 | 20.41 | 8.87 | 8.15 | 90.4 | 73.8 | 47.947
2006-09-06 | 19:45 | 20.29 | 8.81 | 8.03 | 88.9 | 74 | 48.091
2006-09-06 | 20:00 | 20.16 | 8.74 | 7.92 | 87.4 | 74.1 | 48.192
2006-09-06 | 20:15 | 20.03 | 8.66 | 7.82 | 86.1 | 74.4 | 48.369
2006-09-06 | 20:30 | 19.9 | 8.58 | 7.74 | 85 | 74.6 | 48.498
2006-09-06 | 20:45 | 19.78 | 8.48 | 7.67 | 84 | 74.8 | 48.65
2006-09-06 | 21:00 | 19.66 | 8.38 | 7.61 | 83.1 | 75.1 | 48.783
2006-09-06 | 21:15 | 19.53 | 8.27 | 7.57 | 82.5 | 75.3 | 48.937
2006-09-06 | 21:30 | 19.38 | 8.16 | 7.51 | 81.6 | 75.6 | 49.162
2006-09-06 | 21:45 | 19.24 | 8.06 | 7.5 | 81.2 | 75.8 | 49.252
2006-09-06 | 22:00 | 19.1 | 7.97 | 7.47 | 80.7 | 76 | 49.425
2006-09-06 | 22:15 | 18.97 | 7.9 | 7.46 | 80.3 | 76.2 | 49.537
2006-09-06 | 22:30 | 18.83 | 7.84 | 7.45 | 80 | 76.3 | 49.573
2006-09-06 | 22:45 | 18.68 | 7.79 | 7.44 | 79.7 | 76.4 | 49.674
2006-09-06 | 23:00 | 18.54 | 7.74 | 7.43 | 79.4 | 76.5 | 49.703
2006-09-06 | 23:15 | 18.4 | 7.71 | 7.43 | 79.2 | 76.6 | 49.785
2006-09-06 | 23:30 | 18.27 | 7.67 | 7.45 | 79.1 | 76.4 | 49.671
2006-09-06 | 23:45 | 18.13 | 7.65 | 7.45 | 78.9 | 76.5 | 49.696
2006-09-07 | 0:00 | 18.01 | 7.63 | 7.47 | 78.9 | 76.4 | 49.645
2006-09-07 | 0:15 | 17.89 | 7.62 | 7.48 | 78.8 | 76.3 | 49.585
2006-09-07 | 0:30 | 17.77 | 7.6 | 7.49 | 78.8 | 76.1 | 49.472
2006-09-07 | 0:45 | 17.66 | 7.59 | 7.51 | 78.8 | 76 | 49.395
2006-09-07 | 1:01 | 17.55 | 7.58 | 7.56 | 79.1 | 75.8 | 49.264
2006-09-07 | 1:16 | 17.46 | 7.57 | 7.57 | 79.1 | 75.7 | 49.193
2006-09-07 | 1:31 | 17.36 | 7.56 | 7.59 | 79.1 | 75.5 | 49.071
2006-09-07 | 1:46 | 17.27 | 7.56 | 7.61 | 79.2 | 75.4 | 49.037
2006-09-07 | 2:01 | 17.18 | 7.56 | 7.63 | 79.3 | 75.2 | 48.892
2006-09-07 | 2:16 | 17.11 | 7.55 | 7.66 | 79.4 | 75.1 | 48.792
2006-09-07 | 2:31 | 17.04 | 7.55 | 7.66 | 79.4 | 74.9 | 48.689
2006-09-07 | 2:46 | 16.97 | 7.54 | 7.69 | 79.5 | 74.7 | 48.565
2006-09-07 | 3:01 | 16.92 | 7.54 | 7.69 | 79.5 | 74.6 | 48.502
2006-09-07 | 3:16 | 16.86 | 7.54 | 7.7 | 79.5 | 74.5 | 48.415
2006-09-07 | 3:31 | 16.83 | 7.53 | 7.71 | 79.5 | 74.4 | 48.347
2006-09-07 | 3:46 | 16.79 | 7.53 | 7.72 | 79.5 | 74.2 | 48.226
2006-09-07 | 4:01 | 16.76 | 7.53 | 7.73 | 79.6 | 74.1 | 48.189
2006-09-07 | 4:16 | 16.71 | 7.53 | 7.72 | 79.4 | 74.1 | 48.135
2006-09-07 | 4:31 | 16.7 | 7.52 | 7.74 | 79.6 | 73.8 | 47.985
2006-09-07 | 4:46 | 16.69 | 7.52 | 7.74 | 79.5 | 73.7 | 47.883
2006-09-07 | 5:01 | 16.65 | 7.52 | 7.73 | 79.4 | 73.6 | 47.812
2006-09-07 | 5:16 | 16.62 | 7.52 | 7.74 | 79.4 | 73.4 | 47.73
2006-09-07 | 5:31 | 16.62 | 7.51 | 7.73 | 79.4 | 73.3 | 47.648
2006-09-07 | 5:46 | 16.59 | 7.51 | 7.73 | 79.3 | 73.2 | 47.567
2006-09-07 | 6:01 | 16.59 | 7.51 | 7.74 | 79.4 | 73.1 | 47.522
2006-09-07 | 6:16 | 16.56 | 7.51 | 7.74 | 79.4 | 73 | 47.422
2006-09-07 | 6:31 | 16.55 | 7.51 | 7.73 | 79.3 | 72.9 | 47.354
2006-09-07 | 6:46 | 16.53 | 7.5 | 7.75 | 79.4 | 72.8 | 47.307
2006-09-07 | 7:01 | 16.52 | 7.51 | 7.77 | 79.6 | 72.7 | 47.249
2006-09-07 | 7:16 | 16.53 | 7.52 | 7.83 | 80.2 | 72.6 | 47.191
2006-09-07 | 7:31 | 16.53 | 7.53 | 7.9 | 80.9 | 72.5 | 47.156
2006-09-07 | 7:46 | 16.57 | 7.55 | 7.99 | 81.9 | 72.5 | 47.1
2006-09-07 | 8:01 | 16.62 | 7.58 | 8.1 | 83.1 | 72.4 | 47.052
2006-09-07 | 8:16 | 16.68 | 7.61 | 8.21 | 84.4 | 72.3 | 47.013
2006-09-07 | 8:31 | 16.75 | 7.63 | 8.34 | 85.8 | 72.3 | 46.988
2006-09-07 | 8:46 | 16.83 | 7.68 | 8.47 | 87.4 | 72.3 | 46.969
2006-09-07 | 9:01 | 16.93 | 7.72 | 8.59 | 88.8 | 72.2 | 46.937
2006-09-07 | 9:16 | 17.04 | 7.77 | 8.72 | 90.3 | 72.2 | 46.907
2006-09-07 | 9:31 | 17.17 | 7.82 | 8.85 | 91.9 | 72.2 | 46.898
2006-09-07 | 9:46 | 17.3 | 7.88 | 8.97 | 93.4 | 72.1 | 46.841
2006-09-07 | 10:01 | 17.44 | 7.95 | 9.08 | 94.8 | 72 | 46.816
2006-09-07 | 10:16 | 17.6 | 8.02 | 9.17 | 96.1 | 72 | 46.794
2006-09-07 | 10:31 | 17.78 | 8.09 | 9.27 | 97.5 | 71.9 | 46.753
2006-09-07 | 10:46 | 17.95 | 8.17 | 9.36 | 98.8 | 71.9 | 46.725
2006-09-07 | 11:01 | 18.13 | 8.26 | 9.42 | 99.7 | 71.8 | 46.692
2006-09-07 | 11:16 | 18.31 | 8.33 | 9.49 | 100.9 | 71.9 | 46.71
2006-09-07 | 11:31 | 18.5 | 8.41 | 9.54 | 101.8 | 71.8 | 46.674
2006-09-07 | 11:46 | 18.68 | 8.49 | 9.58 | 102.6 | 71.7 | 46.637
2006-09-07 | 12:01 | 18.86 | 8.56 | 9.62 | 103.4 | 71.8 | 46.676
2006-09-07 | 12:16 | 19.05 | 8.62 | 9.64 | 104 | 71.8 | 46.684
2006-09-07 | 12:31 | 19.23 | 8.68 | 9.66 | 104.7 | 71.8 | 46.694
2006-09-07 | 12:46 | 19.41 | 8.72 | 9.67 | 105.1 | 71.9 | 46.732
2006-09-07 | 13:00 | 19.59 | 8.77 | 9.67 | 105.4 | 71.9 | 46.73
2006-09-07 | 13:15 | 19.75 | 8.81 | 9.65 | 105.6 | 71.9 | 46.756
2006-09-07 | 13:30 | 19.9 | 8.85 | 9.66 | 106 | 72 | 46.775
2006-09-07 | 13:45 | 20.05 | 8.88 | 9.65 | 106.3 | 72.1 | 46.844
2006-09-07 | 14:00 | 20.18 | 8.91 | 9.63 | 106.3 | 72.1 | 46.846
2006-09-07 | 14:15 | 20.3 | 8.94 | 9.62 | 106.4 | 72.1 | 46.86
2006-09-07 | 14:30 | 20.41 | 8.96 | 9.58 | 106.3 | 72.1 | 46.863
2006-09-07 | 14:45 | 20.51 | 8.98 | 9.57 | 106.3 | 72.2 | 46.907
2006-09-07 | 15:00 | 20.6 | 9 | 9.53 | 106.1 | 72.1 | 46.896
2006-09-07 | 15:15 | 20.68 | 9.02 | 9.49 | 105.8 | 72.2 | 46.914
2006-09-07 | 15:30 | 20.75 | 9.03 | 9.46 | 105.6 | 72.1 | 46.887
2006-09-07 | 15:45 | 20.81 | 9.04 | 9.41 | 105.2 | 72.2 | 46.914
2006-09-07 | 16:00 | 20.86 | 9.05 | 9.37 | 104.8 | 72.2 | 46.914
2006-09-07 | 16:15 | 20.91 | 9.06 | 9.31 | 104.2 | 72.2 | 46.925
2006-09-07 | 16:30 | 20.92 | 9.06 | 9.25 | 103.6 | 72.3 | 46.974
2006-09-07 | 16:45 | 20.92 | 9.06 | 9.18 | 102.8 | 72.3 | 46.965
2006-09-07 | 17:00 | 20.92 | 9.06 | 9.12 | 102.1 | 72.3 | 46.995
2006-09-07 | 17:15 | 20.91 | 9.06 | 9.03 | 101.2 | 72.3 | 47.025
2006-09-07 | 17:30 | 20.87 | 9.05 | 8.96 | 100.3 | 72.4 | 47.053
2006-09-07 | 17:45 | 20.81 | 9.04 | 8.89 | 99.4 | 72.5 | 47.101
2006-09-07 | 18:00 | 20.74 | 9.03 | 8.8 | 98.2 | 72.5 | 47.136
2006-09-07 | 18:15 | 20.65 | 9 | 8.7 | 96.9 | 72.6 | 47.199
2006-09-07 | 18:30 | 20.55 | 8.98 | 8.57 | 95.3 | 72.7 | 47.256
2006-09-07 | 18:45 | 20.45 | 8.93 | 8.45 | 93.8 | 72.8 | 47.35
2006-09-07 | 19:00 | 20.32 | 8.89 | 8.34 | 92.3 | 72.9 | 47.393
2006-09-07 | 19:15 | 20.19 | 8.83 | 8.23 | 90.8 | 73.1 | 47.491
2006-09-07 | 19:30 | 20.06 | 8.77 | 8.11 | 89.3 | 73.2 | 47.551
2006-09-07 | 19:45 | 19.92 | 8.69 | 7.99 | 87.8 | 73.3 | 47.639
2006-09-07 | 20:00 | 19.79 | 8.6 | 7.87 | 86.2 | 73.5 | 47.792
2006-09-07 | 20:15 | 19.65 | 8.51 | 7.8 | 85.3 | 73.6 | 47.864
2006-09-07 | 20:30 | 19.51 | 8.41 | 7.74 | 84.3 | 73.9 | 48.019
2006-09-07 | 20:45 | 19.37 | 8.31 | 7.69 | 83.5 | 74.1 | 48.16
2006-09-07 | 21:00 | 19.22 | 8.2 | 7.63 | 82.6 | 74.3 | 48.276
2006-09-07 | 21:15 | 19.08 | 8.1 | 7.59 | 82 | 74.5 | 48.434
2006-09-07 | 21:30 | 18.92 | 8.01 | 7.57 | 81.5 | 74.6 | 48.491
2006-09-07 | 21:45 | 18.77 | 7.93 | 7.54 | 80.9 | 74.8 | 48.622
2006-09-07 | 22:00 | 18.62 | 7.87 | 7.53 | 80.6 | 74.9 | 48.713
2006-09-07 | 22:15 | 18.49 | 7.82 | 7.51 | 80.2 | 75 | 48.781
2006-09-07 | 22:30 | 18.34 | 7.77 | 7.5 | 79.8 | 75.1 | 48.84
2006-09-07 | 22:45 | 18.2 | 7.73 | 7.49 | 79.5 | 75.3 | 48.915
2006-09-07 | 23:00 | 18.07 | 7.7 | 7.51 | 79.5 | 75.1 | 48.836
2006-09-07 | 23:15 | 17.95 | 7.67 | 7.51 | 79.2 | 75.2 | 48.858
2006-09-07 | 23:30 | 17.83 | 7.65 | 7.52 | 79.1 | 75 | 48.766
2006-09-07 | 23:45 | 17.72 | 7.63 | 7.52 | 79 | 74.9 | 48.689
2006-09-08 | 0:00 | 17.61 | 7.62 | 7.54 | 79 | 74.8 | 48.636
2006-09-08 | 0:15 | 17.52 | 7.6 | 7.54 | 78.9 | 74.8 | 48.608
2006-09-08 | 0:30 | 17.42 | 7.59 | 7.55 | 78.8 | 74.7 | 48.551
2006-09-08 | 0:45 | 17.33 | 7.58 | 7.57 | 78.9 | 74.5 | 48.457
2006-09-08 | 1:01 | 17.25 | 7.57 | 7.58 | 78.9 | 74.4 | 48.386
2006-09-08 | 1:16 | 17.17 | 7.57 | 7.59 | 78.8 | 74.4 | 48.352
2006-09-08 | 1:31 | 17.1 | 7.56 | 7.6 | 78.8 | 74.4 | 48.385
2006-09-08 | 1:46 | 17.03 | 7.55 | 7.61 | 78.8 | 74.3 | 48.266
2006-09-08 | 2:01 | 16.98 | 7.55 | 7.62 | 78.8 | 74.2 | 48.244
2006-09-08 | 2:16 | 16.91 | 7.54 | 7.62 | 78.7 | 74.2 | 48.257
2006-09-08 | 2:31 | 16.87 | 7.54 | 7.64 | 78.8 | 74.1 | 48.149
2006-09-08 | 2:46 | 16.83 | 7.53 | 7.64 | 78.8 | 74 | 48.127
2006-09-08 | 3:01 | 16.78 | 7.53 | 7.65 | 78.8 | 74 | 48.077
2006-09-08 | 3:16 | 16.74 | 7.53 | 7.65 | 78.7 | 73.9 | 48.049
2006-09-08 | 3:31 | 16.72 | 7.52 | 7.66 | 78.8 | 73.8 | 47.975
2006-09-08 | 3:46 | 16.68 | 7.52 | 7.66 | 78.7 | 73.8 | 47.956
2006-09-08 | 4:01 | 16.66 | 7.52 | 7.66 | 78.7 | 73.7 | 47.882
2006-09-08 | 4:16 | 16.64 | 7.52 | 7.66 | 78.6 | 73.6 | 47.842
2006-09-08 | 4:31 | 16.63 | 7.51 | 7.67 | 78.7 | 73.5 | 47.791
2006-09-08 | 4:46 | 16.61 | 7.51 | 7.66 | 78.6 | 73.5 | 47.753
2006-09-08 | 5:01 | 16.6 | 7.51 | 7.67 | 78.8 | 73.4 | 47.737
2006-09-08 | 5:16 | 16.59 | 7.51 | 7.66 | 78.6 | 73.4 | 47.717
2006-09-08 | 5:31 | 16.58 | 7.51 | 7.66 | 78.6 | 73.3 | 47.64
2006-09-08 | 5:46 | 16.57 | 7.51 | 7.67 | 78.6 | 73.3 | 47.62
2006-09-08 | 6:01 | 16.56 | 7.51 | 7.66 | 78.5 | 73.2 | 47.598
2006-09-08 | 6:16 | 16.56 | 7.5 | 7.66 | 78.6 | 73.2 | 47.574
2006-09-08 | 6:31 | 16.55 | 7.5 | 7.67 | 78.6 | 73.1 | 47.514
2006-09-08 | 6:46 | 16.54 | 7.5 | 7.68 | 78.7 | 73.1 | 47.492
2006-09-08 | 7:01 | 16.52 | 7.51 | 7.71 | 79 | 73 | 47.428
2006-09-08 | 7:16 | 16.52 | 7.51 | 7.76 | 79.5 | 72.9 | 47.404
2006-09-08 | 7:31 | 16.53 | 7.53 | 7.82 | 80.1 | 72.9 | 47.359
2006-09-08 | 7:46 | 16.55 | 7.54 | 7.92 | 81.2 | 72.7 | 47.273
2006-09-08 | 8:01 | 16.6 | 7.56 | 8.02 | 82.3 | 72.7 | 47.236
2006-09-08 | 8:16 | 16.64 | 7.6 | 8.14 | 83.6 | 72.6 | 47.187
2006-09-08 | 8:31 | 16.69 | 7.63 | 8.27 | 85 | 72.5 | 47.103
2006-09-08 | 8:46 | 16.76 | 7.67 | 8.4 | 86.5 | 72.3 | 47.023
2006-09-08 | 9:01 | 16.84 | 7.71 | 8.51 | 87.8 | 72.3 | 46.98
2006-09-08 | 9:16 | 16.94 | 7.76 | 8.65 | 89.4 | 72.2 | 46.927
2006-09-08 | 9:31 | 17.05 | 7.81 | 8.78 | 91 | 72.1 | 46.852
2006-09-08 | 9:46 | 17.18 | 7.88 | 8.91 | 92.5 | 72 | 46.807
2006-09-08 | 10:01 | 17.31 | 7.93 | 9.02 | 93.9 | 71.9 | 46.738
2006-09-08 | 10:16 | 17.46 | 8.01 | 9.13 | 95.3 | 71.8 | 46.677
2006-09-08 | 10:31 | 17.61 | 8.08 | 9.23 | 96.7 | 71.8 | 46.641
2006-09-08 | 10:46 | 17.76 | 8.17 | 9.31 | 97.9 | 71.7 | 46.598
2006-09-08 | 11:01 | 17.9 | 8.25 | 9.4 | 99.1 | 71.6 | 46.542
2006-09-08 | 11:16 | 18.06 | 8.34 | 9.47 | 100.1 | 71.6 | 46.542
2006-09-08 | 11:31 | 18.21 | 8.42 | 9.51 | 100.9 | 71.5 | 46.492
2006-09-08 | 11:46 | 18.37 | 8.49 | 9.54 | 101.6 | 71.5 | 46.482
2006-09-08 | 12:01 | 18.51 | 8.55 | 9.56 | 102 | 71.5 | 46.491
2006-09-08 | 12:16 | 18.64 | 8.62 | 9.64 | 103.2 | 71.5 | 46.48
2006-09-08 | 12:31 | 18.77 | 8.67 | 9.66 | 103.6 | 71.6 | 46.511
2006-09-08 | 12:46 | 18.92 | 8.73 | 9.68 | 104.2 | 71.5 | 46.495
2006-09-08 | 13:00 | 19.08 | 8.78 | 9.7 | 104.8 | 71.6 | 46.52
2006-09-08 | 13:15 | 19.22 | 8.83 | 9.65 | 104.5 | 71.6 | 46.537
2006-09-08 | 13:30 | 19.36 | 8.87 | 9.72 | 105.6 | 71.7 | 46.588
2006-09-08 | 13:45 | 19.5 | 8.91 | 9.72 | 105.8 | 71.7 | 46.582
2006-09-08 | 14:00 | 19.64 | 8.94 | 9.7 | 106 | 71.7 | 46.605
2006-09-08 | 14:15 | 19.77 | 8.97 | 9.68 | 106 | 71.7 | 46.612
2006-09-08 | 14:30 | 19.88 | 8.99 | 9.69 | 106.3 | 71.8 | 46.677
2006-09-08 | 14:45 | 19.99 | 9.01 | 9.65 | 106.1 | 71.8 | 46.653
2006-09-08 | 15:00 | 20.08 | 9.03 | 9.61 | 105.9 | 71.8 | 46.694
2006-09-08 | 15:15 | 20.15 | 9.05 | 9.6 | 106 | 71.8 | 46.699
2006-09-08 | 15:30 | 20.22 | 9.07 | 9.55 | 105.6 | 71.9 | 46.746
2006-09-08 | 15:45 | 20.28 | 9.08 | 9.5 | 105 | 71.9 | 46.725
2006-09-08 | 16:00 | 20.31 | 9.09 | 9.47 | 104.8 | 71.9 | 46.704
2006-09-08 | 16:15 | 20.36 | 9.1 | 9.4 | 104.2 | 71.9 | 46.724
2006-09-08 | 16:30 | 20.38 | 9.1 | 9.39 | 104 | 71.9 | 46.708
2006-09-08 | 16:45 | 20.4 | 9.12 | 9.35 | 103.6 | 71.8 | 46.686
2006-09-08 | 17:00 | 20.41 | 9.12 | 9.32 | 103.3 | 71.9 | 46.71
2006-09-08 | 17:15 | 20.3 | 9.12 | 9.23 | 102.1 | 71.9 | 46.744
2006-09-08 | 17:30 | 20.29 | 9.12 | 9.15 | 101.3 | 72 | 46.81
2006-09-08 | 17:45 | 20.24 | 9.11 | 9.1 | 100.6 | 72 | 46.805
2006-09-08 | 18:00 | 20.14 | 9.1 | 9.01 | 99.3 | 72.1 | 46.842
2006-09-08 | 18:15 | 20.03 | 9.08 | 8.86 | 97.6 | 72.2 | 46.953
2006-09-08 | 18:30 | 19.91 | 9.04 | 8.71 | 95.6 | 72.3 | 47.005
2006-09-08 | 18:45 | 19.81 | 9.01 | 8.6 | 94.2 | 72.5 | 47.114
2006-09-08 | 19:00 | 19.72 | 8.98 | 8.5 | 92.9 | 72.5 | 47.152
2006-09-08 | 19:15 | 19.63 | 8.94 | 8.38 | 91.5 | 72.7 | 47.234
2006-09-08 | 19:30 | 19.53 | 8.89 | 8.25 | 89.8 | 72.8 | 47.343
2006-09-08 | 19:45 | 19.4 | 8.82 | 8.11 | 88.1 | 73.1 | 47.524
2006-09-08 | 20:00 | 19.27 | 8.75 | 8.01 | 86.8 | 73.3 | 47.646
2006-09-08 | 20:15 | 19.13 | 8.67 | 7.92 | 85.6 | 73.5 | 47.773
2006-09-08 | 20:30 | 19.01 | 8.59 | 7.85 | 84.6 | 73.7 | 47.924
2006-09-08 | 20:45 | 18.89 | 8.51 | 7.79 | 83.8 | 74 | 48.076
2006-09-08 | 21:00 | 18.76 | 8.41 | 7.74 | 83 | 74.1 | 48.197
2006-09-08 | 21:15 | 18.62 | 8.31 | 7.7 | 82.4 | 74.5 | 48.434
2006-09-08 | 21:30 | 18.49 | 8.2 | 7.66 | 81.8 | 74.8 | 48.601
2006-09-08 | 21:45 | 18.36 | 8.1 | 7.64 | 81.3 | 75 | 48.748
2006-09-08 | 22:00 | 18.23 | 8.02 | 7.62 | 80.9 | 75.2 | 48.854
2006-09-08 | 22:15 | 18.11 | 7.94 | 7.6 | 80.4 | 75.4 | 48.985
2006-09-08 | 22:30 | 17.98 | 7.87 | 7.59 | 80.2 | 75.5 | 49.065
2006-09-08 | 22:45 | 17.85 | 7.82 | 7.58 | 79.9 | 75.6 | 49.126
2006-09-08 | 23:00 | 17.73 | 7.78 | 7.58 | 79.6 | 75.6 | 49.154
2006-09-08 | 23:15 | 17.61 | 7.74 | 7.57 | 79.3 | 75.6 | 49.149
2006-09-08 | 23:30 | 17.52 | 7.71 | 7.57 | 79.2 | 75.6 | 49.164
2006-09-08 | 23:45 | 17.42 | 7.69 | 7.59 | 79.2 | 75.5 | 49.079
2006-09-09 | 0:00 | 17.33 | 7.67 | 7.6 | 79.2 | 75.5 | 49.059
2006-09-09 | 0:15 | 17.23 | 7.65 | 7.61 | 79.1 | 75.4 | 48.989
2006-09-09 | 0:30 | 17.14 | 7.63 | 7.62 | 79.1 | 75.3 | 48.923
2006-09-09 | 0:45 | 17.06 | 7.62 | 7.62 | 79 | 75.2 | 48.878
2006-09-09 | 1:01 | 16.97 | 7.61 | 7.65 | 79.1 | 75.1 | 48.786
2006-09-09 | 1:16 | 16.88 | 7.6 | 7.66 | 79 | 75 | 48.72
2006-09-09 | 1:31 | 16.8 | 7.59 | 7.68 | 79.1 | 74.8 | 48.637
2006-09-09 | 1:46 | 16.73 | 7.58 | 7.69 | 79.1 | 74.7 | 48.584
2006-09-09 | 2:01 | 16.66 | 7.58 | 7.71 | 79.2 | 74.6 | 48.505
2006-09-09 | 2:16 | 16.59 | 7.57 | 7.71 | 79.1 | 74.6 | 48.483
2006-09-09 | 2:31 | 16.53 | 7.56 | 7.72 | 79.1 | 74.5 | 48.447
2006-09-09 | 2:46 | 16.48 | 7.56 | 7.73 | 79.1 | 74.4 | 48.349
2006-09-09 | 3:01 | 16.43 | 7.55 | 7.73 | 79 | 74.3 | 48.299
2006-09-09 | 3:16 | 16.38 | 7.55 | 7.73 | 79 | 74.3 | 48.264
2006-09-09 | 3:31 | 16.35 | 7.54 | 7.74 | 79.1 | 74.1 | 48.162
2006-09-09 | 3:46 | 16.32 | 7.54 | 7.75 | 79 | 74 | 48.095
2006-09-09 | 4:01 | 16.29 | 7.54 | 7.75 | 79 | 73.9 | 48.03
2006-09-09 | 4:16 | 16.26 | 7.53 | 7.75 | 79 | 73.8 | 47.953
2006-09-09 | 4:31 | 16.24 | 7.53 | 7.75 | 78.9 | 73.7 | 47.918
2006-09-09 | 4:46 | 16.23 | 7.52 | 7.76 | 79 | 73.6 | 47.844
2006-09-09 | 5:01 | 16.22 | 7.52 | 7.75 | 78.9 | 73.5 | 47.796
2006-09-09 | 5:16 | 16.2 | 7.52 | 7.76 | 78.9 | 73.5 | 47.788
2006-09-09 | 5:31 | 16.19 | 7.52 | 7.76 | 79 | 73.4 | 47.7
2006-09-09 | 5:46 | 16.17 | 7.52 | 7.76 | 78.9 | 73.3 | 47.649
2006-09-09 | 6:01 | 16.16 | 7.51 | 7.76 | 78.9 | 73.3 | 47.626
2006-09-09 | 6:16 | 16.14 | 7.52 | 7.76 | 78.9 | 73.3 | 47.648
2006-09-09 | 6:31 | 16.11 | 7.51 | 7.77 | 78.9 | 73.2 | 47.572
2006-09-09 | 6:46 | 16.1 | 7.51 | 7.78 | 79 | 73.2 | 47.575
2006-09-09 | 7:01 | 16.08 | 7.51 | 7.8 | 79.1 | 73.2 | 47.557
2006-09-09 | 7:16 | 16.06 | 7.52 | 7.84 | 79.5 | 73.1 | 47.493
2006-09-09 | 7:31 | 16.05 | 7.53 | 7.9 | 80.1 | 72.9 | 47.417
2006-09-09 | 7:46 | 16.09 | 7.55 | 8.02 | 81.4 | 72.8 | 47.326
2006-09-09 | 8:01 | 16.16 | 7.59 | 8.17 | 83 | 72.6 | 47.194
2006-09-09 | 8:16 | 16.21 | 7.62 | 8.3 | 84.4 | 72.4 | 47.049
2006-09-09 | 8:31 | 16.22 | 7.64 | 8.36 | 85.1 | 72.4 | 47.053
2006-09-09 | 8:46 | 16.27 | 7.67 | 8.45 | 86.1 | 72.4 | 47.044
2006-09-09 | 9:01 | 16.42 | 7.71 | 8.62 | 88.1 | 72.3 | 46.97
2006-09-09 | 9:16 | 16.51 | 7.77 | 8.78 | 89.9 | 72.2 | 46.911
2006-09-09 | 9:31 | 16.54 | 7.8 | 8.83 | 90.5 | 72.2 | 46.919
2006-09-09 | 9:46 | 16.65 | 7.85 | 8.9 | 91.4 | 72.2 | 46.927
2006-09-09 | 10:01 | 16.82 | 7.91 | 9.05 | 93.3 | 72.1 | 46.855
2006-09-09 | 10:16 | 16.97 | 8 | 9.18 | 95 | 72 | 46.825
2006-09-09 | 10:31 | 17.13 | 8.06 | 9.26 | 96.1 | 72.1 | 46.836
2006-09-09 | 10:46 | 17.28 | 8.14 | 9.35 | 97.3 | 72 | 46.806
2006-09-09 | 11:01 | 17.43 | 8.22 | 9.43 | 98.4 | 71.9 | 46.761
2006-09-09 | 11:16 | 17.6 | 8.3 | 9.5 | 99.5 | 71.9 | 46.745
2006-09-09 | 11:31 | 17.76 | 8.36 | 9.53 | 100.2 | 71.9 | 46.722
2006-09-09 | 11:46 | 17.95 | 8.44 | 9.58 | 101.1 | 71.9 | 46.726
2006-09-09 | 12:01 | 18.14 | 8.5 | 9.62 | 101.9 | 72 | 46.773
2006-09-09 | 12:16 | 18.3 | 8.57 | 9.64 | 102.5 | 71.9 | 46.732
2006-09-09 | 12:31 | 18.46 | 8.63 | 9.66 | 103 | 71.9 | 46.76
2006-09-09 | 12:46 | 18.63 | 8.67 | 9.66 | 103.3 | 72 | 46.787
2006-09-09 | 13:00 | 18.8 | 8.73 | 9.67 | 103.8 | 72 | 46.804
2006-09-09 | 13:15 | 18.95 | 8.77 | 9.68 | 104.3 | 72 | 46.799
2006-09-09 | 13:30 | 19.1 | 8.85 | 9.67 | 104.5 | 71.5 | 46.498
2006-09-09 | 13:45 | 19.24 | 8.89 | 9.64 | 104.5 | 71.5 | 46.49
2006-09-09 | 14:00 | 19.37 | 8.92 | 9.65 | 104.8 | 71.5 | 46.488
2006-09-09 | 14:15 | 19.49 | 8.95 | 9.65 | 105 | 71.5 | 46.49
2006-09-09 | 14:30 | 19.6 | 8.97 | 9.61 | 104.9 | 71.6 | 46.519
2006-09-09 | 14:45 | 19.69 | 8.99 | 9.59 | 104.8 | 71.5 | 46.464
2006-09-09 | 15:00 | 19.79 | 9 | 9.57 | 104.8 | 71.5 | 46.503
2006-09-09 | 15:15 | 19.85 | 9.02 | 9.53 | 104.5 | 71.6 | 46.509
2006-09-09 | 15:30 | 19.92 | 9.04 | 9.5 | 104.3 | 71.5 | 46.497
2006-09-09 | 15:45 | 19.96 | 9.05 | 9.46 | 104 | 71.5 | 46.457
2006-09-09 | 16:00 | 20 | 9.06 | 9.42 | 103.6 | 71.5 | 46.493
2006-09-09 | 16:15 | 20.02 | 9.06 | 9.38 | 103.3 | 71.5 | 46.489
2006-09-09 | 16:30 | 20.03 | 9.06 | 9.31 | 102.5 | 71.5 | 46.462
2006-09-09 | 16:45 | 20.02 | 9.07 | 9.27 | 102 | 71.5 | 46.471
2006-09-09 | 17:00 | 19.99 | 9.08 | 9.23 | 101.5 | 71.5 | 46.494
2006-09-09 | 17:15 | 19.96 | 9.06 | 9.14 | 100.4 | 71.6 | 46.511
2006-09-09 | 17:30 | 19.91 | 9.06 | 9.07 | 99.6 | 71.6 | 46.551
2006-09-09 | 17:45 | 19.83 | 9.05 | 8.98 | 98.5 | 71.6 | 46.557
2006-09-09 | 18:00 | 19.73 | 9.03 | 8.91 | 97.5 | 71.7 | 46.578
2006-09-09 | 18:15 | 19.62 | 9.02 | 8.83 | 96.4 | 71.7 | 46.613
2006-09-09 | 18:30 | 19.5 | 8.99 | 8.72 | 95 | 71.8 | 46.686
2006-09-09 | 18:45 | 19.36 | 8.95 | 8.6 | 93.4 | 71.8 | 46.698
2006-09-09 | 19:00 | 19.22 | 8.9 | 8.47 | 91.8 | 71.9 | 46.766
2006-09-09 | 19:15 | 19.08 | 8.84 | 8.37 | 90.3 | 72.1 | 46.848
2006-09-09 | 19:30 | 18.94 | 8.79 | 8.26 | 89 | 72.2 | 46.908
2006-09-09 | 19:45 | 18.8 | 8.72 | 8.16 | 87.6 | 72.3 | 46.991
2006-09-09 | 20:00 | 18.67 | 8.64 | 8.07 | 86.5 | 72.4 | 47.063
2006-09-09 | 20:15 | 18.53 | 8.55 | 8.01 | 85.5 | 72.6 | 47.168
2006-09-09 | 20:30 | 18.39 | 8.45 | 7.95 | 84.7 | 72.7 | 47.272
2006-09-09 | 20:45 | 18.25 | 8.35 | 7.9 | 83.8 | 72.9 | 47.386
2006-09-09 | 21:00 | 18.11 | 8.24 | 7.86 | 83.3 | 73.2 | 47.559
2006-09-09 | 21:15 | 17.97 | 8.14 | 7.82 | 82.6 | 73.3 | 47.664
2006-09-09 | 21:30 | 17.83 | 8.05 | 7.81 | 82.2 | 73.6 | 47.826
2006-09-09 | 21:45 | 17.69 | 7.97 | 7.78 | 81.7 | 73.7 | 47.92
2006-09-09 | 22:00 | 17.53 | 7.9 | 7.78 | 81.4 | 73.9 | 48.018
2006-09-09 | 22:15 | 17.39 | 7.85 | 7.78 | 81.2 | 74.1 | 48.147
2006-09-09 | 22:30 | 17.25 | 7.81 | 7.78 | 80.9 | 74.1 | 48.18
2006-09-09 | 22:45 | 17.11 | 7.77 | 7.78 | 80.7 | 74.3 | 48.27
2006-09-09 | 23:00 | 16.97 | 7.74 | 7.78 | 80.4 | 74.3 | 48.267
2006-09-09 | 23:15 | 16.84 | 7.71 | 7.79 | 80.3 | 74.2 | 48.257
2006-09-09 | 23:30 | 16.71 | 7.68 | 7.8 | 80.2 | 74.2 | 48.221
2006-09-09 | 23:45 | 16.59 | 7.66 | 7.82 | 80.2 | 74.2 | 48.217
2006-09-10 | 0:00 | 16.48 | 7.65 | 7.83 | 80.2 | 74 | 48.132
2006-09-10 | 0:15 | 16.37 | 7.64 | 7.85 | 80.2 | 74 | 48.073
2006-09-10 | 0:30 | 16.26 | 7.63 | 7.86 | 80.1 | 73.9 | 48.037
2006-09-10 | 0:45 | 16.16 | 7.62 | 7.89 | 80.2 | 73.7 | 47.922
2006-09-10 | 1:01 | 16.06 | 7.61 | 7.91 | 80.2 | 73.6 | 47.858
2006-09-10 | 1:16 | 15.96 | 7.6 | 7.93 | 80.3 | 73.5 | 47.8
2006-09-10 | 1:31 | 15.86 | 7.6 | 7.95 | 80.3 | 73.4 | 47.714
2006-09-10 | 1:46 | 15.77 | 7.59 | 7.97 | 80.3 | 73.3 | 47.642
2006-09-10 | 2:01 | 15.7 | 7.58 | 7.98 | 80.3 | 73.3 | 47.618
2006-09-10 | 2:16 | 15.63 | 7.58 | 8 | 80.4 | 73.1 | 47.516
2006-09-10 | 2:31 | 15.56 | 7.57 | 8.01 | 80.4 | 73 | 47.441
2006-09-10 | 2:46 | 15.51 | 7.57 | 8.03 | 80.5 | 72.9 | 47.37
2006-09-10 | 3:01 | 15.45 | 7.57 | 8.03 | 80.4 | 72.8 | 47.332
2006-09-10 | 3:16 | 15.4 | 7.56 | 8.04 | 80.4 | 72.7 | 47.277
2006-09-10 | 3:31 | 15.35 | 7.56 | 8.05 | 80.4 | 72.7 | 47.235
2006-09-10 | 3:46 | 15.32 | 7.56 | 8.05 | 80.4 | 72.5 | 47.154
2006-09-10 | 4:01 | 15.28 | 7.56 | 8.06 | 80.4 | 72.5 | 47.101
2006-09-10 | 4:16 | 15.26 | 7.55 | 8.06 | 80.4 | 72.4 | 47.037
2006-09-10 | 4:31 | 15.24 | 7.55 | 8.07 | 80.5 | 72.3 | 47.002
2006-09-10 | 4:46 | 15.22 | 7.55 | 8.06 | 80.3 | 72.2 | 46.953
2006-09-10 | 5:01 | 15.19 | 7.54 | 8.07 | 80.3 | 72.2 | 46.927
2006-09-10 | 5:16 | 15.17 | 7.54 | 8.07 | 80.3 | 72.1 | 46.893
2006-09-10 | 5:31 | 15.15 | 7.54 | 8.07 | 80.3 | 72.1 | 46.851
2006-09-10 | 5:46 | 15.13 | 7.54 | 8.07 | 80.3 | 72.1 | 46.833
2006-09-10 | 6:01 | 15.11 | 7.54 | 8.07 | 80.2 | 72 | 46.8
2006-09-10 | 6:16 | 15.1 | 7.53 | 8.07 | 80.2 | 72 | 46.798
2006-09-10 | 6:31 | 15.08 | 7.54 | 8.07 | 80.2 | 72 | 46.775
2006-09-10 | 6:46 | 15.06 | 7.53 | 8.09 | 80.4 | 72 | 46.768
2006-09-10 | 7:01 | 15.04 | 7.54 | 8.11 | 80.5 | 71.9 | 46.738
2006-09-10 | 7:16 | 15.04 | 7.55 | 8.16 | 81 | 71.9 | 46.725
2006-09-10 | 7:31 | 15.04 | 7.56 | 8.23 | 81.7 | 71.9 | 46.733
2006-09-10 | 7:46 | 15.07 | 7.58 | 8.31 | 82.6 | 71.8 | 46.672
2006-09-10 | 8:01 | 15.11 | 7.6 | 8.42 | 83.7 | 71.8 | 46.669
2006-09-10 | 8:16 | 15.17 | 7.63 | 8.52 | 84.8 | 71.8 | 46.65
2006-09-10 | 8:31 | 15.23 | 7.67 | 8.64 | 86.1 | 71.8 | 46.654
2006-09-10 | 8:46 | 15.3 | 7.71 | 8.77 | 87.5 | 71.8 | 46.639
2006-09-10 | 9:01 | 15.38 | 7.75 | 8.87 | 88.6 | 71.7 | 46.622
2006-09-10 | 9:16 | 15.47 | 7.79 | 8.97 | 89.9 | 71.7 | 46.601
2006-09-10 | 9:31 | 15.59 | 7.85 | 9.1 | 91.4 | 71.7 | 46.616
2006-09-10 | 9:46 | 15.71 | 7.91 | 9.21 | 92.7 | 71.7 | 46.586
2006-09-10 | 10:01 | 15.85 | 7.97 | 9.31 | 94 | 71.6 | 46.549
2006-09-10 | 10:16 | 16 | 8.04 | 9.41 | 95.3 | 71.6 | 46.55
2006-09-10 | 10:31 | 16.16 | 8.12 | 9.5 | 96.6 | 71.5 | 46.487
2006-09-10 | 10:46 | 16.34 | 8.19 | 9.57 | 97.7 | 71.5 | 46.45
2006-09-10 | 11:01 | 16.53 | 8.27 | 9.64 | 98.7 | 71.4 | 46.436
2006-09-10 | 11:16 | 16.7 | 8.35 | 9.71 | 99.8 | 71.4 | 46.436
2006-09-10 | 11:31 | 16.88 | 8.41 | 9.72 | 100.4 | 71.4 | 46.433
2006-09-10 | 11:46 | 17.06 | 8.48 | 9.77 | 101.2 | 71.4 | 46.414
2006-09-10 | 12:01 | 17.25 | 8.54 | 9.79 | 101.8 | 71.4 | 46.402
2006-09-10 | 12:16 | 17.42 | 8.58 | 9.8 | 102.3 | 71.3 | 46.366
2006-09-10 | 12:31 | 17.6 | 8.64 | 9.8 | 102.7 | 71.4 | 46.407
2006-09-10 | 12:46 | 17.78 | 8.69 | 9.82 | 103.2 | 71.4 | 46.416
2006-09-10 | 13:00 | 17.93 | 8.73 | 9.81 | 103.5 | 71.3 | 46.355
2006-09-10 | 13:15 | 18.1 | 8.76 | 9.8 | 103.7 | 71.3 | 46.332
2006-09-10 | 13:30 | 18.25 | 8.8 | 9.79 | 103.9 | 71.2 | 46.309
2006-09-10 | 13:45 | 18.4 | 8.83 | 9.79 | 104.2 | 71.2 | 46.281
2006-09-10 | 14:00 | 18.53 | 8.86 | 9.76 | 104.2 | 71.2 | 46.269
2006-09-10 | 14:15 | 18.66 | 8.89 | 9.76 | 104.5 | 71.1 | 46.23
2006-09-10 | 14:30 | 18.76 | 8.91 | 9.72 | 104.3 | 71.1 | 46.224
2006-09-10 | 14:45 | 18.87 | 8.93 | 9.69 | 104.2 | 71.3 | 46.327
2006-09-10 | 15:00 | 18.97 | 8.95 | 9.66 | 104.1 | 71.2 | 46.291
2006-09-10 | 15:15 | 19.06 | 8.97 | 9.63 | 104 | 71.2 | 46.306
2006-09-10 | 15:30 | 19.14 | 8.98 | 9.61 | 103.9 | 71.2 | 46.279
2006-09-10 | 15:45 | 19.2 | 8.99 | 9.54 | 103.3 | 71.2 | 46.255
2006-09-10 | 16:00 | 19.24 | 9 | 9.51 | 103.1 | 71.2 | 46.266
2006-09-10 | 16:15 | 19.27 | 9 | 9.46 | 102.6 | 71.1 | 46.229
2006-09-10 | 16:30 | 19.28 | 9.01 | 9.41 | 102 | 71 | 46.137
2006-09-10 | 16:45 | 19.31 | 9.02 | 9.36 | 101.5 | 71 | 46.152
2006-09-10 | 17:00 | 19.31 | 9.01 | 9.3 | 100.9 | 71 | 46.137
2006-09-10 | 17:15 | 19.28 | 9.01 | 9.25 | 100.2 | 71 | 46.143
2006-09-10 | 17:30 | 19.23 | 9 | 9.16 | 99.2 | 71 | 46.17
2006-09-10 | 17:45 | 19.16 | 8.98 | 9.09 | 98.3 | 71 | 46.181
2006-09-10 | 18:00 | 19.08 | 8.97 | 9.03 | 97.5 | 71.1 | 46.202
2006-09-10 | 18:15 | 19 | 8.94 | 8.92 | 96.2 | 71.1 | 46.23
2006-09-10 | 18:30 | 18.9 | 8.91 | 8.82 | 94.9 | 71.2 | 46.273
2006-09-10 | 18:45 | 18.78 | 8.87 | 8.71 | 93.5 | 71.2 | 46.279
2006-09-10 | 19:00 | 18.67 | 8.82 | 8.6 | 92 | 71.3 | 46.338
2006-09-10 | 19:15 | 18.55 | 8.76 | 8.49 | 90.7 | 71.4 | 46.4
2006-09-10 | 19:30 | 18.43 | 8.68 | 8.38 | 89.3 | 71.4 | 46.441
2006-09-10 | 19:45 | 18.29 | 8.6 | 8.3 | 88.2 | 71.5 | 46.506
2006-09-10 | 20:00 | 18.15 | 8.5 | 8.21 | 87 | 71.7 | 46.62
2006-09-10 | 20:15 | 18.03 | 8.4 | 8.14 | 86 | 71.9 | 46.724
2006-09-10 | 20:30 | 17.89 | 8.31 | 8.1 | 85.3 | 72 | 46.811
2006-09-10 | 20:45 | 17.75 | 8.21 | 8.05 | 84.5 | 72.2 | 46.942
2006-09-10 | 21:00 | 17.6 | 8.11 | 8.01 | 83.9 | 72.4 | 47.032
2006-09-10 | 21:15 | 17.45 | 8.02 | 7.97 | 83.3 | 72.6 | 47.19
2006-09-10 | 21:30 | 17.31 | 7.95 | 7.95 | 82.8 | 72.7 | 47.258
2006-09-10 | 21:45 | 17.17 | 7.89 | 7.93 | 82.4 | 72.9 | 47.353
2006-09-10 | 22:00 | 17.02 | 7.84 | 7.92 | 82 | 73 | 47.471
2006-09-10 | 22:15 | 16.87 | 7.79 | 7.9 | 81.6 | 73.1 | 47.539
2006-09-10 | 22:30 | 16.73 | 7.75 | 7.89 | 81.2 | 73.3 | 47.614
2006-09-10 | 22:45 | 16.59 | 7.72 | 7.91 | 81.1 | 73.3 | 47.655
2006-09-10 | 23:00 | 16.46 | 7.69 | 7.91 | 80.9 | 73.4 | 47.707
2006-09-10 | 23:15 | 16.33 | 7.67 | 7.92 | 80.8 | 73.4 | 47.698
2006-09-10 | 23:30 | 16.22 | 7.65 | 7.93 | 80.7 | 73.4 | 47.702
2006-09-10 | 23:45 | 16.11 | 7.64 | 7.93 | 80.6 | 73.4 | 47.687
2006-09-11 | 0:00 | 16.01 | 7.63 | 7.95 | 80.6 | 73.3 | 47.66
2006-09-11 | 0:15 | 15.91 | 7.61 | 7.97 | 80.6 | 73.3 | 47.671
2006-09-11 | 0:30 | 15.82 | 7.6 | 7.97 | 80.4 | 73.2 | 47.595
2006-09-11 | 0:45 | 15.72 | 7.58 | 7.98 | 80.4 | 73.2 | 47.549
2006-09-11 | 1:01 | 15.64 | 7.59 | 8.01 | 80.5 | 73.1 | 47.507
2006-09-11 | 1:16 | 15.56 | 7.58 | 8.01 | 80.4 | 73 | 47.477
2006-09-11 | 1:31 | 15.48 | 7.58 | 8.05 | 80.6 | 73 | 47.426
2006-09-11 | 1:46 | 15.41 | 7.57 | 8.04 | 80.5 | 72.9 | 47.389
2006-09-11 | 2:01 | 15.34 | 7.57 | 8.07 | 80.7 | 72.8 | 47.35
2006-09-11 | 2:16 | 15.28 | 7.56 | 8.08 | 80.6 | 72.8 | 47.323
2006-09-11 | 2:31 | 15.23 | 7.56 | 8.08 | 80.6 | 72.7 | 47.277
2006-09-11 | 2:46 | 15.17 | 7.56 | 8.1 | 80.7 | 72.7 | 47.25
2006-09-11 | 3:01 | 15.12 | 7.55 | 8.1 | 80.6 | 72.6 | 47.201
2006-09-11 | 3:16 | 15.08 | 7.55 | 8.11 | 80.6 | 72.5 | 47.135
2006-09-11 | 3:31 | 15.04 | 7.55 | 8.12 | 80.6 | 72.4 | 47.078
2006-09-11 | 3:46 | 15.01 | 7.54 | 8.12 | 80.6 | 72.4 | 47.059
2006-09-11 | 4:01 | 14.98 | 7.54 | 8.14 | 80.7 | 72.3 | 46.989
2006-09-11 | 4:16 | 14.95 | 7.54 | 8.14 | 80.6 | 72.2 | 46.935
2006-09-11 | 4:31 | 14.92 | 7.53 | 8.14 | 80.6 | 72.2 | 46.898
2006-09-11 | 4:46 | 14.91 | 7.54 | 8.14 | 80.6 | 72 | 46.826
2006-09-11 | 5:01 | 14.9 | 7.53 | 8.15 | 80.6 | 72 | 46.78
2006-09-11 | 5:16 | 14.89 | 7.53 | 8.14 | 80.5 | 71.9 | 46.718
2006-09-11 | 5:31 | 14.87 | 7.53 | 8.14 | 80.5 | 71.8 | 46.679
2006-09-11 | 5:46 | 14.86 | 7.53 | 8.14 | 80.5 | 71.8 | 46.678
2006-09-11 | 6:01 | 14.86 | 7.52 | 8.14 | 80.5 | 71.7 | 46.61
2006-09-11 | 6:16 | 14.85 | 7.52 | 8.14 | 80.5 | 71.6 | 46.57
2006-09-11 | 6:31 | 14.86 | 7.53 | 8.15 | 80.5 | 71.6 | 46.563
2006-09-11 | 6:46 | 14.86 | 7.53 | 8.15 | 80.6 | 71.5 | 46.495
2006-09-11 | 7:01 | 14.86 | 7.53 | 8.18 | 80.9 | 71.5 | 46.461
2006-09-11 | 7:16 | 14.87 | 7.54 | 8.21 | 81.2 | 71.5 | 46.445
2006-09-11 | 7:31 | 14.89 | 7.55 | 8.27 | 81.8 | 71.4 | 46.396
2006-09-11 | 7:46 | 14.94 | 7.57 | 8.36 | 82.8 | 71.3 | 46.334
2006-09-11 | 8:01 | 14.98 | 7.6 | 8.46 | 83.9 | 71.2 | 46.28
2006-09-11 | 8:16 | 15.03 | 7.63 | 8.57 | 85 | 71.1 | 46.247
2006-09-11 | 8:31 | 15.1 | 7.66 | 8.67 | 86.2 | 71.1 | 46.214
2006-09-11 | 8:46 | 15.18 | 7.7 | 8.8 | 87.6 | 71 | 46.167
2006-09-11 | 9:01 | 15.29 | 7.75 | 8.92 | 89 | 70.9 | 46.106
2006-09-11 | 9:16 | 15.37 | 7.79 | 9.02 | 90.2 | 70.9 | 46.066
2006-09-11 | 9:31 | 15.49 | 7.84 | 9.12 | 91.4 | 70.9 | 46.068
2006-09-11 | 9:46 | 15.62 | 7.9 | 9.22 | 92.6 | 70.8 | 46.003
2006-09-11 | 10:01 | 15.78 | 7.97 | 9.34 | 94.2 | 70.7 | 45.958
2006-09-11 | 10:16 | 15.94 | 8.05 | 9.45 | 95.6 | 70.7 | 45.936
2006-09-11 | 10:31 | 16.06 | 8.13 | 9.52 | 96.6 | 70.6 | 45.898
2006-09-11 | 10:46 | 16.19 | 8.19 | 9.57 | 97.4 | 70.6 | 45.885
2006-09-11 | 11:01 | 16.38 | 8.27 | 9.64 | 98.5 | 70.6 | 45.906
2006-09-11 | 11:16 | 16.58 | 8.36 | 9.71 | 99.6 | 70.5 | 45.843
2006-09-11 | 11:31 | 16.73 | 8.42 | 9.73 | 100.1 | 70.6 | 45.873
2006-09-11 | 11:46 | 16.91 | 8.48 | 9.77 | 100.9 | 70.6 | 45.876
2006-09-11 | 12:01 | 17.1 | 8.54 | 9.8 | 101.6 | 70.6 | 45.901
2006-09-11 | 12:16 | 17.27 | 8.6 | 9.82 | 102.2 | 70.6 | 45.896
2006-09-11 | 12:31 | 17.43 | 8.66 | 9.83 | 102.6 | 70.6 | 45.914
2006-09-11 | 12:46 | 17.62 | 8.7 | 9.84 | 103.1 | 70.6 | 45.919
2006-09-11 | 13:00 | 17.8 | 8.75 | 9.84 | 103.5 | 70.7 | 45.937
2006-09-11 | 13:15 | 17.94 | 8.78 | 9.83 | 103.7 | 70.7 | 45.956
2006-09-11 | 13:30 | 18.09 | 8.81 | 9.81 | 103.8 | 70.8 | 45.991
2006-09-11 | 13:45 | 18.25 | 8.85 | 9.81 | 104.2 | 70.8 | 45.991
2006-09-11 | 14:00 | 18.39 | 8.88 | 9.79 | 104.3 | 70.7 | 45.969
2006-09-11 | 14:15 | 18.52 | 8.9 | 9.77 | 104.3 | 70.7 | 45.964
2006-09-11 | 14:30 | 18.63 | 8.93 | 9.75 | 104.3 | 70.7 | 45.978
2006-09-11 | 14:45 | 18.74 | 8.95 | 9.73 | 104.3 | 70.7 | 45.962
2006-09-11 | 15:00 | 18.84 | 8.96 | 9.69 | 104.2 | 70.7 | 45.958
2006-09-11 | 15:15 | 18.93 | 8.98 | 9.66 | 104 | 70.7 | 45.948
2006-09-11 | 15:30 | 19.01 | 8.99 | 9.62 | 103.8 | 70.7 | 45.969
2006-09-11 | 15:45 | 19.06 | 9.01 | 9.59 | 103.5 | 70.6 | 45.921
2006-09-11 | 16:00 | 19.13 | 9.01 | 9.53 | 103 | 70.6 | 45.919
2006-09-11 | 16:15 | 19.16 | 9.02 | 9.47 | 102.5 | 70.6 | 45.918
2006-09-11 | 16:30 | 19.18 | 9.03 | 9.43 | 102.1 | 70.6 | 45.916
2006-09-11 | 16:45 | 19.19 | 9.03 | 9.37 | 101.4 | 70.7 | 45.928
2006-09-11 | 17:00 | 19.17 | 9.04 | 9.32 | 100.8 | 70.6 | 45.914
2006-09-11 | 17:15 | 19.14 | 9.03 | 9.26 | 100.1 | 70.7 | 45.968
2006-09-11 | 17:30 | 19.08 | 9.02 | 9.18 | 99.2 | 70.7 | 45.956
2006-09-11 | 17:45 | 19.02 | 9.01 | 9.11 | 98.2 | 70.7 | 45.981
2006-09-11 | 18:00 | 18.95 | 9 | 9.04 | 97.3 | 70.8 | 46.036
2006-09-11 | 18:15 | 18.86 | 8.97 | 8.94 | 96.1 | 70.8 | 46.036
2006-09-11 | 18:30 | 18.77 | 8.94 | 8.85 | 94.9 | 70.9 | 46.077
2006-09-11 | 18:45 | 18.66 | 8.9 | 8.73 | 93.4 | 70.9 | 46.11
2006-09-11 | 19:00 | 18.55 | 8.85 | 8.62 | 92.1 | 71.1 | 46.188
2006-09-11 | 19:15 | 18.44 | 8.8 | 8.51 | 90.7 | 71.1 | 46.24
2006-09-11 | 19:30 | 18.34 | 8.73 | 8.4 | 89.3 | 71.3 | 46.323
2006-09-11 | 19:45 | 18.22 | 8.65 | 8.3 | 88.1 | 71.4 | 46.436
2006-09-11 | 20:00 | 18.1 | 8.56 | 8.21 | 86.9 | 71.5 | 46.5
2006-09-11 | 20:15 | 17.97 | 8.48 | 8.15 | 86 | 71.7 | 46.602
2006-09-11 | 20:30 | 17.85 | 8.38 | 8.09 | 85.2 | 71.9 | 46.729
2006-09-11 | 20:45 | 17.73 | 8.27 | 8.03 | 84.3 | 72 | 46.831
2006-09-11 | 21:00 | 17.6 | 8.17 | 7.99 | 83.7 | 72.3 | 46.964
2006-09-11 | 21:15 | 17.47 | 8.08 | 7.96 | 83.2 | 72.4 | 47.091
2006-09-11 | 21:30 | 17.35 | 7.99 | 7.93 | 82.7 | 72.6 | 47.212
2006-09-11 | 21:45 | 17.22 | 7.92 | 7.91 | 82.3 | 72.8 | 47.348
2006-09-11 | 22:00 | 17.09 | 7.85 | 7.9 | 81.9 | 73 | 47.422
2006-09-11 | 22:15 | 16.96 | 7.81 | 7.89 | 81.5 | 73.1 | 47.527
2006-09-11 | 22:30 | 16.83 | 7.76 | 7.89 | 81.3 | 73.2 | 47.563
2006-09-11 | 22:45 | 16.71 | 7.73 | 7.88 | 81.1 | 73.3 | 47.633
2006-09-11 | 23:00 | 16.59 | 7.7 | 7.88 | 80.9 | 73.3 | 47.647
2006-09-11 | 23:15 | 16.49 | 7.67 | 7.89 | 80.8 | 73.3 | 47.626
2006-09-11 | 23:30 | 16.38 | 7.66 | 7.9 | 80.6 | 73.2 | 47.59
2006-09-11 | 23:45 | 16.28 | 7.64 | 7.91 | 80.6 | 73.2 | 47.589
2006-09-12 | 0:00 | 16.19 | 7.62 | 7.92 | 80.5 | 73.1 | 47.529
2006-09-12 | 0:15 | 16.1 | 7.61 | 7.93 | 80.5 | 73 | 47.455
2006-09-12 | 0:30 | 16.01 | 7.6 | 7.93 | 80.4 | 72.9 | 47.412
2006-09-12 | 0:45 | 15.94 | 7.6 | 7.95 | 80.4 | 72.8 | 47.334
2006-09-12 | 1:01 | 15.86 | 7.58 | 7.96 | 80.4 | 72.8 | 47.299
2006-09-12 | 1:16 | 15.79 | 7.57 | 7.96 | 80.3 | 72.6 | 47.219
2006-09-12 | 1:31 | 15.73 | 7.57 | 7.97 | 80.3 | 72.6 | 47.158
2006-09-12 | 1:46 | 15.67 | 7.57 | 7.99 | 80.3 | 72.5 | 47.123
2006-09-12 | 2:01 | 15.62 | 7.56 | 8 | 80.4 | 72.4 | 47.064
2006-09-12 | 2:16 | 15.57 | 7.55 | 8 | 80.4 | 72.3 | 47.006
2006-09-12 | 2:31 | 15.53 | 7.55 | 8 | 80.3 | 72.3 | 46.986
2006-09-12 | 2:46 | 15.49 | 7.55 | 8.01 | 80.3 | 72.1 | 46.896
2006-09-12 | 3:01 | 15.46 | 7.54 | 8.01 | 80.2 | 72.1 | 46.846
2006-09-12 | 3:16 | 15.43 | 7.54 | 8.01 | 80.2 | 72 | 46.778
2006-09-12 | 3:31 | 15.4 | 7.53 | 8.01 | 80.1 | 71.9 | 46.722
2006-09-12 | 3:46 | 15.39 | 7.53 | 8.01 | 80.1 | 71.8 | 46.671
2006-09-12 | 4:01 | 15.37 | 7.53 | 8.01 | 80.1 | 71.7 | 46.586
2006-09-12 | 4:16 | 15.35 | 7.53 | 8.01 | 80.1 | 71.6 | 46.541
2006-09-12 | 4:31 | 15.33 | 7.52 | 8.01 | 80 | 71.6 | 46.513
2006-09-12 | 4:46 | 15.33 | 7.52 | 8.01 | 80 | 71.5 | 46.469
2006-09-12 | 5:01 | 15.31 | 7.52 | 8.01 | 79.9 | 71.5 | 46.45
2006-09-12 | 5:16 | 15.31 | 7.52 | 8 | 79.9 | 71.3 | 46.376
2006-09-12 | 5:31 | 15.3 | 7.52 | 8 | 79.9 | 71.3 | 46.351
2006-09-12 | 5:46 | 15.3 | 7.51 | 7.98 | 79.7 | 71.2 | 46.312
2006-09-12 | 6:01 | 15.28 | 7.51 | 7.99 | 79.7 | 71.3 | 46.338
2006-09-12 | 6:16 | 15.28 | 7.52 | 8 | 79.9 | 71.3 | 46.321
2006-09-12 | 6:31 | 15.28 | 7.51 | 7.99 | 79.7 | 71.2 | 46.311
2006-09-12 | 6:46 | 15.28 | 7.51 | 7.99 | 79.8 | 71.2 | 46.266
2006-09-12 | 7:01 | 15.27 | 7.52 | 8.02 | 80 | 71.2 | 46.277
2006-09-12 | 7:16 | 15.28 | 7.53 | 8.06 | 80.4 | 71.2 | 46.25
2006-09-12 | 7:31 | 15.29 | 7.53 | 8.11 | 80.9 | 71.2 | 46.251
2006-09-12 | 7:46 | 15.31 | 7.55 | 8.18 | 81.7 | 71.1 | 46.229
2006-09-12 | 8:01 | 15.35 | 7.57 | 8.28 | 82.8 | 71.1 | 46.247
2006-09-12 | 8:16 | 15.42 | 7.6 | 8.39 | 84 | 71.1 | 46.243
2006-09-12 | 8:31 | 15.48 | 7.64 | 8.51 | 85.3 | 71.1 | 46.222
2006-09-12 | 8:46 | 15.56 | 7.67 | 8.61 | 86.5 | 71.1 | 46.243
2006-09-12 | 9:01 | 15.65 | 7.72 | 8.72 | 87.7 | 71.2 | 46.248
2006-09-12 | 9:16 | 15.75 | 7.76 | 8.84 | 89.1 | 71.1 | 46.235
2006-09-12 | 9:31 | 15.88 | 7.82 | 8.95 | 90.5 | 71.1 | 46.237
2006-09-12 | 9:46 | 16.02 | 7.87 | 9.05 | 91.8 | 71.1 | 46.221
2006-09-12 | 10:01 | 16.17 | 7.94 | 9.16 | 93.2 | 71.1 | 46.229
2006-09-12 | 10:16 | 16.34 | 8.01 | 9.25 | 94.4 | 71.1 | 46.212
2006-09-12 | 10:31 | 16.51 | 8.08 | 9.33 | 95.6 | 71.1 | 46.212
2006-09-12 | 10:46 | 16.69 | 8.16 | 9.4 | 96.7 | 71.1 | 46.225
2006-09-12 | 11:01 | 16.88 | 8.24 | 9.48 | 97.8 | 71.1 | 46.227
2006-09-12 | 11:16 | 17.06 | 8.31 | 9.52 | 98.7 | 71.1 | 46.216
2006-09-12 | 11:31 | 17.26 | 8.38 | 9.56 | 99.5 | 71.1 | 46.22
2006-09-12 | 11:46 | 17.46 | 8.46 | 9.61 | 100.4 | 71.1 | 46.223
2006-09-12 | 12:01 | 17.67 | 8.52 | 9.63 | 101 | 71.2 | 46.287
2006-09-12 | 12:16 | 17.87 | 8.59 | 9.65 | 101.7 | 71.2 | 46.277
2006-09-12 | 12:31 | 18.07 | 8.63 | 9.65 | 102.1 | 71.3 | 46.32
2006-09-12 | 12:46 | 18.26 | 8.69 | 9.66 | 102.6 | 71.3 | 46.349
2006-09-12 | 13:00 | 18.45 | 8.74 | 9.66 | 103 | 71.4 | 46.398
2006-09-12 | 13:15 | 18.63 | 8.78 | 9.65 | 103.3 | 71.4 | 46.41
2006-09-12 | 13:30 | 18.82 | 8.81 | 9.64 | 103.5 | 71.5 | 46.451
2006-09-12 | 13:45 | 18.97 | 8.85 | 9.62 | 103.6 | 71.5 | 46.468
2006-09-12 | 14:00 | 19.15 | 8.87 | 9.59 | 103.7 | 71.5 | 46.506
2006-09-12 | 14:15 | 19.29 | 8.9 | 9.57 | 103.8 | 71.6 | 46.523
2006-09-12 | 14:30 | 19.43 | 8.93 | 9.54 | 103.7 | 71.6 | 46.552
2006-09-12 | 14:45 | 19.55 | 8.95 | 9.5 | 103.6 | 71.6 | 46.565
2006-09-12 | 15:00 | 19.65 | 8.97 | 9.46 | 103.4 | 71.6 | 46.57
2006-09-12 | 15:15 | 19.74 | 8.98 | 9.41 | 103 | 71.7 | 46.591
2006-09-12 | 15:30 | 19.81 | 9 | 9.39 | 102.9 | 71.7 | 46.619
2006-09-12 | 15:45 | 19.86 | 9.01 | 9.35 | 102.6 | 71.7 | 46.629
2006-09-12 | 16:00 | 19.9 | 9.01 | 9.31 | 102.2 | 71.7 | 46.62
2006-09-12 | 16:15 | 19.94 | 9.03 | 9.26 | 101.7 | 71.7 | 46.606
2006-09-12 | 16:30 | 19.95 | 9.03 | 9.2 | 101.1 | 71.7 | 46.605
2006-09-12 | 16:45 | 19.96 | 9.03 | 9.15 | 100.6 | 71.7 | 46.627
2006-09-12 | 17:00 | 19.94 | 9.04 | 9.08 | 99.8 | 71.8 | 46.677
2006-09-12 | 17:15 | 19.91 | 9.03 | 9.02 | 99 | 71.8 | 46.661
2006-09-12 | 17:30 | 19.87 | 9.02 | 8.95 | 98.2 | 71.8 | 46.696
2006-09-12 | 17:45 | 19.8 | 9.01 | 8.87 | 97.2 | 71.9 | 46.724
2006-09-12 | 18:00 | 19.72 | 9 | 8.77 | 96 | 72 | 46.785
2006-09-12 | 18:15 | 19.61 | 8.97 | 8.68 | 94.8 | 72.1 | 46.838
2006-09-12 | 18:30 | 19.49 | 8.94 | 8.59 | 93.5 | 72.1 | 46.892
2006-09-12 | 18:45 | 19.38 | 8.9 | 8.46 | 91.9 | 72.3 | 46.97
2006-09-12 | 19:00 | 19.25 | 8.85 | 8.36 | 90.6 | 72.4 | 47.035
2006-09-12 | 19:15 | 19.15 | 8.79 | 8.24 | 89.1 | 72.5 | 47.116
2006-09-12 | 19:30 | 19.03 | 8.73 | 8.14 | 87.9 | 72.6 | 47.213
2006-09-12 | 19:45 | 18.91 | 8.65 | 8.03 | 86.4 | 72.8 | 47.327
2006-09-12 | 20:00 | 18.78 | 8.56 | 7.95 | 85.4 | 73 | 47.459
2006-09-12 | 20:15 | 18.65 | 8.47 | 7.88 | 84.4 | 73.2 | 47.554
2006-09-12 | 20:30 | 18.51 | 8.37 | 7.83 | 83.6 | 73.3 | 47.671
2006-09-12 | 20:45 | 18.38 | 8.27 | 7.78 | 82.8 | 73.6 | 47.814
2006-09-12 | 21:00 | 18.24 | 8.17 | 7.74 | 82.2 | 73.8 | 47.965
2006-09-12 | 21:15 | 18.1 | 8.07 | 7.71 | 81.6 | 74 | 48.12
2006-09-12 | 21:30 | 17.96 | 7.99 | 7.68 | 81 | 74.2 | 48.202
2006-09-12 | 21:45 | 17.82 | 7.91 | 7.66 | 80.7 | 74.4 | 48.332
2006-09-12 | 22:00 | 17.67 | 7.85 | 7.65 | 80.2 | 74.5 | 48.426
2006-09-12 | 22:15 | 17.53 | 7.8 | 7.64 | 79.9 | 74.7 | 48.53
2006-09-12 | 22:30 | 17.4 | 7.76 | 7.63 | 79.6 | 74.8 | 48.615
2006-09-12 | 22:45 | 17.26 | 7.72 | 7.63 | 79.4 | 74.9 | 48.676
2006-09-12 | 23:00 | 17.11 | 7.69 | 7.64 | 79.3 | 74.9 | 48.709
2006-09-12 | 23:15 | 16.98 | 7.67 | 7.65 | 79.1 | 74.9 | 48.713
2006-09-12 | 23:30 | 16.85 | 7.64 | 7.66 | 79 | 74.9 | 48.708
2006-09-12 | 23:45 | 16.74 | 7.62 | 7.66 | 78.8 | 74.9 | 48.698
2006-09-13 | 0:00 | 16.62 | 7.61 | 7.68 | 78.8 | 74.9 | 48.653
2006-09-13 | 0:15 | 16.49 | 7.6 | 7.7 | 78.8 | 74.8 | 48.634
2006-09-13 | 0:30 | 16.39 | 7.58 | 7.71 | 78.8 | 74.7 | 48.568
2006-09-13 | 0:45 | 16.28 | 7.58 | 7.74 | 78.9 | 74.6 | 48.517
2006-09-13 | 1:01 | 16.19 | 7.57 | 7.75 | 78.9 | 74.6 | 48.503
2006-09-13 | 1:16 | 16.09 | 7.56 | 7.77 | 78.9 | 74.5 | 48.441
2006-09-13 | 1:31 | 16 | 7.55 | 7.77 | 78.8 | 74.4 | 48.375
2006-09-13 | 1:46 | 15.91 | 7.55 | 7.81 | 78.9 | 74.3 | 48.311
2006-09-13 | 2:01 | 15.82 | 7.55 | 7.82 | 78.9 | 74.3 | 48.293
2006-09-13 | 2:16 | 15.74 | 7.54 | 7.84 | 79 | 74.2 | 48.244
2006-09-13 | 2:31 | 15.67 | 7.54 | 7.84 | 78.9 | 74.1 | 48.187
2006-09-13 | 2:46 | 15.6 | 7.53 | 7.85 | 78.9 | 74.1 | 48.16
2006-09-13 | 3:01 | 15.54 | 7.53 | 7.87 | 78.9 | 74.1 | 48.138
2006-09-13 | 3:16 | 15.48 | 7.53 | 7.88 | 78.9 | 74 | 48.073
2006-09-13 | 3:31 | 15.43 | 7.53 | 7.89 | 79 | 73.9 | 48.043
2006-09-13 | 3:46 | 15.38 | 7.52 | 7.89 | 78.9 | 73.9 | 48.036
2006-09-13 | 4:01 | 15.34 | 7.53 | 7.91 | 79 | 73.8 | 47.957
2006-09-13 | 4:16 | 15.3 | 7.52 | 7.91 | 78.9 | 73.7 | 47.934
2006-09-13 | 4:31 | 15.27 | 7.52 | 7.91 | 79 | 73.6 | 47.872
2006-09-13 | 4:46 | 15.24 | 7.52 | 7.92 | 79 | 73.6 | 47.835
2006-09-13 | 5:01 | 15.22 | 7.52 | 7.94 | 79.1 | 73.5 | 47.784
2006-09-13 | 5:16 | 15.18 | 7.51 | 7.93 | 78.9 | 73.5 | 47.753
2006-09-13 | 5:31 | 15.16 | 7.51 | 7.93 | 78.9 | 73.4 | 47.718
2006-09-13 | 5:46 | 15.14 | 7.51 | 7.93 | 78.9 | 73.3 | 47.661
2006-09-13 | 6:01 | 15.14 | 7.52 | 7.94 | 78.9 | 73.3 | 47.614
2006-09-13 | 6:16 | 15.12 | 7.51 | 7.93 | 78.9 | 73.3 | 47.621
2006-09-13 | 6:31 | 15.11 | 7.51 | 7.93 | 78.9 | 73.2 | 47.551
2006-09-13 | 6:46 | 15.11 | 7.52 | 7.93 | 78.8 | 73.2 | 47.562
2006-09-13 | 7:01 | 15.1 | 7.51 | 7.96 | 79.1 | 73.1 | 47.484
2006-09-13 | 7:16 | 15.1 | 7.52 | 8.01 | 79.6 | 73 | 47.442
2006-09-13 | 7:31 | 15.11 | 7.53 | 8.05 | 80.1 | 72.9 | 47.387
2006-09-13 | 7:46 | 15.15 | 7.55 | 8.12 | 80.8 | 72.8 | 47.342
2006-09-13 | 8:01 | 15.21 | 7.57 | 8.22 | 81.9 | 72.7 | 47.287
2006-09-13 | 8:16 | 15.27 | 7.6 | 8.33 | 83.1 | 72.6 | 47.2
2006-09-13 | 8:31 | 15.35 | 7.63 | 8.44 | 84.3 | 72.6 | 47.161
2006-09-13 | 8:46 | 15.45 | 7.67 | 8.55 | 85.6 | 72.5 | 47.097
2006-09-13 | 9:01 | 15.55 | 7.7 | 8.64 | 86.7 | 72.4 | 47.075
2006-09-13 | 9:16 | 15.68 | 7.74 | 8.74 | 88 | 72.4 | 47.043
2006-09-13 | 9:31 | 15.83 | 7.8 | 8.85 | 89.4 | 72.3 | 46.97
2006-09-13 | 9:46 | 15.99 | 7.85 | 8.96 | 90.7 | 72.1 | 46.895
2006-09-13 | 10:01 | 16.15 | 7.91 | 9.05 | 92 | 72.1 | 46.851
2006-09-13 | 10:16 | 16.33 | 7.98 | 9.14 | 93.3 | 72 | 46.801
2006-09-13 | 10:31 | 16.51 | 8.05 | 9.22 | 94.4 | 72 | 46.785
2006-09-13 | 10:46 | 16.7 | 8.12 | 9.3 | 95.6 | 71.9 | 46.741
2006-09-13 | 11:01 | 16.89 | 8.2 | 9.36 | 96.6 | 71.9 | 46.726
2006-09-13 | 11:16 | 17.09 | 8.27 | 9.4 | 97.4 | 71.9 | 46.714
2006-09-13 | 11:31 | 17.25 | 8.35 | 9.45 | 98.3 | 71.9 | 46.735
2006-09-13 | 11:46 | 17.44 | 8.42 | 9.47 | 98.9 | 71.9 | 46.726
2006-09-13 | 12:01 | 17.63 | 8.48 | 9.49 | 99.5 | 72 | 46.772
2006-09-13 | 12:16 | 17.82 | 8.55 | 9.52 | 100.2 | 72 | 46.772
2006-09-13 | 12:31 | 18.01 | 8.6 | 9.52 | 100.6 | 72.1 | 46.843
2006-09-13 | 12:46 | 18.17 | 8.64 | 9.51 | 100.8 | 72.1 | 46.862
2006-09-13 | 13:00 | 18.31 | 8.68 | 9.52 | 101.2 | 72.2 | 46.915
2006-09-13 | 13:15 | 18.45 | 8.73 | 9.52 | 101.5 | 72.2 | 46.955
2006-09-13 | 13:30 | 18.56 | 8.77 | 9.51 | 101.6 | 72.2 | 46.949
2006-09-13 | 13:45 | 18.7 | 8.8 | 9.5 | 101.7 | 72.3 | 46.98
2006-09-13 | 14:00 | 18.8 | 8.83 | 9.47 | 101.7 | 72.3 | 47.003
2006-09-13 | 14:15 | 18.88 | 8.87 | 9.48 | 102 | 72.4 | 47.04
2006-09-13 | 14:30 | 18.95 | 8.89 | 9.46 | 101.8 | 72.3 | 47.017
2006-09-13 | 14:45 | 19.01 | 8.9 | 9.42 | 101.6 | 72.3 | 47.013
2006-09-13 | 15:00 | 19.04 | 8.93 | 9.43 | 101.7 | 72.3 | 47.013
2006-09-13 | 15:15 | 19.06 | 8.94 | 9.37 | 101.1 | 72.3 | 47.023
2006-09-13 | 15:30 | 19.08 | 8.96 | 9.35 | 101 | 72.3 | 47.001
2006-09-13 | 15:45 | 19.06 | 8.97 | 9.32 | 100.6 | 72.3 | 46.991
2006-09-13 | 16:00 | 19.01 | 8.98 | 9.28 | 100 | 72.3 | 46.997
2006-09-13 | 16:15 | 18.96 | 8.99 | 9.24 | 99.5 | 72.3 | 47.005
2006-09-13 | 16:30 | 18.91 | 8.99 | 9.2 | 99 | 72.4 | 47.034
2006-09-13 | 16:45 | 18.89 | 8.99 | 9.14 | 98.3 | 72.4 | 47.054
2006-09-13 | 17:00 | 18.84 | 9 | 9.09 | 97.7 | 72.5 | 47.104
2006-09-13 | 17:15 | 18.78 | 8.99 | 9.03 | 96.9 | 72.5 | 47.154
2006-09-13 | 17:30 | 18.68 | 8.99 | 8.98 | 96.2 | 72.6 | 47.208
2006-09-13 | 17:45 | 18.56 | 8.97 | 8.9 | 95.1 | 72.7 | 47.255
2006-09-13 | 18:00 | 18.44 | 8.95 | 8.79 | 93.7 | 72.8 | 47.352
2006-09-13 | 18:15 | 18.32 | 8.92 | 8.7 | 92.5 | 73 | 47.433
2006-09-13 | 18:30 | 18.21 | 8.89 | 8.58 | 91 | 73.2 | 47.568
2006-09-13 | 18:45 | 18.1 | 8.83 | 8.45 | 89.5 | 73.3 | 47.664
2006-09-13 | 19:00 | 17.99 | 8.78 | 8.34 | 88.1 | 73.5 | 47.785
2006-09-13 | 19:15 | 17.84 | 8.71 | 8.24 | 86.8 | 73.7 | 47.893
2006-09-13 | 19:30 | 17.74 | 8.64 | 8.15 | 85.6 | 73.9 | 48.059
2006-09-13 | 19:45 | 17.65 | 8.56 | 8.07 | 84.7 | 74.1 | 48.148
2006-09-13 | 20:00 | 17.52 | 8.47 | 8 | 83.7 | 74.3 | 48.315
2006-09-13 | 20:15 | 17.38 | 8.38 | 7.95 | 83 | 74.6 | 48.483
2006-09-13 | 20:30 | 17.27 | 8.29 | 7.9 | 82.2 | 74.8 | 48.605
2006-09-13 | 20:45 | 17.14 | 8.19 | 7.88 | 81.8 | 75 | 48.763
2006-09-13 | 21:00 | 17.02 | 8.11 | 7.86 | 81.4 | 75.2 | 48.87
2006-09-13 | 21:15 | 16.91 | 8.03 | 7.83 | 80.9 | 75.4 | 49.009
2006-09-13 | 21:30 | 16.78 | 7.96 | 7.83 | 80.6 | 75.6 | 49.117
2006-09-13 | 21:45 | 16.67 | 7.9 | 7.81 | 80.2 | 75.8 | 49.291
2006-09-13 | 22:00 | 16.59 | 7.84 | 7.79 | 79.9 | 75.9 | 49.355
2006-09-13 | 22:15 | 16.48 | 7.8 | 7.77 | 79.5 | 76.1 | 49.465
2006-09-13 | 22:30 | 16.4 | 7.77 | 7.77 | 79.4 | 76.1 | 49.493
2006-09-13 | 22:45 | 16.31 | 7.73 | 7.76 | 79.1 | 76.2 | 49.534
2006-09-13 | 23:00 | 16.22 | 7.71 | 7.76 | 79 | 76.1 | 49.496
2006-09-13 | 23:15 | 16.1 | 7.69 | 7.77 | 78.9 | 76.2 | 49.524
2006-09-13 | 23:30 | 16.02 | 7.66 | 7.78 | 78.9 | 76 | 49.401
2006-09-13 | 23:45 | 15.96 | 7.65 | 7.79 | 78.9 | 75.9 | 49.33
2006-09-14 | 0:00 | 15.88 | 7.64 | 7.79 | 78.8 | 75.8 | 49.267
2006-09-14 | 0:15 | 15.81 | 7.63 | 7.8 | 78.7 | 75.8 | 49.243
2006-09-14 | 0:30 | 15.73 | 7.62 | 7.82 | 78.8 | 75.6 | 49.137
2006-09-14 | 0:45 | 15.65 | 7.6 | 7.82 | 78.7 | 75.5 | 49.095
2006-09-14 | 1:01 | 15.59 | 7.6 | 7.83 | 78.7 | 75.4 | 49.021
2006-09-14 | 1:16 | 15.52 | 7.59 | 7.84 | 78.7 | 75.4 | 48.98
2006-09-14 | 1:31 | 15.45 | 7.59 | 7.85 | 78.7 | 75.3 | 48.927
2006-09-14 | 1:46 | 15.39 | 7.59 | 7.87 | 78.7 | 75.3 | 48.94
2006-09-14 | 2:01 | 15.33 | 7.58 | 7.87 | 78.6 | 75.2 | 48.905
2006-09-14 | 2:16 | 15.27 | 7.57 | 7.89 | 78.7 | 75.2 | 48.855
2006-09-14 | 2:31 | 15.2 | 7.57 | 7.89 | 78.6 | 75.1 | 48.834
2006-09-14 | 2:46 | 15.14 | 7.57 | 7.91 | 78.7 | 75.1 | 48.815
2006-09-14 | 3:01 | 15.08 | 7.56 | 7.9 | 78.5 | 75.1 | 48.822
2006-09-14 | 3:16 | 15.04 | 7.56 | 7.91 | 78.5 | 75 | 48.743
2006-09-14 | 3:31 | 14.98 | 7.56 | 7.92 | 78.5 | 74.9 | 48.717
2006-09-14 | 3:46 | 14.93 | 7.55 | 7.93 | 78.6 | 74.9 | 48.69
2006-09-14 | 4:01 | 14.87 | 7.56 | 7.94 | 78.6 | 74.9 | 48.676
2006-09-14 | 4:16 | 14.85 | 7.55 | 7.96 | 78.7 | 74.8 | 48.593
2006-09-14 | 4:31 | 14.81 | 7.55 | 7.97 | 78.7 | 74.7 | 48.551
2006-09-14 | 4:46 | 14.77 | 7.55 | 7.91 | 78.1 | 74.7 | 48.548
2006-09-14 | 5:01 | 14.72 | 7.54 | 7.98 | 78.6 | 74.6 | 48.505
2006-09-14 | 5:16 | 14.67 | 7.54 | 7.98 | 78.6 | 74.5 | 48.444
2006-09-14 | 5:31 | 14.63 | 7.54 | 8 | 78.7 | 74.5 | 48.405
2006-09-14 | 5:46 | 14.6 | 7.54 | 8 | 78.7 | 74.4 | 48.367
2006-09-14 | 6:01 | 14.56 | 7.54 | 8 | 78.6 | 74.4 | 48.391
2006-09-14 | 6:16 | 14.52 | 7.54 | 8.01 | 78.6 | 74.4 | 48.329
2006-09-14 | 6:31 | 14.48 | 7.54 | 8.01 | 78.6 | 74.4 | 48.329
2006-09-14 | 6:46 | 14.45 | 7.54 | 8.03 | 78.7 | 74.3 | 48.306
2006-09-14 | 7:01 | 14.41 | 7.54 | 8.05 | 78.8 | 74.3 | 48.272
2006-09-14 | 7:16 | 14.38 | 7.55 | 8.1 | 79.3 | 74.2 | 48.237
2006-09-14 | 7:31 | 14.34 | 7.57 | 8.16 | 79.7 | 74.2 | 48.216
2006-09-14 | 7:46 | 14.35 | 7.58 | 8.23 | 80.5 | 74.1 | 48.179
2006-09-14 | 8:01 | 14.37 | 7.6 | 8.34 | 81.6 | 74.1 | 48.135
2006-09-14 | 8:16 | 14.38 | 7.64 | 8.44 | 82.6 | 74 | 48.12
2006-09-14 | 8:31 | 14.4 | 7.67 | 8.55 | 83.7 | 74 | 48.086
2006-09-14 | 8:46 | 14.44 | 7.71 | 8.67 | 84.9 | 73.9 | 48.021
2006-09-14 | 9:01 | 14.48 | 7.74 | 8.78 | 86.1 | 73.9 | 48.024
2006-09-14 | 9:16 | 14.53 | 7.79 | 8.89 | 87.3 | 73.8 | 47.992
2006-09-14 | 9:31 | 14.53 | 7.83 | 8.99 | 88.2 | 73.8 | 47.955
2006-09-14 | 9:46 | 14.62 | 7.88 | 9.08 | 89.3 | 73.7 | 47.894
2006-09-14 | 10:01 | 14.71 | 7.94 | 9.2 | 90.7 | 73.6 | 47.847
2006-09-14 | 10:16 | 14.78 | 8 | 9.28 | 91.6 | 73.6 | 47.818
2006-09-14 | 10:31 | 14.84 | 8.05 | 9.36 | 92.6 | 73.5 | 47.789
2006-09-14 | 10:46 | 15 | 8.14 | 9.46 | 93.8 | 73.4 | 47.742
2006-09-14 | 11:01 | 15.14 | 8.21 | 9.54 | 94.9 | 73.4 | 47.732
2006-09-14 | 11:16 | 15.29 | 8.29 | 9.61 | 95.9 | 73.4 | 47.719
2006-09-14 | 11:31 | 15.41 | 8.35 | 9.65 | 96.6 | 73.4 | 47.702
2006-09-14 | 11:46 | 15.53 | 8.41 | 9.68 | 97.1 | 73.4 | 47.718
2006-09-14 | 12:01 | 15.67 | 8.49 | 9.72 | 97.8 | 73.3 | 47.656
2006-09-14 | 12:16 | 15.83 | 8.54 | 9.75 | 98.4 | 73.4 | 47.689
2006-09-14 | 12:31 | 15.97 | 8.6 | 9.77 | 98.9 | 73.3 | 47.677
2006-09-14 | 12:46 | 16.11 | 8.64 | 9.76 | 99.1 | 73.4 | 47.701
2006-09-14 | 13:00 | 16.26 | 8.68 | 9.74 | 99.3 | 73.4 | 47.687
2006-09-14 | 13:15 | 16.39 | 8.72 | 9.76 | 99.7 | 73.4 | 47.68
2006-09-14 | 13:30 | 16.51 | 8.77 | 9.75 | 99.9 | 73.4 | 47.683
2006-09-14 | 13:45 | 16.63 | 8.8 | 9.75 | 100.1 | 73.3 | 47.674
2006-09-14 | 14:00 | 16.73 | 8.82 | 9.74 | 100.2 | 73.4 | 47.689
2006-09-14 | 14:15 | 16.84 | 8.85 | 9.72 | 100.2 | 73.3 | 47.654
2006-09-14 | 14:30 | 16.93 | 8.89 | 9.71 | 100.3 | 73.3 | 47.658
2006-09-14 | 14:45 | 17.02 | 8.9 | 9.68 | 100.2 | 73.3 | 47.615
2006-09-14 | 15:00 | 17.04 | 8.93 | 9.64 | 99.9 | 73.2 | 47.606
2006-09-14 | 15:15 | 16.97 | 8.92 | 9.56 | 98.9 | 73.2 | 47.574
2006-09-14 | 15:30 | 16.87 | 8.92 | 9.49 | 98 | 73.1 | 47.506
2006-09-14 | 15:45 | 16.81 | 8.92 | 9.45 | 97.4 | 73.1 | 47.485
2006-09-14 | 16:00 | 16.77 | 8.92 | 9.4 | 96.8 | 73 | 47.475
2006-09-14 | 16:15 | 16.74 | 8.92 | 9.34 | 96.1 | 73 | 47.474
2006-09-14 | 16:30 | 16.71 | 8.92 | 9.31 | 95.8 | 73 | 47.424
2006-09-14 | 16:45 | 16.68 | 8.92 | 9.27 | 95.2 | 72.9 | 47.379
2006-09-14 | 17:00 | 16.66 | 8.92 | 9.22 | 94.7 | 72.9 | 47.363
2006-09-14 | 17:15 | 16.58 | 8.91 | 9.16 | 93.9 | 72.9 | 47.355
2006-09-14 | 17:30 | 16.45 | 8.89 | 9.07 | 92.8 | 72.9 | 47.394
2006-09-14 | 17:45 | 16.33 | 8.85 | 8.96 | 91.4 | 73 | 47.458
2006-09-14 | 18:00 | 16.22 | 8.82 | 8.9 | 90.6 | 73.1 | 47.491
2006-09-14 | 18:15 | 16.13 | 8.79 | 8.83 | 89.7 | 73.1 | 47.539
2006-09-14 | 18:30 | 16.07 | 8.74 | 8.76 | 88.9 | 73.2 | 47.593
2006-09-14 | 18:45 | 15.97 | 8.7 | 8.7 | 88.1 | 73.3 | 47.643
2006-09-14 | 19:00 | 15.85 | 8.63 | 8.61 | 87 | 73.4 | 47.723
2006-09-14 | 19:15 | 15.73 | 8.55 | 8.53 | 86 | 73.5 | 47.795
2006-09-14 | 19:30 | 15.61 | 8.48 | 8.48 | 85.2 | 73.7 | 47.875
2006-09-14 | 19:45 | 15.49 | 8.39 | 8.43 | 84.5 | 73.8 | 47.946
2006-09-14 | 20:00 | 15.39 | 8.31 | 8.39 | 83.9 | 73.9 | 48.034
2006-09-14 | 20:15 | 15.27 | 8.21 | 8.35 | 83.3 | 74 | 48.127
2006-09-14 | 20:30 | 15.16 | 8.13 | 8.32 | 82.8 | 74.2 | 48.245
2006-09-14 | 20:45 | 15.04 | 8.05 | 8.3 | 82.3 | 74.3 | 48.321
2006-09-14 | 21:00 | 14.94 | 7.99 | 8.28 | 82 | 74.5 | 48.414
2006-09-14 | 21:15 | 14.84 | 7.93 | 8.27 | 81.8 | 74.5 | 48.449
2006-09-14 | 21:30 | 14.75 | 7.89 | 8.26 | 81.5 | 74.6 | 48.483
2006-09-14 | 21:45 | 14.68 | 7.84 | 8.25 | 81.3 | 74.7 | 48.553
2006-09-14 | 22:00 | 14.57 | 7.81 | 8.25 | 81.1 | 74.7 | 48.561
2006-09-14 | 22:15 | 14.49 | 7.78 | 8.25 | 80.9 | 74.8 | 48.588
2006-09-14 | 22:30 | 14.41 | 7.75 | 8.24 | 80.7 | 74.8 | 48.607
2006-09-14 | 22:45 | 14.35 | 7.73 | 8.23 | 80.5 | 74.7 | 48.58
2006-09-14 | 23:00 | 14.28 | 7.71 | 8.24 | 80.5 | 74.7 | 48.534
2006-09-14 | 23:15 | 14.21 | 7.7 | 8.25 | 80.4 | 74.6 | 48.504
2006-09-14 | 23:30 | 14.15 | 7.67 | 8.24 | 80.2 | 74.6 | 48.493
2006-09-14 | 23:45 | 14.09 | 7.67 | 8.26 | 80.3 | 74.5 | 48.426
2006-09-15 | 0:00 | 14.04 | 7.66 | 8.26 | 80.3 | 74.5 | 48.396
2006-09-15 | 0:15 | 13.99 | 7.65 | 8.27 | 80.2 | 74.4 | 48.363
2006-09-15 | 0:30 | 13.95 | 7.64 | 8.28 | 80.2 | 74.4 | 48.339
2006-09-15 | 0:45 | 13.91 | 7.63 | 8.28 | 80.2 | 74.3 | 48.286
2006-09-15 | 1:01 | 13.88 | 7.62 | 8.28 | 80.1 | 74.2 | 48.25
2006-09-15 | 1:16 | 13.84 | 7.62 | 8.29 | 80.2 | 74.2 | 48.221
2006-09-15 | 1:31 | 13.81 | 7.61 | 8.29 | 80.1 | 74.1 | 48.184
2006-09-15 | 1:46 | 13.78 | 7.61 | 8.3 | 80.1 | 74.1 | 48.153
2006-09-15 | 2:01 | 13.76 | 7.6 | 8.29 | 80 | 74.1 | 48.144
2006-09-15 | 2:16 | 13.74 | 7.6 | 8.29 | 80 | 74.1 | 48.136
2006-09-15 | 2:31 | 13.71 | 7.6 | 8.3 | 80 | 73.9 | 48.059
2006-09-15 | 2:46 | 13.69 | 7.58 | 8.29 | 79.9 | 74.1 | 48.162
2006-09-15 | 3:01 | 13.68 | 7.59 | 8.29 | 79.9 | 74.1 | 48.181
2006-09-15 | 3:16 | 13.67 | 7.59 | 8.29 | 79.8 | 74 | 48.1
2006-09-15 | 3:31 | 13.67 | 7.58 | 8.29 | 79.8 | 73.9 | 48.059
2006-09-15 | 3:46 | 13.66 | 7.58 | 8.29 | 79.8 | 73.9 | 48.034
2006-09-15 | 4:01 | 13.66 | 7.58 | 8.27 | 79.6 | 74 | 48.089
2006-09-15 | 4:16 | 13.66 | 7.57 | 8.26 | 79.6 | 73.9 | 48.033
2006-09-15 | 4:31 | 13.67 | 7.57 | 8.27 | 79.6 | 73.8 | 47.991
2006-09-15 | 4:46 | 13.67 | 7.57 | 8.26 | 79.6 | 73.8 | 47.977
2006-09-15 | 5:01 | 13.68 | 7.56 | 8.26 | 79.6 | 73.7 | 47.9
2006-09-15 | 5:16 | 13.68 | 7.56 | 8.25 | 79.5 | 73.7 | 47.895
2006-09-15 | 5:31 | 13.67 | 7.56 | 8.26 | 79.6 | 73.6 | 47.868
2006-09-15 | 5:46 | 13.67 | 7.56 | 8.25 | 79.5 | 73.9 | 48.008
2006-09-15 | 6:01 | 13.67 | 7.56 | 8.23 | 79.3 | 74 | 48.11
2006-09-15 | 6:16 | 13.68 | 7.55 | 8.23 | 79.3 | 74 | 48.07
2006-09-15 | 6:31 | 13.69 | 7.55 | 8.23 | 79.3 | 73.8 | 47.976
2006-09-15 | 6:46 | 13.67 | 7.55 | 8.23 | 79.3 | 73.8 | 47.938
2006-09-15 | 7:01 | 13.66 | 7.55 | 8.23 | 79.3 | 73.8 | 47.972
2006-09-15 | 7:16 | 13.67 | 7.55 | 8.25 | 79.5 | 73.8 | 47.975
2006-09-15 | 7:31 | 13.69 | 7.56 | 8.28 | 79.8 | 73.7 | 47.926
2006-09-15 | 7:46 | 13.71 | 7.57 | 8.35 | 80.5 | 73.6 | 47.853
2006-09-15 | 8:01 | 13.72 | 7.6 | 8.45 | 81.5 | 73.5 | 47.762
2006-09-15 | 8:16 | 13.76 | 7.63 | 8.58 | 82.8 | 73.4 | 47.686
2006-09-15 | 8:31 | 13.82 | 7.68 | 8.73 | 84.3 | 73.3 | 47.632
2006-09-15 | 8:46 | 13.84 | 7.71 | 8.81 | 85.2 | 73.3 | 47.613
2006-09-15 | 9:01 | 13.93 | 7.75 | 8.93 | 86.5 | 73.1 | 47.529
2006-09-15 | 9:16 | 13.98 | 7.81 | 9.07 | 88 | 73 | 47.448
2006-09-15 | 9:31 | 14 | 7.87 | 9.19 | 89.2 | 72.9 | 47.396
2006-09-15 | 9:46 | 14.03 | 7.92 | 9.27 | 90 | 72.9 | 47.377
2006-09-15 | 10:01 | 14.13 | 7.98 | 9.36 | 91.1 | 72.9 | 47.355
2006-09-15 | 10:16 | 14.28 | 8.05 | 9.45 | 92.3 | 72.8 | 47.298
2006-09-15 | 10:31 | 14.4 | 8.11 | 9.52 | 93.2 | 72.7 | 47.267
2006-09-15 | 10:46 | 14.42 | 8.18 | 9.58 | 93.9 | 72.7 | 47.262
2006-09-15 | 11:01 | 14.38 | 8.21 | 9.61 | 94 | 72.6 | 47.218
2006-09-15 | 11:16 | 14.39 | 8.26 | 9.64 | 94.4 | 72.6 | 47.203
2006-09-15 | 11:31 | 14.67 | 8.34 | 9.7 | 95.5 | 72.6 | 47.179
2006-09-15 | 11:46 | 14.86 | 8.41 | 9.76 | 96.5 | 72.5 | 47.136
2006-09-15 | 12:01 | 14.87 | 8.45 | 9.75 | 96.4 | 72.5 | 47.148
2006-09-15 | 12:16 | 14.98 | 8.48 | 9.74 | 96.5 | 72.5 | 47.135
2006-09-15 | 12:31 | 15.15 | 8.55 | 9.77 | 97.2 | 71.5 | 46.495