## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | LC 001.25
Program Name | Water Quality Monitoring Program |  | Description |  | Lake Creek, at Old Pond
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 6920 sonde #06A1700AB
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-07 |  | 2006-08-07 |  | A
Time | Computer set to atomic time |  | 14:20 |  | 14:20 |  | A
Temp | VWR NIST traceable surface water |  | 17.84 |  | 17.86 |  | A
DO | Winkler titration of surface water |  | 10.12 |  | 10.32 |  | A
SpC | 148 µS/mL |  | 59.1 |  | 60.6 |  | A
Ph | pH 7 standard solution |  | 8.5 |  | 8.65 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2006-08-10 11:01 | 16.96 | 7.41 | 7.38 | 76.3 | 53.1 | 34.503
Sonde |  | 16.7 | 7.4 | 8 |  | 55
Grade |  | A | A | B |  | A
Audit 2 | 2006-08-12 11:47 | 15.94 | 7.51 | 6.68 | 67.6 | 52.6 | 34.215
Sonde |  | 15.9 | 7.2 | 6.9 |  | 55
Grade |  | A | B | A |  | A
Audit 3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Sonde
Grade |  | #REF! | #REF! | #REF! |  | #REF!
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
no post retrieval audit info. Probe needed maintenance due to algal fouling of d.o. membrane. passed retrieval audit so data is good, but sonde needed maintenance post deployment |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | LC 001.25
Program Name | Water Quality Monitoring Program |  | Description |  | Lake Creek, at Old Pond
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS 05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-07 |  | 2006-08-07 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | VWR NIST traceable surface water |  | 21.6 |  | 21.1 |  | A
DO | Winkler titration of surface water |  | 9 |  | 8.2 |  | B
SpC | 148 µS/mL |  | 152 |  | 143 |  | A
Ph | pH 7 standard solution |  | 7.1 |  | 7.4 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-08-10 11:01 | 16.96 | 7.41 | 7.38 | 76.3 | 53.1 | 34.503
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-10 |  | 2006-08-10 |  | A
Time | Computer set to atomic time |  | 13:00 |  | 13:00 |  | A
Temp | VWR NIST traceable surface water |  | 19.4 |  | 19 |  | A
DO | Winkler titration of surface water |  | 9.3 |  | 8.4 |  | B
SpC | 148 µS/mL |  | 145 |  | 140 |  | A
Ph | pH 7 standard solution |  | 7.1 |  | 7.5 |  | B
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | LC 001.25
Program Name | Water Quality Monitoring Program |  | Description |  | Lake Creek, at Old Pond
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS 05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-10 |  | 2006-08-10 |  | A
Time | Computer set to atomic time |  | 13:00 |  | 13:00 |  | A
Temp | VWR NIST traceable surface water |  | 19.4 |  | 19 |  | A
DO | Winkler titration of surface water |  | 9.3 |  | 8.4 |  | B
SpC | 148 µS/mL |  | 145 |  | 140 |  | A
Ph | pH 7 standard solution |  | 7.1 |  | 7.5 |  | B
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 51 |  | B
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 478 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-08-12 11:47 | 15.94 | 7.51 | 6.68 | 67.6 | 52.6 | 34.215
1
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-08-15 |  | 2006-08-15 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 22.5 |  | 22.2 |  | A
DO | Winkler titration of surface water |  | 9.5 |  | 9.3 |  | A
SpC | 148 µS/mL |  | 148 |  | 141 |  | A
Ph | pH 7 standard solution |  | 7.2 |  | 7.5 |  | A
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 51 |  | B
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 477 |  | B
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | DO Conc | DO% | SpCond | TDS
|  | °C |  | mg/L | % | uS/cm | mg/L
2006-08-10 | 11:01 | 16.72 | 7.38 | 7.96 | 81.9 | 54.7 | 35.558
2006-08-10 | 11:16 | 16.77 | 7.4 | 8.05 | 82.9 | 54.6 | 35.519
2006-08-10 | 11:31 | 16.77 | 7.42 | 8.06 | 83 | 54.6 | 35.516
2006-08-10 | 11:46 | 16.78 | 7.46 | 8.14 | 83.9 | 54.6 | 35.481
2006-08-10 | 12:01 | 16.84 | 7.55 | 8.62 | 88.9 | 54.3 | 35.283
2006-08-10 | 12:16 | 16.86 | 7.65 | 9.14 | 94.3 | 54.1 | 35.184
2006-08-10 | 12:31 | 16.88 | 7.67 | 9.06 | 93.6 | 54.1 | 35.166
2006-08-10 | 12:46 | 16.86 | 7.68 | 9.29 | 95.9 | 54.1 | 35.163
2006-08-10 | 13:00 | 16.87 | 7.69 | 9.35 | 96.5 | 54.1 | 35.167
2006-08-10 | 13:15 | 16.89 | 7.7 | 9.39 | 96.9 | 54.1 | 35.186
2006-08-10 | 13:30 | 16.9 | 7.72 | 9.47 | 97.8 | 54.2 | 35.253
2006-08-10 | 13:45 | 16.92 | 7.72 | 9.52 | 98.3 | 54.3 | 35.276
2006-08-10 | 14:00 | 16.93 | 7.74 | 9.55 | 98.7 | 54.3 | 35.294
2006-08-10 | 14:15 | 16.94 | 7.74 | 9.5 | 98.2 | 54.3 | 35.275
2006-08-10 | 14:30 | 16.94 | 7.74 | 9.47 | 97.9 | 54.3 | 35.267
2006-08-10 | 14:45 | 16.95 | 7.74 | 9.44 | 97.6 | 54.3 | 35.325
2006-08-10 | 15:00 | 16.96 | 7.63 | 9.47 | 97.9 | 54.4 | 35.337
2006-08-10 | 15:15 | 16.97 | 7.48 | 9.83 | 101.7 | 54.4 | 35.344
2006-08-10 | 15:30 | 16.99 | 7.32 | 9.41 | 97.4 | 54.4 | 35.338
2006-08-10 | 15:45 | 16.98 | 7.32 | 8.83 | 91.3 | 54.3 | 35.27
2006-08-10 | 16:00 | 16.97 | 7.23 | 7.88 | 81.5 | 54.4 | 35.35
2006-08-10 | 16:15 | 17.07 | 7.23 | 7.95 | 82.4 | 54.4 | 35.391
2006-08-10 | 16:30 | 17.03 | 7.2 | 7.83 | 81 | 54.9 | 35.66
2006-08-10 | 16:45 | 17.08 | 7.24 | 7.81 | 81 | 54.8 | 35.636
2006-08-10 | 17:00 | 17.1 | 7.19 | 7.56 | 78.4 | 55.1 | 35.811
2006-08-10 | 17:15 | 17.08 | 7.29 | 8.27 | 85.7 | 55 | 35.781
2006-08-10 | 17:30 | 17.11 | 7.54 | 8.97 | 93 | 55.1 | 35.81
2006-08-10 | 17:45 | 17.09 | 7.34 | 8.39 | 87 | 54.7 | 35.586
2006-08-10 | 18:00 | 17.07 | 7.34 | 8.42 | 87.3 | 54.4 | 35.374
2006-08-10 | 18:15 | 17.11 | 7.31 | 8.17 | 84.7 | 54.6 | 35.497
2006-08-10 | 18:30 | 17.08 | 7.16 | 8.24 | 85.4 | 54.5 | 35.436
2006-08-10 | 18:45 | 17.09 | 7.21 | 8.3 | 86 | 54.7 | 35.537
2006-08-10 | 19:00 | 17.12 | 7.18 | 8.12 | 84.2 | 54.7 | 35.539
2006-08-10 | 19:15 | 17.11 | 7.2 | 8.03 | 83.3 | 54.8 | 35.595
2006-08-10 | 19:30 | 17.1 | 7.25 | 7.68 | 79.7 | 54.8 | 35.641
2006-08-10 | 19:45 | 17.12 | 7.29 | 8.44 | 87.6 | 54.8 | 35.634
2006-08-10 | 20:00 | 17.1 | 7.32 | 8.19 | 84.9 | 54.8 | 35.632
2006-08-10 | 20:15 | 17.11 | 7.27 | 8.66 | 89.8 | 54.8 | 35.632
2006-08-10 | 20:30 | 17.12 | 7.2 | 8.05 | 83.5 | 54.9 | 35.69
2006-08-10 | 20:45 | 17.15 | 7.17 | 7.76 | 80.6 | 55 | 35.758
2006-08-10 | 21:00 | 17.15 | 7.17 | 7.6 | 78.9 | 55 | 35.729
2006-08-10 | 21:15 | 17.14 | 7.16 | 7.59 | 78.7 | 55.1 | 35.822
2006-08-10 | 21:30 | 17.17 | 7.16 | 7.55 | 78.4 | 55.1 | 35.814
2006-08-10 | 21:45 | 17.2 | 7.16 | 7.4 | 76.9 | 55.3 | 35.919
2006-08-10 | 22:00 | 17.2 | 7.17 | 7.34 | 76.3 | 55.6 | 36.162
2006-08-10 | 22:15 | 17.24 | 7.15 | 7.31 | 76 | 55.2 | 35.912
2006-08-10 | 22:30 | 17.38 | 7.23 | 7.33 | 76.4 | 54.9 | 35.69
2006-08-10 | 22:45 | 17.48 | 7.47 | 7.99 | 83.5 | 54.3 | 35.293
2006-08-10 | 23:00 | 17.33 | 7.24 | 7.78 | 81 | 54.4 | 35.35
2006-08-10 | 23:15 | 17.33 | 7.16 | 7.18 | 74.8 | 54.6 | 35.465
2006-08-10 | 23:30 | 17.4 | 7.34 | 7.49 | 78.2 | 54.5 | 35.449
2006-08-10 | 23:45 | 17.42 | 7.29 | 7.5 | 78.3 | 54.5 | 35.438
2006-08-11 | 0:00 | 17.65 | 7.35 | 8.16 | 85.6 | 54.6 | 35.49
2006-08-11 | 0:15 | 17.59 | 7.27 | 8.1 | 84.8 | 54.8 | 35.614
2006-08-11 | 0:30 | 17.67 | 7.29 | 7.77 | 81.5 | 54.7 | 35.538
2006-08-11 | 0:45 | 17.65 | 7.29 | 7.53 | 78.9 | 54.7 | 35.557
2006-08-11 | 1:01 | 17.72 | 7.37 | 7.39 | 77.6 | 54.6 | 35.512
2006-08-11 | 1:16 | 17.71 | 7.44 | 7.85 | 82.4 | 54.6 | 35.519
2006-08-11 | 1:31 | 17.7 | 7.44 | 8.08 | 84.8 | 54.7 | 35.556
2006-08-11 | 1:46 | 17.68 | 7.45 | 8.16 | 85.6 | 54.6 | 35.502
2006-08-11 | 2:01 | 17.6 | 7.44 | 8.12 | 85 | 54.6 | 35.519
2006-08-11 | 2:16 | 17.54 | 7.43 | 8.07 | 84.4 | 54.7 | 35.529
2006-08-11 | 2:31 | 17.49 | 7.42 | 8.02 | 83.8 | 54.8 | 35.592
2006-08-11 | 2:46 | 17.42 | 7.42 | 8.04 | 83.9 | 54.7 | 35.543
2006-08-11 | 3:01 | 17.35 | 7.42 | 8.08 | 84.2 | 54.8 | 35.618
2006-08-11 | 3:16 | 17.3 | 7.41 | 8.08 | 84.1 | 54.8 | 35.611
2006-08-11 | 3:31 | 17.24 | 7.41 | 8.13 | 84.5 | 54.7 | 35.579
2006-08-11 | 3:46 | 17.2 | 7.39 | 8 | 83.1 | 54.8 | 35.587
2006-08-11 | 4:01 | 17.15 | 7.41 | 8.11 | 84.2 | 54.8 | 35.588
2006-08-11 | 4:16 | 17.1 | 7.41 | 8.18 | 84.8 | 54.7 | 35.587
2006-08-11 | 4:31 | 17.07 | 7.41 | 8.11 | 84 | 54.8 | 35.604
2006-08-11 | 4:46 | 17.03 | 7.4 | 8.12 | 84 | 54.8 | 35.595
2006-08-11 | 5:01 | 17 | 7.42 | 8.09 | 83.7 | 54.7 | 35.583
2006-08-11 | 5:16 | 16.94 | 7.4 | 8.07 | 83.4 | 54.8 | 35.592
2006-08-11 | 5:31 | 16.89 | 7.4 | 8.08 | 83.5 | 54.8 | 35.606
2006-08-11 | 5:46 | 16.85 | 7.41 | 8.04 | 82.9 | 54.8 | 35.619
2006-08-11 | 6:01 | 16.81 | 7.4 | 8.07 | 83.2 | 54.8 | 35.633
2006-08-11 | 6:16 | 16.77 | 7.38 | 8.05 | 82.9 | 54.8 | 35.609
2006-08-11 | 6:31 | 16.72 | 7.42 | 8.11 | 83.4 | 54.6 | 35.463
2006-08-11 | 6:46 | 16.68 | 7.41 | 8.14 | 83.6 | 54.5 | 35.426
2006-08-11 | 7:01 | 16.64 | 7.41 | 8.13 | 83.5 | 54.7 | 35.576
2006-08-11 | 7:16 | 16.6 | 7.42 | 8.17 | 83.8 | 54.5 | 35.457
2006-08-11 | 7:31 | 16.55 | 7.4 | 8.13 | 83.4 | 54.7 | 35.536
2006-08-11 | 7:46 | 16.54 | 7.42 | 8.21 | 84.2 | 54.6 | 35.513
2006-08-11 | 8:01 | 16.49 | 7.42 | 8.22 | 84.2 | 54.5 | 35.414
2006-08-11 | 8:16 | 16.46 | 7.39 | 8.14 | 83.3 | 54.5 | 35.442
2006-08-11 | 8:31 | 16.44 | 7.38 | 7.95 | 81.3 | 54.5 | 35.425
2006-08-11 | 8:46 | 16.43 | 7.35 | 7.87 | 80.5 | 54.6 | 35.52
2006-08-11 | 9:01 | 16.41 | 7.3 | 7.59 | 77.6 | 54.7 | 35.534
2006-08-11 | 9:16 | 16.42 | 7.29 | 7.57 | 77.4 | 54.8 | 35.609
2006-08-11 | 9:31 | 16.42 | 7.26 | 7.53 | 77 | 54.8 | 35.611
2006-08-11 | 9:46 | 16.43 | 7.25 | 7.09 | 72.5 | 54.9 | 35.657
2006-08-11 | 10:01 | 16.43 | 7.2 | 6.86 | 70.1 | 54.9 | 35.666
2006-08-11 | 10:16 | 16.43 | 7.21 | 7.06 | 72.2 | 54.9 | 35.694
2006-08-11 | 10:31 | 16.44 | 7.2 | 7 | 71.6 | 54.9 | 35.716
2006-08-11 | 10:46 | 16.43 | 7.2 | 7.02 | 71.7 | 55.1 | 35.798
2006-08-11 | 11:01 | 16.42 | 7.23 | 6.99 | 71.5 | 55.3 | 35.976
2006-08-11 | 11:16 | 16.39 | 7.2 | 6.97 | 71.2 | 55.4 | 36.001
2006-08-11 | 11:31 | 16.39 | 7.2 | 6.94 | 70.9 | 55.4 | 36.014
2006-08-11 | 11:46 | 16.37 | 7.26 | 7.35 | 75.1 | 55.2 | 35.893
2006-08-11 | 12:01 | 16.41 | 7.36 | 7.29 | 74.5 | 55 | 35.743
2006-08-11 | 12:16 | 16.42 | 7.33 | 7.51 | 76.7 | 54.8 | 35.641
2006-08-11 | 12:31 | 16.45 | 7.32 | 7.72 | 78.9 | 54.8 | 35.626
2006-08-11 | 12:46 | 16.48 | 7.47 | 8.14 | 83.3 | 55 | 35.767
2006-08-11 | 13:00 | 16.48 | 7.63 | 8.99 | 92 | 54.6 | 35.467
2006-08-11 | 13:15 | 16.51 | 7.57 | 8.96 | 91.7 | 54.4 | 35.375
2006-08-11 | 13:30 | 16.52 | 7.54 | 8.87 | 90.9 | 54.3 | 35.296
2006-08-11 | 13:45 | 16.5 | 7.55 | 8.99 | 92.1 | 54.3 | 35.288
2006-08-11 | 14:00 | 16.5 | 7.48 | 9.06 | 92.8 | 54.4 | 35.376
2006-08-11 | 14:15 | 16.52 | 7.49 | 8.94 | 91.6 | 54.4 | 35.346
2006-08-11 | 14:30 | 16.53 | 7.53 | 9.25 | 94.8 | 54.4 | 35.375
2006-08-11 | 14:45 | 16.54 | 7.52 | 9.17 | 94 | 54.5 | 35.438
2006-08-11 | 15:00 | 16.55 | 7.57 | 9.1 | 93.3 | 54.3 | 35.274
2006-08-11 | 15:15 | 16.55 | 7.51 | 9.04 | 92.7 | 54.3 | 35.285
2006-08-11 | 15:30 | 16.55 | 7.47 | 9.2 | 94.3 | 54 | 35.132
2006-08-11 | 15:45 | 16.58 | 7.45 | 9.09 | 93.2 | 54 | 35.102
2006-08-11 | 16:00 | 16.58 | 7.45 | 9.06 | 93 | 53.9 | 35.03
2006-08-11 | 16:15 | 16.58 | 7.49 | 9.23 | 94.7 | 53.9 | 35.059
2006-08-11 | 16:30 | 16.58 | 7.46 | 9.15 | 93.9 | 53.8 | 34.996
2006-08-11 | 16:45 | 16.6 | 7.43 | 8.98 | 92.2 | 53.8 | 35.001
2006-08-11 | 17:00 | 16.59 | 7.5 | 9.49 | 97.3 | 53.8 | 34.975
2006-08-11 | 17:15 | 16.59 | 7.57 | 9.64 | 98.9 | 53.8 | 34.959
2006-08-11 | 17:30 | 16.6 | 7.69 | 9.23 | 94.7 | 53.9 | 35.013
2006-08-11 | 17:45 | 16.62 | 7.83 | 9.38 | 96.2 | 53.8 | 34.979
2006-08-11 | 18:00 | 16.62 | 7.64 | 9.38 | 96.3 | 53.8 | 34.971
2006-08-11 | 18:15 | 16.66 | 7.42 | 9.03 | 92.8 | 53.7 | 34.889
2006-08-11 | 18:30 | 16.72 | 7.64 | 9.82 | 101 | 53.7 | 34.887
2006-08-11 | 18:45 | 16.69 | 7.49 | 9.3 | 95.6 | 53.6 | 34.846
2006-08-11 | 19:00 | 16.68 | 7.53 | 9.25 | 95.1 | 53.6 | 34.861
2006-08-11 | 19:15 | 16.64 | 7.85 | 9.67 | 99.3 | 53.1 | 34.497
2006-08-11 | 19:30 | 16.64 | 7.72 | 9.58 | 98.4 | 53.1 | 34.503
2006-08-11 | 19:45 | 16.64 | 7.63 | 9.6 | 98.5 | 53.2 | 34.555
2006-08-11 | 20:00 | 16.66 | 7.62 | 9.51 | 97.7 | 53.3 | 34.642
2006-08-11 | 20:15 | 16.68 | 7.62 | 9.5 | 97.6 | 53.2 | 34.581
2006-08-11 | 20:30 | 16.67 | 7.59 | 9.57 | 98.3 | 53.3 | 34.651
2006-08-11 | 20:45 | 16.68 | 7.57 | 9.43 | 96.9 | 53.4 | 34.714
2006-08-11 | 21:00 | 16.68 | 7.51 | 9.41 | 96.7 | 53.4 | 34.739
2006-08-11 | 21:15 | 16.69 | 7.5 | 9.36 | 96.2 | 53.4 | 34.715
2006-08-11 | 21:30 | 16.7 | 7.48 | 9.27 | 95.4 | 53.4 | 34.732
2006-08-11 | 21:45 | 16.69 | 7.46 | 9.3 | 95.6 | 53.5 | 34.768
2006-08-11 | 22:00 | 16.72 | 7.46 | 9.24 | 95.1 | 53.6 | 34.816
2006-08-11 | 22:15 | 16.73 | 7.45 | 9.09 | 93.5 | 53.5 | 34.805
2006-08-11 | 22:30 | 16.74 | 7.41 | 8.81 | 90.6 | 53.5 | 34.765
2006-08-11 | 22:45 | 16.76 | 7.41 | 8.64 | 89 | 53.5 | 34.79
2006-08-11 | 23:00 | 16.74 | 7.36 | 8.55 | 88 | 53.6 | 34.811
2006-08-11 | 23:15 | 16.78 | 7.37 | 8.51 | 87.6 | 53.6 | 34.856
2006-08-11 | 23:30 | 16.78 | 7.39 | 8.52 | 87.7 | 53.7 | 34.903
2006-08-11 | 23:45 | 16.76 | 7.38 | 8.44 | 86.9 | 53.8 | 34.972
2006-08-12 | 0:00 | 16.77 | 7.38 | 8.34 | 85.9 | 53.9 | 35.003
2006-08-12 | 0:15 | 16.78 | 7.35 | 8.18 | 84.3 | 53.8 | 34.992
2006-08-12 | 0:30 | 16.8 | 7.33 | 8.12 | 83.7 | 53.9 | 35.034
2006-08-12 | 0:45 | 16.8 | 7.32 | 7.99 | 82.3 | 53.9 | 35.038
2006-08-12 | 1:01 | 16.8 | 7.32 | 8.19 | 84.4 | 53.9 | 35.054
2006-08-12 | 1:16 | 16.82 | 7.36 | 8.73 | 90 | 54 | 35.083
2006-08-12 | 1:31 | 16.83 | 7.39 | 8.77 | 90.4 | 53.9 | 35.017
2006-08-12 | 1:46 | 16.84 | 7.43 | 8.78 | 90.6 | 53.9 | 35.027
2006-08-12 | 2:01 | 16.83 | 7.54 | 9.22 | 95.1 | 54 | 35.075
2006-08-12 | 2:16 | 16.82 | 7.54 | 9 | 92.8 | 54 | 35.088
2006-08-12 | 2:31 | 16.82 | 7.5 | 8.77 | 90.4 | 54 | 35.105
2006-08-12 | 2:46 | 16.83 | 7.46 | 8.59 | 88.6 | 54.1 | 35.162
2006-08-12 | 3:01 | 16.82 | 7.41 | 8.34 | 86 | 54.3 | 35.274
2006-08-12 | 3:16 | 16.79 | 7.28 | 8.01 | 82.5 | 54.4 | 35.334
2006-08-12 | 3:31 | 16.76 | 7.27 | 7.59 | 78.1 | 54.7 | 35.554
2006-08-12 | 3:46 | 16.71 | 7.31 | 7.63 | 78.5 | 54.8 | 35.627
2006-08-12 | 4:01 | 16.64 | 7.34 | 7.65 | 78.5 | 54.9 | 35.656
2006-08-12 | 4:16 | 16.58 | 7.22 | 7.4 | 75.9 | 54.9 | 35.672
2006-08-12 | 4:31 | 16.55 | 7.27 | 7.64 | 78.3 | 54.9 | 35.706
2006-08-12 | 4:46 | 16.51 | 7.3 | 7.87 | 80.6 | 55 | 35.756
2006-08-12 | 5:01 | 16.46 | 7.34 | 8.17 | 83.6 | 55.3 | 35.915
2006-08-12 | 5:16 | 16.42 | 7.34 | 8.14 | 83.2 | 55.1 | 35.823
2006-08-12 | 5:31 | 16.38 | 7.3 | 7.55 | 77.1 | 55 | 35.753
2006-08-12 | 5:46 | 16.33 | 7.3 | 8.01 | 81.7 | 55 | 35.769
2006-08-12 | 6:01 | 16.29 | 7.35 | 8.16 | 83.2 | 55.1 | 35.824
2006-08-12 | 6:16 | 16.24 | 7.37 | 8.28 | 84.3 | 54.9 | 35.68
2006-08-12 | 6:31 | 16.19 | 7.36 | 8.27 | 84.1 | 55 | 35.731
2006-08-12 | 6:46 | 16.14 | 7.35 | 8.32 | 84.5 | 54.8 | 35.604
2006-08-12 | 7:01 | 16.09 | 7.39 | 8.27 | 83.9 | 54.4 | 35.346
2006-08-12 | 7:16 | 16.04 | 7.33 | 8.06 | 81.7 | 54.2 | 35.249
2006-08-12 | 7:31 | 16 | 7.42 | 8.2 | 83.1 | 54.1 | 35.19
2006-08-12 | 7:46 | 15.96 | 7.4 | 8.09 | 81.9 | 54.3 | 35.29
2006-08-12 | 8:01 | 15.93 | 7.4 | 8.15 | 82.4 | 54.2 | 35.22
2006-08-12 | 8:16 | 15.9 | 7.42 | 8.16 | 82.5 | 54.2 | 35.238
2006-08-12 | 8:31 | 15.87 | 7.43 | 8.28 | 83.7 | 54.1 | 35.151
2006-08-12 | 8:46 | 15.84 | 7.32 | 7.89 | 79.6 | 54.2 | 35.248
2006-08-12 | 9:01 | 15.82 | 7.29 | 7.69 | 77.7 | 54.3 | 35.305
2006-08-12 | 9:16 | 15.83 | 7.29 | 7.73 | 78.1 | 54.4 | 35.329
2006-08-12 | 9:31 | 15.83 | 7.31 | 7.58 | 76.6 | 54.3 | 35.265
2006-08-12 | 9:46 | 15.83 | 7.27 | 7.28 | 73.5 | 54.4 | 35.359
2006-08-12 | 10:01 | 15.83 | 7.21 | 7.18 | 72.5 | 54.6 | 35.466
2006-08-12 | 10:16 | 15.84 | 7.15 | 6.73 | 67.9 | 54.7 | 35.537
2006-08-12 | 10:31 | 15.84 | 7.12 | 6.59 | 66.6 | 54.8 | 35.613
2006-08-12 | 10:46 | 15.85 | 7.13 | 6.56 | 66.3 | 54.9 | 35.659
2006-08-12 | 11:01 | 15.85 | 7.12 | 6.54 | 66.1 | 54.8 | 35.628
2006-08-12 | 11:16 | 15.87 | 7.15 | 6.98 | 70.5 | 54.8 | 35.632
2006-08-12 | 11:31 | 15.85 | 7.15 | 6.87 | 69.4 | 54.9 | 35.69
2006-08-12 | 11:46 | 15.87 | 7.16 | 6.94 | 70.1 | 54.9 | 35.677