## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek City Park gauge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 6920 # 06A2389AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-18 |  | 2006-09-18 |  | A
Time | Computer set to atomic time |  | 14:32 |  | 14:32 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 14.5 |  | 14.7 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7 |  | A
DO | Winkler titration of surface water |  | 10.3 |  | 10.1 |  | A
SpC | 148.0 µS/mL |  | 148 |  | 141 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2006-09-19 15:00 | 11.11 | 7.64 | 10.68 | 97.1 | 32.3 | 20.98
Field Equipment |  | 11.09 | 7.66 | 10.35 | 94.1 | 33.8 | 21.961
Grade |  | A | A | B |  | A
Audit 2 | 2006-09-21 12:03 | 8.39 |  | 11.52 | 98.2 | 29.8 | 19.364
Field Equipment |  | 8.29 | 7.68 | 3.79 | 32.2 | 30.8 | 20.025
Grade |  | A |  | 0 |  | A
Audit 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Field Equipment
Grade |  | A | A | A |  | #DIV/0!
Post-sampling Field Equipment Check
Field equipment ID | YSI 6920 # 06A2389AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-10-05 |  | 2006-10-05 |  | A
Time | Computer set to atomic time |  | 17:19 |  | 17:17 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 15.6 |  | 16.7 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7 |  | A
DO | Winkler titration of surface water |  | 10.5 |  |  |  | 0
SpC | 148.0 µS/mL |  | 148 |  | 141 |  | A
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
|  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek City Park gauge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-16 |  | 2006-09-16 |  | A
Time | Computer set to atomic time |  | 13:55 |  | 13:56 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 15.5 |  | 14.8 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.1 |  | A
DO | Winkler titration of surface water |  | 10.5 |  | 10.7 |  | A
SpC | 148.0 µS/mL |  | 148 |  | 145 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
Field equipment ID | Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
YSI 556 MPS #05G1083AM | 2006-09-19 15:00 | 11.11 | 7.64 | 10.68 | 97.1 | 32.3 | 20.98
Hach 2100P #50828
Post-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-19 |  | 2006-09-19 |  | A
Time | Computer set to atomic time |  | 16:49 |  | 16:48 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 14.5 |  | 14 |  | A
Ph | pH 7 standard solution |  | 7 |  | 8 |  | override
DO | Winkler titration of surface water |  | 10.5 |  | 10.9 |  | B
SpC | 148.0 µS/mL |  | 148 |  | 138 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
calibrated pH in field because both sondes read a higer pH than the YSI 556. Is is perhaps the ionic strength of the river water? Temperature? Also, I calibrated the DO in the field to mg/L since the other sonde and the YSI were stable and precise. |  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek City Park gauge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-20 |  | 2006-09-20 |  | A
Time | Computer set to atomic time |  | 8:53 |  | 8:52 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 14.5 |  | 14 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7 |  | A
DO | Winkler titration of surface water |  | 10.5 |  | 10.9 |  | B
SpC | 148.0 µS/mL |  | 148 |  | 138 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
Field equipment ID | Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
YSI 556 MPS #05G1083AM | 2006-09-21 12:03 | 8.39 | 6.7 | 11.52 | 98.2 | 29.8 | 19.364
Hach 2100P #50828
Post-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-25 |  | 2006-09-25 |  | A
Time | Computer set to atomic time |  | 12:03 |  | 12:01 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 13.1 |  | 12.8 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.2 |  | A
DO | Winkler titration of surface water |  | 10.6 |  | 11.5 |  | B
SpC | 148.0 µS/mL |  | 148 |  | 157 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
|  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1
Formulas:
Upper Deschutes Watershed Council = '[2]month year'!D9

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek City Park gauge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-30 |  | 2006-09-30 |  | A
Time | Computer set to atomic time |  | 14:16 |  | 14:14 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 15.5 |  | 15.4 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.2 |  | A
DO | Winkler titration of surface water |  | 10.5 |  | 10.8 |  | A
SpC | 148.0 µS/mL |  | 148 |  | 139 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
Field equipment ID | Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
YSI 556 MPS #05G1083AM
Hach 2100P #50828
Post-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-10-05 |  | 2006-10-05 |  | A
Time | Computer set to atomic time |  | 17:19 |  | 17:17 |  | A
Temp | VWR NIST traceable #51063402 surface water |  | 15.6 |  | 15 |  | A
Ph | pH 7 standard solution |  | 7 |  | 7.1 |  | A
DO | Winkler titration of surface water |  | 10.5 |  | 10.5 |  | A
SpC | 148.0 µS/mL |  | 148 |  | 139 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
|  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1

## data
Date | Time | Temp | pH | DO Conc | DO% | SpCond | TDS
|  | °C |  | mg/L | % | uS/cm | mg/L
2006-09-19 | 15:00 | 11.09 | 7.66 |  |  | 33.8 | 21.961
2006-09-19 | 15:15 | 11.21 | 7.67 |  |  | 33.8 | 21.968
2006-09-19 | 15:30 | 11.32 | 7.69 |  |  | 33.8 | 21.984
2006-09-19 | 15:45 | 11.39 | 7.71 |  |  | 33.8 | 21.968
2006-09-19 | 16:00 | 11.44 | 7.71 |  |  | 33.8 | 21.972
2006-09-19 | 16:15 | 11.32 | 7.71 |  |  | 33.9 | 22.006
2006-09-19 | 16:30 | 11.06 | 7.72 |  |  | 33.8 | 21.997
2006-09-19 | 16:45 | 10.88 | 7.72 |  |  | 33.8 | 21.997
2006-09-19 | 17:00 | 10.65 | 7.71 |  |  | 33.9 | 22.012
2006-09-19 | 17:15 | 10.49 | 7.71 |  |  | 33.8 | 21.989
2006-09-19 | 17:30 | 10.24 | 7.71 |  |  | 33.8 | 21.997
2006-09-19 | 17:45 | 10.01 | 7.7 |  |  | 33.8 | 21.945
2006-09-19 | 18:00 | 9.79 | 7.69 |  |  | 33.9 | 22.003
2006-09-19 | 18:15 | 9.6 | 7.69 |  |  | 33.8 | 21.996
2006-09-19 | 18:30 | 9.4 | 7.69 |  |  | 33.8 | 21.995
2006-09-19 | 18:45 | 9.2 | 7.68 |  |  | 33.9 | 22.004
2006-09-19 | 19:00 | 9.02 | 7.68 |  |  | 33.8 | 21.981
2006-09-19 | 19:15 | 8.84 | 7.67 |  |  | 33.8 | 21.979
2006-09-19 | 19:30 | 8.65 | 7.67 |  |  | 33.8 | 21.97
2006-09-19 | 19:45 | 8.44 | 7.66 |  |  | 33.8 | 21.951
2006-09-19 | 20:00 | 8.26 | 7.65 |  |  | 33.7 | 21.932
2006-09-19 | 20:15 | 8.09 | 7.65 |  |  | 33.6 | 21.857
2006-09-19 | 20:30 | 7.95 | 7.65 |  |  | 33.7 | 21.882
2006-09-19 | 20:45 | 7.83 | 7.65 |  |  | 33.6 | 21.834
2006-09-19 | 21:00 | 7.71 | 7.65 |  |  | 33.7 | 21.884
2006-09-19 | 21:15 | 7.61 | 7.64 |  |  | 33.6 | 21.857
2006-09-19 | 21:30 | 7.51 | 7.64 |  |  | 33.6 | 21.849
2006-09-19 | 21:45 | 7.43 | 7.63 |  |  | 33.6 | 21.838
2006-09-19 | 22:00 | 7.35 | 7.63 |  |  | 33.6 | 21.859
2006-09-19 | 22:15 | 7.27 | 7.63 |  |  | 33.5 | 21.801
2006-09-19 | 22:30 | 7.19 | 7.63 |  |  | 33.5 | 21.781
2006-09-19 | 22:45 | 7.1 | 7.62 |  |  | 33.5 | 21.774
2006-09-19 | 23:00 | 6.99 | 7.62 |  |  | 33.4 | 21.708
2006-09-19 | 23:15 | 6.88 | 7.62 |  |  | 33.4 | 21.698
2006-09-19 | 23:30 | 6.8 | 7.62 |  |  | 33.3 | 21.677
2006-09-19 | 23:45 | 6.74 | 7.62 |  |  | 33.4 | 21.739
2006-09-20 | 0:00 | 6.68 | 7.61 |  |  | 33.4 | 21.72
2006-09-20 | 0:15 | 6.62 | 7.61 |  |  | 33.3 | 21.653
2006-09-20 | 0:30 | 6.55 | 7.61 |  |  | 33.3 | 21.638
2006-09-20 | 0:45 | 6.48 | 7.61 |  |  | 33.4 | 21.701
2006-09-20 | 1:01 | 6.4 | 7.61 |  |  | 33.3 | 21.623
2006-09-20 | 1:16 | 6.3 | 7.61 |  |  | 33.4 | 21.701
2006-09-20 | 1:31 | 6.2 | 7.61 |  |  | 33.3 | 21.676
2006-09-20 | 1:46 | 6.1 | 7.61 |  |  | 33.4 | 21.7
2006-09-20 | 2:01 | 5.99 | 7.61 |  |  | 33.4 | 21.713
2006-09-20 | 2:16 | 5.89 | 7.61 |  |  | 33.5 | 21.765
2006-09-20 | 2:31 | 5.78 | 7.61 |  |  | 33.4 | 21.705
2006-09-20 | 2:46 | 5.68 | 7.61 |  |  | 33.4 | 21.716
2006-09-20 | 3:01 | 5.59 | 7.6 |  |  | 33.5 | 21.785
2006-09-20 | 3:16 | 5.51 | 7.61 |  |  | 33.4 | 21.687
2006-09-20 | 3:31 | 5.43 | 7.61 |  |  | 33.6 | 21.825
2006-09-20 | 3:46 | 5.36 | 7.61 |  |  | 33.5 | 21.769
2006-09-20 | 4:01 | 5.29 | 7.61 |  |  | 33.6 | 21.823
2006-09-20 | 4:16 | 5.22 | 7.61 |  |  | 33.8 | 21.993
2006-09-20 | 4:31 | 5.16 | 7.61 |  |  | 33.9 | 22.049
2006-09-20 | 4:46 | 5.1 | 7.61 |  |  | 34 | 22.098
2006-09-20 | 5:01 | 5.04 | 7.61 |  |  | 34 | 22.1
2006-09-20 | 5:16 | 4.99 | 7.62 |  |  | 34 | 22.085
2006-09-20 | 5:31 | 4.94 | 7.62 |  |  | 34 | 22.119
2006-09-20 | 5:46 | 4.89 | 7.62 |  |  | 34.1 | 22.152
2006-09-20 | 6:01 | 4.84 | 7.62 |  |  | 34.3 | 22.267
2006-09-20 | 6:16 | 4.78 | 7.62 |  |  | 34.3 | 22.298
2006-09-20 | 6:31 | 4.73 | 7.62 |  |  | 34.4 | 22.336
2006-09-20 | 6:46 | 4.66 | 7.62 |  |  | 34.5 | 22.411
2006-09-20 | 7:01 | 4.6 | 7.61 |  |  | 34.3 | 22.313
2006-09-20 | 7:16 | 4.53 | 7.61 |  |  | 34.5 | 22.409
2006-09-20 | 7:31 | 4.48 | 7.62 |  |  | 34.4 | 22.355
2006-09-20 | 7:46 | 4.45 | 7.62 |  |  | 34.5 | 22.438
2006-09-20 | 8:01 | 4.45 | 7.63 |  |  | 34.5 | 22.439
2006-09-20 | 8:16 | 4.45 | 7.63 |  |  | 34.5 | 22.442
2006-09-20 | 8:31 | 4.44 | 7.63 |  |  | 34.4 | 22.367
2006-09-20 | 8:46 | 4.47 | 7.64 |  |  | 34.4 | 22.376
2006-09-20 | 9:01 | 4.53 | 7.64 |  |  | 34.5 | 22.446
2006-09-20 | 9:16 | 4.63 | 7.65 |  |  | 34.4 | 22.388
2006-09-20 | 9:31 | 4.74 | 7.66 |  |  | 34.4 | 22.382
2006-09-20 | 9:46 | 4.86 | 7.67 |  |  | 34.4 | 22.379
2006-09-20 | 10:01 | 4.98 | 7.67 |  |  | 34.5 | 22.431
2006-09-20 | 10:16 | 5.15 | 7.68 |  |  | 34.5 | 22.456
2006-09-20 | 10:31 | 5.34 | 7.69 |  |  | 34.4 | 22.391
2006-09-20 | 10:46 | 5.49 | 7.69 |  |  | 34.3 | 22.272
2006-09-20 | 11:01 | 5.75 | 7.7 |  |  | 34.2 | 22.212
2006-09-20 | 11:16 | 5.97 | 7.7 |  |  | 34.2 | 22.222
2006-09-20 | 11:31 | 6.19 | 7.7 |  |  | 34.1 | 22.148
2006-09-20 | 11:46 | 6.46 | 7.7 |  |  | 34 | 22.098
2006-09-20 | 12:01 | 6.66 | 7.7 |  |  | 34 | 22.068
2006-09-20 | 12:16 | 6.89 | 7.71 |  |  | 33.9 | 22.033
2006-09-20 | 12:31 | 7.09 | 7.71 |  |  | 33.7 | 21.927
2006-09-20 | 12:46 | 7.38 | 7.71 |  |  | 33.7 | 21.916
2006-09-20 | 13:00 | 7.6 | 7.72 |  |  | 33.7 | 21.896
2006-09-20 | 13:15 | 7.71 | 7.72 |  |  | 33.7 | 21.923
2006-09-20 | 13:30 | 7.82 | 7.71 |  |  | 33.7 | 21.925
2006-09-20 | 13:45 | 7.86 | 7.72 |  |  | 33.7 | 21.925
2006-09-20 | 14:00 | 7.95 | 7.72 |  |  | 33.7 | 21.931
2006-09-20 | 14:15 | 8.1 | 7.72 |  |  | 33.7 | 21.925
2006-09-20 | 14:30 | 8.15 | 7.72 |  |  | 33.7 | 21.905
2006-09-20 | 14:45 | 8.13 | 7.72 |  |  | 33.6 | 21.832
2006-09-20 | 15:00 | 8.11 | 7.71 |  |  | 33.6 | 21.832
2006-09-20 | 15:15 | 8.06 | 7.71 |  |  | 33.6 | 21.834
2006-09-20 | 15:30 | 8 | 7.7 |  |  | 33.6 | 21.839
2006-09-20 | 15:45 | 7.94 | 7.69 |  |  | 33.6 | 21.833
2006-09-20 | 16:00 | 7.89 | 7.69 |  |  | 33.7 | 21.879
2006-09-20 | 16:15 | 7.84 | 7.69 |  |  | 33.5 | 21.774
2006-09-20 | 16:30 | 7.81 | 7.69 |  |  | 33.5 | 21.752
2006-09-20 | 16:45 | 7.75 | 7.69 |  |  | 33.5 | 21.8
2006-09-20 | 17:00 | 7.72 | 7.69 |  |  | 33.4 | 21.722
2006-09-20 | 17:15 | 7.68 | 7.69 |  |  | 33.4 | 21.7
2006-09-20 | 17:30 | 7.65 | 7.69 |  |  | 33.4 | 21.692
2006-09-20 | 17:45 | 7.61 | 7.68 |  |  | 33.4 | 21.741
2006-09-20 | 18:00 | 7.58 | 7.68 |  |  | 33.4 | 21.742
2006-09-20 | 18:15 | 7.53 | 7.68 |  |  | 33.5 | 21.802
2006-09-20 | 18:30 | 7.47 | 7.67 |  |  | 33.3 | 21.672
2006-09-20 | 18:45 | 7.41 | 7.67 |  |  | 33.5 | 21.744
2006-09-20 | 19:00 | 7.34 | 7.66 |  |  | 33.4 | 21.68
2006-09-20 | 19:15 | 7.27 | 7.66 |  |  | 33.5 | 21.744
2006-09-20 | 19:30 | 7.21 | 7.65 |  |  | 33.3 | 21.667
2006-09-20 | 19:45 | 7.15 | 7.65 |  |  | 33.3 | 21.666
2006-09-20 | 20:00 | 7.1 | 7.64 |  |  | 33.5 | 21.764
2006-09-20 | 20:15 | 7.06 | 7.64 |  |  | 33.4 | 21.712
2006-09-20 | 20:30 | 7.02 | 7.64 |  |  | 33.3 | 21.653
2006-09-20 | 20:45 | 6.98 | 7.63 |  |  | 33.3 | 21.635
2006-09-20 | 21:00 | 6.95 | 7.63 |  |  | 33.4 | 21.7
2006-09-20 | 21:15 | 6.93 | 7.63 |  |  | 33.3 | 21.671
2006-09-20 | 21:30 | 6.91 | 7.63 |  |  | 33.2 | 21.584
2006-09-20 | 21:45 | 6.91 | 7.63 |  |  | 33.2 | 21.557
2006-09-20 | 22:00 | 6.91 | 7.63 |  |  | 33.3 | 21.655
2006-09-20 | 22:15 | 6.91 | 7.63 |  |  | 33.3 | 21.614
2006-09-20 | 22:30 | 6.93 | 7.63 |  |  | 33.1 | 21.531
2006-09-20 | 22:45 | 6.93 | 7.63 |  |  | 33.1 | 21.52
2006-09-20 | 23:00 | 6.93 | 7.63 |  |  | 33.1 | 21.496
2006-09-20 | 23:15 | 6.93 | 7.63 |  |  | 33.2 | 21.554
2006-09-20 | 23:30 | 6.93 | 7.63 |  |  | 33.2 | 21.586
2006-09-20 | 23:45 | 6.92 | 7.63 |  |  | 33.1 | 21.534
2006-09-21 | 0:00 | 6.92 | 7.63 |  |  | 33.1 | 21.515
2006-09-21 | 0:15 | 6.92 | 7.63 |  |  | 33 | 21.439
2006-09-21 | 0:30 | 6.91 | 7.63 |  |  | 33.1 | 21.542
2006-09-21 | 0:45 | 6.9 | 7.63 |  |  | 33 | 21.478
2006-09-21 | 1:01 | 6.88 | 7.63 |  |  | 32.9 | 21.389
2006-09-21 | 1:16 | 6.87 | 7.63 |  |  | 33 | 21.438
2006-09-21 | 1:31 | 6.86 | 7.62 |  |  | 33 | 21.422
2006-09-21 | 1:46 | 6.85 | 7.63 |  |  | 32.8 | 21.341
2006-09-21 | 2:01 | 6.85 | 7.63 |  |  | 32.9 | 21.374
2006-09-21 | 2:16 | 6.86 | 7.63 |  |  | 32.9 | 21.363
2006-09-21 | 2:31 | 6.87 | 7.63 |  |  | 32.8 | 21.289
2006-09-21 | 2:46 | 6.89 | 7.63 |  |  | 32.7 | 21.266
2006-09-21 | 3:01 | 6.91 | 7.63 |  |  | 32.7 | 21.267
2006-09-21 | 3:16 | 6.91 | 7.63 |  |  | 32.8 | 21.317
2006-09-21 | 3:31 | 6.91 | 7.63 |  |  | 32.7 | 21.236
2006-09-21 | 3:46 | 6.92 | 7.62 |  |  | 32.6 | 21.215
2006-09-21 | 4:01 | 6.92 | 7.63 |  |  | 32.6 | 21.19
2006-09-21 | 4:16 | 6.91 | 7.62 |  |  | 32.5 | 21.154
2006-09-21 | 4:31 | 6.91 | 7.62 |  |  | 32.4 | 21.085
2006-09-21 | 4:46 | 6.9 | 7.62 |  |  | 32.5 | 21.12
2006-09-21 | 5:01 | 6.88 | 7.62 |  |  | 32.3 | 20.991
2006-09-21 | 5:16 | 6.87 | 7.62 |  |  | 32.1 | 20.843
2006-09-21 | 5:31 | 6.84 | 7.61 |  |  | 32.2 | 20.902
2006-09-21 | 5:46 | 6.81 | 7.61 |  |  | 31.9 | 20.725
2006-09-21 | 6:01 | 6.78 | 7.61 |  |  | 31.8 | 20.659
2006-09-21 | 6:16 | 6.74 | 7.61 |  |  | 31.9 | 20.727
2006-09-21 | 6:31 | 6.71 | 7.61 |  |  | 31.8 | 20.666
2006-09-21 | 6:46 | 6.67 | 7.61 |  |  | 31.8 | 20.676
2006-09-21 | 7:01 | 6.63 | 7.61 |  |  | 31.6 | 20.544
2006-09-21 | 7:16 | 6.59 | 7.61 |  |  | 31.7 | 20.626
2006-09-21 | 7:31 | 6.56 | 7.61 |  |  | 31.7 | 20.6
2006-09-21 | 7:46 | 6.53 | 7.62 |  |  | 31.6 | 20.524
2006-09-21 | 8:01 | 6.52 | 7.61 |  |  | 31.5 | 20.506
2006-09-21 | 8:16 | 6.52 | 7.63 |  |  | 31.4 | 20.435
2006-09-21 | 8:31 | 6.54 | 7.63 |  |  | 31.4 | 20.416
2006-09-21 | 8:46 | 6.56 | 7.64 |  |  | 31.5 | 20.446
2006-09-21 | 9:01 | 6.61 | 7.64 |  |  | 31.5 | 20.446
2006-09-21 | 9:16 | 6.67 | 7.65 |  |  | 31.3 | 20.362
2006-09-21 | 9:31 | 6.74 | 7.65 |  |  | 31.4 | 20.431
2006-09-21 | 9:46 | 6.83 | 7.66 |  |  | 31.4 | 20.431
2006-09-21 | 10:01 | 6.9 | 7.66 |  |  | 31.3 | 20.36
2006-09-21 | 10:16 | 7.03 | 7.67 |  |  | 31.1 | 20.244
2006-09-21 | 10:31 | 7.17 | 7.67 |  |  | 31.1 | 20.218
2006-09-21 | 10:46 | 7.36 | 7.67 |  |  | 31.1 | 20.226
2006-09-21 | 11:01 | 7.52 | 7.67 |  |  | 31 | 20.174
2006-09-21 | 11:16 | 7.66 | 7.67 |  |  | 30.8 | 20.031
2006-09-21 | 11:31 | 7.86 | 7.67 |  |  | 30.9 | 20.071
2006-09-21 | 11:46 | 8.08 | 7.68 |  |  | 30.7 | 19.927
2006-09-21 | 12:01 | 8.29 | 7.68 |  |  | 30.8 | 20.025
2006-09-21 | 12:16 | 8.49 | 7.68 |  |  | 30.7 | 19.975
2006-09-21 | 12:31 | 8.67 | 7.69 |  |  | 30.8 | 20.002
2006-09-21 | 12:46 | 8.85 | 7.68 |  |  | 30.8 | 20.051
2006-09-21 | 13:00 | 9.02 | 7.69 |  |  | 30.7 | 19.964
2006-09-21 | 13:15 | 9.11 | 7.69 |  |  | 30.7 | 19.947
2006-09-21 | 13:30 | 9.25 | 7.69 |  |  | 30.7 | 19.95
2006-09-21 | 13:45 | 9.38 | 7.69 |  |  | 30.8 | 20.016
2006-09-21 | 14:00 | 9.55 | 7.7 |  |  | 30.7 | 19.944
2006-09-21 | 14:15 | 9.67 | 7.69 |  |  | 30.8 | 20.036
2006-09-21 | 14:30 | 9.87 | 7.69 |  |  | 31 | 20.137
2006-09-21 | 14:45 | 9.88 | 7.7 |  |  | 31 | 20.166
2006-09-21 | 15:00 | 9.92 | 7.69 |  |  | 31.1 | 20.19
2006-09-21 | 15:15 | 9.89 | 7.69 |  |  | 31.1 | 20.23
2006-09-21 | 15:30 | 9.79 | 7.69 |  |  | 31.1 | 20.23
2006-09-21 | 15:45 | 9.71 | 7.69 |  |  | 31.2 | 20.276
2006-09-21 | 16:00 | 9.61 | 7.68 |  |  | 31.4 | 20.419
2006-09-21 | 16:15 | 9.52 | 7.69 |  |  | 31.3 | 20.352
2006-09-21 | 16:30 | 9.39 | 7.68 |  |  | 31.4 | 20.39
2006-09-21 | 16:45 | 9.3 | 7.68 |  |  | 31.4 | 20.439
2006-09-21 | 17:00 | 9.29 | 7.68 |  |  | 31.4 | 20.408
2006-09-21 | 17:15 | 9.17 | 7.68 |  |  | 31.5 | 20.496
2006-09-21 | 17:30 | 9 | 7.68 |  |  | 31.5 | 20.495
2006-09-21 | 17:45 | 8.84 | 7.67 |  |  | 31.6 | 20.562
2006-09-21 | 18:00 | 8.69 | 7.67 |  |  | 31.8 | 20.68
2006-09-21 | 18:15 | 8.52 | 7.66 |  |  | 32.4 | 21.052
2006-09-21 | 18:30 | 8.33 | 7.65 |  |  | 32.5 | 21.131
2006-09-21 | 18:45 | 8.16 | 7.65 |  |  | 32.5 | 21.123
2006-09-21 | 19:00 | 7.99 | 7.64 |  |  | 32.5 | 21.109
2006-09-21 | 19:15 | 7.81 | 7.64 |  |  | 32.1 | 20.89
2006-09-21 | 19:30 | 7.62 | 7.63 |  |  | 32.2 | 20.942
2006-09-21 | 19:45 | 7.46 | 7.63 |  |  | 32.1 | 20.847
2006-09-21 | 20:00 | 7.3 | 7.63 |  |  | 32.1 | 20.875
2006-09-21 | 20:15 | 7.16 | 7.63 |  |  | 32.2 | 20.958
2006-09-21 | 20:30 | 7.03 | 7.62 |  |  | 32.2 | 20.929
2006-09-21 | 20:45 | 6.92 | 7.62 |  |  | 32.4 | 21.052
2006-09-21 | 21:00 | 6.8 | 7.62 |  |  | 32.4 | 21.073
2006-09-21 | 21:15 | 6.69 | 7.61 |  |  | 32.3 | 20.98
2006-09-21 | 21:30 | 6.6 | 7.62 |  |  | 32.4 | 21.044
2006-09-21 | 21:45 | 6.54 | 7.61 |  |  | 32.5 | 21.114
2006-09-21 | 22:00 | 6.45 | 7.61 |  |  | 32.4 | 21.054
2006-09-21 | 22:15 | 6.42 | 7.61 |  |  | 32.3 | 21.002
2006-09-21 | 22:30 | 6.38 | 7.61 |  |  | 32.3 | 21.013
2006-09-21 | 22:45 | 6.34 | 7.61 |  |  | 32.4 | 21.075
2006-09-21 | 23:00 | 6.3 | 7.61 |  |  | 32.2 | 20.921
2006-09-21 | 23:15 | 6.25 | 7.61 |  |  | 32.2 | 20.925
2006-09-21 | 23:30 | 6.19 | 7.61 |  |  | 32.1 | 20.846
2006-09-21 | 23:45 | 6.11 | 7.61 |  |  | 32.1 | 20.859
2006-09-22 | 0:00 | 6.04 | 7.61 |  |  | 32.1 | 20.867
2006-09-22 | 0:15 | 5.97 | 7.61 |  |  | 32 | 20.828
2006-09-22 | 0:30 | 5.9 | 7.61 |  |  | 31.9 | 20.762
2006-09-22 | 0:45 | 5.82 | 7.61 |  |  | 31.8 | 20.686
2006-09-22 | 1:01 | 5.75 | 7.6 |  |  | 31.7 | 20.624
2006-09-22 | 1:16 | 5.68 | 7.6 |  |  | 31.7 | 20.636
2006-09-22 | 1:31 | 5.61 | 7.6 |  |  | 31.4 | 20.415
2006-09-22 | 1:46 | 5.55 | 7.6 |  |  | 31.6 | 20.535
2006-09-22 | 2:01 | 5.48 | 7.6 |  |  | 31.5 | 20.487
2006-09-22 | 2:16 | 5.4 | 7.59 |  |  | 31.7 | 20.632
2006-09-22 | 2:31 | 5.32 | 7.59 |  |  | 32.1 | 20.86
2006-09-22 | 2:46 | 5.23 | 7.59 |  |  | 32 | 20.82
2006-09-22 | 3:01 | 5.14 | 7.59 |  |  | 32.1 | 20.854
2006-09-22 | 3:16 | 5.05 | 7.59 |  |  | 32.2 | 20.919
2006-09-22 | 3:31 | 4.96 | 7.59 |  |  | 32.3 | 21.012
2006-09-22 | 3:46 | 4.86 | 7.59 |  |  | 32.4 | 21.09
2006-09-22 | 4:01 | 4.75 | 7.59 |  |  | 32.5 | 21.108
2006-09-22 | 4:16 | 4.63 | 7.58 |  |  | 32.3 | 20.997
2006-09-22 | 4:31 | 4.53 | 7.59 |  |  | 32.6 | 21.219
2006-09-22 | 4:46 | 4.43 | 7.59 |  |  | 32.7 | 21.279
2006-09-22 | 5:01 | 4.32 | 7.58 |  |  | 32.7 | 21.276
2006-09-22 | 5:16 | 4.22 | 7.59 |  |  | 32.9 | 21.402
2006-09-22 | 5:31 | 4.13 | 7.59 |  |  | 33 | 21.457
2006-09-22 | 5:46 | 4.03 | 7.59 |  |  | 33.4 | 21.683
2006-09-22 | 6:01 | 3.94 | 7.59 |  |  | 33.3 | 21.637
2006-09-22 | 6:16 | 3.85 | 7.59 |  |  | 33.3 | 21.654
2006-09-22 | 6:31 | 3.76 | 7.59 |  |  | 33.4 | 21.714
2006-09-22 | 6:46 | 3.68 | 7.59 |  |  | 33.5 | 21.756
2006-09-22 | 7:01 | 3.6 | 7.59 |  |  | 33.5 | 21.77
2006-09-22 | 7:16 | 3.54 | 7.59 |  |  | 33.7 | 21.918
2006-09-22 | 7:31 | 3.47 | 7.58 |  |  | 33.7 | 21.886
2006-09-22 | 7:46 | 3.41 | 7.58 |  |  | 33.5 | 21.781
2006-09-22 | 8:01 | 3.37 | 7.59 |  |  | 33.6 | 21.841
2006-09-22 | 8:16 | 3.35 | 7.59 |  |  | 33.5 | 21.794
2006-09-22 | 8:31 | 3.33 | 7.59 |  |  | 33.4 | 21.723
2006-09-22 | 8:46 | 3.34 | 7.59 |  |  | 33.3 | 21.669
2006-09-22 | 9:01 | 3.38 | 7.61 |  |  | 33.3 | 21.645
2006-09-22 | 9:16 | 3.44 | 7.61 |  |  | 33.4 | 21.716
2006-09-22 | 9:31 | 3.52 | 7.62 |  |  | 33.1 | 21.541
2006-09-22 | 9:46 | 3.62 | 7.62 |  |  | 33.3 | 21.667
2006-09-22 | 10:01 | 3.75 | 7.63 |  |  | 33.3 | 21.676
2006-09-22 | 10:16 | 3.89 | 7.64 |  |  | 33 | 21.47
2006-09-22 | 10:31 | 4.05 | 7.64 |  |  | 33 | 21.482
2006-09-22 | 10:46 | 4.24 | 7.64 |  |  | 32.9 | 21.405
2006-09-22 | 11:01 | 4.46 | 7.65 |  |  | 33.3 | 21.634
2006-09-22 | 11:16 | 4.69 | 7.65 |  |  | 32.9 | 21.406
2006-09-22 | 11:31 | 4.94 | 7.66 |  |  | 33 | 21.454
2006-09-22 | 11:46 | 5.22 | 7.66 |  |  | 33.1 | 21.544
2006-09-22 | 12:01 | 5.56 | 7.66 |  |  | 32.9 | 21.37
2006-09-22 | 12:16 | 5.92 | 7.66 |  |  | 32.5 | 21.135
2006-09-22 | 12:31 | 6.31 | 7.67 |  |  | 32.6 | 21.199
2006-09-22 | 12:46 | 6.74 | 7.67 |  |  | 32.3 | 21.019
2006-09-22 | 13:00 | 7.16 | 7.67 |  |  | 32.1 | 20.893
2006-09-22 | 13:15 | 7.55 | 7.67 |  |  | 32.1 | 20.857
2006-09-22 | 13:30 | 7.93 | 7.68 |  |  | 31.8 | 20.688
2006-09-22 | 13:45 | 8.29 | 7.68 |  |  | 31.8 | 20.663
2006-09-22 | 14:00 | 8.63 | 7.69 |  |  | 31.6 | 20.557
2006-09-22 | 14:15 | 8.94 | 7.69 |  |  | 31.5 | 20.468
2006-09-22 | 14:30 | 9.22 | 7.69 |  |  | 31.3 | 20.367
2006-09-22 | 14:45 | 9.45 | 7.69 |  |  | 31.2 | 20.253
2006-09-22 | 15:00 | 9.63 | 7.69 |  |  | 31.1 | 20.187
2006-09-22 | 15:15 | 9.75 | 7.7 |  |  | 31.3 | 20.342
2006-09-22 | 15:30 | 9.81 | 7.7 |  |  | 31.1 | 20.213
2006-09-22 | 15:45 | 9.87 | 7.69 |  |  | 31.5 | 20.475
2006-09-22 | 16:00 | 9.89 | 7.7 |  |  | 31.3 | 20.314
2006-09-22 | 16:15 | 9.88 | 7.7 |  |  | 31.2 | 20.256
2006-09-22 | 16:30 | 9.79 | 7.69 |  |  | 31.1 | 20.214
2006-09-22 | 16:45 | 9.63 | 7.69 |  |  | 31.3 | 20.326
2006-09-22 | 17:00 | 9.46 | 7.69 |  |  | 31.2 | 20.279
2006-09-22 | 17:15 | 9.23 | 7.68 |  |  | 31 | 20.147
2006-09-22 | 17:30 | 8.98 | 7.68 |  |  | 31 | 20.171
2006-09-22 | 17:45 | 8.74 | 7.67 |  |  | 31.1 | 20.195
2006-09-22 | 18:00 | 8.52 | 7.67 |  |  | 31 | 20.178
2006-09-22 | 18:15 | 8.3 | 7.66 |  |  | 31.1 | 20.185
2006-09-22 | 18:30 | 8.1 | 7.66 |  |  | 31 | 20.129
2006-09-22 | 18:45 | 7.92 | 7.65 |  |  | 30.9 | 20.07
2006-09-22 | 19:00 | 7.75 | 7.64 |  |  | 30.9 | 20.083
2006-09-22 | 19:15 | 7.58 | 7.64 |  |  | 30.7 | 19.973
2006-09-22 | 19:30 | 7.41 | 7.63 |  |  | 30.6 | 19.908
2006-09-22 | 19:45 | 7.24 | 7.62 |  |  | 30.7 | 19.941
2006-09-22 | 20:00 | 7.06 | 7.61 |  |  | 30.9 | 20.102
2006-09-22 | 20:15 | 6.88 | 7.6 |  |  | 30.8 | 20.051
2006-09-22 | 20:30 | 6.71 | 7.6 |  |  | 30.6 | 19.919
2006-09-22 | 20:45 | 6.54 | 7.6 |  |  | 30.7 | 19.974
2006-09-22 | 21:00 | 6.39 | 7.6 |  |  | 30.8 | 20.03
2006-09-22 | 21:15 | 6.25 | 7.6 |  |  | 30.8 | 20.038
2006-09-22 | 21:30 | 6.12 | 7.59 |  |  | 30.8 | 19.989
2006-09-22 | 21:45 | 6 | 7.59 |  |  | 30.6 | 19.902
2006-09-22 | 22:00 | 5.89 | 7.59 |  |  | 30.9 | 20.104
2006-09-22 | 22:15 | 5.79 | 7.59 |  |  | 30.9 | 20.106
2006-09-22 | 22:30 | 5.68 | 7.59 |  |  | 30.7 | 19.946
2006-09-22 | 22:45 | 5.58 | 7.59 |  |  | 30.8 | 20.05
2006-09-22 | 23:00 | 5.49 | 7.58 |  |  | 30.7 | 19.973
2006-09-22 | 23:15 | 5.41 | 7.58 |  |  | 30.7 | 19.985
2006-09-22 | 23:30 | 5.33 | 7.58 |  |  | 30.5 | 19.85
2006-09-22 | 23:45 | 5.25 | 7.58 |  |  | 30.8 | 20.016
2006-09-23 | 0:00 | 5.19 | 7.58 |  |  | 30.7 | 19.987
2006-09-23 | 0:15 | 5.13 | 7.58 |  |  | 30.6 | 19.922
2006-09-23 | 0:30 | 5.08 | 7.58 |  |  | 30.5 | 19.793
2006-09-23 | 0:45 | 5.02 | 7.58 |  |  | 30.5 | 19.804
2006-09-23 | 1:01 | 4.98 | 7.58 |  |  | 30.7 | 19.927
2006-09-23 | 1:16 | 4.93 | 7.57 |  |  | 30.6 | 19.885
2006-09-23 | 1:31 | 4.88 | 7.57 |  |  | 30.6 | 19.879
2006-09-23 | 1:46 | 4.83 | 7.57 |  |  | 31 | 20.181
2006-09-23 | 2:01 | 4.78 | 7.57 |  |  | 31 | 20.125
2006-09-23 | 2:16 | 4.73 | 7.57 |  |  | 31.2 | 20.294
2006-09-23 | 2:31 | 4.67 | 7.57 |  |  | 31.4 | 20.421
2006-09-23 | 2:46 | 4.61 | 7.57 |  |  | 31.4 | 20.417
2006-09-23 | 3:01 | 4.55 | 7.57 |  |  | 31.1 | 20.206
2006-09-23 | 3:16 | 4.49 | 7.57 |  |  | 31.3 | 20.338
2006-09-23 | 3:31 | 4.43 | 7.57 |  |  | 31.6 | 20.518
2006-09-23 | 3:46 | 4.37 | 7.57 |  |  | 31.7 | 20.626
2006-09-23 | 4:01 | 4.31 | 7.57 |  |  | 31.7 | 20.605
2006-09-23 | 4:16 | 4.24 | 7.57 |  |  | 31.7 | 20.599
2006-09-23 | 4:31 | 4.19 | 7.57 |  |  | 31.9 | 20.736
2006-09-23 | 4:46 | 4.13 | 7.57 |  |  | 32 | 20.819
2006-09-23 | 5:01 | 4.07 | 7.57 |  |  | 31.9 | 20.732
2006-09-23 | 5:16 | 4.01 | 7.57 |  |  | 32.1 | 20.845
2006-09-23 | 5:31 | 3.95 | 7.57 |  |  | 32.1 | 20.88
2006-09-23 | 5:46 | 3.89 | 7.58 |  |  | 31.8 | 20.665
2006-09-23 | 6:01 | 3.84 | 7.58 |  |  | 32.2 | 20.954
2006-09-23 | 6:16 | 3.79 | 7.58 |  |  | 32.5 | 21.103
2006-09-23 | 6:31 | 3.74 | 7.58 |  |  | 32.4 | 21.066
2006-09-23 | 6:46 | 3.68 | 7.57 |  |  | 32.3 | 20.982
2006-09-23 | 7:01 | 3.63 | 7.57 |  |  | 32.6 | 21.204
2006-09-23 | 7:16 | 3.58 | 7.57 |  |  | 32.3 | 20.994
2006-09-23 | 7:31 | 3.54 | 7.57 |  |  | 31.9 | 20.751
2006-09-23 | 7:46 | 3.5 | 7.58 |  |  | 31.6 | 20.549
2006-09-23 | 8:01 | 3.47 | 7.58 |  |  | 32.6 | 21.217
2006-09-23 | 8:16 | 3.45 | 7.58 |  |  | 32.5 | 21.153
2006-09-23 | 8:31 | 3.45 | 7.59 |  |  | 32.7 | 21.268
2006-09-23 | 8:46 | 3.47 | 7.59 |  |  | 32.8 | 21.312
2006-09-23 | 9:01 | 3.5 | 7.6 |  |  | 33.1 | 21.492
2006-09-23 | 9:16 | 3.56 | 7.6 |  |  | 33 | 21.459
2006-09-23 | 9:31 | 3.65 | 7.61 |  |  | 32.4 | 21.083
2006-09-23 | 9:46 | 3.74 | 7.61 |  |  | 32.1 | 20.847
2006-09-23 | 10:01 | 3.86 | 7.62 |  |  | 31.9 | 20.764
2006-09-23 | 10:16 | 4 | 7.62 |  |  | 31.6 | 20.566
2006-09-23 | 10:31 | 4.15 | 7.63 |  |  | 31.5 | 20.463
2006-09-23 | 10:46 | 4.34 | 7.64 |  |  | 33 | 21.465
2006-09-23 | 11:01 | 4.56 | 7.65 |  |  | 32.8 | 21.308
2006-09-23 | 11:16 | 4.8 | 7.65 |  |  | 32.6 | 21.213
2006-09-23 | 11:31 | 5.06 | 7.66 |  |  | 32.1 | 20.862
2006-09-23 | 11:46 | 5.35 | 7.66 |  |  | 32.2 | 20.903
2006-09-23 | 12:01 | 5.67 | 7.67 |  |  | 32 | 20.77
2006-09-23 | 12:16 | 6.02 | 7.67 |  |  | 31.8 | 20.661
2006-09-23 | 12:31 | 6.43 | 7.67 |  |  | 32.7 | 21.25
2006-09-23 | 12:46 | 6.83 | 7.68 |  |  | 32.8 | 21.319
2006-09-23 | 13:00 | 7.25 | 7.68 |  |  | 33.2 | 21.555
2006-09-23 | 13:15 | 7.65 | 7.69 |  |  | 33.3 | 21.626
2006-09-23 | 13:30 | 8.05 | 7.69 |  |  | 33.1 | 21.545
2006-09-23 | 13:45 | 8.43 | 7.69 |  |  | 33.1 | 21.522
2006-09-23 | 14:00 | 8.78 | 7.69 |  |  | 32.8 | 21.289
2006-09-23 | 14:15 | 9.1 | 7.7 |  |  | 33.1 | 21.51
2006-09-23 | 14:30 | 9.4 | 7.7 |  |  | 33 | 21.451
2006-09-23 | 14:45 | 9.66 | 7.7 |  |  | 32.8 | 21.346
2006-09-23 | 15:00 | 9.86 | 7.7 |  |  | 32.6 | 21.206
2006-09-23 | 15:15 | 10.01 | 7.7 |  |  | 32.4 | 21.084
2006-09-23 | 15:30 | 10.11 | 7.7 |  |  | 32.6 | 21.182
2006-09-23 | 15:45 | 10.2 | 7.7 |  |  | 32.5 | 21.133
2006-09-23 | 16:00 | 10.23 | 7.7 |  |  | 32.6 | 21.185
2006-09-23 | 16:15 | 10.22 | 7.7 |  |  | 32.3 | 20.98
2006-09-23 | 16:30 | 10.16 | 7.7 |  |  | 32.4 | 21.039
2006-09-23 | 16:45 | 10.03 | 7.69 |  |  | 32.1 | 20.834
2006-09-23 | 17:00 | 9.86 | 7.69 |  |  | 31.9 | 20.75
2006-09-23 | 17:15 | 9.66 | 7.68 |  |  | 31.6 | 20.54
2006-09-23 | 17:30 | 9.41 | 7.68 |  |  | 31.4 | 20.421
2006-09-23 | 17:45 | 9.2 | 7.67 |  |  | 31.2 | 20.273
2006-09-23 | 18:00 | 9 | 7.66 |  |  | 31.2 | 20.254
2006-09-23 | 18:15 | 8.81 | 7.65 |  |  | 31.1 | 20.211
2006-09-23 | 18:30 | 8.62 | 7.65 |  |  | 30.8 | 20.048
2006-09-23 | 18:45 | 8.45 | 7.65 |  |  | 30.7 | 19.97
2006-09-23 | 19:00 | 8.28 | 7.64 |  |  | 30.7 | 19.984
2006-09-23 | 19:15 | 8.12 | 7.63 |  |  | 30.6 | 19.909
2006-09-23 | 19:30 | 7.96 | 7.62 |  |  | 30.4 | 19.778
2006-09-23 | 19:45 | 7.79 | 7.62 |  |  | 30.4 | 19.738
2006-09-23 | 20:00 | 7.62 | 7.61 |  |  | 30.3 | 19.69
2006-09-23 | 20:15 | 7.44 | 7.6 |  |  | 30.3 | 19.669
2006-09-23 | 20:30 | 7.28 | 7.6 |  |  | 31.5 | 20.451
2006-09-23 | 20:45 | 7.14 | 7.6 |  |  | 31.2 | 20.258
2006-09-23 | 21:00 | 7.01 | 7.59 |  |  | 32 | 20.808
2006-09-23 | 21:15 | 6.89 | 7.6 |  |  | 32.4 | 21.047
2006-09-23 | 21:30 | 6.77 | 7.6 |  |  | 32.3 | 20.995
2006-09-23 | 21:45 | 6.66 | 7.59 |  |  | 32.2 | 20.906
2006-09-23 | 22:00 | 6.56 | 7.59 |  |  | 32.3 | 20.97
2006-09-23 | 22:15 | 6.45 | 7.59 |  |  | 32 | 20.781
2006-09-23 | 22:30 | 6.35 | 7.58 |  |  | 31.8 | 20.69
2006-09-23 | 22:45 | 6.27 | 7.59 |  |  | 31.8 | 20.66
2006-09-23 | 23:00 | 6.19 | 7.58 |  |  | 31.6 | 20.522
2006-09-23 | 23:15 | 6.12 | 7.58 |  |  | 31.3 | 20.324
2006-09-23 | 23:30 | 6.05 | 7.58 |  |  | 31.3 | 20.313
2006-09-23 | 23:45 | 5.99 | 7.58 |  |  | 31 | 20.141
2006-09-24 | 0:00 | 5.93 | 7.57 |  |  | 30.6 | 19.916
2006-09-24 | 0:15 | 5.88 | 7.57 |  |  | 30.5 | 19.82
2006-09-24 | 0:30 | 5.83 | 7.57 |  |  | 30.3 | 19.703
2006-09-24 | 0:45 | 5.77 | 7.57 |  |  | 30.3 | 19.69
2006-09-24 | 1:01 | 5.72 | 7.57 |  |  | 30.2 | 19.619
2006-09-24 | 1:16 | 5.68 | 7.57 |  |  | 30.1 | 19.564
2006-09-24 | 1:31 | 5.62 | 7.56 |  |  | 30.1 | 19.554
2006-09-24 | 1:46 | 5.58 | 7.56 |  |  | 30.2 | 19.629
2006-09-24 | 2:01 | 5.54 | 7.56 |  |  | 30.1 | 19.577
2006-09-24 | 2:16 | 5.49 | 7.55 |  |  | 30.1 | 19.583
2006-09-24 | 2:31 | 5.43 | 7.56 |  |  | 30 | 19.519
2006-09-24 | 2:46 | 5.37 | 7.56 |  |  | 29.9 | 19.433
2006-09-24 | 3:01 | 5.3 | 7.56 |  |  | 29.9 | 19.466
2006-09-24 | 3:16 | 5.22 | 7.56 |  |  | 29.8 | 19.384
2006-09-24 | 3:31 | 5.14 | 7.56 |  |  | 29.7 | 19.329
2006-09-24 | 3:46 | 5.06 | 7.56 |  |  | 29.8 | 19.361
2006-09-24 | 4:01 | 4.96 | 7.56 |  |  | 29.6 | 19.213
2006-09-24 | 4:16 | 4.87 | 7.56 |  |  | 29.5 | 19.145
2006-09-24 | 4:31 | 4.8 | 7.56 |  |  | 30.8 | 20.005
2006-09-24 | 4:46 | 4.71 | 7.55 |  |  | 30.4 | 19.759
2006-09-24 | 5:01 | 4.64 | 7.56 |  |  | 32.4 | 21.058
2006-09-24 | 5:16 | 4.56 | 7.56 |  |  | 31.3 | 20.36
2006-09-24 | 5:31 | 4.47 | 7.56 |  |  | 32.1 | 20.885
2006-09-24 | 5:46 | 4.4 | 7.56 |  |  | 31.6 | 20.563
2006-09-24 | 6:01 | 4.33 | 7.55 |  |  | 31.5 | 20.474
2006-09-24 | 6:16 | 4.27 | 7.56 |  |  | 31.4 | 20.44
2006-09-24 | 6:31 | 4.22 | 7.56 |  |  | 31.1 | 20.199
2006-09-24 | 6:46 | 4.15 | 7.56 |  |  | 31.1 | 20.2
2006-09-24 | 7:01 | 4.09 | 7.56 |  |  | 31.1 | 20.242
2006-09-24 | 7:16 | 4.04 | 7.56 |  |  | 31.1 | 20.213
2006-09-24 | 7:31 | 4 | 7.56 |  |  | 31.1 | 20.194
2006-09-24 | 7:46 | 3.96 | 7.57 |  |  | 30.8 | 20.022
2006-09-24 | 8:01 | 3.93 | 7.57 |  |  | 30.8 | 20.001
2006-09-24 | 8:16 | 3.91 | 7.57 |  |  | 30.6 | 19.9
2006-09-24 | 8:31 | 3.9 | 7.57 |  |  | 30.7 | 19.987
2006-09-24 | 8:46 | 3.92 | 7.58 |  |  | 30.6 | 19.871
2006-09-24 | 9:01 | 3.97 | 7.58 |  |  | 30.5 | 19.856
2006-09-24 | 9:16 | 4.04 | 7.58 |  |  | 30.3 | 19.706
2006-09-24 | 9:31 | 4.13 | 7.59 |  |  | 31.9 | 20.75
2006-09-24 | 9:46 | 4.24 | 7.59 |  |  | 31.5 | 20.485
2006-09-24 | 10:01 | 4.37 | 7.61 |  |  | 31.4 | 20.379
2006-09-24 | 10:16 | 4.52 | 7.62 |  |  | 31.1 | 20.235
2006-09-24 | 10:31 | 4.68 | 7.62 |  |  | 31.2 | 20.282
2006-09-24 | 10:46 | 4.87 | 7.63 |  |  | 31 | 20.18
2006-09-24 | 11:01 | 5.1 | 7.63 |  |  | 30.7 | 19.97
2006-09-24 | 11:16 | 5.35 | 7.64 |  |  | 30.6 | 19.911
2006-09-24 | 11:31 | 5.61 | 7.64 |  |  | 30.6 | 19.866
2006-09-24 | 11:46 | 5.89 | 7.65 |  |  | 30.3 | 19.723
2006-09-24 | 12:01 | 6.21 | 7.66 |  |  | 30.3 | 19.699
2006-09-24 | 12:16 | 6.56 | 7.66 |  |  | 30.2 | 19.6
2006-09-24 | 12:31 | 6.96 | 7.67 |  |  | 30 | 19.523
2006-09-24 | 12:46 | 7.39 | 7.67 |  |  | 29.9 | 19.452
2006-09-24 | 13:00 | 7.81 | 7.67 |  |  | 30.7 | 19.947
2006-09-24 | 13:15 | 8.22 | 7.68 |  |  | 30.6 | 19.911
2006-09-24 | 13:30 | 8.62 | 7.69 |  |  | 31.1 | 20.23
2006-09-24 | 13:45 | 9 | 7.69 |  |  | 31.1 | 20.226
2006-09-24 | 14:00 | 9.35 | 7.69 |  |  | 31.1 | 20.205
2006-09-24 | 14:15 | 9.67 | 7.7 |  |  | 31.3 | 20.368
2006-09-24 | 14:30 | 9.97 | 7.7 |  |  | 31.3 | 20.37
2006-09-24 | 14:45 | 10.21 | 7.7 |  |  | 31.1 | 20.222
2006-09-24 | 15:00 | 10.41 | 7.71 |  |  | 31 | 20.16
2006-09-24 | 15:15 | 10.56 | 7.71 |  |  | 31.3 | 20.326
2006-09-24 | 15:30 | 10.65 | 7.71 |  |  | 32.1 | 20.85
2006-09-24 | 15:45 | 10.72 | 7.71 |  |  | 31.8 | 20.686
2006-09-24 | 16:00 | 10.76 | 7.71 |  |  | 31.6 | 20.532
2006-09-24 | 16:15 | 10.74 | 7.71 |  |  | 31.2 | 20.294
2006-09-24 | 16:30 | 10.67 | 7.7 |  |  | 31.1 | 20.206
2006-09-24 | 16:45 | 10.53 | 7.7 |  |  | 33.6 | 21.84
2006-09-24 | 17:00 | 10.36 | 7.69 |  |  | 33.4 | 21.714
2006-09-24 | 17:15 | 10.16 | 7.69 |  |  | 33.3 | 21.629
2006-09-24 | 17:30 | 9.94 | 7.68 |  |  | 34 | 22.121
2006-09-24 | 17:45 | 9.74 | 7.67 |  |  | 33.6 | 21.817
2006-09-24 | 18:00 | 9.56 | 7.66 |  |  | 33.6 | 21.827
2006-09-24 | 18:15 | 9.38 | 7.66 |  |  | 33.2 | 21.566
2006-09-24 | 18:30 | 9.21 | 7.65 |  |  | 33.2 | 21.563
2006-09-24 | 18:45 | 9.05 | 7.64 |  |  | 33.2 | 21.608
2006-09-24 | 19:00 | 8.9 | 7.64 |  |  | 33.1 | 21.491
2006-09-24 | 19:15 | 8.75 | 7.63 |  |  | 32.9 | 21.385
2006-09-24 | 19:30 | 8.6 | 7.62 |  |  | 32.7 | 21.273
2006-09-24 | 19:45 | 8.43 | 7.61 |  |  | 32.5 | 21.141
2006-09-24 | 20:00 | 8.26 | 7.59 |  |  | 32.3 | 21.012
2006-09-24 | 20:15 | 8.08 | 7.59 |  |  | 32.3 | 20.984
2006-09-24 | 20:30 | 7.92 | 7.6 |  |  | 32.2 | 20.934
2006-09-24 | 20:45 | 7.78 | 7.6 |  |  | 31.9 | 20.714
2006-09-24 | 21:00 | 7.65 | 7.59 |  |  | 31.7 | 20.61
2006-09-24 | 21:15 | 7.54 | 7.6 |  |  | 31.6 | 20.548
2006-09-24 | 21:30 | 7.44 | 7.59 |  |  | 31.4 | 20.417
2006-09-24 | 21:45 | 7.32 | 7.59 |  |  | 31.4 | 20.419
2006-09-24 | 22:00 | 7.21 | 7.59 |  |  | 31.3 | 20.36
2006-09-24 | 22:15 | 7.11 | 7.59 |  |  | 32.5 | 21.132
2006-09-24 | 22:30 | 7.02 | 7.58 |  |  | 32.5 | 21.123
2006-09-24 | 22:45 | 6.93 | 7.59 |  |  | 32.9 | 21.391
2006-09-24 | 23:00 | 6.85 | 7.58 |  |  | 32.7 | 21.234
2006-09-24 | 23:15 | 6.76 | 7.58 |  |  | 32.6 | 21.176
2006-09-24 | 23:30 | 6.69 | 7.58 |  |  | 32.2 | 20.948
2006-09-24 | 23:45 | 6.63 | 7.58 |  |  | 32.3 | 20.966
2006-09-25 | 0:00 | 6.57 | 7.58 |  |  | 32 | 20.812
2006-09-25 | 0:15 | 6.52 | 7.58 |  |  | 31.8 | 20.669
2006-09-25 | 0:30 | 6.47 | 7.58 |  |  | 31.7 | 20.633
2006-09-25 | 0:45 | 6.43 | 7.58 |  |  | 31.5 | 20.489
2006-09-25 | 1:01 | 6.4 | 7.57 |  |  | 31.3 | 20.361
2006-09-25 | 1:16 | 6.36 | 7.57 |  |  | 31.2 | 20.257
2006-09-25 | 1:31 | 6.33 | 7.56 |  |  | 31 | 20.157
2006-09-25 | 1:46 | 6.3 | 7.56 |  |  | 31 | 20.128
2006-09-25 | 2:01 | 6.25 | 7.56 |  |  | 30.7 | 19.967
2006-09-25 | 2:16 | 6.21 | 7.56 |  |  | 30.6 | 19.922
2006-09-25 | 2:31 | 6.16 | 7.56 |  |  | 30.6 | 19.912
2006-09-25 | 2:46 | 6.1 | 7.55 |  |  | 30.7 | 19.928
2006-09-25 | 3:01 | 6.04 | 7.55 |  |  | 30.5 | 19.83
2006-09-25 | 3:16 | 5.97 | 7.56 |  |  | 30.5 | 19.832
2006-09-25 | 3:31 | 5.9 | 7.56 |  |  | 30.6 | 19.865
2006-09-25 | 3:46 | 5.83 | 7.56 |  |  | 30.3 | 19.714
2006-09-25 | 4:01 | 5.76 | 7.56 |  |  | 30.2 | 19.612
2006-09-25 | 4:16 | 5.68 | 7.56 |  |  | 30.1 | 19.563
2006-09-25 | 4:31 | 5.61 | 7.56 |  |  | 30.7 | 19.935
2006-09-25 | 4:46 | 5.54 | 7.56 |  |  | 30.6 | 19.883
2006-09-25 | 5:01 | 5.47 | 7.56 |  |  | 30.8 | 20.013
2006-09-25 | 5:16 | 5.39 | 7.56 |  |  | 30.9 | 20.053
2006-09-25 | 5:31 | 5.31 | 7.56 |  |  | 30.8 | 20.033
2006-09-25 | 5:46 | 5.24 | 7.56 |  |  | 30.9 | 20.054
2006-09-25 | 6:01 | 5.17 | 7.55 |  |  | 30.7 | 19.954
2006-09-25 | 6:16 | 5.09 | 7.56 |  |  | 30.9 | 20.104
2006-09-25 | 6:31 | 5.03 | 7.56 |  |  | 30.8 | 20.022
2006-09-25 | 6:46 | 4.96 | 7.56 |  |  | 30.7 | 19.977
2006-09-25 | 7:01 | 4.89 | 7.56 |  |  | 30.9 | 20.069
2006-09-25 | 7:16 | 4.83 | 7.56 |  |  | 30.9 | 20.106
2006-09-25 | 7:31 | 4.77 | 7.55 |  |  | 30.9 | 20.106
2006-09-25 | 7:46 | 4.72 | 7.56 |  |  | 31 | 20.136
2006-09-25 | 8:01 | 4.68 | 7.56 |  |  | 30.9 | 20.083
2006-09-25 | 8:16 | 4.65 | 7.57 |  |  | 30.8 | 20.039
2006-09-25 | 8:31 | 4.64 | 7.57 |  |  | 30.9 | 20.102
2006-09-25 | 8:46 | 4.66 | 7.58 |  |  | 30.9 | 20.077
2006-09-25 | 9:01 | 4.7 | 7.58 |  |  | 30.8 | 20.014
2006-09-25 | 9:16 | 4.76 | 7.58 |  |  | 31 | 20.118
2006-09-25 | 9:31 | 4.84 | 7.59 |  |  | 30.8 | 20.024
2006-09-25 | 9:46 | 4.95 | 7.6 |  |  | 30.9 | 20.062
2006-09-25 | 10:01 | 5.07 | 7.61 |  |  | 30.7 | 19.972
2006-09-25 | 10:16 | 5.21 | 7.62 |  |  | 30.7 | 19.942
2006-09-25 | 10:31 | 5.38 | 7.62 |  |  | 30.6 | 19.909
2006-09-25 | 10:46 | 5.57 | 7.63 |  |  | 30.4 | 19.766
2006-09-25 | 11:01 | 5.79 | 7.63 |  |  | 30.2 | 19.635
2006-09-25 | 11:16 | 6.04 | 7.64 |  |  | 30.8 | 20.032
2006-09-25 | 11:31 | 6.31 | 7.64 |  |  | 30.6 | 19.92
2006-09-25 | 11:46 | 6.59 | 7.65 |  |  | 30.7 | 19.973
2006-09-25 | 12:01 | 6.89 | 7.65 |  |  | 30.5 | 19.845
2006-09-25 | 12:16 | 7.24 | 7.66 |  |  | 30.5 | 19.807
2006-09-25 | 12:31 | 7.65 | 7.66 |  |  | 30.2 | 19.658
2006-09-25 | 12:46 | 8.06 | 7.66 |  |  | 30.3 | 19.674
2006-09-25 | 13:00 | 8.49 | 7.67 |  |  | 30.2 | 19.605
2006-09-25 | 13:15 | 8.91 | 7.67 |  |  | 30.4 | 19.739
2006-09-25 | 13:30 | 9.32 | 7.67 |  |  | 30.2 | 19.616
2006-09-25 | 13:45 | 9.73 | 7.68 |  |  | 30.3 | 19.711
2006-09-25 | 14:00 | 10.09 | 7.68 |  |  | 30.2 | 19.607
2006-09-25 | 14:15 | 10.42 | 7.69 |  |  | 30.4 | 19.768
2006-09-25 | 14:30 | 10.73 | 7.69 |  |  | 30.2 | 19.651
2006-09-25 | 14:45 | 10.99 | 7.69 |  |  | 30.2 | 19.653
2006-09-25 | 15:00 | 11.19 | 7.71 |  |  | 30.2 | 19.599
2006-09-25 | 15:15 | 11.34 | 7.7 |  |  | 30.4 | 19.76
2006-09-25 | 15:30 | 11.45 | 7.7 |  |  | 30.5 | 19.805
2006-09-25 | 15:45 | 11.52 | 7.7 |  |  | 30.6 | 19.862
2006-09-25 | 16:00 | 11.55 | 7.7 |  |  | 30.6 | 19.88
2006-09-25 | 16:15 | 11.54 | 7.7 |  |  | 30.4 | 19.769
2006-09-25 | 16:30 | 11.47 | 7.69 |  |  | 30.5 | 19.824
2006-09-25 | 16:45 | 11.33 | 7.69 |  |  | 30.6 | 19.868
2006-09-25 | 17:00 | 11.16 | 7.68 |  |  | 30.4 | 19.763
2006-09-25 | 17:15 | 10.95 | 7.67 |  |  | 30.7 | 19.932
2006-09-25 | 17:30 | 10.71 | 7.66 |  |  | 30.5 | 19.814
2006-09-25 | 17:45 | 10.49 | 7.65 |  |  | 30.5 | 19.838
2006-09-25 | 18:00 | 10.28 | 7.64 |  |  | 30.6 | 19.872
2006-09-25 | 18:15 | 10.07 | 7.64 |  |  | 30.5 | 19.823
2006-09-25 | 18:30 | 9.88 | 7.63 |  |  | 30.4 | 19.743
2006-09-25 | 18:45 | 9.7 | 7.62 |  |  | 30.4 | 19.787
2006-09-25 | 19:00 | 9.53 | 7.62 |  |  | 30.3 | 19.701
2006-09-25 | 19:15 | 9.36 | 7.61 |  |  | 30.3 | 19.726
2006-09-25 | 19:30 | 9.18 | 7.6 |  |  | 30.4 | 19.767
2006-09-25 | 19:45 | 9.01 | 7.58 |  |  | 30.4 | 19.757
2006-09-25 | 20:00 | 8.82 | 7.58 |  |  | 30.3 | 19.68
2006-09-25 | 20:15 | 8.63 | 7.58 |  |  | 30.3 | 19.664
2006-09-25 | 20:30 | 8.46 | 7.57 |  |  | 30.2 | 19.61
2006-09-25 | 20:45 | 8.3 | 7.57 |  |  | 30.2 | 19.658
2006-09-25 | 21:00 | 8.16 | 7.57 |  |  | 30.5 | 19.806
2006-09-25 | 21:15 | 8.02 | 7.57 |  |  | 30.4 | 19.739
2006-09-25 | 21:30 | 7.89 | 7.57 |  |  | 30.6 | 19.917
2006-09-25 | 21:45 | 7.77 | 7.57 |  |  | 30.3 | 19.726
2006-09-25 | 22:00 | 7.67 | 7.57 |  |  | 30.4 | 19.768
2006-09-25 | 22:15 | 7.57 | 7.56 |  |  | 30.3 | 19.669
2006-09-25 | 22:30 | 7.47 | 7.56 |  |  | 30.3 | 19.699
2006-09-25 | 22:45 | 7.39 | 7.56 |  |  | 30.3 | 19.697
2006-09-25 | 23:00 | 7.29 | 7.56 |  |  | 30.3 | 19.688
2006-09-25 | 23:15 | 7.21 | 7.55 |  |  | 30.1 | 19.584
2006-09-25 | 23:30 | 7.14 | 7.55 |  |  | 30.2 | 19.641
2006-09-25 | 23:45 | 7.08 | 7.54 |  |  | 30.2 | 19.601
2006-09-26 | 0:00 | 7.02 | 7.54 |  |  | 30 | 19.489
2006-09-26 | 0:15 | 6.98 | 7.54 |  |  | 29.9 | 19.428
2006-09-26 | 0:30 | 6.95 | 7.53 |  |  | 29.9 | 19.43
2006-09-26 | 0:45 | 6.91 | 7.54 |  |  | 30 | 19.515
2006-09-26 | 1:01 | 6.88 | 7.54 |  |  | 29.8 | 19.371
2006-09-26 | 1:16 | 6.85 | 7.54 |  |  | 29.9 | 19.446
2006-09-26 | 1:31 | 6.83 | 7.55 |  |  | 29.9 | 19.42
2006-09-26 | 1:46 | 6.8 | 7.55 |  |  | 29.9 | 19.452
2006-09-26 | 2:01 | 6.77 | 7.54 |  |  | 30.1 | 19.554
2006-09-26 | 2:16 | 6.73 | 7.54 |  |  | 30.1 | 19.556
2006-09-26 | 2:31 | 6.69 | 7.54 |  |  | 30.1 | 19.578
2006-09-26 | 2:46 | 6.65 | 7.54 |  |  | 30 | 19.517
2006-09-26 | 3:01 | 6.6 | 7.54 |  |  | 29.7 | 19.336
2006-09-26 | 3:16 | 6.54 | 7.54 |  |  | 30.1 | 19.583
2006-09-26 | 3:31 | 6.49 | 7.53 |  |  | 30.2 | 19.61
2006-09-26 | 3:46 | 6.43 | 7.53 |  |  | 30 | 19.483
2006-09-26 | 4:01 | 6.37 | 7.54 |  |  | 30.3 | 19.68
2006-09-26 | 4:16 | 6.32 | 7.54 |  |  | 30.3 | 19.673
2006-09-26 | 4:31 | 6.26 | 7.54 |  |  | 30.2 | 19.655
2006-09-26 | 4:46 | 6.2 | 7.54 |  |  | 30.5 | 19.85
2006-09-26 | 5:01 | 6.16 | 7.54 |  |  | 30.3 | 19.673
2006-09-26 | 5:16 | 6.11 | 7.53 |  |  | 30.5 | 19.852
2006-09-26 | 5:31 | 6.05 | 7.55 |  |  | 30.5 | 19.825
2006-09-26 | 5:46 | 5.99 | 7.55 |  |  | 30.6 | 19.915
2006-09-26 | 6:01 | 5.94 | 7.55 |  |  | 30.6 | 19.882
2006-09-26 | 6:16 | 5.89 | 7.55 |  |  | 30.5 | 19.829
2006-09-26 | 6:31 | 5.83 | 7.55 |  |  | 30.6 | 19.862
2006-09-26 | 6:46 | 5.77 | 7.56 |  |  | 30.3 | 19.689
2006-09-26 | 7:01 | 5.71 | 7.55 |  |  | 30.4 | 19.785
2006-09-26 | 7:16 | 5.66 | 7.55 |  |  | 30.4 | 19.781
2006-09-26 | 7:31 | 5.6 | 7.54 |  |  | 30.4 | 19.774
2006-09-26 | 7:46 | 5.55 | 7.55 |  |  | 30.4 | 19.75
2006-09-26 | 8:01 | 5.52 | 7.55 |  |  | 30.3 | 19.705
2006-09-26 | 8:16 | 5.49 | 7.56 |  |  | 30.3 | 19.724
2006-09-26 | 8:31 | 5.47 | 7.56 |  |  | 30.4 | 19.737
2006-09-26 | 8:46 | 5.48 | 7.57 |  |  | 30.4 | 19.733
2006-09-26 | 9:01 | 5.51 | 7.58 |  |  | 30.2 | 19.617
2006-09-26 | 9:16 | 5.57 | 7.58 |  |  | 30.5 | 19.809
2006-09-26 | 9:31 | 5.64 | 7.59 |  |  | 30.5 | 19.843
2006-09-26 | 9:46 | 5.73 | 7.59 |  |  | 30.4 | 19.786
2006-09-26 | 10:01 | 5.84 | 7.6 |  |  | 30.3 | 19.688
2006-09-26 | 10:16 | 5.96 | 7.6 |  |  | 30.3 | 19.692
2006-09-26 | 10:31 | 6.12 | 7.61 |  |  | 30.2 | 19.661
2006-09-26 | 10:46 | 6.3 | 7.62 |  |  | 30.3 | 19.675
2006-09-26 | 11:01 | 6.51 | 7.63 |  |  | 30 | 19.511
2006-09-26 | 11:16 | 6.76 | 7.64 |  |  | 29.9 | 19.467
2006-09-26 | 11:31 | 7.01 | 7.64 |  |  | 29.9 | 19.406
2006-09-26 | 11:46 | 7.3 | 7.64 |  |  | 29.8 | 19.343
2006-09-26 | 12:01 | 7.62 | 7.65 |  |  | 29.7 | 19.334
2006-09-26 | 12:16 | 7.97 | 7.65 |  |  | 29.6 | 19.219
2006-09-26 | 12:31 | 8.38 | 7.66 |  |  | 29.6 | 19.24
2006-09-26 | 12:46 | 8.82 | 7.67 |  |  | 29.4 | 19.132
2006-09-26 | 13:00 | 9.26 | 7.67 |  |  | 29.4 | 19.14
2006-09-26 | 13:15 | 9.69 | 7.67 |  |  | 29.5 | 19.175
2006-09-26 | 13:30 | 10.1 | 7.67 |  |  | 29.4 | 19.118
2006-09-26 | 13:45 | 10.52 | 7.68 |  |  | 29.3 | 19.036
2006-09-26 | 14:00 | 10.89 | 7.69 |  |  | 29.4 | 19.09
2006-09-26 | 14:15 | 11.23 | 7.69 |  |  | 30.4 | 19.775
2006-09-26 | 14:30 | 11.55 | 7.69 |  |  | 30.4 | 19.77
2006-09-26 | 14:45 | 11.81 | 7.7 |  |  | 30.4 | 19.748
2006-09-26 | 15:00 | 12.03 | 7.7 |  |  | 30.8 | 20.038
2006-09-26 | 15:15 | 12.19 | 7.7 |  |  | 30.9 | 20.093
2006-09-26 | 15:30 | 12.29 | 7.7 |  |  | 31.1 | 20.218
2006-09-26 | 15:45 | 12.36 | 7.7 |  |  | 31 | 20.176
2006-09-26 | 16:00 | 12.4 | 7.7 |  |  | 31 | 20.143
2006-09-26 | 16:15 | 12.4 | 7.7 |  |  | 31 | 20.171
2006-09-26 | 16:30 | 12.34 | 7.69 |  |  | 31.1 | 20.246
2006-09-26 | 16:45 | 12.21 | 7.68 |  |  | 31.2 | 20.266
2006-09-26 | 17:00 | 12.04 | 7.68 |  |  | 31 | 20.14
2006-09-26 | 17:15 | 11.84 | 7.67 |  |  | 30.9 | 20.104
2006-09-26 | 17:30 | 11.6 | 7.66 |  |  | 31.1 | 20.221
2006-09-26 | 17:45 | 11.39 | 7.65 |  |  | 31.3 | 20.34
2006-09-26 | 18:00 | 11.18 | 7.65 |  |  | 31.2 | 20.29
2006-09-26 | 18:15 | 10.99 | 7.64 |  |  | 31.1 | 20.217
2006-09-26 | 18:30 | 10.8 | 7.63 |  |  | 31.1 | 20.193
2006-09-26 | 18:45 | 10.62 | 7.62 |  |  | 31.1 | 20.186
2006-09-26 | 19:00 | 10.44 | 7.61 |  |  | 30.9 | 20.083
2006-09-26 | 19:15 | 10.27 | 7.6 |  |  | 30.8 | 20.033
2006-09-26 | 19:30 | 10.09 | 7.6 |  |  | 30.8 | 20.003
2006-09-26 | 19:45 | 9.91 | 7.58 |  |  | 30.6 | 19.915
2006-09-26 | 20:00 | 9.73 | 7.58 |  |  | 30.7 | 19.97
2006-09-26 | 20:15 | 9.54 | 7.57 |  |  | 30.7 | 19.943
2006-09-26 | 20:30 | 9.37 | 7.57 |  |  | 30.5 | 19.824
2006-09-26 | 20:45 | 9.21 | 7.56 |  |  | 30.4 | 19.766
2006-09-26 | 21:00 | 9.05 | 7.56 |  |  | 30.3 | 19.701
2006-09-26 | 21:15 | 8.91 | 7.56 |  |  | 30.2 | 19.635
2006-09-26 | 21:30 | 8.78 | 7.56 |  |  | 30.2 | 19.617
2006-09-26 | 21:45 | 8.65 | 7.56 |  |  | 30.1 | 19.59
2006-09-26 | 22:00 | 8.55 | 7.56 |  |  | 30.2 | 19.618
2006-09-26 | 22:15 | 8.44 | 7.56 |  |  | 30.1 | 19.563
2006-09-26 | 22:30 | 8.31 | 7.54 |  |  | 29.9 | 19.435
2006-09-26 | 22:45 | 8.21 | 7.55 |  |  | 29.8 | 19.358
2006-09-26 | 23:00 | 8.12 | 7.55 |  |  | 29.8 | 19.361
2006-09-26 | 23:15 | 8.03 | 7.55 |  |  | 29.7 | 19.308
2006-09-26 | 23:30 | 7.96 | 7.55 |  |  | 29.5 | 19.194
2006-09-26 | 23:45 | 7.89 | 7.55 |  |  | 29.4 | 19.078
2006-09-27 | 0:00 | 7.84 | 7.55 |  |  | 29.3 | 19.032
2006-09-27 | 0:15 | 7.79 | 7.55 |  |  | 29.3 | 19.056
2006-09-27 | 0:30 | 7.75 | 7.54 |  |  | 29.2 | 18.977
2006-09-27 | 0:45 | 7.72 | 7.54 |  |  | 29.1 | 18.918
2006-09-27 | 1:01 | 7.69 | 7.54 |  |  | 29 | 18.863
2006-09-27 | 1:16 | 7.66 | 7.53 |  |  | 29 | 18.843
2006-09-27 | 1:31 | 7.62 | 7.53 |  |  | 28.8 | 18.727
2006-09-27 | 1:46 | 7.6 | 7.53 |  |  | 28.8 | 18.7
2006-09-27 | 2:01 | 7.58 | 7.54 |  |  | 28.6 | 18.579
2006-09-27 | 2:16 | 7.55 | 7.55 |  |  | 28.6 | 18.612
2006-09-27 | 2:31 | 7.51 | 7.55 |  |  | 30.9 | 20.082
2006-09-27 | 2:46 | 7.45 | 7.54 |  |  | 30.7 | 19.932
2006-09-27 | 3:01 | 7.38 | 7.53 |  |  | 31.2 | 20.255
2006-09-27 | 3:16 | 7.31 | 7.53 |  |  | 31 | 20.133
2006-09-27 | 3:31 | 7.24 | 7.52 |  |  | 30.9 | 20.082
2006-09-27 | 3:46 | 7.16 | 7.54 |  |  | 30.8 | 20.042
2006-09-27 | 4:01 | 7.09 | 7.54 |  |  | 30.7 | 19.966
2006-09-27 | 4:16 | 7.02 | 7.53 |  |  | 30.5 | 19.811
2006-09-27 | 4:31 | 6.94 | 7.54 |  |  | 30.5 | 19.799
2006-09-27 | 4:46 | 6.88 | 7.54 |  |  | 30.5 | 19.831
2006-09-27 | 5:01 | 6.83 | 7.53 |  |  | 30.3 | 19.685
2006-09-27 | 5:16 | 6.78 | 7.52 |  |  | 30.2 | 19.654
2006-09-27 | 5:31 | 6.72 | 7.53 |  |  | 30.3 | 19.69
2006-09-27 | 5:46 | 6.66 | 7.54 |  |  | 30.2 | 19.626
2006-09-27 | 6:01 | 6.6 | 7.54 |  |  | 30.1 | 19.594
2006-09-27 | 6:16 | 6.54 | 7.54 |  |  | 30 | 19.526
2006-09-27 | 6:31 | 6.49 | 7.54 |  |  | 31.6 | 20.553
2006-09-27 | 6:46 | 6.43 | 7.54 |  |  | 31.9 | 20.713
2006-09-27 | 7:01 | 6.38 | 7.54 |  |  | 31.6 | 20.532
2006-09-27 | 7:16 | 6.32 | 7.55 |  |  | 31.6 | 20.509
2006-09-27 | 7:31 | 6.26 | 7.55 |  |  | 31.3 | 20.363
2006-09-27 | 7:46 | 6.22 | 7.55 |  |  | 31.2 | 20.306
2006-09-27 | 8:01 | 6.18 | 7.55 |  |  | 31.2 | 20.301
2006-09-27 | 8:16 | 6.15 | 7.56 |  |  | 31.1 | 20.193
2006-09-27 | 8:31 | 6.14 | 7.56 |  |  | 31 | 20.177
2006-09-27 | 8:46 | 6.16 | 7.57 |  |  | 30.9 | 20.098
2006-09-27 | 9:01 | 6.19 | 7.58 |  |  | 30.9 | 20.108
2006-09-27 | 9:16 | 6.25 | 7.58 |  |  | 30.9 | 20.054
2006-09-27 | 9:31 | 6.33 | 7.59 |  |  | 30.8 | 19.995
2006-09-27 | 9:46 | 6.43 | 7.59 |  |  | 30.8 | 20.039
2006-09-27 | 10:01 | 6.53 | 7.6 |  |  | 30.7 | 19.954
2006-09-27 | 10:16 | 6.66 | 7.6 |  |  | 31.3 | 20.35
2006-09-27 | 10:31 | 6.81 | 7.62 |  |  | 32.1 | 20.865
2006-09-27 | 10:46 | 6.98 | 7.62 |  |  | 31.8 | 20.651
2006-09-27 | 11:01 | 7.19 | 7.64 |  |  | 31.7 | 20.637
2006-09-27 | 11:16 | 7.42 | 7.64 |  |  | 31.5 | 20.466
2006-09-27 | 11:31 | 7.66 | 7.65 |  |  | 31.4 | 20.394
2006-09-27 | 11:46 | 7.93 | 7.66 |  |  | 31.2 | 20.251
2006-09-27 | 12:01 | 8.21 | 7.67 |  |  | 31.4 | 20.384
2006-09-27 | 12:16 | 8.53 | 7.67 |  |  | 31.2 | 20.294
2006-09-27 | 12:31 | 8.89 | 7.68 |  |  | 31.2 | 20.274
2006-09-27 | 12:46 | 9.27 | 7.69 |  |  | 31.1 | 20.247
2006-09-27 | 13:00 | 9.65 | 7.69 |  |  | 31.1 | 20.232
2006-09-27 | 13:15 | 10.02 | 7.7 |  |  | 31.1 | 20.207
2006-09-27 | 13:30 | 10.38 | 7.7 |  |  | 31.2 | 20.262
2006-09-27 | 13:45 | 10.72 | 7.71 |  |  | 31.1 | 20.184
2006-09-27 | 14:00 | 11.04 | 7.71 |  |  | 31.1 | 20.24
2006-09-27 | 14:15 | 11.32 | 7.72 |  |  | 31.2 | 20.258
2006-09-27 | 14:30 | 11.58 | 7.72 |  |  | 31.1 | 20.209
2006-09-27 | 14:45 | 11.8 | 7.72 |  |  | 31.1 | 20.226
2006-09-27 | 15:00 | 11.98 | 7.72 |  |  | 31.2 | 20.252
2006-09-27 | 15:15 | 12.11 | 7.73 |  |  | 31.2 | 20.254
2006-09-27 | 15:30 | 12.2 | 7.73 |  |  | 31.2 | 20.256
2006-09-27 | 15:45 | 12.27 | 7.72 |  |  | 31.2 | 20.248
2006-09-27 | 16:00 | 12.3 | 7.72 |  |  | 31.2 | 20.248
2006-09-27 | 16:15 | 12.3 | 7.71 |  |  | 31.1 | 20.226
2006-09-27 | 16:30 | 12.26 | 7.71 |  |  | 31.1 | 20.213
2006-09-27 | 16:45 | 12.15 | 7.7 |  |  | 31.1 | 20.193
2006-09-27 | 17:00 | 12.02 | 7.69 |  |  | 31 | 20.166
2006-09-27 | 17:15 | 11.85 | 7.69 |  |  | 31 | 20.164
2006-09-27 | 17:30 | 11.66 | 7.68 |  |  | 31 | 20.138
2006-09-27 | 17:45 | 11.47 | 7.67 |  |  | 30.9 | 20.056
2006-09-27 | 18:00 | 11.28 | 7.66 |  |  | 30.8 | 20.036
2006-09-27 | 18:15 | 11.1 | 7.65 |  |  | 31 | 20.122
2006-09-27 | 18:30 | 10.92 | 7.64 |  |  | 30.8 | 20.047
2006-09-27 | 18:45 | 10.75 | 7.63 |  |  | 30.8 | 20.018
2006-09-27 | 19:00 | 10.58 | 7.62 |  |  | 30.8 | 20.015
2006-09-27 | 19:15 | 10.41 | 7.61 |  |  | 30.8 | 19.99
2006-09-27 | 19:30 | 10.25 | 7.59 |  |  | 30.7 | 19.94
2006-09-27 | 19:45 | 10.07 | 7.58 |  |  | 30.8 | 19.995
2006-09-27 | 20:00 | 9.9 | 7.57 |  |  | 30.6 | 19.91
2006-09-27 | 20:15 | 9.73 | 7.57 |  |  | 30.5 | 19.809
2006-09-27 | 20:30 | 9.57 | 7.56 |  |  | 30.5 | 19.82
2006-09-27 | 20:45 | 9.4 | 7.56 |  |  | 30.5 | 19.835
2006-09-27 | 21:00 | 9.26 | 7.56 |  |  | 30.3 | 19.672
2006-09-27 | 21:15 | 9.12 | 7.55 |  |  | 30.7 | 19.927
2006-09-27 | 21:30 | 9 | 7.56 |  |  | 30.1 | 19.575
2006-09-27 | 21:45 | 8.87 | 7.55 |  |  | 30 | 19.496
2006-09-27 | 22:00 | 8.75 | 7.55 |  |  | 29.9 | 19.457
2006-09-27 | 22:15 | 8.64 | 7.55 |  |  | 29.9 | 19.426
2006-09-27 | 22:30 | 8.54 | 7.54 |  |  | 30 | 19.486
2006-09-27 | 22:45 | 8.44 | 7.54 |  |  | 29.8 | 19.389
2006-09-27 | 23:00 | 8.35 | 7.53 |  |  | 29.8 | 19.386
2006-09-27 | 23:15 | 8.26 | 7.54 |  |  | 29.6 | 19.223
2006-09-27 | 23:30 | 8.19 | 7.55 |  |  | 29.5 | 19.143
2006-09-27 | 23:45 | 8.11 | 7.55 |  |  | 29.5 | 19.161
2006-09-28 | 0:00 | 8.05 | 7.55 |  |  | 29.3 | 19.073
2006-09-28 | 0:15 | 7.99 | 7.54 |  |  | 29.2 | 18.967
2006-09-28 | 0:30 | 7.94 | 7.55 |  |  | 29.1 | 18.935
2006-09-28 | 0:45 | 7.9 | 7.54 |  |  | 29 | 18.865
2006-09-28 | 1:01 | 7.85 | 7.54 |  |  | 28.7 | 18.623
2006-09-28 | 1:16 | 7.81 | 7.53 |  |  | 28.7 | 18.647
2006-09-28 | 1:31 | 7.76 | 7.53 |  |  | 28.5 | 18.528
2006-09-28 | 1:46 | 7.71 | 7.52 |  |  | 28.4 | 18.441
2006-09-28 | 2:01 | 7.67 | 7.53 |  |  | 28.9 | 18.764
2006-09-28 | 2:16 | 7.63 | 7.53 |  |  | 28.6 | 18.593
2006-09-28 | 2:31 | 7.57 | 7.52 |  |  | 28.6 | 18.564
2006-09-28 | 2:46 | 7.52 | 7.53 |  |  | 28.3 | 18.385
2006-09-28 | 3:01 | 7.45 | 7.53 |  |  | 28.1 | 18.279
2006-09-28 | 3:16 | 7.38 | 7.53 |  |  | 28.4 | 18.433
2006-09-28 | 3:31 | 7.31 | 7.52 |  |  | 28.3 | 18.398
2006-09-28 | 3:46 | 7.24 | 7.52 |  |  | 28.3 | 18.381
2006-09-28 | 4:01 | 7.17 | 7.52 |  |  | 28.2 | 18.328
2006-09-28 | 4:16 | 7.1 | 7.52 |  |  | 28 | 18.214
2006-09-28 | 4:31 | 7.04 | 7.52 |  |  | 28 | 18.216
2006-09-28 | 4:46 | 6.97 | 7.53 |  |  | 28 | 18.2
2006-09-28 | 5:01 | 6.91 | 7.53 |  |  | 28.1 | 18.289
2006-09-28 | 5:16 | 6.84 | 7.53 |  |  | 27.9 | 18.141
2006-09-28 | 5:31 | 6.77 | 7.53 |  |  | 28.5 | 18.539
2006-09-28 | 5:46 | 6.71 | 7.53 |  |  | 28.5 | 18.527
2006-09-28 | 6:01 | 6.64 | 7.53 |  |  | 28.4 | 18.492
2006-09-28 | 6:16 | 6.57 | 7.53 |  |  | 28.8 | 18.717
2006-09-28 | 6:31 | 6.5 | 7.53 |  |  | 28.8 | 18.726
2006-09-28 | 6:46 | 6.43 | 7.53 |  |  | 28.8 | 18.746
2006-09-28 | 7:01 | 6.36 | 7.53 |  |  | 29.1 | 18.902
2006-09-28 | 7:16 | 6.28 | 7.53 |  |  | 29.1 | 18.93
2006-09-28 | 7:31 | 6.2 | 7.54 |  |  | 29.1 | 18.928
2006-09-28 | 7:46 | 6.14 | 7.54 |  |  | 29.1 | 18.888
2006-09-28 | 8:01 | 6.08 | 7.55 |  |  | 29.2 | 18.97
2006-09-28 | 8:16 | 6.02 | 7.55 |  |  | 29.2 | 18.98
2006-09-28 | 8:31 | 5.99 | 7.56 |  |  | 29.3 | 19.067
2006-09-28 | 8:46 | 5.98 | 7.57 |  |  | 29.3 | 19.056
2006-09-28 | 9:01 | 6 | 7.57 |  |  | 29.3 | 19.049
2006-09-28 | 9:16 | 6.04 | 7.58 |  |  | 29.4 | 19.096
2006-09-28 | 9:31 | 6.12 | 7.59 |  |  | 29.4 | 19.098
2006-09-28 | 9:46 | 6.21 | 7.59 |  |  | 29.5 | 19.171
2006-09-28 | 10:01 | 6.31 | 7.6 |  |  | 29.5 | 19.144
2006-09-28 | 10:16 | 6.43 | 7.6 |  |  | 29.6 | 19.22
2006-09-28 | 10:31 | 6.59 | 7.61 |  |  | 29.5 | 19.144
2006-09-28 | 10:46 | 6.76 | 7.61 |  |  | 29.4 | 19.083
2006-09-28 | 11:01 | 6.97 | 7.62 |  |  | 29.5 | 19.145
2006-09-28 | 11:16 | 7.2 | 7.63 |  |  | 29.4 | 19.089
2006-09-28 | 11:31 | 7.46 | 7.64 |  |  | 29.6 | 19.229
2006-09-28 | 11:46 | 7.74 | 7.64 |  |  | 29.4 | 19.125
2006-09-28 | 12:01 | 8.05 | 7.65 |  |  | 29.5 | 19.164
2006-09-28 | 12:16 | 8.4 | 7.66 |  |  | 29.4 | 19.132
2006-09-28 | 12:31 | 8.8 | 7.66 |  |  | 29.7 | 19.298
2006-09-28 | 12:46 | 9.24 | 7.67 |  |  | 29.7 | 19.278
2006-09-28 | 13:00 | 9.68 | 7.67 |  |  | 29.6 | 19.263
2006-09-28 | 13:15 | 10.09 | 7.68 |  |  | 29.6 | 19.248
2006-09-28 | 13:30 | 10.49 | 7.68 |  |  | 29.5 | 19.207
2006-09-28 | 13:45 | 10.88 | 7.69 |  |  | 29.5 | 19.186
2006-09-28 | 14:00 | 11.23 | 7.69 |  |  | 29.7 | 19.308
2006-09-28 | 14:15 | 11.54 | 7.7 |  |  | 29.8 | 19.375
2006-09-28 | 14:30 | 11.83 | 7.71 |  |  | 29.8 | 19.342
2006-09-28 | 14:45 | 12.06 | 7.71 |  |  | 30.1 | 19.533
2006-09-28 | 15:00 | 12.25 | 7.7 |  |  | 30.2 | 19.647
2006-09-28 | 15:15 | 12.37 | 7.71 |  |  | 30.3 | 19.67
2006-09-28 | 15:30 | 12.44 | 7.71 |  |  | 30.3 | 19.673
2006-09-28 | 15:45 | 12.48 | 7.71 |  |  | 30.2 | 19.605
2006-09-28 | 16:00 | 12.5 | 7.71 |  |  | 30.2 | 19.605
2006-09-28 | 16:15 | 12.47 | 7.7 |  |  | 30.3 | 19.726
2006-09-28 | 16:30 | 12.38 | 7.7 |  |  | 30.2 | 19.618
2006-09-28 | 16:45 | 12.25 | 7.7 |  |  | 30.4 | 19.792
2006-09-28 | 17:00 | 12.06 | 7.69 |  |  | 30.5 | 19.81
2006-09-28 | 17:15 | 11.85 | 7.68 |  |  | 30.5 | 19.795
2006-09-28 | 17:30 | 11.61 | 7.67 |  |  | 30.4 | 19.78
2006-09-28 | 17:45 | 11.38 | 7.66 |  |  | 30.3 | 19.699
2006-09-28 | 18:00 | 11.15 | 7.65 |  |  | 30.4 | 19.787
2006-09-28 | 18:15 | 10.94 | 7.65 |  |  | 30.4 | 19.779
2006-09-28 | 18:30 | 10.73 | 7.63 |  |  | 30.3 | 19.72
2006-09-28 | 18:45 | 10.54 | 7.63 |  |  | 30.2 | 19.623
2006-09-28 | 19:00 | 10.34 | 7.62 |  |  | 30.2 | 19.617
2006-09-28 | 19:15 | 10.15 | 7.6 |  |  | 30.3 | 19.704
2006-09-28 | 19:30 | 9.95 | 7.6 |  |  | 30.2 | 19.641
2006-09-28 | 19:45 | 9.75 | 7.59 |  |  | 30.2 | 19.613
2006-09-28 | 20:00 | 9.56 | 7.57 |  |  | 30.1 | 19.564
2006-09-28 | 20:15 | 9.37 | 7.57 |  |  | 30.1 | 19.588
2006-09-28 | 20:30 | 9.18 | 7.57 |  |  | 30.1 | 19.539
2006-09-28 | 20:45 | 9 | 7.56 |  |  | 29.9 | 19.466
2006-09-28 | 21:00 | 8.84 | 7.56 |  |  | 30 | 19.489
2006-09-28 | 21:15 | 8.68 | 7.56 |  |  | 29.8 | 19.397
2006-09-28 | 21:30 | 8.53 | 7.56 |  |  | 29.9 | 19.42
2006-09-28 | 21:45 | 8.39 | 7.56 |  |  | 29.7 | 19.321
2006-09-28 | 22:00 | 8.25 | 7.56 |  |  | 29.8 | 19.384
2006-09-28 | 22:15 | 8.13 | 7.55 |  |  | 29.8 | 19.398
2006-09-28 | 22:30 | 8.02 | 7.55 |  |  | 29.7 | 19.281
2006-09-28 | 22:45 | 7.9 | 7.54 |  |  | 29.6 | 19.259
2006-09-28 | 23:00 | 7.8 | 7.54 |  |  | 29.8 | 19.347
2006-09-28 | 23:15 | 7.7 | 7.53 |  |  | 29.7 | 19.276
2006-09-28 | 23:30 | 7.61 | 7.54 |  |  | 29.7 | 19.287
2006-09-28 | 23:45 | 7.53 | 7.54 |  |  | 29.6 | 19.26
2006-09-29 | 0:00 | 7.46 | 7.54 |  |  | 29.6 | 19.237
2006-09-29 | 0:15 | 7.4 | 7.54 |  |  | 29.5 | 19.143
2006-09-29 | 0:30 | 7.34 | 7.53 |  |  | 29.4 | 19.124
2006-09-29 | 0:45 | 7.29 | 7.53 |  |  | 29.4 | 19.11
2006-09-29 | 1:01 | 7.24 | 7.53 |  |  | 29.4 | 19.093
2006-09-29 | 1:16 | 7.2 | 7.54 |  |  | 29.5 | 19.152
2006-09-29 | 1:31 | 7.15 | 7.54 |  |  | 29.4 | 19.124
2006-09-29 | 1:46 | 7.1 | 7.54 |  |  | 29.3 | 19.025
2006-09-29 | 2:01 | 7.03 | 7.54 |  |  | 29.4 | 19.085
2006-09-29 | 2:16 | 6.97 | 7.54 |  |  | 29.3 | 19.036
2006-09-29 | 2:31 | 6.9 | 7.53 |  |  | 29.2 | 18.99
2006-09-29 | 2:46 | 6.83 | 7.52 |  |  | 29.2 | 18.958
2006-09-29 | 3:01 | 6.75 | 7.52 |  |  | 29 | 18.877
2006-09-29 | 3:16 | 6.68 | 7.52 |  |  | 29.1 | 18.89
2006-09-29 | 3:31 | 6.6 | 7.52 |  |  | 29 | 18.849
2006-09-29 | 3:46 | 6.52 | 7.52 |  |  | 29 | 18.84
2006-09-29 | 4:01 | 6.46 | 7.52 |  |  | 28.9 | 18.781
2006-09-29 | 4:16 | 6.39 | 7.53 |  |  | 28.8 | 18.691
2006-09-29 | 4:31 | 6.32 | 7.53 |  |  | 28.7 | 18.674
2006-09-29 | 4:46 | 6.26 | 7.53 |  |  | 28.7 | 18.659
2006-09-29 | 5:01 | 6.2 | 7.53 |  |  | 28.7 | 18.657
2006-09-29 | 5:16 | 6.15 | 7.53 |  |  | 28.7 | 18.681
2006-09-29 | 5:31 | 6.09 | 7.53 |  |  | 28.8 | 18.701
2006-09-29 | 5:46 | 6.03 | 7.53 |  |  | 28.9 | 18.758
2006-09-29 | 6:01 | 5.96 | 7.53 |  |  | 28.9 | 18.785
2006-09-29 | 6:16 | 5.9 | 7.53 |  |  | 28.9 | 18.794
2006-09-29 | 6:31 | 5.83 | 7.53 |  |  | 29.1 | 18.919
2006-09-29 | 6:46 | 5.76 | 7.53 |  |  | 29 | 18.882
2006-09-29 | 7:01 | 5.69 | 7.53 |  |  | 29.1 | 18.9
2006-09-29 | 7:16 | 5.62 | 7.53 |  |  | 29.2 | 19.008
2006-09-29 | 7:31 | 5.55 | 7.54 |  |  | 29.3 | 19.073
2006-09-29 | 7:46 | 5.49 | 7.54 |  |  | 29.3 | 19.055
2006-09-29 | 8:01 | 5.45 | 7.54 |  |  | 29.4 | 19.111
2006-09-29 | 8:16 | 5.42 | 7.55 |  |  | 29.4 | 19.078
2006-09-29 | 8:31 | 5.4 | 7.56 |  |  | 29.5 | 19.155
2006-09-29 | 8:46 | 5.4 | 7.56 |  |  | 29.6 | 19.212
2006-09-29 | 9:01 | 5.41 | 7.58 |  |  | 29.6 | 19.23
2006-09-29 | 9:16 | 5.46 | 7.59 |  |  | 29.6 | 19.264
2006-09-29 | 9:31 | 5.53 | 7.59 |  |  | 29.7 | 19.297
2006-09-29 | 9:46 | 5.63 | 7.59 |  |  | 29.6 | 19.22
2006-09-29 | 10:01 | 5.74 | 7.6 |  |  | 29.7 | 19.334
2006-09-29 | 10:16 | 5.86 | 7.61 |  |  | 29.6 | 19.224
2006-09-29 | 10:31 | 6.01 | 7.61 |  |  | 29.5 | 19.206
2006-09-29 | 10:46 | 6.18 | 7.62 |  |  | 29.6 | 19.246
2006-09-29 | 11:01 | 6.37 | 7.63 |  |  | 29.4 | 19.123
2006-09-29 | 11:16 | 6.6 | 7.63 |  |  | 29.4 | 19.085
2006-09-29 | 11:31 | 6.84 | 7.64 |  |  | 29.3 | 19.068
2006-09-29 | 11:46 | 7.11 | 7.64 |  |  | 29.3 | 19.051
2006-09-29 | 12:01 | 7.4 | 7.65 |  |  | 29.3 | 19.031
2006-09-29 | 12:16 | 7.74 | 7.65 |  |  | 29.3 | 19.07
2006-09-29 | 12:31 | 8.13 | 7.66 |  |  | 29.3 | 19.074
2006-09-29 | 12:46 | 8.54 | 7.66 |  |  | 29.3 | 19.019
2006-09-29 | 13:00 | 8.97 | 7.67 |  |  | 29.3 | 19.016
2006-09-29 | 13:15 | 9.38 | 7.68 |  |  | 29.4 | 19.085
2006-09-29 | 13:30 | 9.79 | 7.68 |  |  | 29.4 | 19.09
2006-09-29 | 13:45 | 10.19 | 7.68 |  |  | 29.3 | 19.051
2006-09-29 | 14:00 | 10.55 | 7.68 |  |  | 29.3 | 19.039
2006-09-29 | 14:15 | 10.9 | 7.69 |  |  | 29.2 | 18.983
2006-09-29 | 14:30 | 11.23 | 7.69 |  |  | 29.2 | 18.964
2006-09-29 | 14:45 | 11.52 | 7.69 |  |  | 28.8 | 18.737
2006-09-29 | 15:00 | 11.77 | 7.69 |  |  | 28.6 | 18.581
2006-09-29 | 15:15 | 11.95 | 7.69 |  |  | 28.6 | 18.6
2006-09-29 | 15:30 | 12.09 | 7.69 |  |  | 28.5 | 18.543
2006-09-29 | 15:45 | 12.18 | 7.69 |  |  | 28.5 | 18.503
2006-09-29 | 16:00 | 12.23 | 7.69 |  |  | 28.3 | 18.374
2006-09-29 | 16:15 | 12.24 | 7.68 |  |  | 28.3 | 18.417
2006-09-29 | 16:30 | 12.21 | 7.68 |  |  | 28.3 | 18.415
2006-09-29 | 16:45 | 12.12 | 7.67 |  |  | 28.2 | 18.302
2006-09-29 | 17:00 | 11.98 | 7.67 |  |  | 28.2 | 18.336
2006-09-29 | 17:15 | 11.81 | 7.66 |  |  | 28.3 | 18.382
2006-09-29 | 17:30 | 11.61 | 7.65 |  |  | 28.1 | 18.27
2006-09-29 | 17:45 | 11.43 | 7.64 |  |  | 28.4 | 18.431
2006-09-29 | 18:00 | 11.24 | 7.64 |  |  | 28.1 | 18.279
2006-09-29 | 18:15 | 11.04 | 7.62 |  |  | 28.2 | 18.31
2006-09-29 | 18:30 | 10.86 | 7.62 |  |  | 28.1 | 18.246
2006-09-29 | 18:45 | 10.68 | 7.61 |  |  | 28 | 18.209
2006-09-29 | 19:00 | 10.5 | 7.59 |  |  | 28 | 18.172
2006-09-29 | 19:15 | 10.32 | 7.58 |  |  | 28 | 18.192
2006-09-29 | 19:30 | 10.15 | 7.57 |  |  | 27.9 | 18.121
2006-09-29 | 19:45 | 9.97 | 7.56 |  |  | 27.9 | 18.126
2006-09-29 | 20:00 | 9.8 | 7.55 |  |  | 28 | 18.17
2006-09-29 | 20:15 | 9.64 | 7.55 |  |  | 27.8 | 18.09
2006-09-29 | 20:30 | 9.47 | 7.55 |  |  | 27.8 | 18.059
2006-09-29 | 20:45 | 9.33 | 7.54 |  |  | 27.8 | 18.06
2006-09-29 | 21:00 | 9.2 | 7.54 |  |  | 27.8 | 18.053
2006-09-29 | 21:15 | 9.08 | 7.54 |  |  | 27.6 | 17.932
2006-09-29 | 21:30 | 8.97 | 7.54 |  |  | 27.5 | 17.894
2006-09-29 | 21:45 | 8.85 | 7.54 |  |  | 27.6 | 17.948
2006-09-29 | 22:00 | 8.75 | 7.54 |  |  | 27.4 | 17.827
2006-09-29 | 22:15 | 8.64 | 7.53 |  |  | 27.5 | 17.865
2006-09-29 | 22:30 | 8.54 | 7.53 |  |  | 27.4 | 17.79
2006-09-29 | 22:45 | 8.44 | 7.53 |  |  | 27.3 | 17.757
2006-09-29 | 23:00 | 8.33 | 7.52 |  |  | 27.3 | 17.745
2006-09-29 | 23:15 | 8.23 | 7.52 |  |  | 27.4 | 17.807
2006-09-29 | 23:30 | 8.14 | 7.52 |  |  | 27.4 | 17.814
2006-09-29 | 23:45 | 8.06 | 7.52 |  |  | 27.4 | 17.794
2006-09-30 | 0:00 | 7.99 | 7.52 |  |  | 27.2 | 17.688
2006-09-30 | 0:15 | 7.93 | 7.52 |  |  | 27.3 | 17.735
2006-09-30 | 0:30 | 7.87 | 7.52 |  |  | 27.3 | 17.748
2006-09-30 | 0:45 | 7.82 | 7.51 |  |  | 27.2 | 17.655
2006-09-30 | 1:01 | 7.76 | 7.5 |  |  | 27.2 | 17.69
2006-09-30 | 1:16 | 7.71 | 7.5 |  |  | 27.2 | 17.708
2006-09-30 | 1:31 | 7.66 | 7.5 |  |  | 27.2 | 17.69
2006-09-30 | 1:46 | 7.61 | 7.5 |  |  | 27.1 | 17.586
2006-09-30 | 2:01 | 7.55 | 7.5 |  |  | 27.1 | 17.643
2006-09-30 | 2:16 | 7.49 | 7.5 |  |  | 27 | 17.548
2006-09-30 | 2:31 | 7.42 | 7.5 |  |  | 26.9 | 17.503
2006-09-30 | 2:46 | 7.36 | 7.5 |  |  | 26.9 | 17.517
2006-09-30 | 3:01 | 7.29 | 7.49 |  |  | 26.7 | 17.387
2006-09-30 | 3:16 | 7.21 | 7.49 |  |  | 26.6 | 17.317
2006-09-30 | 3:31 | 7.13 | 7.49 |  |  | 26.6 | 17.29
2006-09-30 | 3:46 | 7.06 | 7.49 |  |  | 26.5 | 17.202
2006-09-30 | 4:01 | 6.98 | 7.49 |  |  | 26.5 | 17.21
2006-09-30 | 4:16 | 6.9 | 7.48 |  |  | 26.2 | 17.054
2006-09-30 | 4:31 | 6.83 | 7.49 |  |  | 26.3 | 17.068
2006-09-30 | 4:46 | 6.75 | 7.49 |  |  | 26.1 | 16.991
2006-09-30 | 5:01 | 6.67 | 7.48 |  |  | 26 | 16.914
2006-09-30 | 5:16 | 6.59 | 7.48 |  |  | 26 | 16.891
2006-09-30 | 5:31 | 6.52 | 7.48 |  |  | 26 | 16.886
2006-09-30 | 5:46 | 6.45 | 7.48 |  |  | 26 | 16.871
2006-09-30 | 6:01 | 6.37 | 7.48 |  |  | 26 | 16.874
2006-09-30 | 6:16 | 6.3 | 7.48 |  |  | 26 | 16.903
2006-09-30 | 6:31 | 6.24 | 7.48 |  |  | 26.1 | 16.995
2006-09-30 | 6:46 | 6.17 | 7.48 |  |  | 26.1 | 16.983
2006-09-30 | 7:01 | 6.1 | 7.48 |  |  | 26.3 | 17.082
2006-09-30 | 7:16 | 6.04 | 7.49 |  |  | 26.3 | 17.086
2006-09-30 | 7:31 | 5.98 | 7.49 |  |  | 26.4 | 17.148
2006-09-30 | 7:46 | 5.93 | 7.49 |  |  | 26.6 | 17.314
2006-09-30 | 8:01 | 5.89 | 7.49 |  |  | 26.7 | 17.363
2006-09-30 | 8:16 | 5.86 | 7.5 |  |  | 26.7 | 17.381
2006-09-30 | 8:31 | 5.85 | 7.51 |  |  | 26.8 | 17.388
2006-09-30 | 8:46 | 5.86 | 7.52 |  |  | 26.9 | 17.472
2006-09-30 | 9:01 | 5.9 | 7.54 |  |  | 26.9 | 17.457
2006-09-30 | 9:16 | 5.95 | 7.55 |  |  | 26.9 | 17.506
2006-09-30 | 9:31 | 6.04 | 7.56 |  |  | 27 | 17.549
2006-09-30 | 9:46 | 6.16 | 7.57 |  |  | 27.1 | 17.584
2006-09-30 | 10:01 | 6.27 | 7.57 |  |  | 27.1 | 17.628
2006-09-30 | 10:16 | 6.41 | 7.58 |  |  | 27.3 | 17.724
2006-09-30 | 10:31 | 6.57 | 7.59 |  |  | 27.3 | 17.76
2006-09-30 | 10:46 | 6.75 | 7.61 |  |  | 27.4 | 17.825
2006-09-30 | 11:01 | 6.96 | 7.62 |  |  | 27.6 | 17.95
2006-09-30 | 11:16 | 7.2 | 7.62 |  |  | 27.4 | 17.829
2006-09-30 | 11:31 | 7.45 | 7.63 |  |  | 27.5 | 17.878
2006-09-30 | 11:46 | 7.72 | 7.64 |  |  | 27.7 | 18.01
2006-09-30 | 12:01 | 8.01 | 7.65 |  |  | 27.7 | 18.007
2006-09-30 | 12:16 | 8.32 | 7.66 |  |  | 27.7 | 18.003
2006-09-30 | 12:31 | 8.68 | 7.66 |  |  | 27.7 | 18.036
2006-09-30 | 12:46 | 9.07 | 7.67 |  |  | 27.8 | 18.072
2006-09-30 | 13:00 | 9.46 | 7.68 |  |  | 27.8 | 18.088
2006-09-30 | 13:15 | 9.82 | 7.67 |  |  | 28 | 18.198
2006-09-30 | 13:30 | 10.19 | 7.68 |  |  | 28.1 | 18.276
2006-09-30 | 13:45 | 10.56 | 7.68 |  |  | 28.1 | 18.283
2006-09-30 | 14:00 | 10.87 | 7.69 |  |  | 28.2 | 18.317
2006-09-30 | 14:15 | 11.16 | 7.69 |  |  | 28.2 | 18.362
2006-09-30 | 14:30 | 11.42 | 7.69 |  |  | 28.3 | 18.418
2006-09-30 | 14:45 | 11.63 | 7.7 |  |  | 28.3 | 18.389
2006-09-30 | 15:00 | 11.81 | 7.7 |  |  | 28.4 | 18.455
2006-09-30 | 15:15 | 11.93 | 7.7 |  |  | 28.4 | 18.435
2006-09-30 | 15:30 | 12 | 7.7 |  |  | 28.5 | 18.551
2006-09-30 | 15:45 | 12.04 | 7.7 |  |  | 28.5 | 18.499