## cont multiparameter audit
Organization | City Of Bend |  | Station ID |  | DR 160.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Tumalo Boulder Field
Contact Name | Drexell Barnes |  | Latitude
Number | 541-322-6368 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | 04F11310AC
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2006-06-06 12:00 | 9.74 | 6.88 | 10.1 | 0 | 33.1 | 0
Sonde |  | 9.93 | 6.84 | 10.01 | 88.6 | 34.9
Grade |  | A | A | A |  | A
Audit 2 | 2006-06-15 10:30 | 10.27 | 6.49 | 9.99 | 0 | 40.4 | 0
Sonde |  | 10.24 | 6.88 | 9.95 | 88.7 | 42.6
Grade |  | A | B | A |  | A
Audit 3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Sonde
Grade |  | #REF! | #REF! | #REF! |  | #REF!
Post-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | City Of Bend |  | Station ID |  | DR 160.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Tumalo Boulder Field
Contact Name | Drexell Barnes |  | Latitude |  | 0
Number | 541-322-6368 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-06-06 12:00 | 9.74 | 6.88 | 10.1 |  | 33.1
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | City Of Bend |  | Station ID |  | DR 160.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, d/s Tumalo Boulder Field
Contact Name | Drexell Barnes |  | Latitude |  | 0
Number | 541-322-6368 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-06-15 10:30 | 10.27 | 6.49 | 9.99 |  | 40.4
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | [5 empty] | A
Time | Computer set to atomic time | [5 empty] | A
Temp | VWR NIST traceable surface water | [5 empty] | A
DO | Winkler titration of surface water | [5 empty] | A
SpC | 148 µS/mL | [5 empty] | #DIV/0!
Ph | pH 7 standard solution | [5 empty] | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## data
Date | Time | Temp | pH | DO Conc | DO% | SpCond | TDS
|  | °C |  | mg/L | % | uS/cm | ug/L
2006-06-06 | 12:01 | 9.93 | 6.84 | 10.01 | 88.6 | 34.9 | 22.681
2006-06-06 | 12:16 | 10.13 | 6.86 | 9.83 | 87.3 | 35 | 22.724
2006-06-06 | 12:31 | 10.36 | 6.83 | 9.79 | 87.5 | 35 | 22.727
2006-06-06 | 12:46 | 10.59 | 6.87 | 9.7 | 87.1 | 35 | 22.767
2006-06-06 | 13:00 | 10.81 | 6.9 | 9.61 | 86.7 | 35 | 22.778
2006-06-06 | 13:15 | 11.06 | 6.92 | 9.51 | 86.4 | 35.2 | 22.863
2006-06-06 | 13:30 | 11.29 | 6.95 | 9.43 | 86.1 | 35.2 | 22.879
2006-06-06 | 13:45 | 11.53 | 6.97 | 9.36 | 85.9 | 35.3 | 22.933
2006-06-06 | 14:00 | 11.76 | 7.01 | 9.28 | 85.6 | 35.4 | 23.014
2006-06-06 | 14:15 | 11.98 | 7.04 | 9.22 | 85.5 | 35.5 | 23.046
2006-06-06 | 14:30 | 12.2 | 7.07 | 9.15 | 85.3 | 35.6 | 23.117
2006-06-06 | 14:45 | 12.42 | 7.09 | 9.09 | 85.2 | 35.7 | 23.179
2006-06-06 | 15:00 | 12.6 | 7.11 | 9.04 | 85.1 | 35.7 | 23.229
2006-06-06 | 15:15 | 12.79 | 7.12 | 8.99 | 84.9 | 35.8 | 23.244
2006-06-06 | 15:30 | 12.95 | 7.13 | 8.94 | 84.7 | 35.8 | 23.302
2006-06-06 | 15:45 | 13.05 | 7.13 | 8.91 | 84.6 | 35.9 | 23.314
2006-06-06 | 16:00 | 13.15 | 7.13 | 8.87 | 84.4 | 35.9 | 23.341
2006-06-06 | 16:15 | 13.21 | 7.12 | 8.84 | 84.3 | 35.9 | 23.359
2006-06-06 | 16:30 | 13.26 | 7.14 | 8.82 | 84.2 | 35.9 | 23.355
2006-06-06 | 16:45 | 13.26 | 7.12 | 8.81 | 84.1 | 35.9 | 23.35
2006-06-06 | 17:00 | 13.21 | 7.09 | 8.81 | 84 | 35.9 | 23.335
2006-06-06 | 17:15 | 13.18 | 7.07 | 8.8 | 83.8 | 35.8 | 23.29
2006-06-06 | 17:30 | 13.15 | 7.05 | 8.79 | 83.7 | 35.8 | 23.27
2006-06-06 | 17:45 | 13.12 | 7.02 | 8.78 | 83.5 | 35.7 | 23.206
2006-06-06 | 18:00 | 13.08 | 6.98 | 8.77 | 83.3 | 35.7 | 23.176
2006-06-06 | 18:15 | 13.06 | 6.95 | 8.75 | 83.2 | 35.6 | 23.143
2006-06-06 | 18:30 | 13.03 | 6.91 | 8.74 | 83 | 35.6 | 23.108
2006-06-06 | 18:45 | 13.02 | 6.88 | 8.73 | 82.9 | 35.5 | 23.051
2006-06-06 | 19:00 | 13 | 6.85 | 8.72 | 82.7 | 35.4 | 22.983
2006-06-06 | 19:15 | 12.94 | 6.83 | 8.73 | 82.8 | 35.2 | 22.895
2006-06-06 | 19:30 | 12.87 | 6.79 | 8.74 | 82.7 | 35.1 | 22.816
2006-06-06 | 19:45 | 12.8 | 6.78 | 8.75 | 82.7 | 34.9 | 22.71
2006-06-06 | 20:00 | 12.69 | 6.76 | 8.77 | 82.7 | 34.8 | 22.596
2006-06-06 | 20:15 | 12.58 | 6.74 | 8.79 | 82.7 | 34.6 | 22.502
2006-06-06 | 20:30 | 12.43 | 6.72 | 8.82 | 82.7 | 34.5 | 22.415
2006-06-06 | 20:45 | 12.27 | 6.7 | 8.86 | 82.7 | 34.4 | 22.347
2006-06-06 | 21:00 | 12.1 | 6.68 | 8.89 | 82.7 | 34.3 | 22.292
2006-06-06 | 21:15 | 11.9 | 6.67 | 8.94 | 82.8 | 34.2 | 22.2
2006-06-06 | 21:30 | 11.71 | 6.65 | 8.98 | 82.8 | 34.1 | 22.133
2006-06-06 | 21:45 | 11.53 | 6.63 | 9.03 | 82.9 | 33.9 | 22.059
2006-06-06 | 22:00 | 11.37 | 6.62 | 9.07 | 82.9 | 33.8 | 21.991
2006-06-06 | 22:15 | 11.23 | 6.6 | 9.1 | 83 | 33.8 | 21.941
2006-06-06 | 22:30 | 11.08 | 6.6 | 9.14 | 83 | 33.7 | 21.875
2006-06-06 | 22:45 | 10.93 | 6.58 | 9.17 | 83 | 33.6 | 21.828
2006-06-06 | 23:00 | 10.78 | 6.56 | 9.21 | 83.1 | 33.5 | 21.765
2006-06-06 | 23:15 | 10.63 | 6.57 | 9.25 | 83.1 | 33.4 | 21.718
2006-06-06 | 23:30 | 10.47 | 6.55 | 9.29 | 83.2 | 33.3 | 21.646
2006-06-06 | 23:45 | 10.33 | 6.55 | 9.32 | 83.2 | 33.3 | 21.614
2006-06-07 | 0:00 | 10.21 | 6.54 | 9.35 | 83.2 | 33.3 | 21.615
2006-06-07 | 0:15 | 10.08 | 6.53 | 9.38 | 83.3 | 33.2 | 21.608
2006-06-07 | 0:30 | 9.99 | 6.53 | 9.41 | 83.3 | 33.4 | 21.678
2006-06-07 | 0:45 | 9.9 | 6.53 | 9.43 | 83.4 | 33.4 | 21.74
2006-06-07 | 1:01 | 9.81 | 6.52 | 9.45 | 83.4 | 33.5 | 21.8
2006-06-07 | 1:16 | 9.73 | 6.52 | 9.47 | 83.4 | 33.6 | 21.84
2006-06-07 | 1:31 | 9.65 | 6.53 | 9.5 | 83.4 | 33.7 | 21.883
2006-06-07 | 1:46 | 9.59 | 6.53 | 9.52 | 83.5 | 33.8 | 21.943
2006-06-07 | 2:01 | 9.51 | 6.52 | 9.54 | 83.5 | 33.8 | 21.961
2006-06-07 | 2:16 | 9.41 | 6.52 | 9.56 | 83.5 | 33.8 | 21.949
2006-06-07 | 2:31 | 9.32 | 6.52 | 9.58 | 83.5 | 33.8 | 21.953
2006-06-07 | 2:46 | 9.23 | 6.52 | 9.6 | 83.5 | 33.8 | 21.959
2006-06-07 | 3:01 | 9.15 | 6.51 | 9.63 | 83.6 | 33.8 | 21.963
2006-06-07 | 3:16 | 9.05 | 6.51 | 9.66 | 83.6 | 33.7 | 21.919
2006-06-07 | 3:31 | 8.97 | 6.51 | 9.68 | 83.7 | 33.7 | 21.924
2006-06-07 | 3:46 | 8.92 | 6.51 | 9.68 | 83.6 | 33.8 | 21.995
2006-06-07 | 4:01 | 8.87 | 6.51 | 9.7 | 83.7 | 33.9 | 22.034
2006-06-07 | 4:16 | 8.83 | 6.5 | 9.72 | 83.7 | 34 | 22.083
2006-06-07 | 4:31 | 8.8 | 6.51 | 9.73 | 83.7 | 34.1 | 22.151
2006-06-07 | 4:46 | 8.78 | 6.52 | 9.73 | 83.8 | 34.2 | 22.255
2006-06-07 | 5:01 | 8.76 | 6.51 | 9.74 | 83.7 | 34.4 | 22.355
2006-06-07 | 5:16 | 8.73 | 6.52 | 9.74 | 83.7 | 34.5 | 22.419
2006-06-07 | 5:31 | 8.71 | 6.51 | 9.75 | 83.8 | 34.6 | 22.478
2006-06-07 | 5:46 | 8.68 | 6.51 | 9.76 | 83.8 | 34.7 | 22.536
2006-06-07 | 6:01 | 8.64 | 6.54 | 9.78 | 83.8 | 34.7 | 22.556
2006-06-07 | 6:16 | 8.61 | 6.53 | 9.79 | 83.9 | 34.8 | 22.608
2006-06-07 | 6:31 | 8.59 | 6.54 | 9.8 | 83.9 | 34.9 | 22.66
2006-06-07 | 6:46 | 8.56 | 6.54 | 9.82 | 84 | 34.9 | 22.672
2006-06-07 | 7:01 | 8.55 | 6.55 | 9.83 | 84.1 | 35 | 22.718
2006-06-07 | 7:16 | 8.55 | 6.55 | 9.84 | 84.2 | 35 | 22.771
2006-06-07 | 7:31 | 8.55 | 6.58 | 9.84 | 84.2 | 35.1 | 22.795
2006-06-07 | 7:46 | 8.53 | 6.58 | 9.85 | 84.3 | 35.1 | 22.801
2006-06-07 | 8:01 | 8.53 | 6.6 | 9.86 | 84.3 | 35.1 | 22.806
2006-06-07 | 8:16 | 8.54 | 6.61 | 9.86 | 84.4 | 35.1 | 22.808
2006-06-07 | 8:31 | 8.57 | 6.63 | 9.86 | 84.4 | 35.1 | 22.835
2006-06-07 | 8:46 | 8.62 | 6.64 | 9.86 | 84.5 | 35.3 | 22.948
2006-06-07 | 9:01 | 8.66 | 6.65 | 9.85 | 84.5 | 35.2 | 22.892
2006-06-07 | 9:16 | 8.74 | 6.67 | 9.84 | 84.6 | 35.2 | 22.867
2006-06-07 | 9:31 | 8.84 | 6.69 | 9.83 | 84.7 | 35.2 | 22.906
2006-06-07 | 9:46 | 8.93 | 6.71 | 9.82 | 84.8 | 35.2 | 22.911
2006-06-07 | 10:01 | 9.04 | 6.72 | 9.8 | 84.9 | 35.3 | 22.946
2006-06-07 | 10:16 | 9.18 | 6.73 | 9.78 | 84.9 | 35.4 | 22.992
2006-06-07 | 10:31 | 9.32 | 6.76 | 9.75 | 85 | 35.4 | 23.036
2006-06-07 | 10:46 | 9.48 | 6.77 | 9.72 | 85 | 35.5 | 23.093
2006-06-07 | 11:01 | 9.68 | 6.8 | 9.68 | 85.1 | 35.7 | 23.236
2006-06-07 | 11:16 | 9.88 | 6.83 | 9.64 | 85.2 | 35.9 | 23.345
2006-06-07 | 11:31 | 10.07 | 6.85 | 9.6 | 85.2 | 36 | 23.387
2006-06-07 | 11:46 | 10.26 | 6.88 | 9.56 | 85.2 | 36 | 23.398
2006-06-07 | 12:01 | 10.44 | 6.89 | 9.52 | 85.2 | 36 | 23.378
2006-06-07 | 12:16 | 10.64 | 6.92 | 9.48 | 85.2 | 36 | 23.39
2006-06-07 | 12:31 | 10.85 | 6.95 | 9.43 | 85.2 | 36 | 23.383
2006-06-07 | 12:46 | 11.09 | 6.97 | 9.38 | 85.2 | 36.1 | 23.438
2006-06-07 | 13:00 | 11.33 | 7 | 9.32 | 85.2 | 36.1 | 23.496
2006-06-07 | 13:15 | 11.54 | 7.03 | 9.27 | 85.1 | 36.2 | 23.511
2006-06-07 | 13:30 | 11.67 | 7.04 | 9.23 | 85 | 36.2 | 23.538
2006-06-07 | 13:45 | 11.76 | 7.06 | 9.21 | 85 | 36.3 | 23.575
2006-06-07 | 14:00 | 12.09 | 7.08 | 9.13 | 84.9 | 36.4 | 23.637
2006-06-07 | 14:15 | 12.37 | 7.1 | 9.07 | 84.8 | 36.5 | 23.702
2006-06-07 | 14:30 | 12.59 | 7.11 | 9.01 | 84.7 | 36.5 | 23.72
2006-06-07 | 14:45 | 12.78 | 7.13 | 8.97 | 84.7 | 36.6 | 23.772
2006-06-07 | 15:00 | 12.97 | 7.14 | 8.92 | 84.6 | 36.6 | 23.8
2006-06-07 | 15:15 | 13.11 | 7.17 | 8.89 | 84.6 | 36.6 | 23.809
2006-06-07 | 15:30 | 13.24 | 7.17 | 8.86 | 84.5 | 36.7 | 23.83
2006-06-07 | 15:45 | 13.33 | 7.17 | 8.84 | 84.5 | 36.7 | 23.872
2006-06-07 | 16:00 | 13.43 | 7.18 | 8.81 | 84.4 | 36.8 | 23.894
2006-06-07 | 16:15 | 13.49 | 7.18 | 8.78 | 84.3 | 36.7 | 23.881
2006-06-07 | 16:30 | 13.57 | 7.18 | 8.76 | 84.2 | 36.8 | 23.928
2006-06-07 | 16:45 | 13.64 | 7.16 | 8.73 | 84.1 | 36.8 | 23.931
2006-06-07 | 17:00 | 13.66 | 7.14 | 8.72 | 84 | 36.8 | 23.951
2006-06-07 | 17:15 | 13.62 | 7.12 | 8.7 | 83.7 | 36.9 | 23.961
2006-06-07 | 17:30 | 13.64 | 7.1 | 8.69 | 83.6 | 36.9 | 23.958
2006-06-07 | 17:45 | 13.61 | 7.07 | 8.68 | 83.5 | 36.9 | 23.975
2006-06-07 | 18:00 | 13.56 | 7.05 | 8.69 | 83.5 | 36.8 | 23.907
2006-06-07 | 18:15 | 13.48 | 7.02 | 8.7 | 83.4 | 36.7 | 23.85
2006-06-07 | 18:30 | 13.36 | 7 | 8.71 | 83.3 | 36.5 | 23.755
2006-06-07 | 18:45 | 13.26 | 6.97 | 8.73 | 83.3 | 36.4 | 23.676
2006-06-07 | 19:00 | 13.18 | 6.96 | 8.74 | 83.3 | 36.3 | 23.575
2006-06-07 | 19:15 | 13.1 | 6.92 | 8.75 | 83.2 | 36.1 | 23.438
2006-06-07 | 19:30 | 13.03 | 6.89 | 8.76 | 83.2 | 35.8 | 23.296
2006-06-07 | 19:45 | 12.93 | 6.86 | 8.77 | 83.1 | 35.6 | 23.136
2006-06-07 | 20:00 | 12.81 | 6.84 | 8.8 | 83.1 | 35.3 | 22.951
2006-06-07 | 20:15 | 12.71 | 6.81 | 8.81 | 83.1 | 35.1 | 22.806
2006-06-07 | 20:30 | 12.6 | 6.79 | 8.83 | 83.1 | 34.9 | 22.686
2006-06-07 | 20:45 | 12.49 | 6.77 | 8.86 | 83.1 | 34.7 | 22.567
2006-06-07 | 21:00 | 12.36 | 6.74 | 8.88 | 83.1 | 34.6 | 22.467
2006-06-07 | 21:15 | 12.22 | 6.71 | 8.91 | 83.1 | 34.4 | 22.368
2006-06-07 | 21:30 | 12.09 | 6.7 | 8.94 | 83.2 | 34.3 | 22.281
2006-06-07 | 21:45 | 11.96 | 6.67 | 8.98 | 83.2 | 34.2 | 22.224
2006-06-07 | 22:00 | 11.83 | 6.66 | 9.01 | 83.3 | 34.1 | 22.169
2006-06-07 | 22:15 | 11.68 | 6.64 | 9.05 | 83.3 | 34 | 22.089
2006-06-07 | 22:30 | 11.54 | 6.63 | 9.08 | 83.4 | 33.9 | 22.017
2006-06-07 | 22:45 | 11.38 | 6.62 | 9.12 | 83.4 | 33.7 | 21.904
2006-06-07 | 23:00 | 11.24 | 6.62 | 9.16 | 83.5 | 33.6 | 21.864
2006-06-07 | 23:15 | 11.13 | 6.59 | 9.19 | 83.6 | 33.7 | 21.888
2006-06-07 | 23:30 | 11 | 6.6 | 9.22 | 83.6 | 33.7 | 21.906
2006-06-07 | 23:45 | 10.87 | 6.59 | 9.26 | 83.7 | 33.7 | 21.895
2006-06-08 | 0:00 | 10.7 | 6.58 | 9.3 | 83.8 | 33.6 | 21.871
2006-06-08 | 0:15 | 10.54 | 6.57 | 9.34 | 83.8 | 33.6 | 21.866
2006-06-08 | 0:30 | 10.36 | 6.57 | 9.39 | 83.9 | 33.6 | 21.833
2006-06-08 | 0:45 | 10.19 | 6.56 | 9.44 | 84 | 33.6 | 21.808
2006-06-08 | 1:01 | 10.03 | 6.55 | 9.48 | 84 | 33.5 | 21.807
2006-06-08 | 1:16 | 9.87 | 6.56 | 9.52 | 84.1 | 33.6 | 21.811
2006-06-08 | 1:31 | 9.73 | 6.54 | 9.56 | 84.1 | 33.6 | 21.829
2006-06-08 | 1:46 | 9.58 | 6.54 | 9.59 | 84.2 | 33.6 | 21.84
2006-06-08 | 2:01 | 9.45 | 6.54 | 9.63 | 84.2 | 33.6 | 21.841
2006-06-08 | 2:16 | 9.32 | 6.54 | 9.66 | 84.2 | 33.6 | 21.836
2006-06-08 | 2:31 | 9.19 | 6.53 | 9.69 | 84.2 | 33.5 | 21.805
2006-06-08 | 2:46 | 9.08 | 6.53 | 9.73 | 84.3 | 33.6 | 21.842
2006-06-08 | 3:01 | 8.99 | 6.53 | 9.76 | 84.4 | 33.7 | 21.883
2006-06-08 | 3:16 | 8.89 | 6.53 | 9.78 | 84.4 | 33.7 | 21.895
2006-06-08 | 3:31 | 8.78 | 6.53 | 9.81 | 84.4 | 33.7 | 21.878
2006-06-08 | 3:46 | 8.7 | 6.53 | 9.83 | 84.4 | 33.7 | 21.921
2006-06-08 | 4:01 | 8.61 | 6.53 | 9.86 | 84.5 | 33.7 | 21.919
2006-06-08 | 4:16 | 8.53 | 6.53 | 9.88 | 84.5 | 33.8 | 21.938
2006-06-08 | 4:31 | 8.45 | 6.53 | 9.9 | 84.5 | 33.8 | 21.955
2006-06-08 | 4:46 | 8.37 | 6.53 | 9.92 | 84.5 | 33.8 | 21.972
2006-06-08 | 5:01 | 8.3 | 6.53 | 9.94 | 84.5 | 33.9 | 22.023
2006-06-08 | 5:16 | 8.24 | 6.53 | 9.96 | 84.6 | 34 | 22.073
2006-06-08 | 5:31 | 8.17 | 6.54 | 9.98 | 84.6 | 34 | 22.097
2006-06-08 | 5:46 | 8.14 | 6.54 | 9.99 | 84.7 | 34.2 | 22.212
2006-06-08 | 6:01 | 8.08 | 6.55 | 10.01 | 84.7 | 34.2 | 22.232
2006-06-08 | 6:16 | 8.05 | 6.54 | 10.02 | 84.8 | 34.3 | 22.306
2006-06-08 | 6:31 | 8.01 | 6.55 | 10.04 | 84.8 | 34.4 | 22.349
2006-06-08 | 6:46 | 7.97 | 6.56 | 10.06 | 84.9 | 34.4 | 22.369
2006-06-08 | 7:01 | 7.96 | 6.56 | 10.07 | 84.9 | 34.5 | 22.441
2006-06-08 | 7:16 | 7.94 | 6.58 | 10.09 | 85 | 34.5 | 22.457
2006-06-08 | 7:31 | 7.93 | 6.59 | 10.09 | 85.1 | 34.6 | 22.499
2006-06-08 | 7:46 | 7.93 | 6.6 | 10.1 | 85.1 | 34.7 | 22.55
2006-06-08 | 8:01 | 7.94 | 6.6 | 10.1 | 85.2 | 34.8 | 22.612
2006-06-08 | 8:16 | 7.95 | 6.61 | 10.11 | 85.3 | 34.9 | 22.664
2006-06-08 | 8:31 | 7.97 | 6.62 | 10.11 | 85.3 | 34.9 | 22.696
2006-06-08 | 8:46 | 8 | 6.64 | 10.11 | 85.3 | 35 | 22.756
2006-06-08 | 9:01 | 8.06 | 6.63 | 10.1 | 85.4 | 35.1 | 22.835
2006-06-08 | 9:16 | 8.14 | 6.66 | 10.09 | 85.5 | 35.2 | 22.891
2006-06-08 | 9:31 | 8.23 | 6.68 | 10.08 | 85.6 | 35.4 | 23.007
2006-06-08 | 9:46 | 8.38 | 6.68 | 10.05 | 85.7 | 35.7 | 23.178
2006-06-08 | 10:01 | 8.51 | 6.7 | 10.03 | 85.8 | 35.8 | 23.259
2006-06-08 | 10:16 | 8.67 | 6.73 | 10 | 85.8 | 36 | 23.37
2006-06-08 | 10:31 | 8.82 | 6.75 | 9.97 | 85.9 | 36.1 | 23.463
2006-06-08 | 10:46 | 8.98 | 6.77 | 9.94 | 85.9 | 36.2 | 23.525
2006-06-08 | 11:01 | 9.12 | 6.79 | 9.91 | 86 | 36.2 | 23.512
2006-06-08 | 11:16 | 9.29 | 6.81 | 9.87 | 86 | 36.2 | 23.53
2006-06-08 | 11:31 | 9.46 | 6.82 | 9.84 | 86 | 36.2 | 23.515
2006-06-08 | 11:46 | 9.65 | 6.85 | 9.8 | 86.1 | 36.2 | 23.547
2006-06-08 | 12:01 | 9.85 | 6.88 | 9.75 | 86.1 | 36.3 | 23.564
2006-06-08 | 12:16 | 10.07 | 6.89 | 9.7 | 86.1 | 36.3 | 23.606
2006-06-08 | 12:31 | 10.3 | 6.92 | 9.65 | 86.1 | 36.4 | 23.661
2006-06-08 | 12:46 | 10.54 | 6.95 | 9.59 | 86.1 | 36.5 | 23.731
2006-06-08 | 13:00 | 10.79 | 6.98 | 9.53 | 86 | 36.7 | 23.832
2006-06-08 | 13:15 | 11.01 | 7.01 | 9.48 | 86 | 36.7 | 23.869
2006-06-08 | 13:30 | 11.24 | 7.03 | 9.42 | 85.9 | 36.8 | 23.916
2006-06-08 | 13:45 | 11.46 | 7.05 | 9.37 | 85.9 | 36.8 | 23.94
2006-06-08 | 14:00 | 11.67 | 7.08 | 9.32 | 85.8 | 36.9 | 23.96
2006-06-08 | 14:15 | 11.87 | 7.09 | 9.27 | 85.8 | 36.8 | 23.943
2006-06-08 | 14:30 | 12.07 | 7.12 | 9.22 | 85.7 | 36.9 | 23.98
2006-06-08 | 14:45 | 12.25 | 7.13 | 9.17 | 85.6 | 36.9 | 23.984
2006-06-08 | 15:00 | 12.04 | 7.07 | 9.2 | 85.5 | 35.3 | 22.942
2006-06-08 | 15:15 | 12.66 | 7.17 | 9.07 | 85.4 | 37.3 | 24.265
2006-06-08 | 15:30 | 12.74 | 7.17 | 9.04 | 85.3 | 37.2 | 24.186
2006-06-08 | 15:45 | 12.86 | 7.17 | 9.01 | 85.3 | 37.3 | 24.222
2006-06-08 | 16:00 | 12.92 | 7.18 | 8.99 | 85.2 | 37.2 | 24.212
2006-06-08 | 16:15 | 12.99 | 7.18 | 8.97 | 85.1 | 37.3 | 24.252
2006-06-08 | 16:30 | 13.05 | 7.17 | 8.95 | 85 | 37.4 | 24.29
2006-06-08 | 16:45 | 13.06 | 7.16 | 8.93 | 84.9 | 37.4 | 24.293
2006-06-08 | 17:00 | 13.06 | 7.14 | 8.92 | 84.8 | 37.4 | 24.326
2006-06-08 | 17:15 | 13.05 | 7.12 | 8.91 | 84.7 | 37.4 | 24.328
2006-06-08 | 17:30 | 13.03 | 7.09 | 8.9 | 84.5 | 37.5 | 24.348
2006-06-08 | 17:45 | 12.99 | 7.05 | 8.9 | 84.4 | 37.5 | 24.351
2006-06-08 | 18:00 | 12.93 | 7.03 | 8.89 | 84.3 | 37.5 | 24.344
2006-06-08 | 18:15 | 12.85 | 6.99 | 8.89 | 84.1 | 37.4 | 24.31
2006-06-08 | 18:30 | 12.78 | 6.96 | 8.89 | 84 | 37.3 | 24.276
2006-06-08 | 18:45 | 12.72 | 6.93 | 8.89 | 83.8 | 37.3 | 24.267
2006-06-08 | 19:00 | 12.67 | 6.9 | 8.89 | 83.8 | 37.3 | 24.219
2006-06-08 | 19:15 | 12.61 | 6.88 | 8.9 | 83.7 | 37.1 | 24.141
2006-06-08 | 19:30 | 12.53 | 6.86 | 8.92 | 83.7 | 37 | 24.071
2006-06-08 | 19:45 | 12.4 | 6.82 | 8.94 | 83.7 | 36.9 | 23.953
2006-06-08 | 20:00 | 12.25 | 6.8 | 8.98 | 83.8 | 36.7 | 23.829
2006-06-08 | 20:15 | 12.1 | 6.78 | 9.01 | 83.8 | 36.5 | 23.739
2006-06-08 | 20:30 | 11.96 | 6.76 | 9.04 | 83.8 | 36.4 | 23.631
2006-06-08 | 20:45 | 11.83 | 6.74 | 9.07 | 83.9 | 36.2 | 23.545
2006-06-08 | 21:00 | 11.69 | 6.73 | 9.1 | 83.9 | 36.1 | 23.446
2006-06-08 | 21:15 | 11.56 | 6.71 | 9.14 | 83.9 | 35.9 | 23.337
2006-06-08 | 21:30 | 11.46 | 6.7 | 9.16 | 83.9 | 35.8 | 23.246
2006-06-08 | 21:45 | 11.39 | 6.69 | 9.18 | 84 | 35.7 | 23.173
2006-06-08 | 22:00 | 11.33 | 6.67 | 9.19 | 84 | 35.5 | 23.096
2006-06-08 | 22:15 | 11.25 | 6.66 | 9.21 | 84 | 35.4 | 23.036
2006-06-08 | 22:30 | 11.15 | 6.65 | 9.24 | 84.1 | 35.4 | 22.99
2006-06-08 | 22:45 | 11.05 | 6.65 | 9.27 | 84.1 | 35.4 | 22.983
2006-06-08 | 23:00 | 10.95 | 6.64 | 9.29 | 84.1 | 35.4 | 23.002
2006-06-08 | 23:15 | 10.83 | 6.65 | 9.32 | 84.2 | 35.4 | 23.032
2006-06-08 | 23:30 | 10.68 | 6.62 | 9.36 | 84.2 | 35.4 | 23.012
2006-06-08 | 23:45 | 10.54 | 6.62 | 9.4 | 84.3 | 35.4 | 23.021
2006-06-09 | 0:00 | 10.39 | 6.62 | 9.44 | 84.4 | 35.4 | 23.006
2006-06-09 | 0:15 | 10.23 | 6.6 | 9.48 | 84.5 | 35.4 | 23
2006-06-09 | 0:30 | 10.08 | 6.61 | 9.52 | 84.5 | 35.4 | 23.002
2006-06-09 | 0:45 | 9.94 | 6.59 | 9.56 | 84.5 | 35.5 | 23.047
2006-06-09 | 1:01 | 9.81 | 6.6 | 9.59 | 84.6 | 35.5 | 23.087
2006-06-09 | 1:16 | 9.67 | 6.58 | 9.63 | 84.7 | 35.6 | 23.128
2006-06-09 | 1:31 | 9.54 | 6.61 | 9.67 | 84.7 | 35.7 | 23.183
2006-06-09 | 1:46 | 9.41 | 6.6 | 9.71 | 84.8 | 35.7 | 23.215
2006-06-09 | 2:01 | 9.3 | 6.6 | 9.74 | 84.8 | 35.8 | 23.302
2006-06-09 | 2:16 | 9.21 | 6.59 | 9.76 | 84.9 | 36 | 23.395
2006-06-09 | 2:31 | 9.11 | 6.57 | 9.79 | 84.9 | 36.1 | 23.472
2006-06-09 | 2:46 | 9.02 | 6.58 | 9.81 | 84.9 | 36.3 | 23.563
2006-06-09 | 3:01 | 8.94 | 6.57 | 9.84 | 85 | 36.4 | 23.643
2006-06-09 | 3:16 | 8.85 | 6.59 | 9.86 | 85 | 36.4 | 23.689
2006-06-09 | 3:31 | 8.77 | 6.58 | 9.88 | 85 | 36.6 | 23.758
2006-06-09 | 3:46 | 8.7 | 6.58 | 9.9 | 85 | 36.7 | 23.837
2006-06-09 | 4:01 | 8.65 | 6.59 | 9.92 | 85.1 | 36.8 | 23.933
2006-06-09 | 4:16 | 8.61 | 6.58 | 9.93 | 85.1 | 37 | 24.04
2006-06-09 | 4:31 | 8.57 | 6.59 | 9.94 | 85.1 | 37.1 | 24.134
2006-06-09 | 4:46 | 8.5 | 6.58 | 9.96 | 85.1 | 37.2 | 24.161
2006-06-09 | 5:01 | 8.42 | 6.59 | 9.98 | 85.1 | 37.2 | 24.176
2006-06-09 | 5:16 | 8.34 | 6.58 | 10 | 85.1 | 37.2 | 24.163
2006-06-09 | 5:31 | 8.27 | 6.58 | 10.02 | 85.2 | 37.2 | 24.166
2006-06-09 | 5:46 | 8.22 | 6.59 | 10.04 | 85.2 | 37.3 | 24.222
2006-06-09 | 6:01 | 8.17 | 6.59 | 10.05 | 85.3 | 37.4 | 24.288
2006-06-09 | 6:16 | 8.14 | 6.59 | 10.07 | 85.3 | 37.5 | 24.353
2006-06-09 | 6:31 | 8.11 | 6.6 | 10.08 | 85.3 | 37.5 | 24.387
2006-06-09 | 6:46 | 8.07 | 6.61 | 10.09 | 85.4 | 37.5 | 24.404
2006-06-09 | 7:01 | 8.04 | 6.62 | 10.11 | 85.5 | 37.6 | 24.447
2006-06-09 | 7:16 | 8.04 | 6.63 | 10.12 | 85.5 | 37.7 | 24.513
2006-06-09 | 7:31 | 8.04 | 6.63 | 10.12 | 85.6 | 37.8 | 24.589
2006-06-09 | 7:46 | 8.05 | 6.65 | 10.13 | 85.6 | 37.9 | 24.64
2006-06-09 | 8:01 | 8.09 | 6.66 | 10.12 | 85.7 | 38.1 | 24.759
2006-06-09 | 8:16 | 8.13 | 6.66 | 10.12 | 85.7 | 38.3 | 24.875
2006-06-09 | 8:31 | 8.19 | 6.67 | 10.11 | 85.7 | 38.5 | 25.011
2006-06-09 | 8:46 | 8.25 | 6.69 | 10.1 | 85.8 | 38.6 | 25.093
2006-06-09 | 9:01 | 8.29 | 6.69 | 10.1 | 85.9 | 38.6 | 25.092
2006-06-09 | 9:16 | 8.36 | 6.72 | 10.1 | 86 | 38.6 | 25.098
2006-06-09 | 9:31 | 8.44 | 6.72 | 10.09 | 86.1 | 38.7 | 25.146
2006-06-09 | 9:46 | 8.52 | 6.74 | 10.08 | 86.2 | 38.6 | 25.078
2006-06-09 | 10:01 | 8.62 | 6.76 | 10.06 | 86.3 | 38.5 | 25.046
2006-06-09 | 10:16 | 8.75 | 6.78 | 10.04 | 86.3 | 38.6 | 25.08
2006-06-09 | 10:31 | 8.89 | 6.8 | 10.01 | 86.4 | 38.6 | 25.115
2006-06-09 | 10:46 | 9.05 | 6.83 | 9.98 | 86.4 | 38.7 | 25.176
2006-06-09 | 11:01 | 9.23 | 6.84 | 9.94 | 86.5 | 38.8 | 25.235
2006-06-09 | 11:16 | 9.4 | 6.86 | 9.9 | 86.5 | 38.9 | 25.282
2006-06-09 | 11:31 | 9.6 | 6.89 | 9.86 | 86.5 | 39 | 25.343
2006-06-09 | 11:46 | 9.81 | 6.92 | 9.82 | 86.6 | 39.1 | 25.428
2006-06-09 | 12:01 | 10.03 | 6.94 | 9.77 | 86.6 | 39.2 | 25.509
2006-06-09 | 12:16 | 10.25 | 6.96 | 9.72 | 86.6 | 39.3 | 25.561
2006-06-09 | 12:31 | 10.46 | 6.99 | 9.67 | 86.6 | 39.4 | 25.581
2006-06-09 | 12:46 | 10.65 | 7.02 | 9.62 | 86.5 | 39.3 | 25.566
2006-06-09 | 13:00 | 10.81 | 7.04 | 9.58 | 86.5 | 39.3 | 25.564
2006-06-09 | 13:15 | 10.93 | 7.06 | 9.53 | 86.3 | 39.3 | 25.555
2006-06-09 | 13:30 | 11.16 | 7.08 | 9.48 | 86.3 | 39.3 | 25.562
2006-06-09 | 13:45 | 11.44 | 7.1 | 9.42 | 86.3 | 39.5 | 25.682
2006-06-09 | 14:00 | 11.67 | 7.13 | 9.36 | 86.2 | 39.6 | 25.749
2006-06-09 | 14:15 | 11.88 | 7.15 | 9.29 | 86 | 39.9 | 25.954
2006-06-09 | 14:30 | 12.08 | 7.16 | 9.26 | 86.1 | 39.9 | 25.925
2006-06-09 | 14:45 | 12.31 | 7.17 | 9.2 | 86 | 40 | 25.971
2006-06-09 | 15:00 | 12.51 | 7.19 | 9.15 | 85.9 | 40.1 | 26.051
2006-06-09 | 15:15 | 12.69 | 7.2 | 9.1 | 85.8 | 40.1 | 26.074
2006-06-09 | 15:30 | 12.87 | 7.2 | 9.06 | 85.7 | 40.2 | 26.103
2006-06-09 | 15:45 | 12.97 | 7.21 | 9.02 | 85.5 | 40.3 | 26.183
2006-06-09 | 16:00 | 13.02 | 7.2 | 8.98 | 85.2 | 40.4 | 26.246
2006-06-09 | 16:15 | 13.12 | 7.2 | 8.95 | 85.2 | 40.4 | 26.229
2006-06-09 | 16:30 | 13.29 | 7.22 | 8.93 | 85.3 | 40.4 | 26.266
2006-06-09 | 16:45 | 13.37 | 7.22 | 8.91 | 85.2 | 40.4 | 26.272
2006-06-09 | 17:00 | 13.41 | 7.22 | 8.89 | 85.1 | 40.5 | 26.321
2006-06-09 | 17:15 | 13.36 | 7.19 | 8.87 | 84.9 | 40.5 | 26.349
2006-06-09 | 17:30 | 13.25 | 7.15 | 8.86 | 84.6 | 40.5 | 26.346
2006-06-09 | 17:45 | 13.22 | 7.13 | 8.87 | 84.6 | 40.6 | 26.376
2006-06-09 | 18:00 | 13.17 | 7.11 | 8.87 | 84.5 | 40.5 | 26.356
2006-06-09 | 18:15 | 13.15 | 7.07 | 8.85 | 84.3 | 40.6 | 26.358
2006-06-09 | 18:30 | 13.09 | 7.03 | 8.85 | 84.1 | 40.5 | 26.356
2006-06-09 | 18:45 | 12.98 | 6.99 | 8.86 | 84.1 | 40.5 | 26.306
2006-06-09 | 19:00 | 12.85 | 6.96 | 8.89 | 84.1 | 40.3 | 26.163
2006-06-09 | 19:15 | 12.71 | 6.92 | 8.91 | 84 | 40.1 | 26.033
2006-06-09 | 19:30 | 12.58 | 6.9 | 8.94 | 84 | 39.8 | 25.899
2006-06-09 | 19:45 | 12.45 | 6.88 | 8.96 | 84 | 39.7 | 25.79
2006-06-09 | 20:00 | 12.37 | 6.86 | 8.98 | 84 | 39.6 | 25.72
2006-06-09 | 20:15 | 12.29 | 6.84 | 8.99 | 84 | 39.5 | 25.647
2006-06-09 | 20:30 | 12.21 | 6.82 | 9.01 | 84 | 39.4 | 25.586
2006-06-09 | 20:45 | 12.13 | 6.81 | 9.02 | 84 | 39.3 | 25.534
2006-06-09 | 21:00 | 12.06 | 6.78 | 9.04 | 84 | 39.2 | 25.454
2006-06-09 | 21:15 | 12 | 6.76 | 9.05 | 84 | 39.1 | 25.421
2006-06-09 | 21:30 | 11.94 | 6.76 | 9.07 | 84 | 39 | 25.335
2006-06-09 | 21:45 | 11.9 | 6.74 | 9.07 | 84 | 39 | 25.327
2006-06-09 | 22:00 | 11.86 | 6.74 | 9.08 | 84 | 38.9 | 25.301
2006-06-09 | 22:15 | 11.83 | 6.73 | 9.09 | 84 | 38.9 | 25.265
2006-06-09 | 22:30 | 11.8 | 6.71 | 9.09 | 84 | 38.8 | 25.198
2006-06-09 | 22:45 | 11.76 | 6.71 | 9.1 | 84 | 38.7 | 25.16
2006-06-09 | 23:00 | 11.71 | 6.7 | 9.12 | 84 | 38.7 | 25.133
2006-06-09 | 23:15 | 11.63 | 6.69 | 9.13 | 84 | 38.6 | 25.122
2006-06-09 | 23:30 | 11.53 | 6.69 | 9.16 | 84.1 | 38.6 | 25.111
2006-06-09 | 23:45 | 11.42 | 6.67 | 9.19 | 84.1 | 38.6 | 25.088
2006-06-10 | 0:00 | 11.3 | 6.67 | 9.22 | 84.2 | 38.6 | 25.084
2006-06-10 | 0:15 | 11.18 | 6.66 | 9.25 | 84.2 | 38.6 | 25.06
2006-06-10 | 0:30 | 11.05 | 6.65 | 9.28 | 84.3 | 38.5 | 25.045
2006-06-10 | 0:45 | 10.92 | 6.65 | 9.31 | 84.3 | 38.5 | 25.024
2006-06-10 | 1:01 | 10.79 | 6.64 | 9.35 | 84.4 | 38.5 | 25.004
2006-06-10 | 1:16 | 10.65 | 6.64 | 9.39 | 84.4 | 38.4 | 24.973
2006-06-10 | 1:31 | 10.52 | 6.63 | 9.42 | 84.4 | 38.4 | 24.981
2006-06-10 | 1:46 | 10.38 | 6.63 | 9.45 | 84.5 | 38.4 | 24.975
2006-06-10 | 2:01 | 10.25 | 6.62 | 9.49 | 84.6 | 38.4 | 24.984
2006-06-10 | 2:16 | 10.12 | 6.63 | 9.52 | 84.6 | 38.5 | 24.997
2006-06-10 | 2:31 | 10 | 6.62 | 9.55 | 84.6 | 38.4 | 24.989
2006-06-10 | 2:46 | 9.89 | 6.62 | 9.58 | 84.7 | 38.5 | 25.009
2006-06-10 | 3:01 | 9.8 | 6.62 | 9.61 | 84.7 | 38.5 | 25.034
2006-06-10 | 3:16 | 9.71 | 6.61 | 9.63 | 84.7 | 38.5 | 25.057
2006-06-10 | 3:31 | 9.64 | 6.62 | 9.65 | 84.8 | 38.6 | 25.1
2006-06-10 | 3:46 | 9.57 | 6.61 | 9.67 | 84.8 | 38.7 | 25.141
2006-06-10 | 4:01 | 9.5 | 6.61 | 9.69 | 84.8 | 38.8 | 25.197
2006-06-10 | 4:16 | 9.44 | 6.62 | 9.7 | 84.8 | 38.8 | 25.236
2006-06-10 | 4:31 | 9.39 | 6.61 | 9.72 | 84.8 | 38.9 | 25.294
2006-06-10 | 4:46 | 9.34 | 6.61 | 9.73 | 84.8 | 39 | 25.349
2006-06-10 | 5:01 | 9.28 | 6.61 | 9.74 | 84.9 | 39.1 | 25.394
2006-06-10 | 5:16 | 9.25 | 6.62 | 9.75 | 84.9 | 39.2 | 25.476
2006-06-10 | 5:31 | 9.22 | 6.61 | 9.76 | 84.9 | 39.3 | 25.535
2006-06-10 | 5:46 | 9.21 | 6.61 | 9.77 | 85 | 39.4 | 25.607
2006-06-10 | 6:01 | 9.2 | 6.62 | 9.78 | 85 | 39.5 | 25.68
2006-06-10 | 6:16 | 9.18 | 6.62 | 9.79 | 85 | 39.5 | 25.693
2006-06-10 | 6:31 | 9.18 | 6.63 | 9.79 | 85.1 | 39.7 | 25.776
2006-06-10 | 6:46 | 9.18 | 6.64 | 9.8 | 85.1 | 39.7 | 25.835
2006-06-10 | 7:01 | 9.18 | 6.65 | 9.81 | 85.2 | 39.8 | 25.881
2006-06-10 | 7:16 | 9.19 | 6.65 | 9.81 | 85.3 | 39.9 | 25.92
2006-06-10 | 7:31 | 9.21 | 6.67 | 9.82 | 85.4 | 40 | 25.973
2006-06-10 | 7:46 | 9.25 | 6.69 | 9.82 | 85.5 | 40.1 | 26.048
2006-06-10 | 8:01 | 9.29 | 6.71 | 9.82 | 85.6 | 40.2 | 26.099
2006-06-10 | 8:16 | 9.33 | 6.73 | 9.82 | 85.6 | 40.2 | 26.15
2006-06-10 | 8:31 | 9.37 | 6.74 | 9.82 | 85.7 | 40.3 | 26.19
2006-06-10 | 8:46 | 9.43 | 6.76 | 9.82 | 85.8 | 40.4 | 26.264
2006-06-10 | 9:01 | 9.51 | 6.77 | 9.81 | 85.9 | 40.5 | 26.331
2006-06-10 | 9:16 | 9.58 | 6.79 | 9.8 | 86 | 40.6 | 26.402
2006-06-10 | 9:31 | 9.65 | 6.8 | 9.79 | 86.1 | 40.7 | 26.46
2006-06-10 | 9:46 | 9.73 | 6.83 | 9.79 | 86.2 | 40.8 | 26.55
2006-06-10 | 10:01 | 9.78 | 6.84 | 9.77 | 86.1 | 40.9 | 26.589
2006-06-10 | 10:16 | 9.84 | 6.85 | 9.75 | 86.1 | 40.9 | 26.582
2006-06-10 | 10:31 | 10.03 | 6.88 | 9.73 | 86.2 | 40.9 | 26.589
2006-06-10 | 10:46 | 10.13 | 6.9 | 9.73 | 86.4 | 40.9 | 26.588
2006-06-10 | 11:01 | 10.28 | 6.91 | 9.7 | 86.5 | 40.9 | 26.597
2006-06-10 | 11:16 | 10.43 | 6.94 | 9.67 | 86.5 | 40.9 | 26.616
2006-06-10 | 11:31 | 10.52 | 6.95 | 9.64 | 86.5 | 40.9 | 26.575
2006-06-10 | 11:46 | 10.71 | 6.97 | 9.61 | 86.6 | 40.9 | 26.566
2006-06-10 | 12:01 | 10.87 | 7 | 9.58 | 86.6 | 40.9 | 26.557
2006-06-10 | 12:16 | 11.08 | 7.03 | 9.54 | 86.6 | 40.9 | 26.573
2006-06-10 | 12:31 | 11.27 | 7.05 | 9.49 | 86.6 | 40.9 | 26.604
2006-06-10 | 12:46 | 11.42 | 7.07 | 9.45 | 86.5 | 41 | 26.631
2006-06-10 | 13:00 | 11.64 | 7.1 | 9.4 | 86.5 | 41 | 26.633
2006-06-10 | 13:15 | 11.87 | 7.12 | 9.36 | 86.6 | 41 | 26.629
2006-06-10 | 13:30 | 12.07 | 7.15 | 9.31 | 86.6 | 41 | 26.618
2006-06-10 | 13:45 | 12.25 | 7.18 | 9.27 | 86.5 | 40.9 | 26.609
2006-06-10 | 14:00 | 12.36 | 7.19 | 9.23 | 86.3 | 41 | 26.635
2006-06-10 | 14:15 | 12.51 | 7.2 | 9.2 | 86.3 | 41 | 26.669
2006-06-10 | 14:30 | 12.66 | 7.21 | 9.15 | 86.2 | 41.1 | 26.703
2006-06-10 | 14:45 | 12.82 | 7.22 | 9.11 | 86.2 | 41.1 | 26.746
2006-06-10 | 15:00 | 12.95 | 7.23 | 9.07 | 86 | 41.2 | 26.794
2006-06-10 | 15:15 | 13.06 | 7.23 | 9.03 | 85.8 | 41.3 | 26.851
2006-06-10 | 15:30 | 13.21 | 7.25 | 9.01 | 85.9 | 41.3 | 26.866
2006-06-10 | 15:45 | 13.27 | 7.26 | 8.97 | 85.7 | 41.4 | 26.92
2006-06-10 | 16:00 | 13.2 | 7.23 | 8.96 | 85.5 | 41.4 | 26.926
2006-06-10 | 16:15 | 13.14 | 7.23 | 8.97 | 85.4 | 41.4 | 26.927
2006-06-10 | 16:30 | 13.03 | 7.2 | 8.98 | 85.3 | 41.5 | 26.943
2006-06-10 | 16:45 | 12.98 | 7.18 | 8.97 | 85.1 | 41.5 | 26.974
2006-06-10 | 17:00 | 13.08 | 7.18 | 8.97 | 85.2 | 41.5 | 27.002
2006-06-10 | 17:15 | 13.13 | 7.15 | 8.93 | 85 | 41.6 | 27.04
2006-06-10 | 17:30 | 13.15 | 7.13 | 8.89 | 84.7 | 41.6 | 27.072
2006-06-10 | 17:45 | 13.26 | 7.11 | 8.88 | 84.8 | 41.7 | 27.08
2006-06-10 | 18:00 | 13.36 | 7.11 | 8.86 | 84.8 | 41.7 | 27.091
2006-06-10 | 18:15 | 13.36 | 7.07 | 8.84 | 84.6 | 41.7 | 27.122
2006-06-10 | 18:30 | 13.33 | 7.05 | 8.83 | 84.4 | 41.7 | 27.12
2006-06-10 | 18:45 | 13.29 | 7.02 | 8.82 | 84.3 | 41.7 | 27.097
2006-06-10 | 19:00 | 13.18 | 6.99 | 8.84 | 84.2 | 41.7 | 27.111
2006-06-10 | 19:15 | 13.08 | 6.95 | 8.85 | 84.1 | 41.6 | 27.062
2006-06-10 | 19:30 | 12.99 | 6.93 | 8.86 | 84.1 | 41.5 | 26.975
2006-06-10 | 19:45 | 12.91 | 6.9 | 8.88 | 84.1 | 41.3 | 26.865
2006-06-10 | 20:00 | 12.83 | 6.87 | 8.89 | 84 | 41.1 | 26.729
2006-06-10 | 20:15 | 12.71 | 6.85 | 8.91 | 84 | 40.9 | 26.556
2006-06-10 | 20:30 | 12.57 | 6.82 | 8.94 | 84 | 40.6 | 26.398
2006-06-10 | 20:45 | 12.44 | 6.8 | 8.96 | 84 | 40.4 | 26.248
2006-06-10 | 21:00 | 12.32 | 6.78 | 8.99 | 84 | 40.2 | 26.107
2006-06-10 | 21:15 | 12.21 | 6.76 | 9.01 | 84.1 | 39.9 | 25.953
2006-06-10 | 21:30 | 12.1 | 6.74 | 9.04 | 84 | 39.7 | 25.797
2006-06-10 | 21:45 | 11.99 | 6.73 | 9.07 | 84.1 | 39.5 | 25.673
2006-06-10 | 22:00 | 11.87 | 6.71 | 9.09 | 84.1 | 39.3 | 25.54
2006-06-10 | 22:15 | 11.78 | 6.7 | 9.11 | 84.2 | 39.2 | 25.478
2006-06-10 | 22:30 | 11.68 | 6.69 | 9.14 | 84.2 | 39.1 | 25.387
2006-06-10 | 22:45 | 11.58 | 6.68 | 9.16 | 84.2 | 38.9 | 25.278
2006-06-10 | 23:00 | 11.48 | 6.68 | 9.19 | 84.2 | 38.8 | 25.194
2006-06-10 | 23:15 | 11.38 | 6.67 | 9.21 | 84.3 | 38.6 | 25.106
2006-06-10 | 23:30 | 11.27 | 6.66 | 9.24 | 84.3 | 38.5 | 25.017
2006-06-10 | 23:45 | 11.17 | 6.65 | 9.27 | 84.4 | 38.4 | 24.952
2006-06-11 | 0:00 | 11.06 | 6.64 | 9.3 | 84.4 | 38.3 | 24.894
2006-06-11 | 0:15 | 10.95 | 6.64 | 9.32 | 84.5 | 38.2 | 24.835
2006-06-11 | 0:30 | 10.85 | 6.64 | 9.35 | 84.5 | 38.1 | 24.794
2006-06-11 | 0:45 | 10.75 | 6.63 | 9.38 | 84.5 | 38.1 | 24.791
2006-06-11 | 1:01 | 10.65 | 6.63 | 9.4 | 84.6 | 38.1 | 24.796
2006-06-11 | 1:16 | 10.56 | 6.63 | 9.43 | 84.6 | 38.2 | 24.845
2006-06-11 | 1:31 | 10.47 | 6.62 | 9.45 | 84.6 | 38.3 | 24.868
2006-06-11 | 1:46 | 10.37 | 6.62 | 9.48 | 84.7 | 38.3 | 24.886
2006-06-11 | 2:01 | 10.28 | 6.62 | 9.5 | 84.7 | 38.3 | 24.921
2006-06-11 | 2:16 | 10.19 | 6.61 | 9.53 | 84.8 | 38.4 | 24.954
2006-06-11 | 2:31 | 10.1 | 6.62 | 9.55 | 84.8 | 38.4 | 24.98
2006-06-11 | 2:46 | 10 | 6.61 | 9.57 | 84.8 | 38.5 | 25.007
2006-06-11 | 3:01 | 9.91 | 6.61 | 9.6 | 84.9 | 38.5 | 25.04
2006-06-11 | 3:16 | 9.83 | 6.61 | 9.62 | 84.9 | 38.6 | 25.066
2006-06-11 | 3:31 | 9.75 | 6.61 | 9.64 | 84.9 | 38.6 | 25.095
2006-06-11 | 3:46 | 9.67 | 6.61 | 9.67 | 84.9 | 38.6 | 25.112
2006-06-11 | 4:01 | 9.6 | 6.6 | 9.68 | 85 | 38.7 | 25.155
2006-06-11 | 4:16 | 9.53 | 6.6 | 9.7 | 85 | 38.7 | 25.18
2006-06-11 | 4:31 | 9.48 | 6.59 | 9.71 | 85 | 38.8 | 25.248
2006-06-11 | 4:46 | 9.45 | 6.61 | 9.72 | 85 | 39 | 25.36
2006-06-11 | 5:01 | 9.43 | 6.61 | 9.73 | 85 | 39.2 | 25.474
2006-06-11 | 5:16 | 9.4 | 6.61 | 9.74 | 85.1 | 39.3 | 25.574
2006-06-11 | 5:31 | 9.35 | 6.62 | 9.76 | 85.1 | 39.4 | 25.628
2006-06-11 | 5:46 | 9.3 | 6.61 | 9.77 | 85.1 | 39.5 | 25.673
2006-06-11 | 6:01 | 9.26 | 6.61 | 9.79 | 85.2 | 39.5 | 25.704
2006-06-11 | 6:16 | 9.23 | 6.62 | 9.8 | 85.3 | 39.6 | 25.736
2006-06-11 | 6:31 | 9.21 | 6.63 | 9.82 | 85.4 | 39.7 | 25.781
2006-06-11 | 6:46 | 9.2 | 6.64 | 9.83 | 85.4 | 39.7 | 25.836
2006-06-11 | 7:01 | 9.22 | 6.65 | 9.84 | 85.5 | 39.9 | 25.919
2006-06-11 | 7:16 | 9.23 | 6.66 | 9.85 | 85.7 | 40 | 25.971
2006-06-11 | 7:31 | 9.27 | 6.68 | 9.85 | 85.8 | 40.1 | 26.054
2006-06-11 | 7:46 | 9.31 | 6.69 | 9.85 | 85.8 | 40.2 | 26.138
2006-06-11 | 8:01 | 9.34 | 6.71 | 9.86 | 86 | 40.2 | 26.158
2006-06-11 | 8:16 | 9.37 | 6.73 | 9.86 | 86.1 | 40.2 | 26.154
2006-06-11 | 8:31 | 9.43 | 6.77 | 9.86 | 86.2 | 40.3 | 26.189
2006-06-11 | 8:46 | 9.5 | 6.77 | 9.86 | 86.3 | 40.4 | 26.246
2006-06-11 | 9:01 | 9.56 | 6.79 | 9.85 | 86.3 | 40.5 | 26.354
2006-06-11 | 9:16 | 9.61 | 6.81 | 9.83 | 86.3 | 40.7 | 26.474
2006-06-11 | 9:31 | 9.7 | 6.83 | 9.83 | 86.5 | 40.7 | 26.465
2006-06-11 | 9:46 | 9.83 | 6.85 | 9.81 | 86.6 | 40.7 | 26.445
2006-06-11 | 10:01 | 9.89 | 6.85 | 9.79 | 86.6 | 40.6 | 26.386
2006-06-11 | 10:16 | 9.99 | 6.86 | 9.78 | 86.7 | 40.5 | 26.317
2006-06-11 | 10:31 | 10.04 | 6.89 | 9.77 | 86.6 | 40.5 | 26.304
2006-06-11 | 10:46 | 10.1 | 6.89 | 9.75 | 86.6 | 40.4 | 26.292
2006-06-11 | 11:01 | 10.19 | 6.92 | 9.73 | 86.6 | 40.5 | 26.326
2006-06-11 | 11:16 | 10.31 | 6.94 | 9.71 | 86.7 | 40.6 | 26.368
2006-06-11 | 11:31 | 10.45 | 6.96 | 9.69 | 86.7 | 40.6 | 26.388
2006-06-11 | 11:46 | 10.6 | 6.98 | 9.66 | 86.7 | 40.6 | 26.414
2006-06-11 | 12:01 | 10.75 | 7 | 9.63 | 86.9 | 40.7 | 26.425
2006-06-11 | 12:16 | 10.96 | 7.02 | 9.6 | 87 | 40.7 | 26.428
2006-06-11 | 12:31 | 11.14 | 7.04 | 9.55 | 86.9 | 40.7 | 26.472
2006-06-11 | 12:46 | 11.37 | 7.07 | 9.5 | 86.9 | 40.7 | 26.467
2006-06-11 | 13:00 | 11.53 | 7.1 | 9.45 | 86.7 | 40.8 | 26.504
2006-06-11 | 13:15 | 11.82 | 7.13 | 9.4 | 86.9 | 40.8 | 26.54
2006-06-11 | 13:30 | 11.99 | 7.15 | 9.35 | 86.7 | 40.9 | 26.572
2006-06-11 | 13:45 | 12.26 | 7.19 | 9.29 | 86.7 | 41 | 26.633
2006-06-11 | 14:00 | 12.51 | 7.21 | 9.24 | 86.7 | 41 | 26.676
2006-06-11 | 14:15 | 12.72 | 7.23 | 9.18 | 86.6 | 41.1 | 26.726
2006-06-11 | 14:30 | 12.95 | 7.25 | 9.13 | 86.5 | 41.2 | 26.765
2006-06-11 | 14:45 | 13.18 | 7.26 | 9.07 | 86.4 | 41.3 | 26.837
2006-06-11 | 15:00 | 13.37 | 7.29 | 9.02 | 86.3 | 41.4 | 26.886
2006-06-11 | 15:15 | 13.55 | 7.3 | 8.98 | 86.2 | 41.5 | 26.952
2006-06-11 | 15:30 | 13.71 | 7.31 | 8.93 | 86.1 | 41.6 | 27.012
2006-06-11 | 15:45 | 13.85 | 7.31 | 8.9 | 86 | 41.6 | 27.06
2006-06-11 | 16:00 | 13.93 | 7.31 | 8.87 | 85.9 | 41.7 | 27.104
2006-06-11 | 16:15 | 14.01 | 7.3 | 8.84 | 85.8 | 41.8 | 27.153
2006-06-11 | 16:30 | 14.09 | 7.29 | 8.81 | 85.6 | 41.8 | 27.191
2006-06-11 | 16:45 | 14.1 | 7.28 | 8.79 | 85.5 | 41.9 | 27.228
2006-06-11 | 17:00 | 14.07 | 7.24 | 8.76 | 85.1 | 41.9 | 27.248
2006-06-11 | 17:15 | 14.12 | 7.23 | 8.76 | 85.2 | 41.9 | 27.256
2006-06-11 | 17:30 | 14.03 | 7.18 | 8.72 | 84.7 | 42 | 27.291
2006-06-11 | 17:45 | 13.99 | 7.15 | 8.7 | 84.4 | 41.9 | 27.265
2006-06-11 | 18:00 | 13.95 | 7.11 | 8.7 | 84.4 | 41.9 | 27.249
2006-06-11 | 18:15 | 13.91 | 7.08 | 8.7 | 84.3 | 41.9 | 27.215
2006-06-11 | 18:30 | 13.85 | 7.04 | 8.71 | 84.2 | 41.8 | 27.152
2006-06-11 | 18:45 | 13.76 | 7 | 8.71 | 84.1 | 41.7 | 27.095
2006-06-11 | 19:00 | 13.69 | 6.96 | 8.72 | 84 | 41.6 | 27.057
2006-06-11 | 19:15 | 13.63 | 6.95 | 8.74 | 84.2 | 41.5 | 26.964
2006-06-11 | 19:30 | 13.57 | 6.92 | 8.75 | 84.1 | 41.4 | 26.927
2006-06-11 | 19:45 | 13.53 | 6.9 | 8.76 | 84.1 | 41.3 | 26.866
2006-06-11 | 20:00 | 13.49 | 6.89 | 8.77 | 84.1 | 41.3 | 26.816
2006-06-11 | 20:15 | 13.46 | 6.86 | 8.76 | 84 | 41.3 | 26.822
2006-06-11 | 20:30 | 13.43 | 6.85 | 8.77 | 84 | 41.3 | 26.838
2006-06-11 | 20:45 | 13.4 | 6.81 | 8.77 | 84 | 41.4 | 26.888
2006-06-11 | 21:00 | 13.35 | 6.81 | 8.78 | 83.9 | 41.4 | 26.934
2006-06-11 | 21:15 | 13.26 | 6.8 | 8.79 | 83.9 | 41.4 | 26.937
2006-06-11 | 21:30 | 13.15 | 6.77 | 8.82 | 84 | 41.3 | 26.852
2006-06-11 | 21:45 | 13.02 | 6.75 | 8.85 | 84 | 41.1 | 26.747
2006-06-11 | 22:00 | 12.91 | 6.72 | 8.88 | 84.1 | 41 | 26.679
2006-06-11 | 22:15 | 12.81 | 6.73 | 8.9 | 84.1 | 41 | 26.651
2006-06-11 | 22:30 | 12.71 | 6.71 | 8.92 | 84.2 | 40.9 | 26.607
2006-06-11 | 22:45 | 12.61 | 6.69 | 8.95 | 84.2 | 40.8 | 26.533
2006-06-11 | 23:00 | 12.51 | 6.68 | 8.97 | 84.3 | 40.7 | 26.48
2006-06-11 | 23:15 | 12.41 | 6.68 | 9 | 84.3 | 40.6 | 26.415
2006-06-11 | 23:30 | 12.28 | 6.67 | 9.03 | 84.3 | 40.5 | 26.308
2006-06-11 | 23:45 | 12.13 | 6.66 | 9.07 | 84.4 | 40.3 | 26.191
2006-06-12 | 0:00 | 11.98 | 6.65 | 9.1 | 84.4 | 40.2 | 26.098
2006-06-12 | 0:15 | 11.83 | 6.64 | 9.14 | 84.5 | 40 | 26.017
2006-06-12 | 0:30 | 11.7 | 6.64 | 9.18 | 84.6 | 39.9 | 25.957
2006-06-12 | 0:45 | 11.57 | 6.64 | 9.21 | 84.6 | 39.8 | 25.891
2006-06-12 | 1:01 | 11.44 | 6.63 | 9.24 | 84.7 | 39.7 | 25.817
2006-06-12 | 1:16 | 11.34 | 6.62 | 9.27 | 84.7 | 39.7 | 25.778
2006-06-12 | 1:31 | 11.23 | 6.62 | 9.3 | 84.7 | 39.5 | 25.701
2006-06-12 | 1:46 | 11.14 | 6.62 | 9.32 | 84.8 | 39.4 | 25.63
2006-06-12 | 2:01 | 11.04 | 6.62 | 9.35 | 84.8 | 39.3 | 25.56
2006-06-12 | 2:16 | 10.95 | 6.6 | 9.37 | 84.9 | 39.2 | 25.503
2006-06-12 | 2:31 | 10.85 | 6.61 | 9.39 | 84.9 | 39.2 | 25.448
2006-06-12 | 2:46 | 10.77 | 6.61 | 9.42 | 84.9 | 39.1 | 25.41
2006-06-12 | 3:01 | 10.67 | 6.6 | 9.44 | 85 | 39 | 25.351
2006-06-12 | 3:16 | 10.58 | 6.61 | 9.47 | 85 | 38.9 | 25.289
2006-06-12 | 3:31 | 10.49 | 6.61 | 9.49 | 85 | 38.9 | 25.274
2006-06-12 | 3:46 | 10.41 | 6.61 | 9.51 | 85 | 38.8 | 25.233
2006-06-12 | 4:01 | 10.32 | 6.61 | 9.53 | 85.1 | 38.8 | 25.204
2006-06-12 | 4:16 | 10.25 | 6.6 | 9.55 | 85.1 | 38.8 | 25.207
2006-06-12 | 4:31 | 10.17 | 6.59 | 9.57 | 85.1 | 38.7 | 25.186
2006-06-12 | 4:46 | 10.1 | 6.6 | 9.59 | 85.2 | 38.7 | 25.179
2006-06-12 | 5:01 | 10.03 | 6.59 | 9.61 | 85.2 | 38.7 | 25.15
2006-06-12 | 5:16 | 9.95 | 6.59 | 9.63 | 85.2 | 38.7 | 25.129
2006-06-12 | 5:31 | 9.88 | 6.59 | 9.65 | 85.3 | 38.6 | 25.091
2006-06-12 | 5:46 | 9.78 | 6.59 | 9.68 | 85.3 | 38.5 | 25.001
2006-06-12 | 6:01 | 9.69 | 6.58 | 9.71 | 85.4 | 38.4 | 24.934
2006-06-12 | 6:16 | 9.63 | 6.59 | 9.73 | 85.4 | 38.4 | 24.938
2006-06-12 | 6:31 | 9.58 | 6.59 | 9.75 | 85.5 | 38.4 | 24.942
2006-06-12 | 6:46 | 9.54 | 6.58 | 9.77 | 85.6 | 38.4 | 24.965
2006-06-12 | 7:01 | 9.52 | 6.6 | 9.78 | 85.7 | 38.5 | 25.008
2006-06-12 | 7:16 | 9.51 | 6.61 | 9.8 | 85.8 | 38.5 | 25.038
2006-06-12 | 7:31 | 9.51 | 6.63 | 9.81 | 85.9 | 38.6 | 25.068
2006-06-12 | 7:46 | 9.51 | 6.65 | 9.82 | 86 | 38.6 | 25.075
2006-06-12 | 8:01 | 9.49 | 6.65 | 9.82 | 86 | 38.6 | 25.11
2006-06-12 | 8:16 | 9.51 | 6.67 | 9.84 | 86.2 | 38.6 | 25.119
2006-06-12 | 8:31 | 9.52 | 6.68 | 9.84 | 86.2 | 38.7 | 25.162
2006-06-12 | 8:46 | 9.6 | 6.71 | 9.84 | 86.3 | 38.8 | 25.24
2006-06-12 | 9:01 | 9.64 | 6.73 | 9.83 | 86.4 | 38.9 | 25.291
2006-06-12 | 9:16 | 9.71 | 6.74 | 9.84 | 86.6 | 39 | 25.34
2006-06-12 | 9:31 | 9.78 | 6.76 | 9.83 | 86.6 | 39.1 | 25.388
2006-06-12 | 9:46 | 9.88 | 6.79 | 9.82 | 86.7 | 39.2 | 25.458
2006-06-12 | 10:01 | 9.98 | 6.79 | 9.8 | 86.8 | 39.2 | 25.468
2006-06-12 | 10:16 | 10.07 | 6.81 | 9.78 | 86.8 | 39.3 | 25.517
2006-06-12 | 10:31 | 10.2 | 6.83 | 9.76 | 86.9 | 39.3 | 25.546
2006-06-12 | 10:46 | 10.29 | 6.85 | 9.73 | 86.8 | 39.3 | 25.57
2006-06-12 | 11:01 | 10.38 | 6.86 | 9.72 | 86.9 | 39.4 | 25.59
2006-06-12 | 11:16 | 10.48 | 6.88 | 9.69 | 86.8 | 39.4 | 25.623
2006-06-12 | 11:31 | 10.57 | 6.89 | 9.67 | 86.8 | 39.4 | 25.616
2006-06-12 | 11:46 | 10.64 | 6.9 | 9.64 | 86.7 | 39.4 | 25.625
2006-06-12 | 12:01 | 10.75 | 6.91 | 9.62 | 86.8 | 39.5 | 25.662
2006-06-12 | 12:16 | 10.91 | 6.93 | 9.59 | 86.8 | 39.5 | 25.706
2006-06-12 | 12:31 | 11.05 | 6.93 | 9.54 | 86.6 | 39.6 | 25.765
2006-06-12 | 12:46 | 11.18 | 6.95 | 9.52 | 86.7 | 39.7 | 25.809
2006-06-12 | 13:00 | 11.34 | 6.96 | 9.48 | 86.6 | 39.8 | 25.863
2006-06-12 | 13:15 | 11.64 | 6.99 | 9.44 | 86.9 | 39.9 | 25.904
2006-06-12 | 13:30 | 11.84 | 7 | 9.38 | 86.7 | 40 | 25.973
2006-06-12 | 13:45 | 12.07 | 7.03 | 9.31 | 86.5 | 40.1 | 26.033
2006-06-12 | 14:00 | 12.11 | 6.99 | 9.25 | 86 | 40.1 | 26.047
2006-06-12 | 14:15 | 12.16 | 6.98 | 9.22 | 85.8 | 40.1 | 26.044
2006-06-12 | 14:30 | 12.17 | 6.96 | 9.18 | 85.5 | 40.2 | 26.102
2006-06-12 | 14:45 | 12.17 | 6.91 | 9.15 | 85.3 | 39.7 | 25.802
2006-06-12 | 15:00 | 12.17 | 6.89 | 9.15 | 85.2 | 39.9 | 25.925
2006-06-12 | 15:15 | 12.33 | 6.88 | 9.14 | 85.4 | 40.1 | 26.07
2006-06-12 | 15:30 | 12.36 | 6.88 | 9.15 | 85.6 | 39.9 | 25.936
2006-06-12 | 15:45 | 12.32 | 6.87 | 9.18 | 85.8 | 39.5 | 25.678
2006-06-12 | 16:00 | 12.37 | 6.88 | 9.2 | 86.1 | 39 | 25.337
2006-06-12 | 16:15 | 12.47 | 6.9 | 9.19 | 86.2 | 39.2 | 25.471
2006-06-12 | 16:30 | 12.63 | 6.91 | 9.16 | 86.2 | 39.6 | 25.763
2006-06-12 | 16:45 | 12.7 | 6.91 | 9.12 | 86 | 40.1 | 26.042
2006-06-12 | 17:00 | 12.66 | 6.9 | 9.11 | 85.8 | 40.1 | 26.075
2006-06-12 | 17:15 | 12.56 | 6.86 | 9.11 | 85.6 | 40 | 25.987
2006-06-12 | 17:30 | 12.5 | 6.85 | 9.12 | 85.5 | 40 | 25.976
2006-06-12 | 17:45 | 12.44 | 6.82 | 9.1 | 85.2 | 40.1 | 26.038
2006-06-12 | 18:00 | 12.4 | 6.77 | 9.09 | 85.1 | 39.6 | 25.769
2006-06-12 | 18:15 | 12.4 | 6.75 | 9.08 | 85 | 39.7 | 25.773
2006-06-12 | 18:30 | 12.37 | 6.73 | 9.08 | 85 | 39.5 | 25.649
2006-06-12 | 18:45 | 12.32 | 6.75 | 9.1 | 85.1 | 39.1 | 25.392
2006-06-12 | 19:00 | 12.3 | 6.75 | 9.11 | 85.1 | 38.9 | 25.254
2006-06-12 | 19:15 | 12.27 | 6.74 | 9.13 | 85.2 | 38.7 | 25.124
2006-06-12 | 19:30 | 12.27 | 6.74 | 9.12 | 85.2 | 38.8 | 25.201
2006-06-12 | 19:45 | 12.22 | 6.73 | 9.13 | 85.1 | 38.9 | 25.271
2006-06-12 | 20:00 | 12.15 | 6.71 | 9.15 | 85.2 | 39 | 25.327
2006-06-12 | 20:15 | 12.05 | 6.7 | 9.16 | 85.1 | 39 | 25.325
2006-06-12 | 20:30 | 11.95 | 6.69 | 9.18 | 85.1 | 39 | 25.319
2006-06-12 | 20:45 | 11.84 | 6.67 | 9.21 | 85.1 | 38.8 | 25.25
2006-06-12 | 21:00 | 11.74 | 6.67 | 9.23 | 85.1 | 38.7 | 25.157
2006-06-12 | 21:15 | 11.63 | 6.65 | 9.26 | 85.2 | 38.4 | 24.986
2006-06-12 | 21:30 | 11.52 | 6.65 | 9.29 | 85.2 | 38.3 | 24.886
2006-06-12 | 21:45 | 11.41 | 6.64 | 9.32 | 85.3 | 38.1 | 24.792
2006-06-12 | 22:00 | 11.32 | 6.63 | 9.35 | 85.4 | 38 | 24.712
2006-06-12 | 22:15 | 11.24 | 6.63 | 9.37 | 85.4 | 37.9 | 24.621
2006-06-12 | 22:30 | 11.16 | 6.63 | 9.39 | 85.5 | 37.7 | 24.489
2006-06-12 | 22:45 | 11.09 | 6.61 | 9.41 | 85.5 | 37.5 | 24.406
2006-06-12 | 23:00 | 10.99 | 6.62 | 9.43 | 85.5 | 38.1 | 24.757
2006-06-12 | 23:15 | 10.89 | 6.62 | 9.46 | 85.6 | 38 | 24.689
2006-06-12 | 23:30 | 10.8 | 6.6 | 9.49 | 85.6 | 38 | 24.69
2006-06-12 | 23:45 | 10.71 | 6.6 | 9.51 | 85.7 | 37.9 | 24.645
2006-06-13 | 0:00 | 10.63 | 6.6 | 9.54 | 85.7 | 37.8 | 24.581
2006-06-13 | 0:15 | 10.54 | 6.59 | 9.56 | 85.8 | 37.7 | 24.526
2006-06-13 | 0:30 | 10.45 | 6.59 | 9.58 | 85.8 | 37.6 | 24.444
2006-06-13 | 0:45 | 10.37 | 6.58 | 9.61 | 85.8 | 37.6 | 24.426
2006-06-13 | 1:01 | 10.29 | 6.58 | 9.63 | 85.9 | 37.6 | 24.412
2006-06-13 | 1:16 | 10.19 | 6.57 | 9.66 | 85.9 | 37.5 | 24.407
2006-06-13 | 1:31 | 10.11 | 6.57 | 9.68 | 86 | 37.5 | 24.401
2006-06-13 | 1:46 | 10.03 | 6.57 | 9.7 | 86 | 37.5 | 24.38
2006-06-13 | 2:01 | 9.94 | 6.57 | 9.72 | 86 | 37.5 | 24.35
2006-06-13 | 2:16 | 9.88 | 6.57 | 9.75 | 86.1 | 37.5 | 24.384
2006-06-13 | 2:31 | 9.8 | 6.56 | 9.76 | 86.1 | 37.5 | 24.382
2006-06-13 | 2:46 | 9.73 | 6.57 | 9.78 | 86.1 | 37.5 | 24.37
2006-06-13 | 3:01 | 9.67 | 6.55 | 9.8 | 86.1 | 37.7 | 24.485
2006-06-13 | 3:16 | 9.59 | 6.55 | 9.82 | 86.2 | 37.6 | 24.458
2006-06-13 | 3:31 | 9.53 | 6.56 | 9.84 | 86.2 | 37.5 | 24.403
2006-06-13 | 3:46 | 9.49 | 6.56 | 9.85 | 86.2 | 37.6 | 24.434
2006-06-13 | 4:01 | 9.43 | 6.57 | 9.87 | 86.3 | 37.6 | 24.414
2006-06-13 | 4:16 | 9.4 | 6.57 | 9.88 | 86.3 | 37.7 | 24.473
2006-06-13 | 4:31 | 9.36 | 6.57 | 9.89 | 86.3 | 37.7 | 24.506
2006-06-13 | 4:46 | 9.34 | 6.56 | 9.89 | 86.3 | 37.8 | 24.554
2006-06-13 | 5:01 | 9.32 | 6.56 | 9.9 | 86.3 | 37.8 | 24.571
2006-06-13 | 5:16 | 9.28 | 6.57 | 9.91 | 86.4 | 37.8 | 24.59
2006-06-13 | 5:31 | 9.25 | 6.57 | 9.93 | 86.4 | 37.8 | 24.586
2006-06-13 | 5:46 | 9.22 | 6.57 | 9.94 | 86.4 | 37.9 | 24.61
2006-06-13 | 6:01 | 9.18 | 6.57 | 9.94 | 86.4 | 37.9 | 24.63
2006-06-13 | 6:16 | 9.17 | 6.57 | 9.95 | 86.4 | 38 | 24.694
2006-06-13 | 6:31 | 9.15 | 6.58 | 9.97 | 86.5 | 37.9 | 24.645
2006-06-13 | 6:46 | 9.13 | 6.58 | 9.97 | 86.5 | 38.1 | 24.752
2006-06-13 | 7:01 | 9.12 | 6.59 | 9.97 | 86.5 | 38.2 | 24.808
2006-06-13 | 7:16 | 9.11 | 6.59 | 9.99 | 86.6 | 38.2 | 24.841
2006-06-13 | 7:31 | 9.11 | 6.6 | 9.99 | 86.7 | 38.3 | 24.912
2006-06-13 | 7:46 | 9.14 | 6.61 | 10 | 86.8 | 38.4 | 24.975
2006-06-13 | 8:01 | 9.17 | 6.63 | 10 | 86.9 | 38.5 | 25.041
2006-06-13 | 8:16 | 9.22 | 6.66 | 10 | 87 | 38.7 | 25.147
2006-06-13 | 8:31 | 9.28 | 6.69 | 10.01 | 87.2 | 38.8 | 25.203
2006-06-13 | 8:46 | 9.32 | 6.69 | 10 | 87.2 | 38.9 | 25.276
2006-06-13 | 9:01 | 9.39 | 6.72 | 9.99 | 87.3 | 39 | 25.371
2006-06-13 | 9:16 | 9.48 | 6.74 | 9.99 | 87.4 | 39.1 | 25.424
2006-06-13 | 9:31 | 9.55 | 6.76 | 9.97 | 87.4 | 39.2 | 25.504
2006-06-13 | 9:46 | 9.62 | 6.77 | 9.96 | 87.4 | 39.3 | 25.516
2006-06-13 | 10:01 | 9.66 | 6.77 | 9.92 | 87.2 | 39.4 | 25.585
2006-06-13 | 10:16 | 9.73 | 6.77 | 9.92 | 87.3 | 39.4 | 25.631
2006-06-13 | 10:31 | 9.86 | 6.8 | 9.91 | 87.5 | 39.6 | 25.717
2006-06-13 | 10:46 | 9.99 | 6.82 | 9.88 | 87.6 | 39.8 | 25.843
2006-06-13 | 11:01 | 10.08 | 6.83 | 9.87 | 87.6 | 39.9 | 25.947
2006-06-13 | 11:16 | 10.16 | 6.84 | 9.85 | 87.6 | 40.1 | 26.045
2006-06-13 | 11:31 | 10.29 | 6.86 | 9.81 | 87.5 | 40 | 26.019
2006-06-13 | 11:46 | 10.37 | 6.87 | 9.79 | 87.5 | 40.1 | 26.048
2006-06-13 | 12:01 | 10.42 | 6.88 | 9.79 | 87.6 | 40.2 | 26.116
2006-06-13 | 12:16 | 10.45 | 6.87 | 9.73 | 87.1 | 40.4 | 26.241
2006-06-13 | 12:31 | 10.48 | 6.86 | 9.72 | 87.1 | 40.4 | 26.251
2006-06-13 | 12:46 | 10.48 | 6.87 | 9.73 | 87.1 | 40.5 | 26.294
2006-06-13 | 13:00 | 10.53 | 6.87 | 9.73 | 87.3 | 40.4 | 26.291
2006-06-13 | 13:15 | 10.6 | 6.86 | 9.73 | 87.4 | 40.3 | 26.205
2006-06-13 | 13:30 | 10.65 | 6.88 | 9.75 | 87.7 | 39.6 | 25.756
2006-06-13 | 13:45 | 10.65 | 6.87 | 9.73 | 87.6 | 39.1 | 25.406
2006-06-13 | 14:00 | 10.57 | 6.86 | 9.73 | 87.4 | 38.7 | 25.153
2006-06-13 | 14:15 | 10.67 | 6.86 | 9.69 | 87.2 | 39.2 | 25.48
2006-06-13 | 14:30 | 10.78 | 6.88 | 9.67 | 87.2 | 39.8 | 25.887
2006-06-13 | 14:45 | 10.86 | 6.91 | 9.67 | 87.4 | 40 | 26.01
2006-06-13 | 15:00 | 10.91 | 6.92 | 9.65 | 87.3 | 40.1 | 26.059
2006-06-13 | 15:15 | 10.97 | 6.92 | 9.64 | 87.4 | 40.2 | 26.152
2006-06-13 | 15:30 | 10.96 | 6.91 | 9.62 | 87.2 | 40.3 | 26.177
2006-06-13 | 15:45 | 10.96 | 6.9 | 9.6 | 87 | 40.4 | 26.269
2006-06-13 | 16:00 | 11 | 6.88 | 9.57 | 86.8 | 40.6 | 26.384
2006-06-13 | 16:15 | 11.05 | 6.86 | 9.56 | 86.8 | 40.5 | 26.307
2006-06-13 | 16:30 | 11.12 | 6.85 | 9.52 | 86.5 | 40.8 | 26.522
2006-06-13 | 16:45 | 11.16 | 6.84 | 9.5 | 86.5 | 40.8 | 26.532
2006-06-13 | 17:00 | 11.18 | 6.83 | 9.48 | 86.4 | 41 | 26.679
2006-06-13 | 17:15 | 11.2 | 6.82 | 9.47 | 86.3 | 40.9 | 26.583
2006-06-13 | 17:30 | 11.22 | 6.81 | 9.49 | 86.5 | 40.3 | 26.184
2006-06-13 | 17:45 | 11.22 | 6.81 | 9.5 | 86.6 | 40.1 | 26.039
2006-06-13 | 18:00 | 11.21 | 6.79 | 9.5 | 86.6 | 39.8 | 25.889
2006-06-13 | 18:15 | 11.22 | 6.81 | 9.51 | 86.7 | 39.6 | 25.765
2006-06-13 | 18:30 | 11.24 | 6.82 | 9.52 | 86.8 | 39.5 | 25.657
2006-06-13 | 18:45 | 11.38 | 6.83 | 9.48 | 86.8 | 39.9 | 25.944
2006-06-13 | 19:00 | 11.47 | 6.84 | 9.46 | 86.7 | 40.3 | 26.179
2006-06-13 | 19:15 | 11.48 | 6.83 | 9.45 | 86.7 | 40.4 | 26.285
2006-06-13 | 19:30 | 11.47 | 6.83 | 9.45 | 86.6 | 40.5 | 26.311
2006-06-13 | 19:45 | 11.46 | 6.83 | 9.45 | 86.6 | 40.5 | 26.322
2006-06-13 | 20:00 | 11.45 | 6.82 | 9.44 | 86.5 | 40.4 | 26.289
2006-06-13 | 20:15 | 11.4 | 6.8 | 9.44 | 86.4 | 40.3 | 26.218
2006-06-13 | 20:30 | 11.34 | 6.78 | 9.45 | 86.3 | 40.2 | 26.142
2006-06-13 | 20:45 | 11.28 | 6.76 | 9.45 | 86.3 | 40.1 | 26.09
2006-06-13 | 21:00 | 11.21 | 6.75 | 9.46 | 86.2 | 40 | 26.01
2006-06-13 | 21:15 | 11.15 | 6.73 | 9.47 | 86.1 | 39.9 | 25.929
2006-06-13 | 21:30 | 11.12 | 6.73 | 9.46 | 86.1 | 39.9 | 25.904
2006-06-13 | 21:45 | 11.08 | 6.71 | 9.47 | 86 | 39.8 | 25.851
2006-06-13 | 22:00 | 11.05 | 6.71 | 9.47 | 86 | 39.7 | 25.787
2006-06-13 | 22:15 | 11.01 | 6.7 | 9.48 | 86 | 39.6 | 25.753
2006-06-13 | 22:30 | 10.96 | 6.69 | 9.49 | 86 | 39.5 | 25.679
2006-06-13 | 22:45 | 10.91 | 6.69 | 9.51 | 86.1 | 39.4 | 25.628
2006-06-13 | 23:00 | 10.83 | 6.67 | 9.53 | 86.1 | 39.3 | 25.567
2006-06-13 | 23:15 | 10.76 | 6.67 | 9.55 | 86.1 | 39.3 | 25.52
2006-06-13 | 23:30 | 10.68 | 6.67 | 9.57 | 86.1 | 39.2 | 25.49
2006-06-13 | 23:45 | 10.6 | 6.66 | 9.59 | 86.2 | 39.1 | 25.437
2006-06-14 | 0:00 | 10.53 | 6.65 | 9.61 | 86.2 | 39.1 | 25.428
2006-06-14 | 0:15 | 10.45 | 6.64 | 9.63 | 86.2 | 39.1 | 25.429
2006-06-14 | 0:30 | 10.37 | 6.64 | 9.65 | 86.2 | 39.1 | 25.423
2006-06-14 | 0:45 | 10.28 | 6.64 | 9.68 | 86.3 | 39.1 | 25.394
2006-06-14 | 1:01 | 10.17 | 6.64 | 9.7 | 86.3 | 39 | 25.349
2006-06-14 | 1:16 | 10.08 | 6.63 | 9.74 | 86.4 | 39 | 25.329
2006-06-14 | 1:31 | 9.99 | 6.64 | 9.76 | 86.4 | 39 | 25.341
2006-06-14 | 1:46 | 9.9 | 6.63 | 9.78 | 86.4 | 39 | 25.33
2006-06-14 | 2:01 | 9.82 | 6.63 | 9.8 | 86.5 | 39 | 25.352
2006-06-14 | 2:16 | 9.73 | 6.64 | 9.83 | 86.5 | 39 | 25.361
2006-06-14 | 2:31 | 9.63 | 6.64 | 9.86 | 86.6 | 39 | 25.35
2006-06-14 | 2:46 | 9.55 | 6.64 | 9.88 | 86.6 | 39 | 25.38
2006-06-14 | 3:01 | 9.47 | 6.63 | 9.9 | 86.7 | 39.1 | 25.395
2006-06-14 | 3:16 | 9.39 | 6.63 | 9.93 | 86.7 | 39.1 | 25.417
2006-06-14 | 3:31 | 9.33 | 6.63 | 9.94 | 86.7 | 39.2 | 25.453
2006-06-14 | 3:46 | 9.27 | 6.63 | 9.96 | 86.7 | 39.3 | 25.514
2006-06-14 | 4:01 | 9.22 | 6.63 | 9.97 | 86.7 | 39.3 | 25.561
2006-06-14 | 4:16 | 9.18 | 6.63 | 9.99 | 86.8 | 39.4 | 25.599
2006-06-14 | 4:31 | 9.13 | 6.64 | 10 | 86.8 | 39.4 | 25.631
2006-06-14 | 4:46 | 9.08 | 6.63 | 10.02 | 86.8 | 39.5 | 25.658
2006-06-14 | 5:01 | 9.04 | 6.63 | 10.03 | 86.9 | 39.5 | 25.682
2006-06-14 | 5:16 | 8.99 | 6.62 | 10.04 | 86.9 | 39.6 | 25.717
2006-06-14 | 5:31 | 8.95 | 6.63 | 10.06 | 86.9 | 39.6 | 25.75
2006-06-14 | 5:46 | 8.91 | 6.63 | 10.07 | 86.9 | 39.7 | 25.789
2006-06-14 | 6:01 | 8.88 | 6.63 | 10.08 | 87 | 39.7 | 25.821
2006-06-14 | 6:16 | 8.85 | 6.62 | 10.1 | 87 | 39.8 | 25.858
2006-06-14 | 6:31 | 8.82 | 6.64 | 10.11 | 87.1 | 39.9 | 25.904
2006-06-14 | 6:46 | 8.82 | 6.64 | 10.12 | 87.1 | 40 | 25.977
2006-06-14 | 7:01 | 8.81 | 6.65 | 10.13 | 87.2 | 40 | 26.032
2006-06-14 | 7:16 | 8.82 | 6.65 | 10.13 | 87.3 | 40.2 | 26.11
2006-06-14 | 7:31 | 8.82 | 6.67 | 10.14 | 87.4 | 40.2 | 26.145
2006-06-14 | 7:46 | 8.84 | 6.69 | 10.15 | 87.4 | 40.3 | 26.203
2006-06-14 | 8:01 | 8.85 | 6.7 | 10.15 | 87.5 | 40.4 | 26.254
2006-06-14 | 8:16 | 8.86 | 6.7 | 10.16 | 87.6 | 40.4 | 26.271
2006-06-14 | 8:31 | 8.9 | 6.71 | 10.15 | 87.6 | 40.5 | 26.338
2006-06-14 | 8:46 | 8.95 | 6.73 | 10.15 | 87.7 | 40.6 | 26.414
2006-06-14 | 9:01 | 9.01 | 6.74 | 10.15 | 87.8 | 40.7 | 26.459
2006-06-14 | 9:16 | 9.09 | 6.75 | 10.14 | 87.9 | 40.7 | 26.449
2006-06-14 | 9:31 | 9.17 | 6.78 | 10.14 | 88.1 | 40.7 | 26.439
2006-06-14 | 9:46 | 9.23 | 6.79 | 10.13 | 88.1 | 40.6 | 26.417
2006-06-14 | 10:01 | 9.33 | 6.81 | 10.12 | 88.3 | 40.5 | 26.317
2006-06-14 | 10:16 | 9.45 | 6.84 | 10.1 | 88.3 | 40.4 | 26.279
2006-06-14 | 10:31 | 9.59 | 6.86 | 10.08 | 88.4 | 40.5 | 26.332
2006-06-14 | 10:46 | 9.77 | 6.88 | 10.04 | 88.5 | 40.6 | 26.404
2006-06-14 | 11:01 | 9.96 | 6.9 | 10.01 | 88.6 | 40.8 | 26.506
2006-06-14 | 11:16 | 10.12 | 6.92 | 9.97 | 88.6 | 40.9 | 26.571
2006-06-14 | 11:31 | 10.3 | 6.95 | 9.94 | 88.6 | 41 | 26.644
2006-06-14 | 11:46 | 10.55 | 6.98 | 9.88 | 88.7 | 41.2 | 26.761
2006-06-14 | 12:01 | 10.78 | 7.01 | 9.83 | 88.7 | 41.4 | 26.916
2006-06-14 | 12:16 | 10.99 | 7.04 | 9.78 | 88.6 | 41.6 | 27.023
2006-06-14 | 12:31 | 11.13 | 7.06 | 9.74 | 88.6 | 41.7 | 27.098
2006-06-14 | 12:46 | 11.12 | 7.06 | 9.71 | 88.3 | 41.8 | 27.146
2006-06-14 | 13:00 | 11.16 | 7.04 | 9.67 | 88 | 41.9 | 27.218
2006-06-14 | 13:15 | 11.25 | 7.06 | 9.66 | 88.1 | 41.9 | 27.234
2006-06-14 | 13:30 | 11.42 | 7.09 | 9.62 | 88.1 | 41.9 | 27.263
2006-06-14 | 13:45 | 11.63 | 7.1 | 9.58 | 88.2 | 42 | 27.313
2006-06-14 | 14:00 | 11.69 | 7.11 | 9.54 | 87.9 | 42.1 | 27.362
2006-06-14 | 14:15 | 11.7 | 7.1 | 9.53 | 87.9 | 42.2 | 27.435
2006-06-14 | 14:30 | 11.8 | 7.12 | 9.52 | 87.9 | 42.2 | 27.424
2006-06-14 | 14:45 | 11.85 | 7.12 | 9.49 | 87.8 | 42.2 | 27.447
2006-06-14 | 15:00 | 11.85 | 7.1 | 9.47 | 87.6 | 42.3 | 27.494
2006-06-14 | 15:15 | 11.89 | 7.08 | 9.45 | 87.5 | 42.3 | 27.526
2006-06-14 | 15:30 | 12.02 | 7.09 | 9.44 | 87.7 | 42.4 | 27.585
2006-06-14 | 15:45 | 12.11 | 7.08 | 9.4 | 87.5 | 42.5 | 27.63
2006-06-14 | 16:00 | 12.17 | 7.07 | 9.37 | 87.2 | 42.6 | 27.676
2006-06-14 | 16:15 | 12.21 | 7.05 | 9.33 | 87 | 42.6 | 27.696
2006-06-14 | 16:30 | 12.26 | 7.04 | 9.31 | 86.9 | 42.6 | 27.715
2006-06-14 | 16:45 | 12.36 | 7.02 | 9.27 | 86.8 | 42.7 | 27.739
2006-06-14 | 17:00 | 12.41 | 7 | 9.26 | 86.7 | 42.7 | 27.78
2006-06-14 | 17:15 | 12.42 | 6.98 | 9.24 | 86.6 | 42.8 | 27.831
2006-06-14 | 17:30 | 12.44 | 6.96 | 9.23 | 86.5 | 42.9 | 27.872
2006-06-14 | 17:45 | 12.47 | 6.95 | 9.22 | 86.5 | 42.9 | 27.881
2006-06-14 | 18:00 | 12.49 | 6.93 | 9.2 | 86.4 | 42.9 | 27.887
2006-06-14 | 18:15 | 12.53 | 6.93 | 9.21 | 86.4 | 42.9 | 27.876
2006-06-14 | 18:30 | 12.54 | 6.93 | 9.21 | 86.5 | 42.9 | 27.855
2006-06-14 | 18:45 | 12.51 | 6.91 | 9.21 | 86.5 | 42.8 | 27.837
2006-06-14 | 19:00 | 12.46 | 6.9 | 9.22 | 86.4 | 42.7 | 27.772
2006-06-14 | 19:15 | 12.41 | 6.9 | 9.22 | 86.4 | 42.7 | 27.742
2006-06-14 | 19:30 | 12.34 | 6.88 | 9.23 | 86.4 | 42.6 | 27.698
2006-06-14 | 19:45 | 12.25 | 6.86 | 9.24 | 86.3 | 42.5 | 27.635
2006-06-14 | 20:00 | 12.15 | 6.84 | 9.27 | 86.3 | 42.4 | 27.537
2006-06-14 | 20:15 | 12.06 | 6.83 | 9.27 | 86.2 | 42.3 | 27.47
2006-06-14 | 20:30 | 11.97 | 6.81 | 9.28 | 86.1 | 42.1 | 27.374
2006-06-14 | 20:45 | 11.89 | 6.8 | 9.3 | 86.1 | 42 | 27.294
2006-06-14 | 21:00 | 11.8 | 6.78 | 9.32 | 86 | 41.8 | 27.191
2006-06-14 | 21:15 | 11.73 | 6.77 | 9.33 | 86 | 41.7 | 27.124
2006-06-14 | 21:30 | 11.66 | 6.75 | 9.34 | 86 | 41.6 | 27.058
2006-06-14 | 21:45 | 11.59 | 6.75 | 9.35 | 86 | 41.5 | 26.984
2006-06-14 | 22:00 | 11.52 | 6.74 | 9.37 | 86 | 41.4 | 26.899
2006-06-14 | 22:15 | 11.46 | 6.72 | 9.39 | 86 | 41.3 | 26.853
2006-06-14 | 22:30 | 11.38 | 6.71 | 9.41 | 86.1 | 41.2 | 26.789
2006-06-14 | 22:45 | 11.3 | 6.72 | 9.43 | 86.1 | 41.1 | 26.722
2006-06-14 | 23:00 | 11.2 | 6.7 | 9.45 | 86.1 | 41 | 26.675
2006-06-14 | 23:15 | 11.1 | 6.7 | 9.48 | 86.2 | 41 | 26.638
2006-06-14 | 23:30 | 10.99 | 6.7 | 9.51 | 86.3 | 40.9 | 26.605
2006-06-14 | 23:45 | 10.89 | 6.68 | 9.55 | 86.3 | 40.9 | 26.557
2006-06-15 | 0:00 | 10.8 | 6.69 | 9.57 | 86.4 | 40.8 | 26.527
2006-06-15 | 0:15 | 10.71 | 6.68 | 9.59 | 86.4 | 40.7 | 26.474
2006-06-15 | 0:30 | 10.65 | 6.67 | 9.61 | 86.5 | 40.7 | 26.471
2006-06-15 | 0:45 | 10.58 | 6.68 | 9.63 | 86.5 | 40.7 | 26.447
2006-06-15 | 1:01 | 10.51 | 6.68 | 9.65 | 86.5 | 40.7 | 26.436
2006-06-15 | 1:16 | 10.45 | 6.68 | 9.67 | 86.5 | 40.7 | 26.426
2006-06-15 | 1:31 | 10.39 | 6.67 | 9.68 | 86.6 | 40.7 | 26.426
2006-06-15 | 1:46 | 10.32 | 6.67 | 9.7 | 86.6 | 40.7 | 26.424
2006-06-15 | 2:01 | 10.27 | 6.66 | 9.72 | 86.7 | 40.7 | 26.453
2006-06-15 | 2:16 | 10.21 | 6.67 | 9.74 | 86.7 | 40.7 | 26.465
2006-06-15 | 2:31 | 10.15 | 6.67 | 9.75 | 86.7 | 40.7 | 26.486
2006-06-15 | 2:46 | 10.09 | 6.67 | 9.77 | 86.8 | 40.8 | 26.489
2006-06-15 | 3:01 | 10.03 | 6.67 | 9.78 | 86.7 | 40.8 | 26.525
2006-06-15 | 3:16 | 9.99 | 6.67 | 9.8 | 86.8 | 40.9 | 26.557
2006-06-15 | 3:31 | 9.94 | 6.67 | 9.81 | 86.8 | 40.9 | 26.589
2006-06-15 | 3:46 | 9.9 | 6.67 | 9.82 | 86.8 | 41 | 26.628
2006-06-15 | 4:01 | 9.84 | 6.66 | 9.84 | 86.8 | 41 | 26.635
2006-06-15 | 4:16 | 9.8 | 6.67 | 9.85 | 86.8 | 41 | 26.653
2006-06-15 | 4:31 | 9.77 | 6.65 | 9.86 | 86.9 | 41 | 26.681
2006-06-15 | 4:46 | 9.74 | 6.66 | 9.87 | 86.9 | 41.1 | 26.727
2006-06-15 | 5:01 | 9.72 | 6.66 | 9.87 | 86.9 | 41.2 | 26.756
2006-06-15 | 5:16 | 9.68 | 6.65 | 9.88 | 86.9 | 41.2 | 26.786
2006-06-15 | 5:31 | 9.66 | 6.65 | 9.89 | 86.9 | 41.3 | 26.874
2006-06-15 | 5:46 | 9.64 | 6.66 | 9.9 | 87 | 41.4 | 26.88
2006-06-15 | 6:01 | 9.62 | 6.67 | 9.91 | 87 | 41.4 | 26.904
2006-06-15 | 6:16 | 9.61 | 6.67 | 9.93 | 87.1 | 41.5 | 26.948
2006-06-15 | 6:31 | 9.6 | 6.67 | 9.93 | 87.2 | 41.5 | 26.993
2006-06-15 | 6:46 | 9.58 | 6.67 | 9.95 | 87.3 | 41.6 | 27.021
2006-06-15 | 7:01 | 9.57 | 6.69 | 9.96 | 87.3 | 41.6 | 27.057
2006-06-15 | 7:16 | 9.57 | 6.7 | 9.98 | 87.5 | 41.7 | 27.103
2006-06-15 | 7:31 | 9.59 | 6.71 | 9.99 | 87.6 | 41.8 | 27.152
2006-06-15 | 7:46 | 9.58 | 6.71 | 9.99 | 87.6 | 41.9 | 27.207
2006-06-15 | 8:01 | 9.6 | 6.74 | 9.99 | 87.7 | 41.9 | 27.263
2006-06-15 | 8:16 | 9.64 | 6.75 | 10 | 87.9 | 42 | 27.295
2006-06-15 | 8:31 | 9.65 | 6.75 | 10.01 | 87.9 | 42.1 | 27.352
2006-06-15 | 8:46 | 9.69 | 6.77 | 10.01 | 88 | 42.1 | 27.397
2006-06-15 | 9:01 | 9.75 | 6.78 | 10.01 | 88.1 | 42.2 | 27.452
2006-06-15 | 9:16 | 9.82 | 6.81 | 10.01 | 88.3 | 42.3 | 27.489
2006-06-15 | 9:31 | 9.88 | 6.82 | 10.01 | 88.5 | 42.3 | 27.523
2006-06-15 | 9:46 | 9.96 | 6.84 | 10.02 | 88.7 | 42.3 | 27.519
2006-06-15 | 10:01 | 10.04 | 6.86 | 10 | 88.6 | 42.4 | 27.59
2006-06-15 | 10:16 | 10.18 | 6.86 | 9.98 | 88.8 | 42.5 | 27.654
2006-06-15 | 10:31 | 10.24 | 6.88 | 9.95 | 88.7 | 42.6 | 27.698