## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 123.00
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River d/s whychus creek
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-12 |  | 2006-09-12 |  | A
Time | Computer set to atomic time |  | 12:05 |  | 12:04 |  | A
Temp (°C) | VWR NIST traceable #51063402 surface water |  | 18.9 |  | 18.4 |  | A
ph (SU) | pH 7 standard solution |  | 7 |  | 7 |  | A
DO (mg/L) | Winkler titration of surface water |  | 9.7 |  | 9.7 |  | A
SpC (uS/ml) | 148.0 µS/mL |  | 148 |  | 143 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 50 |  | 51 |  | B
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 478 |  | 479 |  | B
Field Equipment Grab Sampling Results
Field equipment ID | Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
YSI 556 MPS #05G1083AM | 2006-09-12 16:33 | 13.52 | 8.16 | 10.9 | 104.7 | 120 | 77.997
Hach 2100P #50828 | [7 empty] | 1
Post-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-15 |  | 2006-09-15 |  | A
Time | Computer set to atomic time |  | 9:16 |  | 9:15 |  | A
Temp (°C) | VWR NIST traceable #51063402 surface water |  | 14.1 |  | 13.7 |  | A
ph (SU) | pH 7 standard solution |  | 7 |  | 7.1 |  | A
DO (mg/L) | Winkler titration of surface water |  | 9.2 |  | 9.9 |  | B
SpC (uS/ml) | 148.0 µS/mL |  | 148 |  | 143 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 478 |  | 479 |  | B
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
|  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1