## WC 030.25
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 030.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek OWRD gauge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 x34 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-20 |  | 2006-09-20 |  | A
Time | Computer set to atomic time |  | 8:53 |  | 8:52 |  | A
Temp (°C) | VWR NIST traceable #51063402 surface water |  | 14.5 |  | 14 |  | A
ph (SU) | pH 7 standard solution |  | 7 |  | 7 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.9 |  | B
SpC (uS/ml) | 148.0 µS/mL |  | 148 |  | 138 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 50 |  | 51 |  | B
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 478 |  | 479 |  | B
Field Equipment Grab Sampling Results
Field equipment ID | Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
YSI 556 MPS #05G1083AM | 2006-09-21 11:36 | 5.96 |  | 12.24 | 98.2 | 24.8 | 16.093
Hach 2100P #50828 | [7 empty] | 22
Post-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS #05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-09-25 |  | 2006-09-25 |  | A
Time | Computer set to atomic time |  | 12:03 |  | 12:01 |  | A
Temp (°C) | VWR NIST traceable #51063402 surface water |  | 13.1 |  | 12.8 |  | A
ph (SU) | pH 7 standard solution |  | 7 |  | 7.2 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 11.5 |  | B
SpC (uS/ml) | 148.0 µS/mL |  | 148 |  | 157 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 50 |  | 51 |  | B
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 478 |  | 477 |  | B
Comments: |  |  | Grade | Temp | pH | DO | SpC | Turb
|  |  |  | C | SU | mg/L | µS/mL | NTU
|  |  | A | <1.50 | <=0.30 | <=0.3 | <=10% | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 0.50 | 0.31 to 1.0 | 10.1% to 15% | 1.1 to 2.0
|  |  | Fail | >2.01 | >0.51 | >1.1 | >15.1% | >2.1
Formulas:
Upper Deschutes Watershed Council = '[2]month year'!D7