## WC 009.00
Organization | UDWC |  | Station ID |  | WC 009.00
Program Name | Water Quality Monitoring Program |  | System |  | Whychus Creek
Contact Name | Lesely Jones |  | Description |  | Rimrock Ranch
Number | 541-382-6103 ex 34 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS 05G1083AM and Hach 2100P #50828
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2006-06-09 |  |  |  | F
Time | Computer set to atomic time |  | 17:00 |  |  |  | F
Temp (°C) | VWR NIST traceable surface water |  | 20.7 |  | 20.4 |  | A
DO (mg/L) | Winkler titration of surface water |  | 9.1 |  | 8.2 |  | B
SpC (uS/ml) | 148 µS/mL |  | 161 |  | 149 |  | A
ph (SU) | pH 7 standard solution |  | 7 |  | 7.5 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 478 |  | 478 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2006-06-10 | 13 | 6.2 | 9.9 | 93.9 | 32.1 | 21
9
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 16/12/06 |  |  |  | F
Time | Computer set to atomic time |  | 8:00 |  |  |  | F
Temp (°C) | VWR NIST traceable surface water |  | 21.4 |  | 21.2 |  | A
DO (mg/L) | Winkler titration of surface water |  | 8.7 |  | 8.6 |  | A
SpC (uS/ml) | 148 µS/mL |  | 154 |  | 153 |  | A
ph (SU) | pH 7 standard solution |  | 7.1 |  | 7.4 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 5 |  | 5 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 50 |  | 50 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 478 |  | 476 |  | B
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1
Formulas:
Audit value = '[2]master equipment audit'!D9:E9