## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AA against NIST, 556, & Titration
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-14 |  | 2008-07-14 |  | A
Time | Computer set to atomic time |  | 10:52 |  | 10:52 |  | A
Temp | VWR NIST traceable surface water |  | 7.7 |  | 7.81 |  | A
DO | Winkler titration of surface water |  | 11.6 |  | 11.92 |  | B
SpC | In Situ µS/mL |  | 24 |  | 24 |  | A
Ph | In Situ pH |  | 8.24 |  | 7.23 |  | Fail
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2008-07-14 16:30 | 22.05 | 8.3 | 9.22 | 105.6 | 56 | 39
Sonde | 2008-07-14 16:30 | 22.1 | 8.16 | 9.39 | 107.6 | 55 | 36
Grade |  | A | A | A |  | A
Audit 2 | 2008-07-18 13:15 | 22.09 | 8.17 | 9.51 | 109 | 68 | 48
Sonde | 2008-07-18 13:15 | 22.08 | 8.15 | 9.68 | 110.8 | 66 | 43
Grade |  | A | A | A |  | A
Audit 3 | 2008-07-22 13:00 | 22.43 | 8.24 | 9.5 | 109.6 | 68 | 47
Sonde | 2008-07-22 13:00 | 22.43 | 8.18 | 9.74 | 112.4 | 66 | 43
Grade |  | A | A | A |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AA against NIST, 556, & Titration
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-22 |  | 2008-07-22 |  | A
Time | Computer set to atomic time |  | 13:00 |  | 13:00 |  | A
Temp | VWR NIST traceable surface water |  | 22.3 |  | 22.43 |  | A
DO | Winkler titration of surface water |  | 9.4 |  | 9.74 |  | B
SpC | In Situ µS/mL |  | 68 |  | 66 |  | A
Ph | pH In Situ |  | 8.24 |  | 8.18 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Sonde File: 071408AA |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS & Sonde AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-14 |  | 2008-07-14 |  | A
Time | Computer set to atomic time |  | 10:52 |  | 10:52 |  | A
Temp | VWR NIST traceable surface water |  | 7.7 |  | 7.82 |  | A
DO | Winkler titration of surface water |  | 11.6 |  | 11.87 |  | A
SpC | 148 µS/mL |  | 24 |  | 24 |  | A
Ph | pH 7 standard solution |  | 7.81 |  | 4.95 |  | Fail
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-07-14 16:30 | 22.05 | 8.3 | 9.22 | 105.6 | 56 | 39
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-14 |  | 2008-07-14 |  | A
Time | Computer set to atomic time |  | 16:30 |  | 16:30 |  | A
Temp | In Situ |  | 22.1 |  | 22.05 |  | A
DO | In Situ |  | 9.39 |  | 9.22 |  | A
SpC | In Situ |  | 55 |  | 56 |  | A
Ph | In Situ |  | 8.16 |  | 8.3 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 File: 071408AA |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS & Sonde AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-17 |  | 2008-07-17 |  | A
Time | Computer set to atomic time |  | 14:15 |  | 14:15 |  | A
Temp | VWR NIST traceable surface water |  | 17.6 |  | 17.81 |  | A
DO | Winkler titration of surface water |  | 11.5 |  | 11.2 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 142 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.19 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-07-18 13:15 | 22.09 | 8.17 | 9.51 | 109 | 68 | 48
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-21 |  | 2008-07-21 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | VWR NIST traceable surface water |  | 18.1 |  | 18.34 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 9.45 |  | B
SpC | 148 µS/mL |  | 149.3 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.21 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS & Sonde AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-21 |  | 2008-07-21 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | VWR NIST traceable surface water |  | 18.1 |  | 18.34 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 9.45 |  | B
SpC | 148 µS/mL |  | 149.3 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.21 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-07-22 13:00 | 22.43 | 8.24 | 9.5 | 109.6 | 68 | 47
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-22 |  | 2008-07-22 |  | A
Time | Computer set to atomic time |  | 14:15 |  | 14:15 |  | A
Temp | VWR NIST traceable surface water |  | 18.7 |  | 18.9 |  | A
DO | Winkler titration of surface water |  | 9.1 |  | 9.25 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.16 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | TDS | DO% | Temp | SpCond | DO Conc | pH
M/D/Y |  | mg/L | % | ˚C | uS/cm | mg/L
2008-07-14 | 16:30 | 36 | 107.6 | 22.1 | 55 | 9.39 | 8.16
2008-07-14 | 16:45 | 36 | 106.6 | 22.07 | 55 | 9.3 | 8.23
2008-07-14 | 17:00 | 35 | 105.5 | 22.01 | 55 | 9.23 | 8.12
2008-07-14 | 17:15 | 35 | 104.3 | 21.98 | 55 | 9.13 | 8.07
2008-07-14 | 17:30 | 35 | 103.8 | 21.94 | 54 | 9.09 | 8.04
2008-07-14 | 17:45 | 35 | 103.1 | 21.86 | 54 | 9.04 | 8.02
2008-07-14 | 18:00 | 35 | 102.3 | 21.77 | 55 | 8.98 | 8
2008-07-14 | 18:15 | 35 | 100.5 | 21.62 | 54 | 8.85 | 7.94
2008-07-14 | 18:30 | 35 | 99.5 | 21.51 | 55 | 8.78 | 7.92
2008-07-14 | 18:45 | 35 | 99.2 | 21.43 | 55 | 8.77 | 7.9
2008-07-14 | 19:00 | 35 | 99.1 | 21.34 | 55 | 8.77 | 7.88
2008-07-14 | 19:15 | 36 | 98.7 | 21.23 | 55 | 8.76 | 7.86
2008-07-14 | 19:30 | 36 | 97.5 | 21.11 | 55 | 8.67 | 7.82
2008-07-14 | 19:45 | 36 | 96.4 | 21.01 | 55 | 8.59 | 7.77
2008-07-14 | 20:00 | 36 | 95.6 | 20.91 | 55 | 8.54 | 7.7
2008-07-14 | 20:15 | 36 | 94.7 | 20.8 | 55 | 8.48 | 7.65
2008-07-14 | 20:30 | 36 | 93.3 | 20.69 | 55 | 8.37 | 7.56
2008-07-14 | 20:45 | 36 | 91.9 | 20.58 | 56 | 8.26 | 7.46
2008-07-14 | 21:00 | 36 | 91.1 | 20.45 | 56 | 8.21 | 7.38
2008-07-14 | 21:15 | 36 | 90.4 | 20.35 | 56 | 8.16 | 7.3
2008-07-14 | 21:30 | 37 | 89.4 | 20.27 | 56 | 8.09 | 7.24
2008-07-14 | 21:45 | 37 | 88.4 | 20.17 | 56 | 8.01 | 7.19
2008-07-14 | 22:00 | 37 | 88.1 | 20.09 | 56 | 7.99 | 7.15
2008-07-14 | 22:15 | 37 | 87.6 | 20.03 | 57 | 7.96 | 7.12
2008-07-14 | 22:30 | 37 | 87 | 19.98 | 57 | 7.91 | 7.08
2008-07-14 | 22:45 | 37 | 86.4 | 19.95 | 57 | 7.86 | 7.06
2008-07-14 | 23:00 | 37 | 85.9 | 19.91 | 57 | 7.83 | 7.03
2008-07-14 | 23:15 | 37 | 85.7 | 19.89 | 57 | 7.81 | 7.01
2008-07-14 | 23:30 | 37 | 85.3 | 19.86 | 57 | 7.78 | 6.99
2008-07-14 | 23:45 | 37 | 84.8 | 19.84 | 58 | 7.74 | 6.96
2008-07-15 | 0:00 | 38 | 84.6 | 19.8 | 58 | 7.72 | 6.94
2008-07-15 | 0:15 | 38 | 84.2 | 19.78 | 58 | 7.69 | 6.94
2008-07-15 | 0:30 | 38 | 84.2 | 19.75 | 58 | 7.69 | 6.94
2008-07-15 | 0:45 | 38 | 84 | 19.73 | 58 | 7.68 | 6.91
2008-07-15 | 1:01 | 38 | 83.6 | 19.71 | 58 | 7.65 | 6.91
2008-07-15 | 1:16 | 38 | 83.4 | 19.69 | 58 | 7.63 | 6.9
2008-07-15 | 1:31 | 38 | 83.4 | 19.68 | 59 | 7.63 | 6.89
2008-07-15 | 1:46 | 38 | 83.4 | 19.66 | 59 | 7.64 | 6.89
2008-07-15 | 2:01 | 38 | 83.4 | 19.64 | 59 | 7.64 | 6.88
2008-07-15 | 2:16 | 38 | 83.2 | 19.63 | 59 | 7.63 | 6.88
2008-07-15 | 2:31 | 38 | 83 | 19.61 | 59 | 7.6 | 6.88
2008-07-15 | 2:46 | 39 | 82.8 | 19.59 | 59 | 7.59 | 6.85
2008-07-15 | 3:01 | 39 | 82.8 | 19.57 | 60 | 7.59 | 6.87
2008-07-15 | 3:16 | 39 | 82.8 | 19.56 | 60 | 7.6 | 6.88
2008-07-15 | 3:31 | 39 | 82.7 | 19.54 | 60 | 7.59 | 6.86
2008-07-15 | 3:46 | 39 | 82.6 | 19.52 | 60 | 7.58 | 6.86
2008-07-15 | 4:01 | 39 | 82.5 | 19.49 | 60 | 7.58 | 6.85
2008-07-15 | 4:16 | 39 | 82.6 | 19.47 | 60 | 7.59 | 6.84
2008-07-15 | 4:31 | 39 | 82.8 | 19.44 | 60 | 7.61 | 6.84
2008-07-15 | 4:46 | 39 | 82.6 | 19.41 | 61 | 7.6 | 6.86
2008-07-15 | 5:01 | 39 | 82.5 | 19.38 | 61 | 7.6 | 6.85
2008-07-15 | 5:16 | 40 | 82.5 | 19.35 | 61 | 7.6 | 6.84
2008-07-15 | 5:31 | 40 | 82.4 | 19.31 | 61 | 7.6 | 6.83
2008-07-15 | 5:46 | 40 | 82.7 | 19.27 | 61 | 7.63 | 6.82
2008-07-15 | 6:01 | 40 | 83 | 19.23 | 61 | 7.66 | 6.84
2008-07-15 | 6:16 | 40 | 83.2 | 19.2 | 61 | 7.68 | 6.83
2008-07-15 | 6:31 | 40 | 83.8 | 19.16 | 61 | 7.75 | 6.84
2008-07-15 | 6:46 | 40 | 84.3 | 19.13 | 62 | 7.8 | 6.85
2008-07-15 | 7:01 | 40 | 85 | 19.11 | 62 | 7.87 | 6.87
2008-07-15 | 7:16 | 40 | 86 | 19.09 | 62 | 7.96 | 6.89
2008-07-15 | 7:31 | 40 | 86.8 | 19.09 | 62 | 8.04 | 6.92
2008-07-15 | 7:46 | 40 | 87.7 | 19.1 | 62 | 8.12 | 6.93
2008-07-15 | 8:01 | 40 | 89.1 | 19.13 | 62 | 8.24 | 6.97
2008-07-15 | 8:16 | 40 | 90.9 | 19.19 | 62 | 8.4 | 7.01
2008-07-15 | 8:31 | 40 | 91.6 | 19.21 | 62 | 8.46 | 7.05
2008-07-15 | 8:46 | 40 | 93.1 | 19.31 | 62 | 8.58 | 7.09
2008-07-15 | 9:01 | 40 | 94.4 | 19.4 | 62 | 8.68 | 7.15
2008-07-15 | 9:16 | 40 | 95.8 | 19.53 | 62 | 8.79 | 7.2
2008-07-15 | 9:31 | 40 | 97.1 | 19.66 | 62 | 8.89 | 7.25
2008-07-15 | 9:46 | 40 | 98.8 | 19.79 | 62 | 9.02 | 7.31
2008-07-15 | 10:01 | 40 | 99.6 | 19.89 | 62 | 9.08 | 7.36
2008-07-15 | 10:16 | 40 | 100.7 | 19.99 | 62 | 9.16 | 7.39
2008-07-15 | 10:31 | 40 | 101.5 | 20.07 | 62 | 9.22 | 7.43
2008-07-15 | 10:46 | 40 | 101.5 | 20.11 | 61 | 9.21 | 7.46
2008-07-15 | 11:01 | 40 | 103.3 | 20.33 | 61 | 9.33 | 7.55
2008-07-15 | 11:16 | 40 | 104.3 | 20.52 | 61 | 9.39 | 7.64
2008-07-15 | 11:31 | 40 | 105.2 | 20.65 | 61 | 9.45 | 7.7
2008-07-15 | 11:46 | 39 | 106 | 20.76 | 61 | 9.49 | 7.75
2008-07-15 | 12:01 | 39 | 106.6 | 20.87 | 60 | 9.53 | 7.81
2008-07-15 | 12:16 | 39 | 106.9 | 21.02 | 60 | 9.52 | 7.88
2008-07-15 | 12:31 | 39 | 107.7 | 21.17 | 60 | 9.57 | 7.95
2008-07-15 | 12:46 | 39 | 107.7 | 21.21 | 59 | 9.56 | 7.96
2008-07-15 | 13:00 | 38 | 108.1 | 21.32 | 59 | 9.57 | 8.02
2008-07-15 | 13:15 | 38 | 108 | 21.46 | 59 | 9.55 | 8.07
2008-07-15 | 13:30 | 38 | 108.4 | 21.58 | 58 | 9.56 | 8.09
2008-07-15 | 13:45 | 38 | 108 | 21.63 | 58 | 9.51 | 8.12
2008-07-15 | 14:00 | 37 | 107.8 | 21.61 | 58 | 9.49 | 8.1
2008-07-15 | 14:15 | 37 | 108.1 | 21.75 | 57 | 9.5 | 8.16
2008-07-15 | 14:30 | 37 | 108.3 | 21.91 | 57 | 9.48 | 8.2
2008-07-15 | 14:45 | 37 | 107.8 | 22.03 | 56 | 9.42 | 8.22
2008-07-15 | 15:00 | 36 | 108 | 22.11 | 56 | 9.42 | 8.21
2008-07-15 | 15:15 | 36 | 107.8 | 22.2 | 55 | 9.39 | 8.25
2008-07-15 | 15:30 | 36 | 107.6 | 22.29 | 55 | 9.35 | 8.26
2008-07-15 | 15:45 | 36 | 107 | 22.3 | 55 | 9.3 | 8.24
2008-07-15 | 16:00 | 35 | 106.2 | 22.34 | 55 | 9.23 | 8.24
2008-07-15 | 16:15 | 35 | 105.7 | 22.38 | 54 | 9.18 | 8.25
2008-07-15 | 16:30 | 35 | 104.5 | 22.31 | 54 | 9.08 | 8.18
2008-07-15 | 16:45 | 35 | 103.3 | 22.18 | 54 | 9 | 8.15
2008-07-15 | 17:00 | 35 | 103.3 | 22.21 | 54 | 9 | 8.2
2008-07-15 | 17:15 | 35 | 103.6 | 22.23 | 54 | 9.02 | 8.2
2008-07-15 | 17:30 | 35 | 102.9 | 22.14 | 53 | 8.97 | 8.14
2008-07-15 | 17:45 | 35 | 101.3 | 22 | 53 | 8.86 | 8.1
2008-07-15 | 18:00 | 35 | 100.1 | 21.9 | 53 | 8.77 | 8.09
2008-07-15 | 18:15 | 35 | 99.8 | 21.86 | 53 | 8.75 | 8.09
2008-07-15 | 18:30 | 35 | 99.3 | 21.81 | 53 | 8.72 | 8.09
2008-07-15 | 18:45 | 35 | 98.7 | 21.67 | 53 | 8.69 | 8.01
2008-07-15 | 19:00 | 35 | 97.9 | 21.53 | 53 | 8.64 | 7.97
2008-07-15 | 19:15 | 35 | 96.9 | 21.43 | 53 | 8.57 | 7.94
2008-07-15 | 19:30 | 35 | 96.4 | 21.32 | 53 | 8.54 | 7.9
2008-07-15 | 19:45 | 35 | 95.3 | 21.19 | 53 | 8.46 | 7.84
2008-07-15 | 20:00 | 35 | 94.6 | 21.06 | 54 | 8.42 | 7.77
2008-07-15 | 20:15 | 35 | 93.8 | 20.92 | 54 | 8.37 | 7.68
2008-07-15 | 20:30 | 35 | 92.6 | 20.8 | 54 | 8.29 | 7.61
2008-07-15 | 20:45 | 35 | 91.3 | 20.69 | 54 | 8.19 | 7.51
2008-07-15 | 21:00 | 35 | 90.7 | 20.6 | 54 | 8.15 | 7.43
2008-07-15 | 21:15 | 35 | 89.6 | 20.51 | 54 | 8.07 | 7.35
2008-07-15 | 21:30 | 35 | 89 | 20.43 | 54 | 8.02 | 7.29
2008-07-15 | 21:45 | 36 | 88.5 | 20.34 | 55 | 7.99 | 7.23
2008-07-15 | 22:00 | 36 | 87.7 | 20.26 | 55 | 7.93 | 7.18
2008-07-15 | 22:15 | 36 | 86.9 | 20.18 | 55 | 7.87 | 7.13
2008-07-15 | 22:30 | 36 | 86.4 | 20.11 | 55 | 7.84 | 7.1
2008-07-15 | 22:45 | 36 | 86.2 | 20.06 | 55 | 7.83 | 7.06
2008-07-15 | 23:00 | 36 | 85.6 | 20 | 56 | 7.79 | 7.04
2008-07-15 | 23:15 | 36 | 85.2 | 19.93 | 56 | 7.76 | 7.01
2008-07-15 | 23:30 | 36 | 85 | 19.87 | 56 | 7.74 | 6.99
2008-07-15 | 23:45 | 37 | 84.8 | 19.82 | 56 | 7.74 | 6.97
2008-07-16 | 0:00 | 37 | 84.4 | 19.76 | 56 | 7.71 | 6.96
2008-07-16 | 0:15 | 37 | 84.2 | 19.72 | 56 | 7.69 | 6.94
2008-07-16 | 0:30 | 37 | 84 | 19.69 | 57 | 7.68 | 6.93
2008-07-16 | 0:45 | 37 | 83.8 | 19.65 | 57 | 7.67 | 6.92
2008-07-16 | 1:01 | 37 | 84 | 19.62 | 57 | 7.69 | 6.92
2008-07-16 | 1:16 | 37 | 83.7 | 19.59 | 57 | 7.67 | 6.9
2008-07-16 | 1:31 | 37 | 83.4 | 19.56 | 57 | 7.65 | 6.9
2008-07-16 | 1:46 | 37 | 83.4 | 19.54 | 57 | 7.65 | 6.89
2008-07-16 | 2:01 | 37 | 83.5 | 19.53 | 58 | 7.66 | 6.88
2008-07-16 | 2:16 | 38 | 83.6 | 19.51 | 58 | 7.68 | 6.9
2008-07-16 | 2:31 | 38 | 83.6 | 19.49 | 58 | 7.68 | 6.89
2008-07-16 | 2:46 | 38 | 83.4 | 19.46 | 58 | 7.67 | 6.87
2008-07-16 | 3:01 | 38 | 83.1 | 19.42 | 58 | 7.65 | 6.88
2008-07-16 | 3:16 | 38 | 83.3 | 19.41 | 58 | 7.66 | 6.87
2008-07-16 | 3:31 | 38 | 83.1 | 19.39 | 59 | 7.65 | 6.86
2008-07-16 | 3:46 | 38 | 82.9 | 19.35 | 59 | 7.63 | 6.87
2008-07-16 | 4:01 | 38 | 83 | 19.33 | 59 | 7.65 | 6.86
2008-07-16 | 4:16 | 38 | 82.9 | 19.31 | 59 | 7.64 | 6.85
2008-07-16 | 4:31 | 39 | 82.9 | 19.28 | 59 | 7.64 | 6.85
2008-07-16 | 4:46 | 39 | 82.9 | 19.25 | 59 | 7.65 | 6.87
2008-07-16 | 5:01 | 39 | 83 | 19.22 | 60 | 7.66 | 6.85
2008-07-16 | 5:16 | 39 | 83 | 19.18 | 60 | 7.67 | 6.86
2008-07-16 | 5:31 | 39 | 83.2 | 19.15 | 60 | 7.69 | 6.85
2008-07-16 | 5:46 | 39 | 83.2 | 19.11 | 60 | 7.7 | 6.85
2008-07-16 | 6:01 | 39 | 83.3 | 19.08 | 60 | 7.71 | 6.87
2008-07-16 | 6:16 | 39 | 83.8 | 19.06 | 60 | 7.77 | 6.86
2008-07-16 | 6:31 | 39 | 84.4 | 19.03 | 61 | 7.82 | 6.87
2008-07-16 | 6:46 | 39 | 84.9 | 19.02 | 61 | 7.87 | 6.87
2008-07-16 | 7:01 | 39 | 85.6 | 19.02 | 61 | 7.94 | 6.91
2008-07-16 | 7:16 | 40 | 87 | 19.05 | 61 | 8.06 | 6.94
2008-07-16 | 7:31 | 40 | 87.9 | 19.08 | 61 | 8.14 | 6.96
2008-07-16 | 7:46 | 40 | 89.4 | 19.12 | 61 | 8.27 | 6.99
2008-07-16 | 8:01 | 40 | 90.6 | 19.19 | 61 | 8.37 | 7.03
2008-07-16 | 8:16 | 40 | 91.7 | 19.27 | 62 | 8.46 | 7.08
2008-07-16 | 8:31 | 40 | 93.1 | 19.36 | 62 | 8.57 | 7.1
2008-07-16 | 8:46 | 40 | 94 | 19.46 | 62 | 8.64 | 7.16
2008-07-16 | 9:01 | 40 | 95.3 | 19.57 | 62 | 8.74 | 7.2
2008-07-16 | 9:16 | 40 | 96.6 | 19.71 | 62 | 8.84 | 7.24
2008-07-16 | 9:31 | 40 | 97.8 | 19.85 | 62 | 8.92 | 7.29
2008-07-16 | 9:46 | 41 | 98.9 | 20 | 62 | 8.99 | 7.34
2008-07-16 | 10:01 | 41 | 99.8 | 20.16 | 62 | 9.04 | 7.38
2008-07-16 | 10:16 | 41 | 101.3 | 20.34 | 63 | 9.15 | 7.44
2008-07-16 | 10:31 | 41 | 102.6 | 20.51 | 63 | 9.23 | 7.49
2008-07-16 | 10:46 | 41 | 103.8 | 20.69 | 63 | 9.31 | 7.56
2008-07-16 | 11:01 | 41 | 104.4 | 20.87 | 63 | 9.33 | 7.62
2008-07-16 | 11:16 | 41 | 105.2 | 21.04 | 63 | 9.37 | 7.68
2008-07-16 | 11:31 | 41 | 106.2 | 21.23 | 63 | 9.43 | 7.72
2008-07-16 | 11:46 | 41 | 106.7 | 21.41 | 63 | 9.43 | 7.78
2008-07-16 | 12:01 | 41 | 107.1 | 21.59 | 63 | 9.44 | 7.84
2008-07-16 | 12:16 | 41 | 107.9 | 21.76 | 63 | 9.48 | 7.89
2008-07-16 | 12:31 | 41 | 108.4 | 21.91 | 63 | 9.49 | 7.94
2008-07-16 | 12:46 | 41 | 108.8 | 22.07 | 63 | 9.5 | 7.99
2008-07-16 | 13:00 | 41 | 108.8 | 22.26 | 63 | 9.47 | 8.04
2008-07-16 | 13:15 | 41 | 109.3 | 22.39 | 62 | 9.48 | 8.07
2008-07-16 | 13:30 | 40 | 109.4 | 22.54 | 62 | 9.47 | 8.09
2008-07-16 | 13:45 | 40 | 109.5 | 22.71 | 62 | 9.44 | 8.16
2008-07-16 | 14:00 | 40 | 108.2 | 22.67 | 61 | 9.34 | 8.08
2008-07-16 | 14:15 | 40 | 107.9 | 22.74 | 61 | 9.3 | 8.14
2008-07-16 | 14:30 | 40 | 108.5 | 22.92 | 61 | 9.32 | 8.19
2008-07-16 | 14:45 | 39 | 107.5 | 22.9 | 60 | 9.24 | 8.13
2008-07-16 | 15:00 | 39 | 106.3 | 22.74 | 60 | 9.16 | 8.09
2008-07-16 | 15:15 | 39 | 106.5 | 22.91 | 60 | 9.15 | 8.19
2008-07-16 | 15:30 | 38 | 107.1 | 23.07 | 59 | 9.18 | 8.24
2008-07-16 | 15:45 | 38 | 106.8 | 23.16 | 59 | 9.13 | 8.24
2008-07-16 | 16:00 | 38 | 106.5 | 23.16 | 58 | 9.11 | 8.2
2008-07-16 | 16:15 | 38 | 106.4 | 23.22 | 58 | 9.09 | 8.24
2008-07-16 | 16:30 | 38 | 106 | 23.19 | 58 | 9.06 | 8.19
2008-07-16 | 16:45 | 37 | 105.5 | 23.19 | 58 | 9.02 | 8.19
2008-07-16 | 17:00 | 37 | 104.2 | 23.02 | 57 | 8.93 | 8.13
2008-07-16 | 17:15 | 37 | 103.9 | 22.95 | 57 | 8.92 | 8.13
2008-07-16 | 17:30 | 37 | 103.6 | 22.89 | 57 | 8.91 | 8.13
2008-07-16 | 17:45 | 37 | 103 | 22.79 | 57 | 8.87 | 8.11
2008-07-16 | 18:00 | 37 | 102.3 | 22.73 | 56 | 8.82 | 8.11
2008-07-16 | 18:15 | 37 | 102 | 22.66 | 56 | 8.81 | 8.11
2008-07-16 | 18:30 | 37 | 100.9 | 22.59 | 56 | 8.73 | 8.09
2008-07-16 | 18:45 | 37 | 100.4 | 22.48 | 56 | 8.7 | 8.06
2008-07-16 | 19:00 | 37 | 99.3 | 22.36 | 56 | 8.62 | 8.03
2008-07-16 | 19:15 | 37 | 98.4 | 22.23 | 56 | 8.57 | 8
2008-07-16 | 19:30 | 36 | 96.9 | 22.11 | 56 | 8.46 | 7.94
2008-07-16 | 19:45 | 37 | 96.1 | 21.97 | 56 | 8.41 | 7.9
2008-07-16 | 20:00 | 37 | 95.2 | 21.83 | 56 | 8.35 | 7.84
2008-07-16 | 20:15 | 37 | 94 | 21.71 | 56 | 8.27 | 7.79
2008-07-16 | 20:30 | 37 | 92.9 | 21.6 | 56 | 8.19 | 7.71
2008-07-16 | 20:45 | 37 | 92 | 21.5 | 56 | 8.13 | 7.65
2008-07-16 | 21:00 | 37 | 91.2 | 21.39 | 56 | 8.07 | 7.55
2008-07-16 | 21:15 | 37 | 90.4 | 21.26 | 57 | 8.01 | 7.46
2008-07-16 | 21:30 | 37 | 89.3 | 21.14 | 57 | 7.94 | 7.38
2008-07-16 | 21:45 | 37 | 88.4 | 21.02 | 57 | 7.88 | 7.3
2008-07-16 | 22:00 | 37 | 88 | 20.91 | 57 | 7.86 | 7.24
2008-07-16 | 22:15 | 37 | 87.2 | 20.8 | 57 | 7.81 | 7.19
2008-07-16 | 22:30 | 37 | 86.9 | 20.7 | 57 | 7.79 | 7.15
2008-07-16 | 22:45 | 37 | 86.4 | 20.59 | 58 | 7.76 | 7.1
2008-07-16 | 23:00 | 38 | 86 | 20.5 | 58 | 7.74 | 7.06
2008-07-16 | 23:15 | 38 | 85.5 | 20.41 | 58 | 7.71 | 7.05
2008-07-16 | 23:30 | 38 | 85.2 | 20.32 | 58 | 7.7 | 7.02
2008-07-16 | 23:45 | 38 | 84.8 | 20.24 | 58 | 7.67 | 7
2008-07-17 | 0:00 | 38 | 84.7 | 20.17 | 58 | 7.67 | 6.98
2008-07-17 | 0:15 | 38 | 84.2 | 20.1 | 58 | 7.64 | 6.97
2008-07-17 | 0:30 | 38 | 84 | 20.04 | 58 | 7.63 | 6.94
2008-07-17 | 0:45 | 38 | 84.1 | 19.98 | 58 | 7.65 | 6.93
2008-07-17 | 1:01 | 38 | 84 | 19.93 | 59 | 7.65 | 6.93
2008-07-17 | 1:16 | 38 | 83.7 | 19.87 | 59 | 7.63 | 6.93
2008-07-17 | 1:31 | 38 | 83.4 | 19.81 | 59 | 7.61 | 6.91
2008-07-17 | 1:46 | 38 | 83.3 | 19.77 | 59 | 7.61 | 6.91
2008-07-17 | 2:01 | 38 | 83.3 | 19.72 | 59 | 7.62 | 6.9
2008-07-17 | 2:16 | 38 | 83.3 | 19.67 | 59 | 7.62 | 6.91
2008-07-17 | 2:31 | 38 | 83.3 | 19.61 | 59 | 7.63 | 6.89
2008-07-17 | 2:46 | 39 | 83.4 | 19.57 | 59 | 7.65 | 6.89
2008-07-17 | 3:01 | 39 | 83.1 | 19.53 | 59 | 7.63 | 6.9
2008-07-17 | 3:16 | 39 | 83.1 | 19.5 | 60 | 7.63 | 6.88
2008-07-17 | 3:31 | 39 | 83 | 19.46 | 60 | 7.63 | 6.89
2008-07-17 | 3:46 | 39 | 82.9 | 19.42 | 60 | 7.62 | 6.87
2008-07-17 | 4:01 | 39 | 83.2 | 19.38 | 60 | 7.65 | 6.87
2008-07-17 | 4:16 | 39 | 83.4 | 19.35 | 60 | 7.68 | 6.88
2008-07-17 | 4:31 | 39 | 83.3 | 19.3 | 60 | 7.68 | 6.88
2008-07-17 | 4:46 | 39 | 83 | 19.26 | 60 | 7.66 | 6.86
2008-07-17 | 5:01 | 39 | 83 | 19.23 | 61 | 7.66 | 6.87
2008-07-17 | 5:16 | 39 | 83 | 19.19 | 61 | 7.67 | 6.87
2008-07-17 | 5:31 | 40 | 83.1 | 19.14 | 61 | 7.69 | 6.87
2008-07-17 | 5:46 | 40 | 83.2 | 19.1 | 61 | 7.7 | 6.86
2008-07-17 | 6:01 | 40 | 83.3 | 19.07 | 61 | 7.72 | 6.87
2008-07-17 | 6:16 | 40 | 83.8 | 19.02 | 61 | 7.77 | 6.89
2008-07-17 | 6:31 | 40 | 84.5 | 18.99 | 61 | 7.84 | 6.89
2008-07-17 | 6:46 | 40 | 85.1 | 18.98 | 62 | 7.89 | 6.9
2008-07-17 | 7:01 | 40 | 85.9 | 18.97 | 62 | 7.97 | 6.92
2008-07-17 | 7:16 | 40 | 87.1 | 18.98 | 62 | 8.08 | 6.94
2008-07-17 | 7:31 | 40 | 88.4 | 18.99 | 62 | 8.2 | 6.97
2008-07-17 | 7:46 | 40 | 89.6 | 19.03 | 62 | 8.3 | 7.03
2008-07-17 | 8:01 | 40 | 90.3 | 19.08 | 62 | 8.37 | 7.02
2008-07-17 | 8:16 | 40 | 91.6 | 19.16 | 62 | 8.47 | 7.08
2008-07-17 | 8:31 | 41 | 92.9 | 19.23 | 62 | 8.58 | 7.09
2008-07-17 | 8:46 | 41 | 93.8 | 19.34 | 62 | 8.65 | 7.13
2008-07-17 | 9:01 | 41 | 95.1 | 19.45 | 63 | 8.74 | 7.2
2008-07-17 | 9:16 | 41 | 96.2 | 19.6 | 63 | 8.82 | 7.23
2008-07-17 | 9:31 | 41 | 97.7 | 19.74 | 63 | 8.93 | 7.29
2008-07-17 | 9:46 | 41 | 98.8 | 19.9 | 63 | 9 | 7.34
2008-07-17 | 10:01 | 41 | 100 | 20.07 | 63 | 9.08 | 7.4
2008-07-17 | 10:16 | 41 | 101.7 | 20.24 | 63 | 9.2 | 7.44
2008-07-17 | 10:31 | 41 | 102.6 | 20.41 | 63 | 9.25 | 7.51
2008-07-17 | 10:46 | 41 | 103.9 | 20.6 | 64 | 9.34 | 7.56
2008-07-17 | 11:01 | 41 | 104.9 | 20.81 | 64 | 9.38 | 7.61
2008-07-17 | 11:16 | 41 | 106.1 | 21.02 | 64 | 9.46 | 7.67
2008-07-17 | 11:31 | 42 | 107 | 21.23 | 64 | 9.5 | 7.74
2008-07-17 | 11:46 | 42 | 107.8 | 21.43 | 64 | 9.53 | 7.8
2008-07-17 | 12:01 | 42 | 108.1 | 21.64 | 64 | 9.52 | 7.86
2008-07-17 | 12:16 | 42 | 108.5 | 21.83 | 64 | 9.52 | 7.91
2008-07-17 | 12:31 | 42 | 108.9 | 22.02 | 64 | 9.52 | 7.95
2008-07-17 | 12:46 | 42 | 109.2 | 22.21 | 64 | 9.51 | 8.02
2008-07-17 | 13:00 | 42 | 109.4 | 22.37 | 64 | 9.5 | 8.05
2008-07-17 | 13:15 | 42 | 109.8 | 22.53 | 64 | 9.5 | 8.11
2008-07-17 | 13:30 | 42 | 110 | 22.65 | 65 | 9.5 | 8.11
2008-07-17 | 13:45 | 42 | 110.1 | 22.77 | 64 | 9.48 | 8.14
2008-07-17 | 14:00 | 42 | 110.1 | 22.86 | 64 | 9.47 | 8.2
2008-07-17 | 14:15 | 42 | 109.6 | 22.95 | 64 | 9.41 | 8.22
2008-07-17 | 14:30 | 42 | 109.8 | 23 | 64 | 9.42 | 8.23
2008-07-17 | 14:45 | 41 | 109.7 | 23.05 | 64 | 9.4 | 8.23
2008-07-17 | 15:00 | 41 | 109.4 | 23.11 | 64 | 9.36 | 8.26
2008-07-17 | 15:15 | 41 | 108.8 | 23.15 | 64 | 9.31 | 8.28
2008-07-17 | 15:30 | 41 | 108.7 | 23.19 | 63 | 9.29 | 8.27
2008-07-17 | 15:45 | 41 | 108.5 | 23.2 | 63 | 9.27 | 8.29
2008-07-17 | 16:00 | 41 | 107.7 | 23.22 | 63 | 9.2 | 8.28
2008-07-17 | 16:15 | 40 | 107.4 | 23.22 | 62 | 9.18 | 8.31
2008-07-17 | 16:30 | 40 | 107.1 | 23.18 | 62 | 9.15 | 8.27
2008-07-17 | 16:45 | 40 | 106.5 | 23.14 | 62 | 9.11 | 8.27
2008-07-17 | 17:00 | 40 | 105.4 | 23.08 | 61 | 9.03 | 8.26
2008-07-17 | 17:15 | 39 | 105.1 | 23.02 | 61 | 9.01 | 8.25
2008-07-17 | 17:30 | 39 | 104 | 22.94 | 60 | 8.93 | 8.25
2008-07-17 | 17:45 | 39 | 103.4 | 22.85 | 60 | 8.89 | 8.25
2008-07-17 | 18:00 | 39 | 103 | 22.74 | 60 | 8.88 | 8.21
2008-07-17 | 18:15 | 39 | 101.8 | 22.61 | 59 | 8.8 | 8.19
2008-07-17 | 18:30 | 38 | 101.4 | 22.48 | 59 | 8.79 | 8.17
2008-07-17 | 18:45 | 38 | 100.3 | 22.35 | 59 | 8.71 | 8.15
2008-07-17 | 19:00 | 38 | 99.2 | 22.2 | 59 | 8.64 | 8.12
2008-07-17 | 19:15 | 38 | 98.5 | 22.03 | 58 | 8.61 | 8.08
2008-07-17 | 19:30 | 38 | 97.5 | 21.87 | 58 | 8.55 | 8.04
2008-07-17 | 19:45 | 38 | 96.2 | 21.72 | 58 | 8.46 | 7.98
2008-07-17 | 20:00 | 38 | 95.4 | 21.56 | 58 | 8.42 | 7.92
2008-07-17 | 20:15 | 38 | 94.5 | 21.39 | 58 | 8.36 | 7.85
2008-07-17 | 20:30 | 38 | 93.4 | 21.2 | 58 | 8.3 | 7.78
2008-07-17 | 20:45 | 38 | 92.6 | 21.05 | 58 | 8.25 | 7.7
2008-07-17 | 21:00 | 38 | 91.6 | 20.88 | 58 | 8.19 | 7.61
2008-07-17 | 21:15 | 38 | 90.6 | 20.73 | 58 | 8.12 | 7.51
2008-07-17 | 21:30 | 38 | 90 | 20.57 | 58 | 8.09 | 7.42
2008-07-17 | 21:45 | 38 | 89.2 | 20.42 | 58 | 8.04 | 7.35
2008-07-17 | 22:00 | 38 | 88.7 | 20.29 | 58 | 8.02 | 7.28
2008-07-17 | 22:15 | 38 | 87.9 | 20.15 | 58 | 7.97 | 7.23
2008-07-17 | 22:30 | 38 | 87.7 | 20.02 | 58 | 7.97 | 7.17
2008-07-17 | 22:45 | 38 | 87.3 | 19.89 | 59 | 7.96 | 7.14
2008-07-17 | 23:00 | 38 | 87 | 19.77 | 59 | 7.95 | 7.1
2008-07-17 | 23:15 | 38 | 86.3 | 19.66 | 59 | 7.9 | 7.07
2008-07-17 | 23:30 | 38 | 86.1 | 19.55 | 59 | 7.9 | 7.05
2008-07-17 | 23:45 | 38 | 85.9 | 19.45 | 59 | 7.9 | 7.02
2008-07-18 | 0:00 | 38 | 85.7 | 19.36 | 59 | 7.89 | 7.01
2008-07-18 | 0:15 | 39 | 85.4 | 19.27 | 59 | 7.87 | 6.99
2008-07-18 | 0:30 | 39 | 85 | 19.18 | 59 | 7.86 | 6.98
2008-07-18 | 0:45 | 39 | 84.9 | 19.1 | 59 | 7.86 | 6.97
2008-07-18 | 1:01 | 39 | 84.7 | 19.02 | 60 | 7.85 | 6.96
2008-07-18 | 1:16 | 39 | 84.6 | 18.95 | 60 | 7.85 | 6.95
2008-07-18 | 1:31 | 39 | 84.8 | 18.89 | 60 | 7.89 | 6.94
2008-07-18 | 1:46 | 39 | 84.5 | 18.82 | 60 | 7.87 | 6.94
2008-07-18 | 2:01 | 39 | 84.7 | 18.77 | 60 | 7.89 | 6.93
2008-07-18 | 2:16 | 39 | 84.4 | 18.71 | 60 | 7.87 | 6.93
2008-07-18 | 2:31 | 39 | 84.4 | 18.66 | 60 | 7.88 | 6.92
2008-07-18 | 2:46 | 39 | 84.5 | 18.61 | 60 | 7.9 | 6.92
2008-07-18 | 3:01 | 39 | 84.5 | 18.57 | 60 | 7.9 | 6.93
2008-07-18 | 3:16 | 39 | 84.2 | 18.53 | 60 | 7.89 | 6.91
2008-07-18 | 3:31 | 39 | 84.2 | 18.49 | 60 | 7.89 | 6.91
2008-07-18 | 3:46 | 39 | 84.2 | 18.46 | 61 | 7.9 | 6.92
2008-07-18 | 4:01 | 39 | 84.2 | 18.44 | 61 | 7.9 | 6.92
2008-07-18 | 4:16 | 40 | 84.4 | 18.4 | 61 | 7.92 | 6.91
2008-07-18 | 4:31 | 40 | 84.3 | 18.38 | 61 | 7.92 | 6.93
2008-07-18 | 4:46 | 40 | 84.1 | 18.35 | 61 | 7.91 | 6.91
2008-07-18 | 5:01 | 40 | 84.3 | 18.3 | 61 | 7.93 | 6.91
2008-07-18 | 5:16 | 40 | 84.4 | 18.27 | 61 | 7.95 | 6.92
2008-07-18 | 5:31 | 40 | 84.4 | 18.24 | 61 | 7.95 | 6.9
2008-07-18 | 5:46 | 40 | 84.4 | 18.2 | 62 | 7.95 | 6.91
2008-07-18 | 6:01 | 40 | 84.8 | 18.16 | 62 | 8 | 6.93
2008-07-18 | 6:16 | 40 | 84.9 | 18.14 | 62 | 8.01 | 6.92
2008-07-18 | 6:31 | 40 | 85.6 | 18.13 | 62 | 8.08 | 6.92
2008-07-18 | 6:46 | 40 | 86.5 | 18.13 | 62 | 8.16 | 6.95
2008-07-18 | 7:01 | 40 | 87.1 | 18.14 | 62 | 8.22 | 6.94
2008-07-18 | 7:16 | 41 | 88 | 18.15 | 62 | 8.31 | 6.99
2008-07-18 | 7:31 | 41 | 89.4 | 18.2 | 62 | 8.43 | 7.02
2008-07-18 | 7:46 | 41 | 90.3 | 18.26 | 63 | 8.5 | 7.06
2008-07-18 | 8:01 | 41 | 91.3 | 18.33 | 63 | 8.59 | 7.09
2008-07-18 | 8:16 | 41 | 92.4 | 18.41 | 63 | 8.68 | 7.12
2008-07-18 | 8:31 | 41 | 93.5 | 18.5 | 63 | 8.76 | 7.17
2008-07-18 | 8:46 | 41 | 95 | 18.63 | 63 | 8.88 | 7.2
2008-07-18 | 9:01 | 41 | 96.4 | 18.77 | 63 | 8.98 | 7.24
2008-07-18 | 9:16 | 41 | 97.2 | 18.93 | 64 | 9.03 | 7.28
2008-07-18 | 9:31 | 41 | 98.6 | 19.08 | 64 | 9.13 | 7.33
2008-07-18 | 9:46 | 41 | 99.5 | 19.24 | 64 | 9.18 | 7.38
2008-07-18 | 10:01 | 41 | 101.1 | 19.43 | 64 | 9.3 | 7.44
2008-07-18 | 10:16 | 42 | 102.6 | 19.61 | 64 | 9.41 | 7.49
2008-07-18 | 10:31 | 42 | 103.8 | 19.8 | 64 | 9.48 | 7.56
2008-07-18 | 10:46 | 42 | 104.9 | 20.01 | 64 | 9.54 | 7.61
2008-07-18 | 11:01 | 42 | 106 | 20.22 | 64 | 9.6 | 7.67
2008-07-18 | 11:16 | 42 | 106.8 | 20.42 | 65 | 9.63 | 7.73
2008-07-18 | 11:31 | 42 | 107.7 | 20.63 | 65 | 9.67 | 7.8
2008-07-18 | 11:46 | 42 | 108.5 | 20.83 | 65 | 9.71 | 7.85
2008-07-18 | 12:01 | 42 | 108.8 | 21.05 | 65 | 9.69 | 7.91
2008-07-18 | 12:16 | 43 | 109.2 | 21.27 | 65 | 9.69 | 7.97
2008-07-18 | 12:31 | 43 | 109.7 | 21.47 | 65 | 9.69 | 8.01
2008-07-18 | 12:46 | 43 | 110.3 | 21.69 | 66 | 9.7 | 8.08
2008-07-18 | 13:00 | 43 | 110.3 | 21.89 | 66 | 9.67 | 8.12
2008-07-18 | 13:15 | 43 | 110.8 | 22.08 | 66 | 9.68 | 8.15
2008-07-18 | 13:30 | 43 | 111 | 22.27 | 66 | 9.66 | 8.18
2008-07-18 | 13:45 | 43 | 111.2 | 22.44 | 66 | 9.64 | 8.22
2008-07-18 | 14:00 | 43 | 110.9 | 22.58 | 66 | 9.59 | 8.26
2008-07-18 | 14:15 | 43 | 111 | 22.73 | 67 | 9.57 | 8.28
2008-07-18 | 14:30 | 43 | 110.7 | 22.82 | 67 | 9.52 | 8.3
2008-07-18 | 14:45 | 43 | 110.8 | 22.88 | 67 | 9.52 | 8.31
2008-07-18 | 15:00 | 43 | 110.7 | 22.94 | 67 | 9.51 | 8.34
2008-07-18 | 15:15 | 43 | 110.1 | 23.01 | 67 | 9.44 | 8.35
2008-07-18 | 15:30 | 43 | 109.6 | 23.06 | 67 | 9.39 | 8.36
2008-07-18 | 15:45 | 43 | 109.1 | 23.13 | 67 | 9.34 | 8.36
2008-07-18 | 16:00 | 43 | 109.1 | 23.12 | 67 | 9.34 | 8.38
2008-07-18 | 16:15 | 43 | 108.6 | 23.12 | 67 | 9.29 | 8.37
2008-07-18 | 16:30 | 43 | 108.2 | 23.1 | 66 | 9.26 | 8.38
2008-07-18 | 16:45 | 43 | 107.4 | 23.09 | 66 | 9.19 | 8.38
2008-07-18 | 17:00 | 43 | 106.6 | 23.05 | 66 | 9.14 | 8.38
2008-07-18 | 17:15 | 43 | 106.1 | 23 | 66 | 9.1 | 8.37
2008-07-18 | 17:30 | 43 | 105.5 | 22.93 | 66 | 9.06 | 8.35
2008-07-18 | 17:45 | 43 | 104.9 | 22.84 | 66 | 9.02 | 8.34
2008-07-18 | 18:00 | 43 | 104 | 22.75 | 66 | 8.96 | 8.34
2008-07-18 | 18:15 | 42 | 103.2 | 22.65 | 65 | 8.91 | 8.32
2008-07-18 | 18:30 | 42 | 102.4 | 22.51 | 65 | 8.86 | 8.3
2008-07-18 | 18:45 | 42 | 101.3 | 22.38 | 65 | 8.79 | 8.28
2008-07-18 | 19:00 | 42 | 100.4 | 22.25 | 64 | 8.74 | 8.26
2008-07-18 | 19:15 | 41 | 99.2 | 22.11 | 64 | 8.66 | 8.21
2008-07-18 | 19:30 | 41 | 97.7 | 21.96 | 63 | 8.55 | 8.18
2008-07-18 | 19:45 | 41 | 96.8 | 21.81 | 63 | 8.49 | 8.13
2008-07-18 | 20:00 | 41 | 95.5 | 21.68 | 63 | 8.4 | 8.08
2008-07-18 | 20:15 | 41 | 94.6 | 21.53 | 62 | 8.35 | 8.02
2008-07-18 | 20:30 | 40 | 93.9 | 21.37 | 62 | 8.31 | 7.96
2008-07-18 | 20:45 | 40 | 92.8 | 21.22 | 62 | 8.23 | 7.87
2008-07-18 | 21:00 | 40 | 91.4 | 21.06 | 62 | 8.14 | 7.78
2008-07-18 | 21:15 | 40 | 90.9 | 20.89 | 61 | 8.12 | 7.68
2008-07-18 | 21:30 | 40 | 90.2 | 20.73 | 61 | 8.08 | 7.58
2008-07-18 | 21:45 | 40 | 89.5 | 20.58 | 61 | 8.04 | 7.49
2008-07-18 | 22:00 | 40 | 88.7 | 20.43 | 61 | 7.99 | 7.41
2008-07-18 | 22:15 | 40 | 88.1 | 20.29 | 61 | 7.96 | 7.33
2008-07-18 | 22:30 | 40 | 87.8 | 20.13 | 61 | 7.96 | 7.28
2008-07-18 | 22:45 | 40 | 87.5 | 19.98 | 61 | 7.95 | 7.22
2008-07-18 | 23:00 | 40 | 88.5 | 19.83 | 61 | 8.08 | 7.15
2008-07-18 | 23:15 | 40 | 87.4 | 19.68 | 61 | 7.99 | 7.1
2008-07-18 | 23:30 | 40 | 87.1 | 19.53 | 61 | 7.99 | 7.09
2008-07-18 | 23:45 | 40 | 86.9 | 19.39 | 61 | 8 | 7.05
2008-07-19 | 0:00 | 40 | 86.2 | 19.26 | 61 | 7.96 | 7.03
2008-07-19 | 0:15 | 40 | 86.2 | 19.13 | 61 | 7.98 | 7.01
2008-07-19 | 0:30 | 40 | 86.1 | 19.01 | 61 | 7.98 | 6.99
2008-07-19 | 0:45 | 40 | 85.8 | 18.89 | 61 | 7.98 | 6.99
2008-07-19 | 1:01 | 40 | 85.8 | 18.78 | 61 | 8 | 6.98
2008-07-19 | 1:16 | 40 | 85.8 | 18.68 | 61 | 8.01 | 6.97
2008-07-19 | 1:31 | 40 | 85.7 | 18.58 | 61 | 8.02 | 6.95
2008-07-19 | 1:46 | 40 | 85.6 | 18.49 | 61 | 8.02 | 6.94
2008-07-19 | 2:01 | 40 | 85.5 | 18.4 | 61 | 8.03 | 6.94
2008-07-19 | 2:16 | 40 | 85.6 | 18.32 | 61 | 8.05 | 6.93
2008-07-19 | 2:31 | 40 | 85.3 | 18.24 | 61 | 8.04 | 6.94
2008-07-19 | 2:46 | 40 | 85.5 | 18.17 | 61 | 8.06 | 6.93
2008-07-19 | 3:01 | 40 | 85.5 | 18.11 | 61 | 8.08 | 6.92
2008-07-19 | 3:16 | 40 | 85.2 | 18.05 | 61 | 8.06 | 6.92
2008-07-19 | 3:31 | 40 | 85.4 | 17.99 | 61 | 8.09 | 6.92
2008-07-19 | 3:46 | 40 | 85.2 | 17.94 | 61 | 8.07 | 6.92
2008-07-19 | 4:01 | 40 | 85.2 | 17.91 | 61 | 8.08 | 6.91
2008-07-19 | 4:16 | 40 | 85.1 | 17.87 | 61 | 8.08 | 6.9
2008-07-19 | 4:31 | 40 | 85 | 17.83 | 62 | 8.08 | 6.93
2008-07-19 | 4:46 | 40 | 85.3 | 17.8 | 62 | 8.11 | 6.9
2008-07-19 | 5:01 | 40 | 85.3 | 17.75 | 62 | 8.12 | 6.92
2008-07-19 | 5:16 | 40 | 85.1 | 17.73 | 62 | 8.1 | 6.91
2008-07-19 | 5:31 | 40 | 85.3 | 17.7 | 62 | 8.13 | 6.9
2008-07-19 | 5:46 | 40 | 85.2 | 17.65 | 62 | 8.12 | 6.9
2008-07-19 | 6:01 | 40 | 85.3 | 17.63 | 62 | 8.14 | 6.92
2008-07-19 | 6:16 | 41 | 85.9 | 17.6 | 62 | 8.2 | 6.93
2008-07-19 | 6:31 | 41 | 86.2 | 17.58 | 62 | 8.23 | 6.93
2008-07-19 | 6:46 | 41 | 86.9 | 17.62 | 62 | 8.29 | 6.95
2008-07-19 | 7:01 | 41 | 87.9 | 17.61 | 62 | 8.38 | 6.98
2008-07-19 | 7:16 | 41 | 89 | 17.65 | 63 | 8.48 | 6.99
2008-07-19 | 7:31 | 41 | 89.8 | 17.69 | 63 | 8.55 | 7.03
2008-07-19 | 7:46 | 41 | 90.9 | 17.75 | 63 | 8.65 | 7.06
2008-07-19 | 8:01 | 41 | 92.1 | 17.84 | 63 | 8.75 | 7.08
2008-07-19 | 8:16 | 41 | 93 | 17.93 | 63 | 8.82 | 7.12
2008-07-19 | 8:31 | 41 | 94 | 18.04 | 63 | 8.89 | 7.16
2008-07-19 | 8:46 | 41 | 95.1 | 18.17 | 63 | 8.97 | 7.2
2008-07-19 | 9:01 | 41 | 96.1 | 18.31 | 63 | 9.04 | 7.24
2008-07-19 | 9:16 | 41 | 97.4 | 18.47 | 64 | 9.13 | 7.28
2008-07-19 | 9:31 | 41 | 98.9 | 18.65 | 64 | 9.24 | 7.33
2008-07-19 | 9:46 | 41 | 100.1 | 18.82 | 64 | 9.32 | 7.38
2008-07-19 | 10:01 | 42 | 101.1 | 19 | 64 | 9.38 | 7.43
2008-07-19 | 10:16 | 42 | 102.3 | 19.2 | 64 | 9.45 | 7.49
2008-07-19 | 10:31 | 42 | 103.9 | 19.39 | 64 | 9.56 | 7.55
2008-07-19 | 10:46 | 42 | 104.7 | 19.59 | 64 | 9.6 | 7.6
2008-07-19 | 11:01 | 42 | 105.6 | 19.81 | 64 | 9.64 | 7.67
2008-07-19 | 11:16 | 42 | 106.9 | 20.02 | 65 | 9.71 | 7.72
2008-07-19 | 11:31 | 42 | 107.9 | 20.23 | 65 | 9.76 | 7.78
2008-07-19 | 11:46 | 42 | 108.2 | 20.43 | 65 | 9.76 | 7.84
2008-07-19 | 12:01 | 42 | 109.2 | 20.63 | 65 | 9.8 | 7.91
2008-07-19 | 12:16 | 42 | 109.9 | 20.87 | 65 | 9.82 | 7.95
2008-07-19 | 12:31 | 43 | 110 | 21.07 | 65 | 9.79 | 8
2008-07-19 | 12:46 | 43 | 110.6 | 21.27 | 66 | 9.8 | 8.06
2008-07-19 | 13:00 | 43 | 110.7 | 21.46 | 66 | 9.78 | 8.11
2008-07-19 | 13:15 | 43 | 111.1 | 21.65 | 66 | 9.78 | 8.14
2008-07-19 | 13:30 | 43 | 110.9 | 21.82 | 66 | 9.73 | 8.18
2008-07-19 | 13:45 | 43 | 111.3 | 21.98 | 66 | 9.73 | 8.21
2008-07-19 | 14:00 | 43 | 111.6 | 22.14 | 66 | 9.73 | 8.25
2008-07-19 | 14:15 | 43 | 111.5 | 22.28 | 66 | 9.7 | 8.27
2008-07-19 | 14:30 | 43 | 111.2 | 22.41 | 66 | 9.64 | 8.3
2008-07-19 | 14:45 | 43 | 111.2 | 22.5 | 67 | 9.63 | 8.32
2008-07-19 | 15:00 | 43 | 111.3 | 22.6 | 67 | 9.62 | 8.38
2008-07-19 | 15:15 | 43 | 111.1 | 22.69 | 67 | 9.59 | 8.39
2008-07-19 | 15:30 | 43 | 110.5 | 22.76 | 67 | 9.52 | 8.37
2008-07-19 | 15:45 | 44 | 110 | 22.81 | 67 | 9.47 | 8.39
2008-07-19 | 16:00 | 44 | 109.8 | 22.85 | 67 | 9.45 | 8.4
2008-07-19 | 16:15 | 44 | 109 | 22.88 | 67 | 9.37 | 8.39
2008-07-19 | 16:30 | 44 | 108.4 | 22.89 | 67 | 9.32 | 8.41
2008-07-19 | 16:45 | 44 | 108.1 | 22.9 | 67 | 9.29 | 8.4
2008-07-19 | 17:00 | 44 | 107.7 | 22.88 | 67 | 9.26 | 8.38
2008-07-19 | 17:15 | 44 | 107 | 22.83 | 67 | 9.21 | 8.38
2008-07-19 | 17:30 | 44 | 105.9 | 22.77 | 67 | 9.13 | 8.36
2008-07-19 | 17:45 | 44 | 105.5 | 22.7 | 67 | 9.1 | 8.36
2008-07-19 | 18:00 | 44 | 104.5 | 22.62 | 67 | 9.03 | 8.34
2008-07-19 | 18:15 | 44 | 104.1 | 22.51 | 67 | 9.01 | 8.31
2008-07-19 | 18:30 | 43 | 103.2 | 22.4 | 67 | 8.95 | 8.3
2008-07-19 | 18:45 | 43 | 102.3 | 22.26 | 67 | 8.9 | 8.27
2008-07-19 | 19:00 | 43 | 101.4 | 22.12 | 66 | 8.84 | 8.24
2008-07-19 | 19:15 | 43 | 100.2 | 21.96 | 66 | 8.77 | 8.22
2008-07-19 | 19:30 | 43 | 99.1 | 21.8 | 66 | 8.7 | 8.18
2008-07-19 | 19:45 | 43 | 97.5 | 21.62 | 66 | 8.59 | 8.15
2008-07-19 | 20:00 | 43 | 96.6 | 21.45 | 66 | 8.53 | 8.1
2008-07-19 | 20:15 | 43 | 95.6 | 21.27 | 66 | 8.48 | 8.05
2008-07-19 | 20:30 | 43 | 94.3 | 21.1 | 66 | 8.39 | 7.99
2008-07-19 | 20:45 | 43 | 93.1 | 20.92 | 66 | 8.32 | 7.91
2008-07-19 | 21:00 | 43 | 92.5 | 20.76 | 66 | 8.29 | 7.84
2008-07-19 | 21:15 | 42 | 91.6 | 20.59 | 65 | 8.24 | 7.74
2008-07-19 | 21:30 | 42 | 90.9 | 20.42 | 65 | 8.2 | 7.66
2008-07-19 | 21:45 | 42 | 90.4 | 20.28 | 65 | 8.17 | 7.57
2008-07-19 | 22:00 | 42 | 89.4 | 20.14 | 65 | 8.11 | 7.48
2008-07-19 | 22:15 | 42 | 89.1 | 20 | 65 | 8.1 | 7.41
2008-07-19 | 22:30 | 42 | 88.3 | 19.87 | 65 | 8.04 | 7.35
2008-07-19 | 22:45 | 42 | 87.9 | 19.74 | 65 | 8.03 | 7.28
2008-07-19 | 23:00 | 42 | 87.4 | 19.63 | 65 | 8.01 | 7.23
2008-07-19 | 23:15 | 42 | 87.2 | 19.51 | 65 | 8.01 | 7.19
2008-07-19 | 23:30 | 42 | 86.6 | 19.41 | 65 | 7.97 | 7.16
2008-07-19 | 23:45 | 42 | 86.3 | 19.31 | 65 | 7.96 | 7.13
2008-07-20 | 0:00 | 42 | 86.3 | 19.22 | 65 | 7.97 | 7.11
2008-07-20 | 0:15 | 42 | 86 | 19.13 | 65 | 7.96 | 7.08
2008-07-20 | 0:30 | 42 | 85.7 | 19.05 | 65 | 7.94 | 7.07
2008-07-20 | 0:45 | 42 | 85.6 | 18.96 | 65 | 7.94 | 7.05
2008-07-20 | 1:01 | 42 | 85.5 | 18.89 | 65 | 7.95 | 7.05
2008-07-20 | 1:16 | 42 | 85.5 | 18.82 | 65 | 7.96 | 7.03
2008-07-20 | 1:31 | 42 | 85.4 | 18.76 | 65 | 7.96 | 7.02
2008-07-20 | 1:46 | 42 | 85.5 | 18.7 | 65 | 7.98 | 7.03
2008-07-20 | 2:01 | 42 | 85.4 | 18.63 | 65 | 7.98 | 7
2008-07-20 | 2:16 | 42 | 85.4 | 18.58 | 65 | 7.99 | 7
2008-07-20 | 2:31 | 42 | 85.1 | 18.54 | 65 | 7.97 | 6.99
2008-07-20 | 2:46 | 42 | 85.2 | 18.48 | 65 | 7.98 | 6.98
2008-07-20 | 3:01 | 42 | 85.3 | 18.45 | 65 | 8 | 6.98
2008-07-20 | 3:16 | 42 | 85.3 | 18.41 | 65 | 8 | 6.99
2008-07-20 | 3:31 | 42 | 85 | 18.37 | 65 | 7.98 | 6.98
2008-07-20 | 3:46 | 42 | 85.1 | 18.34 | 65 | 8 | 6.99
2008-07-20 | 4:01 | 42 | 84.9 | 18.31 | 65 | 7.98 | 6.97
2008-07-20 | 4:16 | 42 | 85 | 18.28 | 65 | 8 | 6.97
2008-07-20 | 4:31 | 42 | 84.9 | 18.26 | 65 | 7.99 | 6.97
2008-07-20 | 4:46 | 42 | 84.9 | 18.24 | 65 | 8 | 6.97
2008-07-20 | 5:01 | 43 | 85 | 18.2 | 65 | 8.01 | 6.97
2008-07-20 | 5:16 | 43 | 84.7 | 18.19 | 65 | 7.99 | 6.96
2008-07-20 | 5:31 | 43 | 84.9 | 18.16 | 66 | 8.01 | 6.96
2008-07-20 | 5:46 | 43 | 84.7 | 18.13 | 66 | 8 | 6.96
2008-07-20 | 6:01 | 43 | 85.2 | 18.12 | 66 | 8.04 | 6.96
2008-07-20 | 6:16 | 43 | 85.7 | 18.11 | 66 | 8.09 | 6.97
2008-07-20 | 6:31 | 43 | 85.9 | 18.1 | 66 | 8.12 | 6.96
2008-07-20 | 6:46 | 43 | 86.6 | 18.1 | 66 | 8.18 | 7.01
2008-07-20 | 7:01 | 43 | 87.8 | 18.12 | 66 | 8.29 | 7.02
2008-07-20 | 7:16 | 43 | 88.8 | 18.14 | 66 | 8.38 | 7.04
2008-07-20 | 7:31 | 43 | 89.8 | 18.18 | 66 | 8.47 | 7.05
2008-07-20 | 7:46 | 43 | 91.2 | 18.21 | 66 | 8.6 | 7.11
2008-07-20 | 8:01 | 43 | 92.5 | 18.28 | 66 | 8.7 | 7.14
2008-07-20 | 8:16 | 43 | 93.6 | 18.37 | 65 | 8.8 | 7.17
2008-07-20 | 8:31 | 43 | 94.7 | 18.45 | 66 | 8.88 | 7.21
2008-07-20 | 8:46 | 43 | 96 | 18.56 | 66 | 8.99 | 7.26
2008-07-20 | 9:01 | 43 | 97.5 | 18.68 | 65 | 9.11 | 7.31
2008-07-20 | 9:16 | 43 | 98.9 | 18.8 | 66 | 9.21 | 7.35
2008-07-20 | 9:31 | 43 | 100.1 | 18.95 | 66 | 9.3 | 7.4
2008-07-20 | 9:46 | 43 | 101 | 19.09 | 66 | 9.35 | 7.45
2008-07-20 | 10:01 | 43 | 102.1 | 19.25 | 66 | 9.43 | 7.5
2008-07-20 | 10:16 | 43 | 103.7 | 19.42 | 66 | 9.54 | 7.58
2008-07-20 | 10:31 | 43 | 105 | 19.6 | 66 | 9.63 | 7.65
2008-07-20 | 10:46 | 43 | 106.2 | 19.79 | 66 | 9.69 | 7.71
2008-07-20 | 11:01 | 43 | 107.2 | 19.98 | 66 | 9.75 | 7.77
2008-07-20 | 11:16 | 43 | 107.8 | 20.17 | 66 | 9.77 | 7.83
2008-07-20 | 11:31 | 43 | 108.5 | 20.35 | 66 | 9.79 | 7.9
2008-07-20 | 11:46 | 43 | 109.6 | 20.56 | 66 | 9.86 | 7.97
2008-07-20 | 12:01 | 43 | 110.1 | 20.76 | 66 | 9.86 | 8.02
2008-07-20 | 12:16 | 43 | 110.9 | 20.94 | 66 | 9.9 | 8.08
2008-07-20 | 12:31 | 43 | 111.6 | 21.12 | 66 | 9.93 | 8.14
2008-07-20 | 12:46 | 43 | 112.2 | 21.32 | 66 | 9.94 | 8.19
2008-07-20 | 13:00 | 43 | 112.6 | 21.5 | 66 | 9.94 | 8.23
2008-07-20 | 13:15 | 43 | 112.9 | 21.67 | 66 | 9.93 | 8.26
2008-07-20 | 13:30 | 43 | 112.7 | 21.85 | 66 | 9.89 | 8.32
2008-07-20 | 13:45 | 43 | 113.2 | 22 | 67 | 9.89 | 8.35
2008-07-20 | 14:00 | 43 | 113 | 22.14 | 67 | 9.85 | 8.38
2008-07-20 | 14:15 | 43 | 112.8 | 22.26 | 67 | 9.81 | 8.41
2008-07-20 | 14:30 | 43 | 113.1 | 22.37 | 67 | 9.82 | 8.42
2008-07-20 | 14:45 | 43 | 112.8 | 22.48 | 67 | 9.77 | 8.45
2008-07-20 | 15:00 | 43 | 112.9 | 22.56 | 67 | 9.77 | 8.45
2008-07-20 | 15:15 | 43 | 112.2 | 22.64 | 67 | 9.69 | 8.48
2008-07-20 | 15:30 | 43 | 112.2 | 22.7 | 67 | 9.68 | 8.47
2008-07-20 | 15:45 | 43 | 111.8 | 22.75 | 67 | 9.64 | 8.49
2008-07-20 | 16:00 | 43 | 111.2 | 22.77 | 67 | 9.58 | 8.48
2008-07-20 | 16:15 | 43 | 110.8 | 22.78 | 67 | 9.54 | 8.48
2008-07-20 | 16:30 | 43 | 110.4 | 22.8 | 67 | 9.5 | 8.49
2008-07-20 | 16:45 | 43 | 109.8 | 22.8 | 67 | 9.45 | 8.48
2008-07-20 | 17:00 | 43 | 108.8 | 22.77 | 66 | 9.38 | 8.47
2008-07-20 | 17:15 | 43 | 107.9 | 22.75 | 66 | 9.3 | 8.46
2008-07-20 | 17:30 | 43 | 107 | 22.71 | 66 | 9.23 | 8.44
2008-07-20 | 17:45 | 43 | 106.3 | 22.65 | 66 | 9.18 | 8.43
2008-07-20 | 18:00 | 43 | 105.4 | 22.58 | 66 | 9.11 | 8.42
2008-07-20 | 18:15 | 43 | 104.8 | 22.5 | 66 | 9.08 | 8.38
2008-07-20 | 18:30 | 42 | 104.1 | 22.4 | 65 | 9.03 | 8.37
2008-07-20 | 18:45 | 42 | 103.2 | 22.28 | 65 | 8.98 | 8.34
2008-07-20 | 19:00 | 42 | 102.2 | 22.16 | 65 | 8.91 | 8.3
2008-07-20 | 19:15 | 42 | 101.2 | 22.03 | 65 | 8.84 | 8.28
2008-07-20 | 19:30 | 42 | 99.6 | 21.89 | 64 | 8.73 | 8.24
2008-07-20 | 19:45 | 42 | 98.4 | 21.74 | 64 | 8.65 | 8.18
2008-07-20 | 20:00 | 42 | 97.5 | 21.58 | 64 | 8.6 | 8.13
2008-07-20 | 20:15 | 42 | 96.5 | 21.44 | 64 | 8.53 | 8.07
2008-07-20 | 20:30 | 41 | 95.2 | 21.28 | 64 | 8.44 | 7.98
2008-07-20 | 20:45 | 41 | 94.3 | 21.14 | 64 | 8.38 | 7.89
2008-07-20 | 21:00 | 41 | 93.3 | 20.99 | 64 | 8.32 | 7.8
2008-07-20 | 21:15 | 41 | 92.2 | 20.85 | 64 | 8.24 | 7.69
2008-07-20 | 21:30 | 41 | 91.4 | 20.72 | 64 | 8.19 | 7.59
2008-07-20 | 21:45 | 41 | 90.3 | 20.59 | 64 | 8.12 | 7.5
2008-07-20 | 22:00 | 41 | 89.5 | 20.47 | 64 | 8.06 | 7.42
2008-07-20 | 22:15 | 41 | 89 | 20.34 | 64 | 8.03 | 7.34
2008-07-20 | 22:30 | 41 | 88.4 | 20.22 | 64 | 8 | 7.3
2008-07-20 | 22:45 | 41 | 87.8 | 20.12 | 64 | 7.96 | 7.24
2008-07-20 | 23:00 | 41 | 87.8 | 20.01 | 64 | 7.98 | 7.2
2008-07-20 | 23:15 | 41 | 87.1 | 19.92 | 64 | 7.93 | 7.16
2008-07-20 | 23:30 | 41 | 87.1 | 19.84 | 64 | 7.94 | 7.14
2008-07-20 | 23:45 | 41 | 86.9 | 19.76 | 64 | 7.94 | 7.12
2008-07-21 | 0:00 | 41 | 86.5 | 19.7 | 64 | 7.92 | 7.09
2008-07-21 | 0:15 | 41 | 86.4 | 19.64 | 64 | 7.91 | 7.07
2008-07-21 | 0:30 | 41 | 86 | 19.58 | 64 | 7.88 | 7.06
2008-07-21 | 0:45 | 41 | 85.7 | 19.54 | 64 | 7.86 | 7.04
2008-07-21 | 1:01 | 41 | 85.6 | 19.49 | 64 | 7.86 | 7.03
2008-07-21 | 1:16 | 41 | 85.5 | 19.45 | 64 | 7.86 | 7.02
2008-07-21 | 1:31 | 42 | 85.3 | 19.41 | 64 | 7.85 | 7.01
2008-07-21 | 1:46 | 42 | 85.1 | 19.38 | 64 | 7.84 | 7.01
2008-07-21 | 2:01 | 42 | 85.2 | 19.35 | 64 | 7.85 | 6.99
2008-07-21 | 2:16 | 42 | 85.1 | 19.35 | 64 | 7.84 | 7
2008-07-21 | 2:31 | 42 | 84.9 | 19.34 | 64 | 7.82 | 6.99
2008-07-21 | 2:46 | 42 | 84.8 | 19.34 | 64 | 7.81 | 6.99
2008-07-21 | 3:01 | 42 | 84.7 | 19.33 | 64 | 7.81 | 6.97
2008-07-21 | 3:16 | 42 | 84.6 | 19.32 | 64 | 7.8 | 7
2008-07-21 | 3:31 | 42 | 84.8 | 19.32 | 64 | 7.82 | 7
2008-07-21 | 3:46 | 42 | 84.5 | 19.34 | 65 | 7.79 | 6.98
2008-07-21 | 4:01 | 42 | 84.7 | 19.35 | 65 | 7.8 | 6.97
2008-07-21 | 4:16 | 42 | 84.4 | 19.35 | 65 | 7.77 | 6.98
2008-07-21 | 4:31 | 42 | 84.6 | 19.34 | 65 | 7.79 | 6.97
2008-07-21 | 4:46 | 42 | 84.6 | 19.33 | 65 | 7.8 | 6.97
2008-07-21 | 5:01 | 42 | 84.7 | 19.33 | 65 | 7.8 | 6.97
2008-07-21 | 5:16 | 42 | 84.4 | 19.35 | 65 | 7.78 | 6.97
2008-07-21 | 5:31 | 42 | 84.5 | 19.34 | 65 | 7.78 | 6.96
2008-07-21 | 5:46 | 42 | 84.2 | 19.32 | 65 | 7.76 | 6.96
2008-07-21 | 6:01 | 42 | 84.4 | 19.31 | 65 | 7.78 | 6.96
2008-07-21 | 6:16 | 42 | 84.8 | 19.31 | 65 | 7.82 | 6.97
2008-07-21 | 6:31 | 43 | 85 | 19.29 | 65 | 7.83 | 6.97
2008-07-21 | 6:46 | 43 | 85.4 | 19.28 | 66 | 7.88 | 6.97
2008-07-21 | 7:01 | 43 | 85.8 | 19.27 | 66 | 7.92 | 6.99
2008-07-21 | 7:16 | 43 | 87.2 | 19.28 | 66 | 8.04 | 7
2008-07-21 | 7:31 | 43 | 88.8 | 19.33 | 66 | 8.18 | 7.04
2008-07-21 | 7:46 | 43 | 89.8 | 19.37 | 66 | 8.27 | 7.06
2008-07-21 | 8:01 | 43 | 91.1 | 19.43 | 66 | 8.38 | 7.09
2008-07-21 | 8:16 | 43 | 92.4 | 19.48 | 66 | 8.48 | 7.14
2008-07-21 | 8:31 | 43 | 93.7 | 19.57 | 66 | 8.59 | 7.18
2008-07-21 | 8:46 | 43 | 94.9 | 19.65 | 66 | 8.68 | 7.21
2008-07-21 | 9:01 | 43 | 96.6 | 19.77 | 66 | 8.82 | 7.26
2008-07-21 | 9:16 | 43 | 97.6 | 19.89 | 66 | 8.9 | 7.31
2008-07-21 | 9:31 | 43 | 99.3 | 20.02 | 66 | 9.02 | 7.36
2008-07-21 | 9:46 | 43 | 100.3 | 20.16 | 66 | 9.09 | 7.41
2008-07-21 | 10:01 | 43 | 101.9 | 20.31 | 66 | 9.21 | 7.47
2008-07-21 | 10:16 | 43 | 103.3 | 20.47 | 66 | 9.3 | 7.53
2008-07-21 | 10:31 | 43 | 104.3 | 20.64 | 66 | 9.36 | 7.59
2008-07-21 | 10:46 | 43 | 105.3 | 20.81 | 66 | 9.42 | 7.65
2008-07-21 | 11:01 | 43 | 105.9 | 20.99 | 66 | 9.44 | 7.72
2008-07-21 | 11:16 | 43 | 106.6 | 21.18 | 67 | 9.47 | 7.8
2008-07-21 | 11:31 | 43 | 108.3 | 21.36 | 67 | 9.59 | 7.86
2008-07-21 | 11:46 | 43 | 109.5 | 21.55 | 67 | 9.66 | 7.92
2008-07-21 | 12:01 | 43 | 109.9 | 21.73 | 67 | 9.66 | 7.97
2008-07-21 | 12:16 | 43 | 110.9 | 21.91 | 67 | 9.72 | 8.05
2008-07-21 | 12:31 | 43 | 111.6 | 22.09 | 67 | 9.74 | 8.08
2008-07-21 | 12:46 | 43 | 112 | 22.26 | 67 | 9.75 | 8.14
2008-07-21 | 13:00 | 43 | 112.4 | 22.43 | 67 | 9.75 | 8.19
2008-07-21 | 13:15 | 43 | 112.3 | 22.57 | 67 | 9.71 | 8.22
2008-07-21 | 13:30 | 44 | 112.8 | 22.72 | 67 | 9.72 | 8.27
2008-07-21 | 13:45 | 44 | 112.6 | 22.85 | 67 | 9.68 | 8.29
2008-07-21 | 14:00 | 43 | 112.6 | 22.97 | 67 | 9.66 | 8.33
2008-07-21 | 14:15 | 43 | 112.4 | 23.08 | 67 | 9.63 | 8.35
2008-07-21 | 14:30 | 43 | 112.8 | 23.19 | 67 | 9.64 | 8.36
2008-07-21 | 14:45 | 43 | 111.9 | 23.2 | 67 | 9.56 | 8.36
2008-07-21 | 15:00 | 43 | 111.6 | 23.27 | 67 | 9.53 | 8.39
2008-07-21 | 15:15 | 43 | 111.4 | 23.27 | 66 | 9.51 | 8.37
2008-07-21 | 15:30 | 43 | 111 | 23.33 | 67 | 9.46 | 8.41
2008-07-21 | 15:45 | 43 | 110.5 | 23.33 | 66 | 9.42 | 8.41
2008-07-21 | 16:00 | 43 | 110.1 | 23.31 | 66 | 9.38 | 8.39
2008-07-21 | 16:15 | 43 | 110.3 | 23.45 | 66 | 9.38 | 8.47
2008-07-21 | 16:30 | 43 | 109.5 | 23.4 | 66 | 9.32 | 8.4
2008-07-21 | 16:45 | 43 | 109 | 23.37 | 66 | 9.28 | 8.4
2008-07-21 | 17:00 | 43 | 109.1 | 23.42 | 66 | 9.29 | 8.45
2008-07-21 | 17:15 | 42 | 106.6 | 23.3 | 65 | 9.09 | 8.35
2008-07-21 | 17:30 | 42 | 104.5 | 23.17 | 65 | 8.93 | 8.3
2008-07-21 | 17:45 | 42 | 103.2 | 23.08 | 65 | 8.84 | 8.29
2008-07-21 | 18:00 | 42 | 101.5 | 22.96 | 65 | 8.71 | 8.25
2008-07-21 | 18:15 | 42 | 100.1 | 22.8 | 64 | 8.62 | 8.19
2008-07-21 | 18:30 | 42 | 98.5 | 22.68 | 64 | 8.5 | 8.13
2008-07-21 | 18:45 | 41 | 97.1 | 22.54 | 64 | 8.41 | 8.07
2008-07-21 | 19:00 | 41 | 96 | 22.42 | 64 | 8.33 | 8.02
2008-07-21 | 19:15 | 41 | 96 | 22.34 | 64 | 8.34 | 7.97
2008-07-21 | 19:30 | 41 | 95.7 | 22.23 | 63 | 8.33 | 7.9
2008-07-21 | 19:45 | 41 | 94.8 | 22.08 | 63 | 8.28 | 7.81
2008-07-21 | 20:00 | 41 | 94 | 21.94 | 63 | 8.23 | 7.74
2008-07-21 | 20:15 | 41 | 93.5 | 21.8 | 63 | 8.2 | 7.65
2008-07-21 | 20:30 | 41 | 92.2 | 21.67 | 63 | 8.11 | 7.56
2008-07-21 | 20:45 | 41 | 91 | 21.53 | 63 | 8.03 | 7.46
2008-07-21 | 21:00 | 41 | 90.3 | 21.39 | 63 | 7.99 | 7.38
2008-07-21 | 21:15 | 41 | 89.4 | 21.28 | 63 | 7.93 | 7.32
2008-07-21 | 21:30 | 41 | 89 | 21.17 | 63 | 7.91 | 7.27
2008-07-21 | 21:45 | 41 | 88.3 | 21.06 | 63 | 7.86 | 7.22
2008-07-21 | 22:00 | 41 | 88 | 20.96 | 63 | 7.85 | 7.18
2008-07-21 | 22:15 | 41 | 87.6 | 20.89 | 63 | 7.83 | 7.15
2008-07-21 | 22:30 | 41 | 87.1 | 20.81 | 63 | 7.79 | 7.12
2008-07-21 | 22:45 | 41 | 87 | 20.73 | 63 | 7.8 | 7.1
2008-07-21 | 23:00 | 41 | 86.5 | 20.66 | 63 | 7.77 | 7.09
2008-07-21 | 23:15 | 41 | 86.4 | 20.59 | 63 | 7.76 | 7.07
2008-07-21 | 23:30 | 41 | 85.9 | 20.52 | 63 | 7.73 | 7.05
2008-07-21 | 23:45 | 41 | 85.6 | 20.46 | 63 | 7.71 | 7.04
2008-07-22 | 0:00 | 41 | 85.4 | 20.4 | 63 | 7.7 | 7.03
2008-07-22 | 0:15 | 41 | 85.2 | 20.35 | 63 | 7.69 | 7.02
2008-07-22 | 0:30 | 41 | 85.1 | 20.32 | 63 | 7.69 | 7.01
2008-07-22 | 0:45 | 41 | 85 | 20.29 | 63 | 7.69 | 7
2008-07-22 | 1:01 | 41 | 85 | 20.25 | 63 | 7.69 | 6.99
2008-07-22 | 1:16 | 41 | 84.9 | 20.22 | 63 | 7.69 | 6.98
2008-07-22 | 1:31 | 41 | 84.9 | 20.21 | 63 | 7.69 | 6.98
2008-07-22 | 1:46 | 41 | 84.9 | 20.2 | 63 | 7.69 | 6.97
2008-07-22 | 2:01 | 41 | 84.7 | 20.2 | 63 | 7.67 | 6.99
2008-07-22 | 2:16 | 41 | 84.7 | 20.21 | 63 | 7.67 | 6.97
2008-07-22 | 2:31 | 41 | 84.3 | 20.21 | 63 | 7.64 | 6.98
2008-07-22 | 2:46 | 41 | 84.5 | 20.22 | 63 | 7.65 | 6.96
2008-07-22 | 3:01 | 41 | 84.2 | 20.21 | 63 | 7.62 | 6.97
2008-07-22 | 3:16 | 41 | 84.2 | 20.21 | 63 | 7.62 | 6.96
2008-07-22 | 3:31 | 41 | 84 | 20.19 | 63 | 7.61 | 6.96
2008-07-22 | 3:46 | 41 | 84.2 | 20.19 | 63 | 7.63 | 6.96
2008-07-22 | 4:01 | 41 | 84.1 | 20.2 | 63 | 7.62 | 6.95
2008-07-22 | 4:16 | 41 | 83.9 | 20.2 | 63 | 7.6 | 6.95
2008-07-22 | 4:31 | 41 | 84.3 | 20.19 | 63 | 7.63 | 6.95
2008-07-22 | 4:46 | 41 | 84 | 20.17 | 63 | 7.61 | 6.96
2008-07-22 | 5:01 | 41 | 84.1 | 20.14 | 63 | 7.62 | 6.93
2008-07-22 | 5:16 | 41 | 84.1 | 20.11 | 64 | 7.63 | 6.93
2008-07-22 | 5:31 | 41 | 83.9 | 20.07 | 64 | 7.62 | 6.93
2008-07-22 | 5:46 | 41 | 83.9 | 20.04 | 64 | 7.62 | 6.93
2008-07-22 | 6:01 | 41 | 84.3 | 20.01 | 64 | 7.66 | 6.92
2008-07-22 | 6:16 | 42 | 84.6 | 19.97 | 64 | 7.7 | 6.94
2008-07-22 | 6:31 | 42 | 85.4 | 19.94 | 64 | 7.77 | 6.94
2008-07-22 | 6:46 | 42 | 85.9 | 19.92 | 64 | 7.83 | 6.96
2008-07-22 | 7:01 | 42 | 87.1 | 19.9 | 64 | 7.93 | 6.98
2008-07-22 | 7:16 | 42 | 88.2 | 19.89 | 64 | 8.04 | 7.01
2008-07-22 | 7:31 | 42 | 89.2 | 19.91 | 65 | 8.13 | 7.03
2008-07-22 | 7:46 | 42 | 90.7 | 19.93 | 65 | 8.26 | 7.06
2008-07-22 | 8:01 | 42 | 92.1 | 19.96 | 65 | 8.38 | 7.09
2008-07-22 | 8:16 | 42 | 93.2 | 20.01 | 65 | 8.47 | 7.13
2008-07-22 | 8:31 | 42 | 94.5 | 20.04 | 65 | 8.58 | 7.17
2008-07-22 | 8:46 | 42 | 95.9 | 20.12 | 65 | 8.7 | 7.22
2008-07-22 | 9:01 | 42 | 97.5 | 20.19 | 65 | 8.83 | 7.26
2008-07-22 | 9:16 | 42 | 99 | 20.28 | 65 | 8.95 | 7.32
2008-07-22 | 9:31 | 42 | 100.3 | 20.38 | 65 | 9.05 | 7.37
2008-07-22 | 9:46 | 42 | 101.2 | 20.48 | 65 | 9.12 | 7.43
2008-07-22 | 10:01 | 42 | 102.8 | 20.61 | 65 | 9.24 | 7.49
2008-07-22 | 10:16 | 42 | 104.3 | 20.73 | 65 | 9.34 | 7.54
2008-07-22 | 10:31 | 42 | 105.4 | 20.85 | 65 | 9.42 | 7.6
2008-07-22 | 10:46 | 43 | 106.5 | 20.98 | 65 | 9.5 | 7.66
2008-07-22 | 11:01 | 43 | 107.6 | 21.12 | 65 | 9.57 | 7.73
2008-07-22 | 11:16 | 43 | 108.5 | 21.3 | 66 | 9.62 | 7.81
2008-07-22 | 11:31 | 43 | 109.4 | 21.47 | 65 | 9.66 | 7.87
2008-07-22 | 11:46 | 43 | 109.8 | 21.62 | 66 | 9.67 | 7.93
2008-07-22 | 12:01 | 43 | 110.7 | 21.79 | 66 | 9.71 | 7.99
2008-07-22 | 12:16 | 43 | 110.9 | 21.97 | 66 | 9.71 | 8.04
2008-07-22 | 12:31 | 43 | 111.4 | 22.12 | 66 | 9.72 | 8.1
2008-07-22 | 12:46 | 43 | 111.5 | 22.26 | 66 | 9.7 | 8.12
2008-07-22 | 13:00 | 43 | 112.4 | 22.43 | 66 | 9.74 | 8.18