## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Loner Sonde V2
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-25 |  | 2008-06-25 |  | A
Time | Computer set to atomic time |  | 12:04 |  | 12:04 |  | A
Temp | VWR NIST traceable surface water |  | 14.8 |  | 14.96 |  | A
DO | Winkler titration of surface water |  | 10.5 |  | 10.58 |  | A
SpC | 148 µS/mL |  | 59 |  | 60 |  | A
Ph | pH 7 standard solution |  | 7.78 |  | 6.75 |  | Fail
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2008-06-25 15:00 | 13.9 | 6.5 | 10.1 | 98 | 31 | 22
Sonde | 2008-06-25 15:00 | 13.96 | 7.6 | 9.78 | 95 | 32 | 21
Grade |  | A | Fail | B |  | A
Audit 3 | 2008-06-30 12:30 | 10.73 | 6.54 | 11.26 | 101.5 | 25 | 18
Sonde | 2008-06-30 12:30 | 10.8 | 7.4 | 10.54 | 95 | 26 | 17
Grade |  | A | Fail | B |  | A
Audit 5 | 2008-07-10 16:30 | 15.9 | 6.7 | 9.92 | 100.2 | 31 | 21
Sonde | 2008-07-10 16:30 | 15.97 | 7.5 | 9.4 | 95 | 31 | 20
Grade |  | A | Fail | B |  | A
Post-sampling Field Equipment Check
Field equipment ID | Loner Sonde V2
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-10 |  | 2008-07-10 |  | A
Time | Computer set to atomic time |  | 16:30 |  | 16:30 |  | A
Temp | VWR NIST traceable surface water |  | 15.9 |  | 15.97 |  | A
DO | Winkler titration of surface water |  | 9.6 |  | 9.4 |  | A
SpC | In situ against 556 µS/mL |  | 31 |  | 31 |  | A
Ph | pH In Situ Against 556 |  | 6.7 |  | 7.5 |  | Fail
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Sonde Files: 062508V2 & 062608V2 (due to in situ battery change) *Combined in data tab for this sheet Turbidity Data is not trust worthy do to Optics trobles & high flow |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Loner Sonde V2 (In Situ Audit)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-25 |  | 2008-06-25 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | YSI 556 |  | 13.87 |  | 13.96 |  | A
DO | YSI 556 |  | 10.13 |  | 9.78 |  | B
SpC | YSI 556 |  | 31 |  | 32 |  | A
Ph | YSI 556 |  | 6.49 |  | 7.6 |  | Fail
Turb | Hatch Filed Audit |  | 3.06 |  | 3.4 |  | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-06-25 15:00 | 13.9 | 6.5 | 10.1 | 98 | 31 | 22
3.06
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-26 |  | 2008-06-26 |  | A
Time | Computer set to atomic time |  | 16:15 |  | 16:15 |  | A
Temp | VWR NIST traceable surface water |  | 11.7 |  | 11.92 |  | A
DO | Winkler titration of surface water |  | 10.4 |  | 10.58 |  | A
SpC | 148 µS/mL |  | 147 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.15 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 Sonde Audit File: 062508V2 15:00 |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Loner (V2) Sonde & 556
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-26 |  | 2008-06-26 |  | A
Time | Computer set to atomic time |  | 16:15 |  | 16:15 |  | A
Temp | VWR NIST traceable surface water |  | 11.7 |  | 11.92 |  | A
DO | Winkler titration of surface water |  | 10.4 |  | 10.58 |  | A
SpC | 148 µS/mL |  | 147 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.15 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-06-26 14:15 | 12.57 | 6.47 | 10.74 | 100.9 | 30 | 21
4
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-30 |  | 2008-06-30 |  | A
Time | Computer set to atomic time |  | 14:40 |  | 14:40 |  | A
Temp | VWR NIST traceable surface water |  | 18.2 |  | 18.45 |  | A
DO | Winkler titration of surface water |  | 9.2 |  | 9.34 |  | A
SpC | 148 µS/mL |  | 147 |  | 144 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.23 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 Sonde Audit File: 062608V2 14:15 |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Loner (V2) Sonde & 556
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-26 |  | 2008-06-26 |  | A
Time | Computer set to atomic time |  | 16:15 |  | 16:15 |  | A
Temp | VWR NIST traceable surface water |  | 11.7 |  | 11.92 |  | A
DO | Winkler titration of surface water |  | 10.4 |  | 10.58 |  | A
SpC | 148 µS/mL |  | 147 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.15 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-06-30 12:30 | 10.73 | 6.54 | 11.26 | 101.5 | 25 | 18
25
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-30 |  | 2008-06-30 |  | A
Time | Computer set to atomic time |  | 14:40 |  | 14:40 |  | A
Temp | VWR NIST traceable surface water |  | 18.2 |  | 18.45 |  | A
DO | Winkler titration of surface water |  | 9.2 |  | 9.34 |  | A
SpC | 148 µS/mL |  | 147 |  | 144 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.23 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 Sonde Audit File: 063008V2 12:30 |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Loner (V2) Sonde & 556
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-07 |  | 2008-07-07 |  | A
Time | Computer set to atomic time |  | 12:26 |  | 12:26 |  | A
Temp | VWR NIST traceable surface water |  | 16.3 |  | 16.37 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 10.31 |  | A
SpC | 148 µS/mL |  | 147 |  | 153 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.2 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-07-07 15:30 | 13.94 | 6.14 | 10.36 | 100.4 | 23 | 16
6.63
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-07 |  | 2008-07-07 |  | A
Time | Computer set to atomic time |  | 17:30 |  | 17:30 |  | A
Temp | VWR NIST traceable surface water |  | 12.1 |  | 12.3 |  | A
DO | Winkler titration of surface water |  | 10.2 |  | 10.71 |  | B
SpC | 148 µS/mL |  | 149 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.22 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 Sonde Audit File: 070708V2 15:30 |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 5
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Loner (V2) Sonde & 556
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-10 |  | 2008-07-10 |  | A
Time | Computer set to atomic time |  | 9:59 |  | 9:59 |  | A
Temp | VWR NIST traceable surface water |  | 16.4 |  | 16.35 |  | A
DO | Winkler titration of surface water |  | 10 |  | 9.86 |  | A
SpC | 148 µS/mL |  | 149 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.22 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-07-10 16:30 | 15.9 | 6.7 | 9.92 | 100.2 | 31 | 21
4
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-10 |  | 2008-07-10 |  | A
Time | Computer set to atomic time |  | 18:15 |  | 18:15 |  | A
Temp | VWR NIST traceable surface water |  | 12.6 |  | 12.75 |  | A
DO | Winkler titration of surface water |  | 10.2 |  | 10.48 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 161 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.21 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 File for Sonde Audit: 7-10-08 16:30 |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | ODO Conc | ODO% | Temp | SpCond | pH | Turbidity+
M/D/Y |  | mg/L | % | C | uS/cm |  | NTU
2008-06-25 | 15:00 | 9.78 | 94.8 | 13.96 | 32 | 7.6 | 3.4
2008-06-25 | 15:15 | 9.74 | 94.9 | 14.17 | 32 | 7.61 | 4.4
2008-06-25 | 15:30 | 9.7 | 94.8 | 14.32 | 32 | 7.62 | 3.4
2008-06-25 | 15:45 | 9.66 | 94.7 | 14.48 | 32 | 7.62 | 3.4
2008-06-25 | 16:00 | 9.63 | 94.6 | 14.57 | 32 | 7.63 | 3.1
2008-06-25 | 16:15 | 9.6 | 94.5 | 14.68 | 32 | 7.63 | 3.2
2008-06-25 | 16:30 | 9.57 | 94.4 | 14.77 | 32 | 7.63 | 3.6
2008-06-25 | 16:45 | 9.54 | 94.3 | 14.84 | 32 | 7.63 | 3.2
2008-06-25 | 17:00 | 9.53 | 94.2 | 14.87 | 32 | 7.63 | 2.9
2008-06-25 | 17:15 | 9.5 | 94 | 14.91 | 32 | 7.62 | 3.2
2008-06-25 | 17:30 | 9.49 | 94 | 14.95 | 32 | 7.62 | 3.2
2008-06-25 | 17:45 | 9.47 | 93.8 | 14.95 | 32 | 7.62 | 3.6
2008-06-25 | 18:00 | 9.47 | 93.7 | 14.94 | 32 | 7.61 | 3.3
2008-06-25 | 18:15 | 9.46 | 93.7 | 14.94 | 32 | 7.61 | 3.4
2008-06-25 | 18:30 | 9.46 | 93.7 | 14.93 | 32 | 7.61 | 2.7
2008-06-25 | 18:45 | 9.46 | 93.7 | 14.92 | 33 | 7.6 | 3.6
2008-06-25 | 19:00 | 9.46 | 93.6 | 14.89 | 33 | 7.6 | 3.1
2008-06-25 | 19:15 | 9.47 | 93.6 | 14.86 | 33 | 7.6 | 2.9
2008-06-25 | 19:30 | 9.48 | 93.6 | 14.81 | 33 | 7.6 | 2.7
2008-06-25 | 19:45 | 9.5 | 93.6 | 14.72 | 33 | 7.59 | 5
2008-06-25 | 20:00 | 9.52 | 93.6 | 14.62 | 33 | 7.59 | 2.9
2008-06-25 | 20:15 | 9.54 | 93.6 | 14.51 | 33 | 7.59 | 2.7
2008-06-25 | 20:30 | 9.56 | 93.6 | 14.41 | 33 | 7.58 | 2.5
2008-06-25 | 20:45 | 9.58 | 93.6 | 14.3 | 33 | 7.58 | 2.6
2008-06-25 | 21:00 | 9.59 | 93.5 | 14.21 | 33 | 7.58 | 3.4
2008-06-25 | 21:15 | 9.61 | 93.5 | 14.12 | 33 | 7.57 | 3.4
2008-06-25 | 21:30 | 9.63 | 93.5 | 14.03 | 33 | 7.57 | 3.1
2008-06-25 | 21:45 | 9.65 | 93.5 | 13.94 | 33 | 7.57 | 3
2008-06-25 | 22:00 | 9.67 | 93.5 | 13.86 | 34 | 7.56 | 2.7
2008-06-25 | 22:15 | 9.69 | 93.6 | 13.79 | 34 | 7.56 | 2.9
2008-06-25 | 22:30 | 9.71 | 93.6 | 13.68 | 34 | 7.56 | 3.7
2008-06-25 | 22:45 | 9.73 | 93.5 | 13.58 | 34 | 7.56 | 4.5
2008-06-25 | 23:00 | 9.75 | 93.5 | 13.48 | 34 | 7.56 | 3.6
2008-06-25 | 23:15 | 9.78 | 93.6 | 13.36 | 34 | 7.55 | 4.8
2008-06-25 | 23:30 | 9.81 | 93.6 | 13.24 | 34 | 7.55 | 4.8
2008-06-25 | 23:45 | 9.83 | 93.6 | 13.12 | 34 | 7.55 | 6.5
2008-06-26 | 0:00 | 9.87 | 93.6 | 12.97 | 34 | 7.54 | 5.6
2008-06-26 | 0:15 | 9.9 | 93.6 | 12.81 | 33 | 7.54 | 6.2
2008-06-26 | 0:30 | 9.94 | 93.6 | 12.67 | 33 | 7.54 | 6.1
2008-06-26 | 0:45 | 9.97 | 93.6 | 12.52 | 33 | 7.53 | 7
2008-06-26 | 1:01 | 10.01 | 93.6 | 12.37 | 33 | 7.53 | 7.6
2008-06-26 | 1:16 | 10.04 | 93.6 | 12.24 | 33 | 7.53 | 7.8
2008-06-26 | 1:31 | 10.07 | 93.6 | 12.09 | 33 | 7.52 | 10.1
2008-06-26 | 1:46 | 10.1 | 93.6 | 11.96 | 32 | 7.51 | 8.9
2008-06-26 | 2:01 | 10.13 | 93.6 | 11.84 | 32 | 7.51 | 9
2008-06-26 | 2:16 | 10.15 | 93.6 | 11.74 | 32 | 7.51 | 9.8
2008-06-26 | 2:31 | 10.17 | 93.7 | 11.65 | 32 | 7.51 | 10.3
2008-06-26 | 2:46 | 10.2 | 93.7 | 11.55 | 32 | 7.5 | 10.1
2008-06-26 | 3:01 | 10.23 | 93.7 | 11.42 | 31 | 7.49 | 11.2
2008-06-26 | 3:16 | 10.27 | 93.7 | 11.28 | 31 | 7.49 | 9.7
2008-06-26 | 3:31 | 10.3 | 93.7 | 11.13 | 31 | 7.48 | 12.5
2008-06-26 | 3:46 | 10.34 | 93.7 | 10.99 | 31 | 7.49 | 11.7
2008-06-26 | 4:01 | 10.37 | 93.7 | 10.84 | 31 | 7.47 | 10.7
2008-06-26 | 4:16 | 10.4 | 93.6 | 10.7 | 31 | 7.47 | 9.8
2008-06-26 | 4:31 | 10.44 | 93.7 | 10.55 | 31 | 7.47 | 10.8
2008-06-26 | 4:46 | 10.48 | 93.7 | 10.4 | 30 | 7.47 | 12
2008-06-26 | 5:01 | 10.51 | 93.7 | 10.26 | 30 | 7.46 | 10.4
2008-06-26 | 5:16 | 10.54 | 93.7 | 10.12 | 30 | 7.46 | 10.6
2008-06-26 | 5:31 | 10.58 | 93.7 | 9.99 | 30 | 7.45 | 9.5
2008-06-26 | 5:46 | 10.62 | 93.7 | 9.85 | 30 | 7.45 | 10.6
2008-06-26 | 6:01 | 10.66 | 93.8 | 9.72 | 30 | 7.45 | 10.2
2008-06-26 | 6:16 | 10.69 | 93.8 | 9.58 | 30 | 7.45 | 10.1
2008-06-26 | 6:31 | 10.73 | 93.8 | 9.45 | 30 | 7.45 | 11.5
2008-06-26 | 6:46 | 10.77 | 93.8 | 9.31 | 30 | 7.45 | 10.9
2008-06-26 | 7:01 | 10.81 | 93.9 | 9.19 | 30 | 7.45 | 8.7
2008-06-26 | 7:16 | 10.84 | 93.9 | 9.07 | 30 | 7.45 | 9.4
2008-06-26 | 7:31 | 10.87 | 93.9 | 8.96 | 30 | 7.45 | 8.7
2008-06-26 | 7:46 | 10.9 | 94 | 8.85 | 30 | 7.46 | 9
2008-06-26 | 8:01 | 10.94 | 94 | 8.75 | 30 | 7.46 | 8.9
2008-06-26 | 8:16 | 10.97 | 94.1 | 8.67 | 30 | 7.46 | 8.5
2008-06-26 | 8:31 | 11 | 94.2 | 8.59 | 30 | 7.46 | 9
2008-06-26 | 8:46 | 11.02 | 94.3 | 8.54 | 30 | 7.47 | 8.1
2008-06-26 | 9:01 | 11.03 | 94.3 | 8.52 | 30 | 7.47 | 8.3
2008-06-26 | 9:16 | 11.03 | 94.3 | 8.51 | 30 | 7.48 | 8.5
2008-06-26 | 9:31 | 11.04 | 94.5 | 8.55 | 30 | 7.48 | 8.1
2008-06-26 | 9:46 | 11.03 | 94.5 | 8.62 | 30 | 7.49 | 8.7
2008-06-26 | 10:01 | 11.02 | 94.6 | 8.71 | 30 | 7.49 | 9.9
2008-06-26 | 10:16 | 10.99 | 94.7 | 8.82 | 30 | 7.49 | 7.5
2008-06-26 | 10:31 | 10.97 | 94.8 | 8.97 | 30 | 7.5 | 7.3
2008-06-26 | 10:46 | 10.93 | 94.9 | 9.13 | 30 | 7.5 | 9.2
2008-06-26 | 11:01 | 10.89 | 94.9 | 9.33 | 30 | 7.51 | 7
2008-06-26 | 11:16 | 10.85 | 95.1 | 9.54 | 30 | 7.51 | 6.7
2008-06-26 | 11:31 | 10.79 | 95.1 | 9.77 | 30 | 7.51 | 6.5
2008-06-26 | 11:46 | 10.74 | 95.2 | 10.01 | 30 | 7.51 | 8.1
2008-06-26 | 12:01 | 10.68 | 95.2 | 10.25 | 30 | 7.52 | 5
2008-06-26 | 12:16 | 10.63 | 95.3 | 10.51 | 30 | 7.52 | 6.5
2008-06-26 | 12:31 | 10.56 | 95.3 | 10.78 | 30 | 7.52 | 5.4
2008-06-26 | 12:46 | 10.5 | 95.3 | 11.05 | 30 | 7.52 | 5
2008-06-26 | 13:00 | 10.45 | 95.4 | 11.33 | 30 | 7.52 | 5
2008-06-26 | 13:15 | 10.37 | 95.4 | 11.62 | 30 | 7.53 | 4.7
2008-06-26 | 13:30 | 10.31 | 95.5 | 11.9 | 30 | 7.53 | 4.5
2008-06-26 | 13:45 | 10.24 | 95.4 | 12.17 | 30 | 7.56 | 3.9
Data gap for Battery Change
2008-06-26 | 14:15 | 10.13 | 95.4 | 12.66 | 30 | 7.5 | 3.8
2008-06-26 | 14:30 | 10.08 | 95.4 | 12.9 | 30 | 7.52 | 5.7
2008-06-26 | 14:45 | 10.02 | 95.4 | 13.12 | 30 | 7.54 | 5.4
2008-06-26 | 15:00 | 9.97 | 95.3 | 13.35 | 30 | 7.55 | 5.5
2008-06-26 | 15:15 | 9.92 | 95.3 | 13.56 | 30 | 7.55 | 5.9
2008-06-26 | 15:30 | 9.87 | 95.3 | 13.76 | 30 | 7.55 | 5.1
2008-06-26 | 15:45 | 9.83 | 95.2 | 13.95 | 30 | 7.55 | 6.2
2008-06-26 | 16:00 | 9.79 | 95.2 | 14.09 | 31 | 7.56 | 6.2
2008-06-26 | 16:15 | 9.75 | 95 | 14.2 | 31 | 7.56 | 8.1
2008-06-26 | 16:30 | 9.72 | 95 | 14.33 | 31 | 7.57 | 6.5
2008-06-26 | 16:45 | 9.69 | 94.9 | 14.4 | 31 | 7.57 | 7.4
2008-06-26 | 17:00 | 9.68 | 94.7 | 14.39 | 31 | 7.56 | 6.2
2008-06-26 | 17:15 | 9.65 | 94.7 | 14.48 | 31 | 7.57 | 4.6
2008-06-26 | 17:30 | 9.63 | 94.5 | 14.55 | 31 | 7.56 | 3.7
2008-06-26 | 17:45 | 9.61 | 94.5 | 14.59 | 31 | 7.55 | 4
2008-06-26 | 18:00 | 9.6 | 94.4 | 14.61 | 31 | 7.55 | 5
2008-06-26 | 18:15 | 9.59 | 94.3 | 14.61 | 31 | 7.54 | 5
2008-06-26 | 18:30 | 9.58 | 94.2 | 14.61 | 31 | 7.55 | 3.9
2008-06-26 | 18:45 | 9.58 | 94.2 | 14.61 | 31 | 7.55 | 5.4
2008-06-26 | 19:00 | 9.57 | 94.1 | 14.61 | 31 | 7.55 | 4.3
2008-06-26 | 19:15 | 9.57 | 94.1 | 14.59 | 31 | 7.53 | 4.3
2008-06-26 | 19:30 | 9.58 | 94.1 | 14.54 | 31 | 7.53 | 4.5
2008-06-26 | 19:45 | 9.59 | 94.1 | 14.49 | 31 | 7.53 | 6.5
2008-06-26 | 20:00 | 9.6 | 94.1 | 14.44 | 32 | 7.53 | 6.8
2008-06-26 | 20:15 | 9.62 | 94.1 | 14.37 | 32 | 7.52 | 4.8
2008-06-26 | 20:30 | 9.64 | 94.1 | 14.29 | 32 | 7.52 | 4.5
2008-06-26 | 20:45 | 9.64 | 94 | 14.23 | 32 | 7.52 | 4.4
2008-06-26 | 21:00 | 9.65 | 94 | 14.17 | 32 | 7.51 | 5
2008-06-26 | 21:15 | 9.66 | 94 | 14.1 | 32 | 7.5 | 4.4
2008-06-26 | 21:30 | 9.68 | 94 | 14.04 | 32 | 7.5 | 4.3
2008-06-26 | 21:45 | 9.69 | 94 | 13.97 | 32 | 7.49 | 5.9
2008-06-26 | 22:00 | 9.71 | 94 | 13.89 | 32 | 7.49 | 4.7
2008-06-26 | 22:15 | 9.73 | 94 | 13.82 | 32 | 7.49 | 5.6
2008-06-26 | 22:30 | 9.74 | 94 | 13.74 | 32 | 7.49 | 4.7
2008-06-26 | 22:45 | 9.76 | 94 | 13.67 | 32 | 7.48 | 5.2
2008-06-26 | 23:00 | 9.78 | 94 | 13.57 | 32 | 7.48 | 5.8
2008-06-26 | 23:15 | 9.8 | 94 | 13.47 | 33 | 7.48 | 9.5
2008-06-26 | 23:30 | 9.82 | 94 | 13.36 | 33 | 7.48 | 12
2008-06-26 | 23:45 | 9.85 | 94 | 13.23 | 33 | 7.48 | 6.7
2008-06-27 | 0:00 | 9.88 | 94 | 13.09 | 32 | 7.48 | 17.3
2008-06-27 | 0:15 | 9.92 | 94 | 12.94 | 32 | 7.47 | 17.7
2008-06-27 | 0:30 | 9.95 | 94 | 12.81 | 32 | 7.47 | 340.8
2008-06-27 | 0:45 | 9.98 | 94 | 12.67 | 32 | 7.47 | 9.1
2008-06-27 | 1:01 | 10.01 | 94 | 12.54 | 32 | 7.46 | 14.4
2008-06-27 | 1:16 | 10.03 | 94 | 12.4 | 32 | 7.46 | 12.3
2008-06-27 | 1:31 | 10.07 | 94 | 12.27 | 32 | 7.46 | 13.3
2008-06-27 | 1:46 | 10.1 | 94 | 12.15 | 32 | 7.45 | 17.7
2008-06-27 | 2:01 | 10.12 | 94 | 12.03 | 31 | 7.44 | 17.1
2008-06-27 | 2:16 | 10.16 | 94.1 | 11.91 | 31 | 7.44 | 18.9
2008-06-27 | 2:31 | 10.18 | 94 | 11.8 | 31 | 7.44 | 9
2008-06-27 | 2:46 | 10.21 | 94 | 11.68 | 31 | 7.43 | 14.5
2008-06-27 | 3:01 | 10.23 | 94 | 11.57 | 31 | 7.43 | 12.1
2008-06-27 | 3:16 | 10.26 | 94 | 11.46 | 31 | 7.43 | 10.2
2008-06-27 | 3:31 | 10.29 | 94 | 11.34 | 31 | 7.42 | 13.3
2008-06-27 | 3:46 | 10.32 | 94 | 11.21 | 31 | 7.42 | 11.8
2008-06-27 | 4:01 | 10.35 | 94 | 11.08 | 30 | 7.42 | 10.8
2008-06-27 | 4:16 | 10.38 | 94 | 10.94 | 30 | 7.42 | 9.5
2008-06-27 | 4:31 | 10.42 | 94 | 10.8 | 30 | 7.42 | 11.2
2008-06-27 | 4:46 | 10.45 | 94 | 10.65 | 30 | 7.41 | 9
2008-06-27 | 5:01 | 10.48 | 94 | 10.5 | 30 | 7.41 | 10
2008-06-27 | 5:16 | 10.52 | 94 | 10.35 | 30 | 7.4 | 9.4
2008-06-27 | 5:31 | 10.56 | 94 | 10.21 | 30 | 7.4 | 31.1
2008-06-27 | 5:46 | 10.59 | 94 | 10.06 | 30 | 7.4 | 10
2008-06-27 | 6:01 | 10.63 | 94 | 9.91 | 30 | 7.4 | 12.1
2008-06-27 | 6:16 | 10.68 | 94.1 | 9.77 | 30 | 7.39 | 7.5
2008-06-27 | 6:31 | 10.71 | 94.1 | 9.63 | 30 | 7.39 | 10.8
2008-06-27 | 6:46 | 10.75 | 94.1 | 9.5 | 30 | 7.39 | 7.3
2008-06-27 | 7:01 | 10.78 | 94.1 | 9.38 | 30 | 7.39 | 7.1
2008-06-27 | 7:16 | 10.82 | 94.2 | 9.27 | 30 | 7.39 | 7.6
2008-06-27 | 7:31 | 10.85 | 94.2 | 9.16 | 29 | 7.39 | 6.9
2008-06-27 | 7:46 | 10.88 | 94.3 | 9.06 | 30 | 7.39 | 8.1
2008-06-27 | 8:01 | 10.91 | 94.3 | 8.97 | 30 | 7.39 | 8
2008-06-27 | 8:16 | 10.94 | 94.4 | 8.89 | 30 | 7.4 | 7.3
2008-06-27 | 8:31 | 10.96 | 94.4 | 8.82 | 30 | 7.4 | 6.1
2008-06-27 | 8:46 | 10.98 | 94.5 | 8.78 | 30 | 7.41 | 7.3
2008-06-27 | 9:01 | 10.99 | 94.6 | 8.77 | 30 | 7.42 | 7.8
2008-06-27 | 9:16 | 10.99 | 94.6 | 8.8 | 30 | 7.42 | 10.4
2008-06-27 | 9:31 | 10.99 | 94.7 | 8.85 | 30 | 7.43 | 5.6
2008-06-27 | 9:46 | 10.97 | 94.8 | 8.95 | 30 | 7.44 | 8.3
2008-06-27 | 10:01 | 10.95 | 94.9 | 9.07 | 30 | 7.44 | 8.2
2008-06-27 | 10:16 | 10.92 | 94.9 | 9.22 | 30 | 7.45 | 8.5
2008-06-27 | 10:31 | 10.88 | 95 | 9.4 | 30 | 7.46 | 7.3
2008-06-27 | 10:46 | 10.84 | 95.1 | 9.6 | 30 | 7.46 | 5.7
2008-06-27 | 11:01 | 10.79 | 95.2 | 9.82 | 30 | 7.48 | 5.6
2008-06-27 | 11:16 | 10.74 | 95.3 | 10.07 | 30 | 7.49 | 7.3
2008-06-27 | 11:31 | 10.68 | 95.4 | 10.34 | 30 | 7.48 | 7.9
2008-06-27 | 11:46 | 10.62 | 95.5 | 10.62 | 30 | 7.49 | 7
2008-06-27 | 12:01 | 10.56 | 95.5 | 10.91 | 30 | 7.49 | 7.3
2008-06-27 | 12:16 | 10.48 | 95.5 | 11.22 | 30 | 7.49 | 5.5
2008-06-27 | 12:31 | 10.41 | 95.6 | 11.54 | 30 | 7.51 | 6.5
2008-06-27 | 12:46 | 10.34 | 95.7 | 11.87 | 30 | 7.51 | 9.2
2008-06-27 | 13:00 | 10.27 | 95.7 | 12.2 | 30 | 7.53 | 13.2
2008-06-27 | 13:15 | 10.2 | 95.7 | 12.52 | 30 | 7.52 | 15.5
2008-06-27 | 13:30 | 10.12 | 95.8 | 12.85 | 30 | 7.53 | 15.9
2008-06-27 | 13:45 | 10.06 | 95.8 | 13.17 | 30 | 7.54 | 11.2
2008-06-27 | 14:00 | 9.98 | 95.7 | 13.48 | 30 | 7.54 | 8
2008-06-27 | 14:15 | 9.92 | 95.7 | 13.76 | 30 | 7.55 | 9.2
2008-06-27 | 14:30 | 9.86 | 95.7 | 14.04 | 30 | 7.54 | 13.4
2008-06-27 | 14:45 | 9.8 | 95.7 | 14.29 | 30 | 7.55 | 8.3
2008-06-27 | 15:00 | 9.74 | 95.6 | 14.53 | 30 | 7.55 | 12.2
2008-06-27 | 15:15 | 9.69 | 95.5 | 14.74 | 30 | 7.55 | 17.1
2008-06-27 | 15:30 | 9.64 | 95.5 | 14.93 | 30 | 7.56 | 6.2
2008-06-27 | 15:45 | 9.6 | 95.4 | 15.1 | 30 | 7.56 | 6.2
2008-06-27 | 16:00 | 9.56 | 95.3 | 15.25 | 30 | 7.55 | 5.7
2008-06-27 | 16:15 | 9.52 | 95.2 | 15.37 | 30 | 7.56 | 6.2
2008-06-27 | 16:30 | 9.49 | 95.1 | 15.48 | 31 | 7.56 | 4.7
2008-06-27 | 16:45 | 9.46 | 95 | 15.57 | 31 | 7.56 | 4.5
2008-06-27 | 17:00 | 9.43 | 94.8 | 15.64 | 31 | 7.55 | 9.5
2008-06-27 | 17:15 | 9.41 | 94.7 | 15.68 | 31 | 7.55 | 5
2008-06-27 | 17:30 | 9.39 | 94.6 | 15.72 | 31 | 7.55 | 7.6
2008-06-27 | 17:45 | 9.38 | 94.5 | 15.75 | 31 | 7.54 | 8.7
2008-06-27 | 18:00 | 9.37 | 94.4 | 15.77 | 31 | 7.55 | 4.4
2008-06-27 | 18:15 | 9.36 | 94.4 | 15.78 | 31 | 7.54 | 4
2008-06-27 | 18:30 | 9.35 | 94.3 | 15.79 | 31 | 7.54 | 4
2008-06-27 | 18:45 | 9.34 | 94.2 | 15.8 | 31 | 7.54 | 4.5
2008-06-27 | 19:00 | 9.34 | 94.2 | 15.79 | 31 | 7.53 | 3.7
2008-06-27 | 19:15 | 9.34 | 94.2 | 15.77 | 32 | 7.53 | 5.6
2008-06-27 | 19:30 | 9.34 | 94.1 | 15.73 | 32 | 7.53 | 9.2
2008-06-27 | 19:45 | 9.35 | 94.1 | 15.68 | 32 | 7.53 | 6.9
2008-06-27 | 20:00 | 9.36 | 94.1 | 15.62 | 32 | 7.52 | 7.4
2008-06-27 | 20:15 | 9.37 | 94.1 | 15.55 | 32 | 7.52 | 3.8
2008-06-27 | 20:30 | 9.39 | 94.1 | 15.47 | 32 | 7.52 | 3.5
2008-06-27 | 20:45 | 9.4 | 94 | 15.4 | 32 | 7.51 | 4.1
2008-06-27 | 21:00 | 9.42 | 94 | 15.33 | 32 | 7.5 | 5.3
2008-06-27 | 21:15 | 9.42 | 94 | 15.25 | 32 | 7.5 | 3.4
2008-06-27 | 21:30 | 9.44 | 93.9 | 15.18 | 32 | 7.49 | 5.2
2008-06-27 | 21:45 | 9.46 | 94 | 15.11 | 32 | 7.49 | 4.9
2008-06-27 | 22:00 | 9.48 | 94 | 15.03 | 32 | 7.48 | 8.7
2008-06-27 | 22:15 | 9.49 | 94 | 14.95 | 32 | 7.48 | 7.3
2008-06-27 | 22:30 | 9.51 | 94 | 14.85 | 32 | 7.47 | 9.8
2008-06-27 | 22:45 | 9.53 | 94 | 14.75 | 32 | 7.46 | 35.7
2008-06-27 | 23:00 | 9.56 | 94.1 | 14.63 | 32 | 7.46 | 16.6
2008-06-27 | 23:15 | 9.59 | 94 | 14.5 | 32 | 7.46 | 10.6
2008-06-27 | 23:30 | 9.62 | 94.1 | 14.36 | 32 | 7.45 | 8.7
2008-06-27 | 23:45 | 9.64 | 94 | 14.22 | 32 | 7.45 | 9.4
2008-06-28 | 0:00 | 9.68 | 94.1 | 14.07 | 31 | 7.45 | 11.9
2008-06-28 | 0:15 | 9.72 | 94.1 | 13.92 | 31 | 7.45 | 13.4
2008-06-28 | 0:30 | 9.75 | 94.1 | 13.79 | 31 | 7.44 | 16.3
2008-06-28 | 0:45 | 9.77 | 94.1 | 13.65 | 31 | 7.44 | 37
2008-06-28 | 1:01 | 9.8 | 94.1 | 13.52 | 31 | 7.43 | 22
2008-06-28 | 1:16 | 9.83 | 94.1 | 13.39 | 30 | 7.43 | 26.1
2008-06-28 | 1:31 | 9.85 | 94 | 13.27 | 30 | 7.42 | 29.9
2008-06-28 | 1:46 | 9.88 | 94.1 | 13.15 | 30 | 7.43 | 32.7
2008-06-28 | 2:01 | 9.91 | 94.1 | 13.01 | 30 | 7.42 | 30
2008-06-28 | 2:16 | 9.94 | 94.1 | 12.87 | 30 | 7.41 | 23.7
2008-06-28 | 2:31 | 9.98 | 94.1 | 12.73 | 30 | 7.41 | 21.1
2008-06-28 | 2:46 | 10.01 | 94.1 | 12.57 | 29 | 7.41 | 20.9
2008-06-28 | 3:01 | 10.04 | 94 | 12.41 | 29 | 7.4 | 24.4
2008-06-28 | 3:16 | 10.08 | 94 | 12.24 | 29 | 7.4 | 25.7
2008-06-28 | 3:31 | 10.12 | 94.1 | 12.06 | 29 | 7.39 | 35.1
2008-06-28 | 3:46 | 10.17 | 94.1 | 11.88 | 29 | 7.39 | 25.1
2008-06-28 | 4:01 | 10.21 | 94.1 | 11.69 | 29 | 7.38 | 23.1
2008-06-28 | 4:16 | 10.26 | 94.1 | 11.49 | 29 | 7.38 | 25.6
2008-06-28 | 4:31 | 10.3 | 94.1 | 11.31 | 29 | 7.38 | 26
2008-06-28 | 4:46 | 10.34 | 94 | 11.12 | 28 | 7.37 | 27.5
2008-06-28 | 5:01 | 10.39 | 94 | 10.94 | 28 | 7.37 | 25.8
2008-06-28 | 5:16 | 10.43 | 94.1 | 10.76 | 28 | 7.37 | 24.1
2008-06-28 | 5:31 | 10.47 | 94 | 10.58 | 28 | 7.37 | 30.7
2008-06-28 | 5:46 | 10.52 | 94.1 | 10.42 | 28 | 7.37 | 54.8
2008-06-28 | 6:01 | 10.56 | 94.1 | 10.25 | 28 | 7.37 | 63.9
2008-06-28 | 6:16 | 10.6 | 94.1 | 10.08 | 28 | 7.36 | 38.8
2008-06-28 | 6:31 | 10.65 | 94.2 | 9.93 | 28 | 7.36 | 134.1
2008-06-28 | 6:46 | 10.68 | 94.2 | 9.79 | 28 | 7.36 | 67.3
2008-06-28 | 7:01 | 10.72 | 94.2 | 9.65 | 28 | 7.36 | 40.5
2008-06-28 | 7:16 | 10.76 | 94.2 | 9.52 | 28 | 7.36 | 20.8
2008-06-28 | 7:31 | 10.79 | 94.3 | 9.41 | 28 | 7.36 | 21.1
2008-06-28 | 7:46 | 10.82 | 94.3 | 9.3 | 28 | 7.36 | 87
2008-06-28 | 8:01 | 10.86 | 94.4 | 9.21 | 28 | 7.36 | 31.4
2008-06-28 | 8:16 | 10.87 | 94.3 | 9.12 | 28 | 7.36 | 163.6
2008-06-28 | 8:31 | 10.9 | 94.4 | 9.06 | 28 | 7.37 | 24.1
2008-06-28 | 8:46 | 10.92 | 94.5 | 9.03 | 28 | 7.38 | 28.5
2008-06-28 | 9:01 | 10.93 | 94.6 | 9.02 | 28 | 7.38 | 24.6
2008-06-28 | 9:16 | 10.93 | 94.6 | 9.04 | 28 | 7.39 | 20
2008-06-28 | 9:31 | 10.92 | 94.7 | 9.09 | 28 | 7.4 | 19.3
2008-06-28 | 9:46 | 10.91 | 94.8 | 9.18 | 28 | 7.41 | 14.4
2008-06-28 | 10:01 | 10.89 | 94.9 | 9.29 | 28 | 7.42 | 15.4
2008-06-28 | 10:16 | 10.87 | 95 | 9.43 | 28 | 7.42 | 18.8
2008-06-28 | 10:31 | 10.84 | 95.1 | 9.59 | 28 | 7.43 | 16.7
2008-06-28 | 10:46 | 10.79 | 95.1 | 9.78 | 28 | 7.43 | 26.3
2008-06-28 | 11:01 | 10.74 | 95.1 | 9.99 | 28 | 7.45 | 18.3
2008-06-28 | 11:16 | 10.7 | 95.3 | 10.22 | 28 | 7.45 | 17.7
2008-06-28 | 11:31 | 10.64 | 95.3 | 10.47 | 28 | 7.45 | 14.1
2008-06-28 | 11:46 | 10.58 | 95.4 | 10.73 | 28 | 7.45 | 12
2008-06-28 | 12:01 | 10.52 | 95.5 | 11.02 | 28 | 7.46 | 20.3
2008-06-28 | 12:16 | 10.46 | 95.5 | 11.32 | 28 | 7.47 | 17.3
2008-06-28 | 12:31 | 10.39 | 95.6 | 11.62 | 28 | 7.47 | 18.5
2008-06-28 | 12:46 | 10.32 | 95.6 | 11.93 | 28 | 7.47 | 21.3
2008-06-28 | 13:00 | 10.24 | 95.6 | 12.24 | 28 | 7.48 | 16.8
2008-06-28 | 13:15 | 10.18 | 95.7 | 12.55 | 28 | 7.49 | 14
2008-06-28 | 13:30 | 10.11 | 95.7 | 12.87 | 28 | 7.49 | 9.6
2008-06-28 | 13:45 | 10.04 | 95.6 | 13.17 | 29 | 7.5 | 12.6
2008-06-28 | 14:00 | 9.97 | 95.6 | 13.45 | 29 | 7.5 | 10.9
2008-06-28 | 14:15 | 9.92 | 95.6 | 13.73 | 29 | 7.5 | 9.8
2008-06-28 | 14:30 | 9.86 | 95.6 | 13.99 | 29 | 7.5 | 34.7
2008-06-28 | 14:45 | 9.79 | 95.5 | 14.24 | 29 | 7.5 | 11.3
2008-06-28 | 15:00 | 9.74 | 95.5 | 14.48 | 29 | 7.51 | 12.8
2008-06-28 | 15:15 | 9.69 | 95.4 | 14.69 | 29 | 7.51 | 8.9
2008-06-28 | 15:30 | 9.65 | 95.4 | 14.88 | 29 | 7.51 | 11.5
2008-06-28 | 15:45 | 9.6 | 95.3 | 15.05 | 29 | 7.52 | 10.2
2008-06-28 | 16:00 | 9.56 | 95.3 | 15.21 | 29 | 7.51 | 10.9
2008-06-28 | 16:15 | 9.52 | 95.1 | 15.35 | 29 | 7.51 | 10.1
2008-06-28 | 16:30 | 9.49 | 95.1 | 15.47 | 29 | 7.52 | 39
2008-06-28 | 16:45 | 9.46 | 94.9 | 15.57 | 29 | 7.52 | 10.2
2008-06-28 | 17:00 | 9.43 | 94.8 | 15.65 | 29 | 7.51 | 10.3
2008-06-28 | 17:15 | 9.4 | 94.7 | 15.71 | 29 | 7.51 | 11.7
2008-06-28 | 17:30 | 9.38 | 94.5 | 15.77 | 29 | 7.51 | 10
2008-06-28 | 17:45 | 9.36 | 94.4 | 15.81 | 29 | 7.5 | 13.7
2008-06-28 | 18:00 | 9.35 | 94.4 | 15.84 | 29 | 7.5 | 8.9
2008-06-28 | 18:15 | 9.33 | 94.3 | 15.86 | 30 | 7.5 | 10
2008-06-28 | 18:30 | 9.32 | 94.2 | 15.87 | 30 | 7.49 | 10.3
2008-06-28 | 18:45 | 9.32 | 94.2 | 15.88 | 30 | 7.5 | 9
2008-06-28 | 19:00 | 9.32 | 94.1 | 15.88 | 30 | 7.5 | 9.7
2008-06-28 | 19:15 | 9.32 | 94.1 | 15.85 | 30 | 7.48 | 8.1
2008-06-28 | 19:30 | 9.32 | 94.1 | 15.82 | 30 | 7.49 | 8.2
2008-06-28 | 19:45 | 9.34 | 94.1 | 15.78 | 30 | 7.49 | 9.3
2008-06-28 | 20:00 | 9.34 | 94 | 15.73 | 30 | 7.48 | 9.6
2008-06-28 | 20:15 | 9.36 | 94.1 | 15.66 | 30 | 7.49 | 10.3
2008-06-28 | 20:30 | 9.36 | 94.1 | 15.6 | 30 | 7.48 | 10.9
2008-06-28 | 20:45 | 9.37 | 94 | 15.54 | 30 | 7.47 | 10
2008-06-28 | 21:00 | 9.39 | 94 | 15.47 | 30 | 7.46 | 8.9
2008-06-28 | 21:15 | 9.4 | 94 | 15.4 | 30 | 7.46 | 12.3
2008-06-28 | 21:30 | 9.41 | 94 | 15.33 | 30 | 7.46 | 10.9
2008-06-28 | 21:45 | 9.42 | 93.9 | 15.25 | 30 | 7.44 | 17.3
2008-06-28 | 22:00 | 9.45 | 94 | 15.15 | 30 | 7.44 | 12.7
2008-06-28 | 22:15 | 9.47 | 94 | 15.05 | 30 | 7.44 | 20.6
2008-06-28 | 22:30 | 9.49 | 94 | 14.94 | 30 | 7.43 | 16.9
2008-06-28 | 22:45 | 9.52 | 94 | 14.81 | 30 | 7.42 | 24.6
2008-06-28 | 23:00 | 9.54 | 94 | 14.69 | 30 | 7.42 | 35.4
2008-06-28 | 23:15 | 9.57 | 94 | 14.56 | 30 | 7.41 | 26.2
2008-06-28 | 23:30 | 9.6 | 94 | 14.43 | 29 | 7.41 | 30.8
2008-06-28 | 23:45 | 9.63 | 94.1 | 14.31 | 29 | 7.4 | 29.4
2008-06-29 | 0:00 | 9.66 | 94.2 | 14.2 | 29 | 7.4 | 42.9
2008-06-29 | 0:15 | 9.69 | 94.2 | 14.09 | 29 | 7.39 | 34.6
2008-06-29 | 0:30 | 9.7 | 94.1 | 14 | 29 | 7.39 | 41.4
2008-06-29 | 0:45 | 9.73 | 94.2 | 13.9 | 29 | 7.38 | 43
2008-06-29 | 1:01 | 9.75 | 94.2 | 13.8 | 28 | 7.38 | 38.1
2008-06-29 | 1:16 | 9.77 | 94.2 | 13.7 | 28 | 7.37 | 46.3
2008-06-29 | 1:31 | 9.8 | 94.2 | 13.58 | 28 | 7.38 | 52.6
2008-06-29 | 1:46 | 9.83 | 94.2 | 13.45 | 28 | 7.37 | 70.6
2008-06-29 | 2:01 | 9.86 | 94.2 | 13.3 | 28 | 7.37 | 48.7
2008-06-29 | 2:16 | 9.88 | 94.1 | 13.14 | 28 | 7.36 | 42
2008-06-29 | 2:31 | 9.93 | 94.1 | 12.98 | 28 | 7.36 | 50.4
2008-06-29 | 2:46 | 9.97 | 94.2 | 12.8 | 28 | 7.35 | 50.5
2008-06-29 | 3:01 | 10.01 | 94.2 | 12.6 | 27 | 7.35 | 48.8
2008-06-29 | 3:16 | 10.06 | 94.2 | 12.4 | 27 | 7.35 | 45.6
2008-06-29 | 3:31 | 10.1 | 94.2 | 12.21 | 27 | 7.35 | 62.8
2008-06-29 | 3:46 | 10.14 | 94.1 | 12.01 | 27 | 7.34 | 45.4
2008-06-29 | 4:01 | 10.19 | 94.2 | 11.82 | 27 | 7.33 | 50.4
2008-06-29 | 4:16 | 10.23 | 94.1 | 11.64 | 27 | 7.33 | 70.3
2008-06-29 | 4:31 | 10.28 | 94.2 | 11.46 | 27 | 7.33 | 53.4
2008-06-29 | 4:46 | 10.32 | 94.2 | 11.29 | 27 | 7.33 | 51.2
2008-06-29 | 5:01 | 10.35 | 94.2 | 11.12 | 26 | 7.33 | 48.7
2008-06-29 | 5:16 | 10.39 | 94.2 | 10.97 | 26 | 7.33 | 46.3
2008-06-29 | 5:31 | 10.42 | 94.1 | 10.82 | 26 | 7.32 | 48.8
2008-06-29 | 5:46 | 10.46 | 94.1 | 10.67 | 26 | 7.31 | 42.9
2008-06-29 | 6:01 | 10.5 | 94.1 | 10.52 | 26 | 7.32 | 42.2
2008-06-29 | 6:16 | 10.53 | 94.1 | 10.38 | 26 | 7.32 | 42.6
2008-06-29 | 6:31 | 10.56 | 94.1 | 10.25 | 26 | 7.32 | 60.5
2008-06-29 | 6:46 | 10.59 | 94.1 | 10.13 | 26 | 7.32 | 305.6
2008-06-29 | 7:01 | 10.63 | 94.2 | 10.01 | 26 | 7.31 | 40.4
2008-06-29 | 7:16 | 10.66 | 94.2 | 9.9 | 26 | 7.31 | 40.6
2008-06-29 | 7:31 | 10.68 | 94.2 | 9.79 | 26 | 7.31 | 34.6
2008-06-29 | 7:46 | 10.71 | 94.2 | 9.69 | 26 | 7.31 | 62.8
2008-06-29 | 8:01 | 10.74 | 94.2 | 9.59 | 26 | 7.32 | 46.8
2008-06-29 | 8:16 | 10.77 | 94.3 | 9.51 | 26 | 7.32 | 163.5
2008-06-29 | 8:31 | 10.79 | 94.4 | 9.44 | 26 | 7.33 | 63.4
2008-06-29 | 8:46 | 10.81 | 94.4 | 9.4 | 26 | 7.33 | 44
2008-06-29 | 9:01 | 10.81 | 94.4 | 9.38 | 26 | 7.34 | 66
2008-06-29 | 9:16 | 10.82 | 94.4 | 9.4 | 26 | 7.35 | 59.3
2008-06-29 | 9:31 | 10.8 | 94.4 | 9.44 | 26 | 7.36 | 67.6
2008-06-29 | 9:46 | 10.8 | 94.6 | 9.51 | 26 | 7.37 | 41.5
2008-06-29 | 10:01 | 10.79 | 94.7 | 9.6 | 26 | 7.36 | 64
2008-06-29 | 10:16 | 10.77 | 94.7 | 9.72 | 26 | 7.38 | 37.2
2008-06-29 | 10:31 | 10.73 | 94.8 | 9.86 | 26 | 7.38 | 33
2008-06-29 | 10:46 | 10.7 | 94.9 | 10.04 | 26 | 7.39 | 33.6
2008-06-29 | 11:01 | 10.65 | 94.9 | 10.23 | 26 | 7.4 | 51.3
2008-06-29 | 11:16 | 10.61 | 95 | 10.45 | 26 | 7.4 | 63.3
2008-06-29 | 11:31 | 10.55 | 95 | 10.69 | 26 | 7.41 | 364.8
2008-06-29 | 11:46 | 10.49 | 95.1 | 10.96 | 26 | 7.41 | 29.1
2008-06-29 | 12:01 | 10.42 | 95 | 11.23 | 26 | 7.41 | 36.5
2008-06-29 | 12:16 | 10.36 | 95 | 11.51 | 27 | 7.42 | 34.2
2008-06-29 | 12:31 | 10.3 | 95.2 | 11.8 | 27 | 7.43 | 42.1
2008-06-29 | 12:46 | 10.23 | 95.2 | 12.15 | 26 | 7.44 | 59
2008-06-29 | 13:00 | 10.16 | 95.2 | 12.45 | 27 | 7.44 | 33
2008-06-29 | 13:15 | 10.1 | 95.3 | 12.73 | 27 | 7.44 | 102.6
2008-06-29 | 13:30 | 10.04 | 95.2 | 12.97 | 27 | 7.43 | 39.9
2008-06-29 | 13:45 | 9.98 | 95.3 | 13.24 | 27 | 7.44 | 44.5
2008-06-29 | 14:00 | 9.92 | 95.2 | 13.52 | 27 | 7.45 | 43
2008-06-29 | 14:15 | 9.85 | 95.1 | 13.75 | 27 | 7.47 | 25
2008-06-29 | 14:30 | 9.8 | 94.8 | 13.82 | 27 | 7.46 | 32.7
2008-06-29 | 14:45 | 9.78 | 94.5 | 13.79 | 27 | 7.46 | 26.8
2008-06-29 | 15:00 | 9.76 | 94.3 | 13.79 | 27 | 7.44 | 27.9
2008-06-29 | 15:15 | 9.77 | 94.2 | 13.73 | 27 | 7.45 | 36.3
2008-06-29 | 15:30 | 9.78 | 94.2 | 13.65 | 27 | 7.44 | 29.1
2008-06-29 | 15:45 | 9.82 | 94.3 | 13.53 | 27 | 7.44 | 30.7
2008-06-29 | 16:00 | 9.84 | 94.4 | 13.48 | 27 | 7.44 | 29.7
2008-06-29 | 16:15 | 9.84 | 94.4 | 13.49 | 27 | 7.44 | 32.2
2008-06-29 | 16:30 | 9.83 | 94.2 | 13.48 | 27 | 7.44 | 80
2008-06-29 | 16:45 | 9.82 | 94.3 | 13.52 | 27 | 7.44 | 38.8
2008-06-29 | 17:00 | 9.79 | 94.3 | 13.67 | 27 | 7.44 | 35.3
2008-06-29 | 17:15 | 9.76 | 94.3 | 13.77 | 27 | 7.44 | 31.4
2008-06-29 | 17:30 | 9.75 | 94.2 | 13.78 | 27 | 7.44 | 37.9
2008-06-29 | 17:45 | 9.75 | 94.2 | 13.79 | 27 | 7.44 | 26.6
2008-06-29 | 18:00 | 9.76 | 94.3 | 13.83 | 27 | 7.41 | 32.5
2008-06-29 | 18:15 | 9.75 | 94.2 | 13.82 | 27 | 7.42 | 24.4
2008-06-29 | 18:30 | 9.74 | 94.1 | 13.79 | 27 | 7.4 | 34
2008-06-29 | 18:45 | 9.75 | 94.2 | 13.78 | 27 | 7.4 | 49.5
2008-06-29 | 19:00 | 9.75 | 94.2 | 13.81 | 28 | 7.4 | 37.9
2008-06-29 | 19:15 | 9.76 | 94.2 | 13.75 | 28 | 7.39 | 34.7
2008-06-29 | 19:30 | 9.78 | 94.2 | 13.68 | 28 | 7.39 | 546.2
2008-06-29 | 19:45 | 9.78 | 94.2 | 13.64 | 28 | 7.39 | 49.1
2008-06-29 | 20:00 | 9.79 | 94.2 | 13.62 | 28 | 7.39 | 55.3
2008-06-29 | 20:15 | 9.8 | 94.2 | 13.59 | 28 | 7.39 | -1
2008-06-29 | 20:30 | 9.81 | 94.2 | 13.53 | 28 | 7.37 | 99
2008-06-29 | 20:45 | 9.82 | 94.2 | 13.49 | 28 | 7.37 | 390.6
2008-06-29 | 21:00 | 9.82 | 94.2 | 13.46 | 28 | 7.39 | 88
2008-06-29 | 21:15 | 9.85 | 94.2 | 13.37 | 28 | 7.38 | 467.2
2008-06-29 | 21:30 | 9.88 | 94.2 | 13.21 | 28 | 7.37 | 0
2008-06-29 | 21:45 | 9.91 | 94.2 | 13.05 | 28 | 7.35 | 3.1
2008-06-29 | 22:00 | 9.94 | 94.2 | 12.95 | 28 | 7.33 | 39.6
2008-06-29 | 22:15 | 9.96 | 94.3 | 12.87 | 27 | 7.34 | 495.4
2008-06-29 | 22:30 | 9.98 | 94.2 | 12.77 | 27 | 7.34 | 45.5
2008-06-29 | 22:45 | 9.99 | 94.2 | 12.7 | 27 | 7.34 | 307.3
2008-06-29 | 23:00 | 10 | 94.2 | 12.68 | 27 | 7.34 | 52.4
2008-06-29 | 23:15 | 10 | 94.3 | 12.7 | 27 | 7.33 | 76.1
2008-06-29 | 23:30 | 10 | 94.2 | 12.68 | 27 | 7.33 | 47.3
2008-06-29 | 23:45 | 10.02 | 94.3 | 12.62 | 27 | 7.34 | 47.7
2008-06-30 | 0:00 | 10.04 | 94.3 | 12.52 | 27 | 7.34 | 47.6
2008-06-30 | 0:15 | 10.07 | 94.3 | 12.42 | 27 | 7.32 | 72.3
2008-06-30 | 0:30 | 10.09 | 94.3 | 12.32 | 27 | 7.32 | 85.5
2008-06-30 | 0:45 | 10.12 | 94.4 | 12.21 | 27 | 7.32 | 525.2
2008-06-30 | 1:01 | 10.13 | 94.2 | 12.1 | 27 | 7.32 | 411.3
2008-06-30 | 1:16 | 10.16 | 94.2 | 11.98 | 26 | 7.31 | 171.9
2008-06-30 | 1:31 | 10.19 | 94.2 | 11.85 | 26 | 7.31 | 189.7
2008-06-30 | 1:46 | 10.21 | 94.1 | 11.72 | 26 | 7.32 | 122.3
2008-06-30 | 2:01 | 10.25 | 94.1 | 11.58 | 26 | 7.32 | 349.4
2008-06-30 | 2:16 | 10.27 | 94.1 | 11.42 | 26 | 7.31 | 393.4
2008-06-30 | 2:31 | 10.32 | 94.1 | 11.24 | 26 | 7.31 | 59.4
2008-06-30 | 2:46 | 10.36 | 94.1 | 11.05 | 26 | 7.3 | 79.7
2008-06-30 | 3:01 | 10.4 | 94 | 10.87 | 26 | 7.3 | 132.6
2008-06-30 | 3:16 | 10.46 | 94.2 | 10.7 | 26 | 7.3 | 74.9
2008-06-30 | 3:31 | 10.49 | 94.1 | 10.54 | 26 | 7.3 | 168
2008-06-30 | 3:46 | 10.52 | 94.1 | 10.4 | 26 | 7.3 | 97.8
2008-06-30 | 4:01 | 10.55 | 94 | 10.27 | 26 | 7.3 | 122.6
2008-06-30 | 4:16 | 10.57 | 94 | 10.16 | 26 | 7.3 | 61
2008-06-30 | 4:31 | 10.6 | 94 | 10.05 | 25 | 7.29 | 411.4
2008-06-30 | 4:46 | 10.63 | 94 | 9.93 | 26 | 7.29 | 48.3
2008-06-30 | 5:01 | 10.66 | 94 | 9.81 | 25 | 7.29 | 336.4
2008-06-30 | 5:16 | 10.69 | 94 | 9.69 | 25 | 7.28 | 443.8
2008-06-30 | 5:31 | 10.72 | 94 | 9.57 | 26 | 7.28 | 672.4
2008-06-30 | 5:46 | 10.76 | 94.1 | 9.46 | 26 | 7.26 | 252.2
2008-06-30 | 6:01 | 10.77 | 94 | 9.35 | 25 | 7.26 | 49.1
2008-06-30 | 6:16 | 10.81 | 94.1 | 9.24 | 25 | 7.27 | 178.4
2008-06-30 | 6:31 | 10.84 | 94 | 9.14 | 25 | 7.28 | 133.2
2008-06-30 | 6:46 | 10.86 | 94.1 | 9.04 | 26 | 7.26 | 94.4
2008-06-30 | 7:01 | 10.9 | 94.1 | 8.95 | 26 | 7.27 | 457.7
2008-06-30 | 7:16 | 10.92 | 94.1 | 8.86 | 26 | 7.28 | 491.7
2008-06-30 | 7:31 | 10.93 | 94.1 | 8.79 | 26 | 7.28 | 95.1
2008-06-30 | 7:46 | 10.96 | 94.2 | 8.72 | 26 | 7.29 | 261.1
2008-06-30 | 8:01 | 10.98 | 94.2 | 8.66 | 26 | 7.29 | 70.9
2008-06-30 | 8:16 | 11 | 94.2 | 8.59 | 26 | 7.3 | 374.3
2008-06-30 | 8:31 | 11.02 | 94.2 | 8.53 | 26 | 7.3 | 470.9
2008-06-30 | 8:46 | 11.03 | 94.3 | 8.51 | 26 | 7.32 | 90.2
2008-06-30 | 9:01 | 11.03 | 94.3 | 8.52 | 26 | 7.32 | 322.7
2008-06-30 | 9:16 | 11.03 | 94.3 | 8.53 | 26 | 7.32 | 299.5
2008-06-30 | 9:31 | 11.03 | 94.4 | 8.55 | 26 | 7.31 | 610.4
2008-06-30 | 9:46 | 11.02 | 94.4 | 8.61 | 26 | 7.32 | 270.3
2008-06-30 | 10:01 | 11 | 94.5 | 8.73 | 26 | 7.33 | 195.3
2008-06-30 | 10:16 | 10.97 | 94.6 | 8.86 | 26 | 7.34 | 182.3
2008-06-30 | 10:31 | 10.93 | 94.6 | 9.02 | 26 | 7.35 | 353.4
2008-06-30 | 10:46 | 10.89 | 94.7 | 9.2 | 26 | 7.36 | 964.6
2008-06-30 | 11:01 | 10.86 | 94.8 | 9.39 | 26 | 7.37 | 512.7
2008-06-30 | 11:16 | 10.81 | 94.9 | 9.6 | 26 | 7.36 | 62.3
2008-06-30 | 11:31 | 10.76 | 94.9 | 9.83 | 26 | 7.36 | 356.5
2008-06-30 | 11:46 | 10.7 | 95 | 10.06 | 26 | 7.36 | 258.5
2008-06-30 | 12:01 | 10.64 | 95 | 10.3 | 26 | 7.37 | 100.7
2008-06-30 | 12:16 | 10.59 | 95 | 10.56 | 26 | 7.38 | 515.2
2008-06-30 | 12:31 | 10.54 | 95.1 | 10.8 | 26 | 7.4 | 248.7
2008-06-30 | 12:46 | 10.47 | 95.1 | 11.06 | 26 | 7.4 | 1085.7
2008-06-30 | 13:00 | 10.41 | 95.1 | 11.33 | 26 | 7.42 | 429.6
2008-06-30 | 13:15 | 10.35 | 95.1 | 11.6 | 26 | 7.41 | 383.9
2008-06-30 | 13:30 | 10.28 | 95.1 | 11.87 | 26 | 7.41 | 390.8
2008-06-30 | 13:45 | 10.22 | 95.1 | 12.13 | 26 | 7.41 | 239.1
2008-06-30 | 14:00 | 10.17 | 95.2 | 12.38 | 26 | 7.42 | 208.6
2008-06-30 | 14:15 | 10.11 | 95.1 | 12.63 | 26 | 7.43 | 304.4
2008-06-30 | 14:30 | 10.05 | 95.2 | 12.88 | 26 | 7.44 | 748.9
2008-06-30 | 14:45 | 9.99 | 95.1 | 13.13 | 26 | 7.43 | 684.8
2008-06-30 | 15:00 | 9.94 | 95.1 | 13.35 | 26 | 7.43 | 571.3
2008-06-30 | 15:15 | 9.91 | 95.1 | 13.51 | 27 | 7.45 | 1116.8
2008-06-30 | 15:30 | 9.86 | 94.9 | 13.63 | 27 | 7.44 | 628.1
2008-06-30 | 15:45 | 9.83 | 94.8 | 13.73 | 27 | 7.45 | 444.8
2008-06-30 | 16:00 | 9.79 | 94.8 | 13.89 | 27 | 7.44 | -0.5
2008-06-30 | 16:15 | 9.76 | 94.8 | 14.05 | 27 | 7.45 | 775.8
2008-06-30 | 16:30 | 9.73 | 94.7 | 14.17 | 26 | 7.46 | 454.7
2008-06-30 | 16:45 | 9.71 | 94.7 | 14.22 | 27 | 7.46 | 534.6
2008-06-30 | 17:00 | 9.68 | 94.6 | 14.31 | 26 | 7.45 | 27.9
2008-06-30 | 17:15 | 9.65 | 94.5 | 14.44 | 27 | 7.45 | 982.5
2008-06-30 | 17:30 | 9.62 | 94.5 | 14.55 | 27 | 7.45 | 525.1
2008-06-30 | 17:45 | 9.6 | 94.4 | 14.65 | 27 | 7.45 | 161.2
2008-06-30 | 18:00 | 9.57 | 94.4 | 14.76 | 27 | 7.45 | 330.4
2008-06-30 | 18:15 | 9.55 | 94.4 | 14.85 | 27 | 7.46 | 177.3
2008-06-30 | 18:30 | 9.54 | 94.4 | 14.87 | 27 | 7.46 | 302.6
2008-06-30 | 18:45 | 9.54 | 94.4 | 14.88 | 27 | 7.46 | 386.7
2008-06-30 | 19:00 | 9.53 | 94.4 | 14.92 | 27 | 7.46 | 148.8
2008-06-30 | 19:15 | 9.52 | 94.3 | 14.97 | 27 | 7.45 | 258.8
2008-06-30 | 19:30 | 9.51 | 94.4 | 15.02 | 27 | 7.45 | 14.1
2008-06-30 | 19:45 | 9.5 | 94.3 | 15.05 | 27 | 7.45 | 868.9
2008-06-30 | 20:00 | 9.49 | 94.2 | 15.07 | 28 | 7.44 | 312.6
2008-06-30 | 20:15 | 9.48 | 94.2 | 15.07 | 28 | 7.44 | 216.7
2008-06-30 | 20:30 | 9.49 | 94.2 | 15.06 | 28 | 7.44 | 388.9
2008-06-30 | 20:45 | 9.5 | 94.2 | 15.02 | 28 | 7.43 | 464.8
2008-06-30 | 21:00 | 9.5 | 94.2 | 14.98 | 28 | 7.42 | 277.1
2008-06-30 | 21:15 | 9.51 | 94.2 | 14.95 | 28 | 7.42 | 118.8
2008-06-30 | 21:30 | 9.53 | 94.3 | 14.89 | 28 | 7.4 | 77.5
2008-06-30 | 21:45 | 9.54 | 94.2 | 14.82 | 27 | 7.4 | 172.6
2008-06-30 | 22:00 | 9.56 | 94.2 | 14.73 | 27 | 7.39 | 79.5
2008-06-30 | 22:15 | 9.58 | 94.2 | 14.63 | 27 | 7.39 | 204.2
2008-06-30 | 22:30 | 9.61 | 94.3 | 14.52 | 27 | 7.39 | 134.2
2008-06-30 | 22:45 | 9.62 | 94.2 | 14.4 | 27 | 7.39 | 692
2008-06-30 | 23:00 | 9.65 | 94.2 | 14.28 | 27 | 7.38 | 319.5
2008-06-30 | 23:15 | 9.69 | 94.3 | 14.15 | 27 | 7.38 | 178.3
2008-06-30 | 23:30 | 9.71 | 94.2 | 14.02 | 27 | 7.37 | 652.7
2008-06-30 | 23:45 | 9.74 | 94.3 | 13.89 | 27 | 7.37 | 73.7
2008-07-01 | 0:00 | 9.77 | 94.3 | 13.77 | 27 | 7.36 | 584.9
2008-07-01 | 0:15 | 9.8 | 94.4 | 13.64 | 27 | 7.36 | 117.2
2008-07-01 | 0:30 | 9.83 | 94.3 | 13.49 | 26 | 7.36 | 195.3
2008-07-01 | 0:45 | 9.86 | 94.3 | 13.35 | 26 | 7.34 | 367.8
2008-07-01 | 1:01 | 9.9 | 94.4 | 13.2 | 26 | 7.35 | 243.7
2008-07-01 | 1:16 | 9.93 | 94.4 | 13.05 | 26 | 7.35 | 67.6
2008-07-01 | 1:31 | 9.96 | 94.4 | 12.9 | 26 | 7.35 | 62.8
2008-07-01 | 1:46 | 9.99 | 94.3 | 12.76 | 26 | 7.34 | 59
2008-07-01 | 2:01 | 10.03 | 94.4 | 12.62 | 26 | 7.34 | 116.8
2008-07-01 | 2:16 | 10.06 | 94.4 | 12.47 | 26 | 7.33 | 52.9
2008-07-01 | 2:31 | 10.09 | 94.3 | 12.32 | 26 | 7.32 | 65.4
2008-07-01 | 2:46 | 10.13 | 94.3 | 12.16 | 26 | 7.32 | 66.2
2008-07-01 | 3:01 | 10.17 | 94.3 | 11.99 | 26 | 7.31 | 63.5
2008-07-01 | 3:16 | 10.21 | 94.3 | 11.81 | 26 | 7.32 | 48.6
2008-07-01 | 3:31 | 10.24 | 94.2 | 11.61 | 26 | 7.31 | 172.4
2008-07-01 | 3:46 | 10.3 | 94.2 | 11.41 | 25 | 7.31 | 184.6
2008-07-01 | 4:01 | 10.34 | 94.2 | 11.2 | 25 | 7.3 | 187.1
2008-07-01 | 4:16 | 10.4 | 94.2 | 10.98 | 25 | 7.3 | 430.6
2008-07-01 | 4:31 | 10.45 | 94.3 | 10.78 | 25 | 7.3 | 523.3
2008-07-01 | 4:46 | 10.5 | 94.3 | 10.58 | 26 | 7.31 | 392.3
2008-07-01 | 5:01 | 10.55 | 94.3 | 10.39 | 25 | 7.3 | 1051.7
2008-07-01 | 5:16 | 10.59 | 94.3 | 10.21 | 25 | 7.31 | 157.4
2008-07-01 | 5:31 | 10.63 | 94.3 | 10.05 | 25 | 7.3 | 53.7
2008-07-01 | 5:46 | 10.67 | 94.3 | 9.91 | 25 | 7.3 | 143
2008-07-01 | 6:01 | 10.7 | 94.3 | 9.79 | 26 | 7.3 | 77.3
2008-07-01 | 6:16 | 10.73 | 94.3 | 9.68 | 26 | 7.3 | 46.6
2008-07-01 | 6:31 | 10.75 | 94.3 | 9.59 | 26 | 7.31 | 238.2
2008-07-01 | 6:46 | 10.78 | 94.3 | 9.49 | 26 | 7.3 | 63.7
2008-07-01 | 7:01 | 10.8 | 94.3 | 9.4 | 26 | 7.31 | 56.7
2008-07-01 | 7:16 | 10.82 | 94.3 | 9.32 | 26 | 7.3 | 49.6
2008-07-01 | 7:31 | 10.85 | 94.4 | 9.24 | 26 | 7.31 | 64
2008-07-01 | 7:46 | 10.86 | 94.3 | 9.16 | 26 | 7.31 | 46.1
2008-07-01 | 8:01 | 10.88 | 94.4 | 9.09 | 26 | 7.32 | 50.4
2008-07-01 | 8:16 | 10.91 | 94.4 | 9.03 | 26 | 7.32 | 38.2
2008-07-01 | 8:31 | 10.93 | 94.5 | 8.98 | 26 | 7.33 | 50.2
2008-07-01 | 8:46 | 10.93 | 94.5 | 8.96 | 26 | 7.34 | 59.4
2008-07-01 | 9:01 | 10.94 | 94.5 | 8.95 | 26 | 7.34 | 246.8
2008-07-01 | 9:16 | 10.94 | 94.6 | 8.97 | 26 | 7.34 | 186.4
2008-07-01 | 9:31 | 10.93 | 94.6 | 9 | 26 | 7.34 | 149.5
2008-07-01 | 9:46 | 10.92 | 94.7 | 9.08 | 26 | 7.35 | 38.2
2008-07-01 | 10:01 | 10.91 | 94.7 | 9.14 | 26 | 7.35 | 36.1
2008-07-01 | 10:16 | 10.9 | 94.8 | 9.22 | 26 | 7.36 | 88
2008-07-01 | 10:31 | 10.87 | 94.8 | 9.33 | 26 | 7.36 | 40.2
2008-07-01 | 10:46 | 10.85 | 94.9 | 9.45 | 26 | 7.37 | 33.5
2008-07-01 | 11:01 | 10.81 | 94.9 | 9.62 | 26 | 7.37 | 49.6
2008-07-01 | 11:16 | 10.77 | 94.9 | 9.79 | 26 | 7.37 | 33.6
2008-07-01 | 11:31 | 10.73 | 95 | 9.96 | 26 | 7.38 | 194.2
2008-07-01 | 11:46 | 10.69 | 95 | 10.16 | 25 | 7.4 | 53.2
2008-07-01 | 12:01 | 10.65 | 95.1 | 10.36 | 25 | 7.4 | 36.8
2008-07-01 | 12:16 | 10.6 | 95.1 | 10.57 | 25 | 7.4 | 53.8
2008-07-01 | 12:31 | 10.54 | 95.2 | 10.82 | 25 | 7.41 | 42.6
2008-07-01 | 12:46 | 10.49 | 95.2 | 11.05 | 25 | 7.41 | 52.6
2008-07-01 | 13:00 | 10.42 | 95.2 | 11.32 | 25 | 7.42 | 30.9
2008-07-01 | 13:15 | 10.36 | 95.1 | 11.53 | 25 | 7.42 | 270.6
2008-07-01 | 13:30 | 10.31 | 95.1 | 11.7 | 25 | 7.42 | 36.3
2008-07-01 | 13:45 | 10.27 | 95.2 | 11.93 | 25 | 7.42 | 34.3
2008-07-01 | 14:00 | 10.22 | 95 | 12.09 | 25 | 7.42 | 31.4
2008-07-01 | 14:15 | 10.19 | 94.9 | 12.16 | 25 | 7.43 | 25.1
2008-07-01 | 14:30 | 10.15 | 94.8 | 12.28 | 25 | 7.43 | 28.4
2008-07-01 | 14:45 | 10.12 | 94.8 | 12.42 | 25 | 7.43 | 25.4
2008-07-01 | 15:00 | 10.1 | 94.9 | 12.58 | 25 | 7.44 | 33.8
2008-07-01 | 15:15 | 10.06 | 94.8 | 12.7 | 25 | 7.44 | 24.5
2008-07-01 | 15:30 | 10.04 | 94.9 | 12.84 | 25 | 7.44 | 29.6
2008-07-01 | 15:45 | 9.99 | 94.8 | 12.99 | 25 | 7.44 | 26.4
2008-07-01 | 16:00 | 9.96 | 94.7 | 13.1 | 25 | 7.45 | 20.1
2008-07-01 | 16:15 | 9.93 | 94.8 | 13.26 | 25 | 7.44 | 25.4
2008-07-01 | 16:30 | 9.88 | 94.7 | 13.47 | 25 | 7.45 | 21.9
2008-07-01 | 16:45 | 9.86 | 94.7 | 13.55 | 25 | 7.45 | 20.9
2008-07-01 | 17:00 | 9.84 | 94.6 | 13.58 | 25 | 7.44 | 22.2
2008-07-01 | 17:15 | 9.81 | 94.5 | 13.67 | 26 | 7.44 | 21.9
2008-07-01 | 17:30 | 9.79 | 94.5 | 13.79 | 25 | 7.44 | 19.1
2008-07-01 | 17:45 | 9.77 | 94.4 | 13.81 | 25 | 7.44 | 19.4
2008-07-01 | 18:00 | 9.76 | 94.3 | 13.79 | 25 | 7.43 | 22
2008-07-01 | 18:15 | 9.76 | 94.3 | 13.79 | 25 | 7.42 | 27.1
2008-07-01 | 18:30 | 9.76 | 94.3 | 13.78 | 25 | 7.42 | 20.6
2008-07-01 | 18:45 | 9.78 | 94.3 | 13.73 | 25 | 7.42 | 17.9
2008-07-01 | 19:00 | 9.79 | 94.3 | 13.69 | 25 | 7.43 | 20.1
2008-07-01 | 19:15 | 9.8 | 94.3 | 13.65 | 25 | 7.43 | 18.6
2008-07-01 | 19:30 | 9.79 | 94.2 | 13.65 | 25 | 7.43 | 37.5
2008-07-01 | 19:45 | 9.79 | 94.3 | 13.66 | 26 | 7.42 | 19
2008-07-01 | 20:00 | 9.79 | 94.2 | 13.64 | 26 | 7.42 | 85.8
2008-07-01 | 20:15 | 9.8 | 94.2 | 13.59 | 26 | 7.42 | 21.4
2008-07-01 | 20:30 | 9.82 | 94.2 | 13.51 | 26 | 7.42 | 25
2008-07-01 | 20:45 | 9.82 | 94.2 | 13.46 | 25 | 7.42 | 19.5
2008-07-01 | 21:00 | 9.83 | 94.2 | 13.42 | 26 | 7.41 | 23.1
2008-07-01 | 21:15 | 9.85 | 94.2 | 13.35 | 25 | 7.4 | 471.1
2008-07-01 | 21:30 | 9.88 | 94.3 | 13.26 | 25 | 7.4 | 20.5
2008-07-01 | 21:45 | 9.89 | 94.3 | 13.18 | 25 | 7.39 | 24.8
2008-07-01 | 22:00 | 9.9 | 94.2 | 13.1 | 25 | 7.39 | 24.7
2008-07-01 | 22:15 | 9.93 | 94.2 | 13.01 | 25 | 7.38 | 27.2
2008-07-01 | 22:30 | 9.94 | 94.2 | 12.96 | 25 | 7.37 | 30.7
2008-07-01 | 22:45 | 9.95 | 94.2 | 12.94 | 25 | 7.37 | 30.5
2008-07-01 | 23:00 | 9.95 | 94.2 | 12.9 | 25 | 7.37 | 29.2
2008-07-01 | 23:15 | 9.97 | 94.3 | 12.83 | 25 | 7.37 | 23.8
2008-07-01 | 23:30 | 9.99 | 94.3 | 12.77 | 25 | 7.36 | 46
2008-07-01 | 23:45 | 9.99 | 94.3 | 12.72 | 25 | 7.36 | 42.5
2008-07-02 | 0:00 | 10 | 94.2 | 12.67 | 25 | 7.36 | 28.6
2008-07-02 | 0:15 | 10.01 | 94.2 | 12.62 | 25 | 7.35 | 31.8
2008-07-02 | 0:30 | 10.03 | 94.2 | 12.56 | 25 | 7.35 | 36
2008-07-02 | 0:45 | 10.04 | 94.2 | 12.49 | 25 | 7.35 | 35.3
2008-07-02 | 1:01 | 10.07 | 94.3 | 12.41 | 25 | 7.35 | 34.8
2008-07-02 | 1:16 | 10.08 | 94.3 | 12.33 | 24 | 7.35 | 35
2008-07-02 | 1:31 | 10.11 | 94.3 | 12.22 | 24 | 7.34 | 37.2
2008-07-02 | 1:46 | 10.13 | 94.2 | 12.1 | 24 | 7.34 | 32.4
2008-07-02 | 2:01 | 10.16 | 94.2 | 11.97 | 24 | 7.33 | 87.5
2008-07-02 | 2:16 | 10.2 | 94.2 | 11.82 | 24 | 7.33 | 123.9
2008-07-02 | 2:31 | 10.23 | 94.2 | 11.67 | 24 | 7.33 | 46
2008-07-02 | 2:46 | 10.27 | 94.2 | 11.5 | 24 | 7.32 | 39
2008-07-02 | 3:01 | 10.3 | 94.1 | 11.33 | 24 | 7.32 | 34.7
2008-07-02 | 3:16 | 10.34 | 94.2 | 11.17 | 24 | 7.32 | 38.6
2008-07-02 | 3:31 | 10.38 | 94.1 | 11.02 | 24 | 7.32 | 50.2
2008-07-02 | 3:46 | 10.41 | 94.2 | 10.88 | 24 | 7.32 | 34.9
2008-07-02 | 4:01 | 10.44 | 94.2 | 10.76 | 24 | 7.31 | 41.2
2008-07-02 | 4:16 | 10.47 | 94.2 | 10.65 | 23 | 7.32 | 44.3
2008-07-02 | 4:31 | 10.49 | 94.1 | 10.55 | 23 | 7.31 | 44.5
2008-07-02 | 4:46 | 10.53 | 94.2 | 10.44 | 23 | 7.31 | 44.5
2008-07-02 | 5:01 | 10.55 | 94.2 | 10.33 | 23 | 7.31 | 52.1
2008-07-02 | 5:16 | 10.57 | 94.2 | 10.22 | 23 | 7.31 | 43.6
2008-07-02 | 5:31 | 10.6 | 94.2 | 10.12 | 23 | 7.31 | 45.5
2008-07-02 | 5:46 | 10.62 | 94.2 | 10.02 | 23 | 7.3 | 44.6
2008-07-02 | 6:01 | 10.64 | 94.1 | 9.94 | 23 | 7.31 | 33.6
2008-07-02 | 6:16 | 10.67 | 94.2 | 9.85 | 23 | 7.31 | 57.6
2008-07-02 | 6:31 | 10.69 | 94.2 | 9.78 | 23 | 7.31 | 44.7
2008-07-02 | 6:46 | 10.71 | 94.2 | 9.71 | 23 | 7.31 | 60.7
2008-07-02 | 7:01 | 10.73 | 94.2 | 9.64 | 23 | 7.31 | 42.1
2008-07-02 | 7:16 | 10.75 | 94.2 | 9.58 | 23 | 7.32 | 72.6
2008-07-02 | 7:31 | 10.77 | 94.3 | 9.52 | 23 | 7.32 | 36.5
2008-07-02 | 7:46 | 10.79 | 94.3 | 9.45 | 24 | 7.33 | 35.1
2008-07-02 | 8:01 | 10.79 | 94.3 | 9.4 | 24 | 7.33 | 31.9
2008-07-02 | 8:16 | 10.82 | 94.4 | 9.35 | 24 | 7.35 | 136.8
2008-07-02 | 8:31 | 10.83 | 94.4 | 9.32 | 24 | 7.35 | 31.6
2008-07-02 | 8:46 | 10.84 | 94.4 | 9.31 | 24 | 7.35 | 32.4
2008-07-02 | 9:01 | 10.84 | 94.5 | 9.3 | 24 | 7.36 | 31
2008-07-02 | 9:16 | 10.85 | 94.5 | 9.31 | 24 | 7.36 | 37.2
2008-07-02 | 9:31 | 10.85 | 94.6 | 9.32 | 24 | 7.37 | 78.1
2008-07-02 | 9:46 | 10.85 | 94.7 | 9.36 | 24 | 7.38 | 28.5
2008-07-02 | 10:01 | 10.83 | 94.5 | 9.37 | 24 | 7.38 | 33.8
2008-07-02 | 10:16 | 10.82 | 94.5 | 9.4 | 24 | 7.38 | 29.6
2008-07-02 | 10:31 | 10.81 | 94.5 | 9.44 | 24 | 7.39 | 35
2008-07-02 | 10:46 | 10.81 | 94.8 | 9.54 | 24 | 7.39 | 25.1
2008-07-02 | 11:01 | 10.79 | 94.9 | 9.68 | 24 | 7.4 | 27.3
2008-07-02 | 11:16 | 10.76 | 94.9 | 9.83 | 24 | 7.4 | 30.8
2008-07-02 | 11:31 | 10.72 | 95 | 10.01 | 24 | 7.41 | 25.2
2008-07-02 | 11:46 | 10.66 | 94.9 | 10.2 | 24 | 7.41 | 26
2008-07-02 | 12:01 | 10.6 | 95 | 10.51 | 24 | 7.42 | 26.6
2008-07-02 | 12:16 | 10.54 | 95.1 | 10.79 | 24 | 7.43 | 154.4
2008-07-02 | 12:31 | 10.47 | 95.1 | 11.05 | 24 | 7.43 | 78.5
2008-07-02 | 12:46 | 10.41 | 95 | 11.29 | 24 | 7.44 | 327.9
2008-07-02 | 13:00 | 10.37 | 95.1 | 11.49 | 24 | 7.44 | 295.4
2008-07-02 | 13:15 | 10.31 | 95 | 11.67 | 24 | 7.43 | 137.5
2008-07-02 | 13:30 | 10.27 | 95 | 11.87 | 24 | 7.43 | 533.1
2008-07-02 | 13:45 | 10.22 | 95 | 12.09 | 24 | 7.44 | 3.6
2008-07-02 | 14:00 | 10.17 | 95 | 12.32 | 24 | 7.44 | 35
2008-07-02 | 14:15 | 10.12 | 95 | 12.5 | 24 | 7.44 | 125.4
2008-07-02 | 14:30 | 10.08 | 94.7 | 12.57 | 24 | 7.45 | 656.5
2008-07-02 | 14:45 | 10.06 | 94.6 | 12.62 | 24 | 7.45 | 447.4
2008-07-02 | 15:00 | 10.02 | 94.8 | 12.83 | 24 | 7.45 | 16.1
2008-07-02 | 15:15 | 9.95 | 94.8 | 13.15 | 24 | 7.46 | 439.2
2008-07-02 | 15:30 | 9.9 | 94.8 | 13.41 | 24 | 7.47 | 392
2008-07-02 | 15:45 | 9.85 | 94.9 | 13.67 | 24 | 7.46 | 85.4
2008-07-02 | 16:00 | 9.81 | 94.8 | 13.83 | 24 | 7.46 | 797
2008-07-02 | 16:15 | 9.78 | 94.7 | 13.89 | 24 | 7.45 | 517.8
2008-07-02 | 16:30 | 9.76 | 94.6 | 13.92 | 24 | 7.44 | 364.5
2008-07-02 | 16:45 | 9.75 | 94.4 | 13.92 | 25 | 7.45 | 694.8
2008-07-02 | 17:00 | 9.74 | 94.4 | 13.97 | 25 | 7.45 | 321.1
2008-07-02 | 17:15 | 9.74 | 94.4 | 13.97 | 25 | 7.46 | 568.9
2008-07-02 | 17:30 | 9.73 | 94.4 | 13.95 | 25 | 7.45 | 497.5
2008-07-02 | 17:45 | 9.74 | 94.4 | 13.95 | 25 | 7.46 | 102.1
2008-07-02 | 18:00 | 9.74 | 94.3 | 13.94 | 25 | 7.45 | 110.2
2008-07-02 | 18:15 | 9.72 | 94.2 | 13.94 | 25 | 7.44 | 64.6
2008-07-02 | 18:30 | 9.71 | 94.2 | 14 | 25 | 7.45 | 61.4
2008-07-02 | 18:45 | 9.69 | 94.2 | 14.09 | 25 | 7.45 | 38.6
2008-07-02 | 19:00 | 9.69 | 94.3 | 14.13 | 25 | 7.45 | 26.7
2008-07-02 | 19:15 | 9.67 | 94.2 | 14.18 | 25 | 7.45 | 29.6
2008-07-02 | 19:30 | 9.67 | 94.2 | 14.18 | 25 | 7.44 | 23.3
2008-07-02 | 19:45 | 9.68 | 94.2 | 14.14 | 25 | 7.44 | 24.7
2008-07-02 | 20:00 | 9.69 | 94.2 | 14.1 | 26 | 7.43 | 34.2
2008-07-02 | 20:15 | 9.68 | 94.1 | 14.08 | 26 | 7.41 | 27.5
2008-07-02 | 20:30 | 9.69 | 94.1 | 14.05 | 26 | 7.42 | 67.8
2008-07-02 | 20:45 | 9.69 | 94.1 | 14.04 | 26 | 7.42 | 20.5
2008-07-02 | 21:00 | 9.69 | 94.1 | 14.04 | 26 | 7.41 | 62.8
2008-07-02 | 21:15 | 9.69 | 94 | 14.03 | 26 | 7.41 | 38.6
2008-07-02 | 21:30 | 9.68 | 94 | 14.05 | 26 | 7.4 | 19.8
2008-07-02 | 21:45 | 9.69 | 94.1 | 14.03 | 26 | 7.39 | 19.1
2008-07-02 | 22:00 | 9.71 | 94.1 | 13.96 | 26 | 7.39 | 31.9
2008-07-02 | 22:15 | 9.72 | 94.1 | 13.87 | 26 | 7.38 | 42.5
2008-07-02 | 22:30 | 9.74 | 94.1 | 13.8 | 26 | 7.38 | 27.7
2008-07-02 | 22:45 | 9.76 | 94.1 | 13.72 | 25 | 7.38 | 63.9
2008-07-02 | 23:00 | 9.78 | 94.1 | 13.64 | 25 | 7.38 | 49.2
2008-07-02 | 23:15 | 9.79 | 94.1 | 13.59 | 25 | 7.37 | 89
2008-07-02 | 23:30 | 9.8 | 94.1 | 13.52 | 25 | 7.37 | 86.2
2008-07-02 | 23:45 | 9.82 | 94.1 | 13.45 | 25 | 7.36 | 95.3
2008-07-03 | 0:00 | 9.84 | 94.1 | 13.36 | 25 | 7.36 | 103.5
2008-07-03 | 0:15 | 9.86 | 94.1 | 13.26 | 25 | 7.35 | 65.6
2008-07-03 | 0:30 | 9.89 | 94.2 | 13.13 | 25 | 7.35 | 39.2
2008-07-03 | 0:45 | 9.91 | 94.1 | 13.02 | 25 | 7.35 | 77.1
2008-07-03 | 1:01 | 9.92 | 94 | 12.93 | 25 | 7.34 | 91
2008-07-03 | 1:16 | 9.95 | 94.1 | 12.86 | 24 | 7.34 | 57.9
2008-07-03 | 1:31 | 9.97 | 94.1 | 12.77 | 24 | 7.34 | 90.4
2008-07-03 | 1:46 | 9.99 | 94.1 | 12.67 | 24 | 7.33 | 55.5
2008-07-03 | 2:01 | 10.01 | 94.1 | 12.55 | 24 | 7.33 | 52
2008-07-03 | 2:16 | 10.04 | 94 | 12.41 | 24 | 7.32 | 64
2008-07-03 | 2:31 | 10.08 | 94.1 | 12.26 | 24 | 7.32 | 88.3
2008-07-03 | 2:46 | 10.12 | 94.1 | 12.1 | 24 | 7.32 | 90.2
2008-07-03 | 3:01 | 10.16 | 94.2 | 11.95 | 24 | 7.32 | 245.6
2008-07-03 | 3:16 | 10.19 | 94.1 | 11.79 | 24 | 7.32 | 52.2
2008-07-03 | 3:31 | 10.22 | 94.1 | 11.63 | 24 | 7.33 | 46
2008-07-03 | 3:46 | 10.26 | 94.1 | 11.48 | 24 | 7.32 | 37.7
2008-07-03 | 4:01 | 10.29 | 94.1 | 11.34 | 24 | 7.32 | 40.3
2008-07-03 | 4:16 | 10.33 | 94.1 | 11.19 | 24 | 7.32 | 196.7
2008-07-03 | 4:31 | 10.36 | 94 | 11.05 | 24 | 7.31 | 60.6
2008-07-03 | 4:46 | 10.4 | 94 | 10.89 | 24 | 7.3 | 43.8
2008-07-03 | 5:01 | 10.43 | 94.1 | 10.74 | 24 | 7.3 | 68.1
2008-07-03 | 5:16 | 10.47 | 94 | 10.6 | 24 | 7.3 | 42.2
2008-07-03 | 5:31 | 10.5 | 94 | 10.46 | 24 | 7.3 | 46.6
2008-07-03 | 5:46 | 10.53 | 94 | 10.33 | 24 | 7.3 | 38.9
2008-07-03 | 6:01 | 10.56 | 94 | 10.21 | 24 | 7.3 | 48.7
2008-07-03 | 6:16 | 10.59 | 94 | 10.1 | 24 | 7.3 | 36.8
2008-07-03 | 6:31 | 10.62 | 94.1 | 10.01 | 24 | 7.3 | 39.3
2008-07-03 | 6:46 | 10.64 | 94.1 | 9.92 | 24 | 7.3 | 51.1
2008-07-03 | 7:01 | 10.67 | 94.1 | 9.83 | 24 | 7.31 | 88.8
2008-07-03 | 7:16 | 10.69 | 94.2 | 9.75 | 24 | 7.32 | 41.5
2008-07-03 | 7:31 | 10.71 | 94.2 | 9.69 | 24 | 7.32 | 36.3
2008-07-03 | 7:46 | 10.72 | 94.1 | 9.62 | 24 | 7.33 | 44.9
2008-07-03 | 8:01 | 10.75 | 94.2 | 9.56 | 24 | 7.34 | 43.2
2008-07-03 | 8:16 | 10.76 | 94.2 | 9.5 | 24 | 7.34 | 54.5
2008-07-03 | 8:31 | 10.78 | 94.3 | 9.46 | 24 | 7.35 | 178.2
2008-07-03 | 8:46 | 10.79 | 94.3 | 9.43 | 24 | 7.36 | 92.5
2008-07-03 | 9:01 | 10.8 | 94.3 | 9.43 | 24 | 7.37 | 107.4
2008-07-03 | 9:16 | 10.8 | 94.4 | 9.44 | 24 | 7.37 | 118.4
2008-07-03 | 9:31 | 10.79 | 94.4 | 9.46 | 24 | 7.37 | 144.5
2008-07-03 | 9:46 | 10.78 | 94.4 | 9.52 | 24 | 7.38 | 134.8
2008-07-03 | 10:01 | 10.77 | 94.5 | 9.59 | 24 | 7.38 | 396.1
2008-07-03 | 10:16 | 10.75 | 94.5 | 9.65 | 24 | 7.39 | 427.9
2008-07-03 | 10:31 | 10.74 | 94.5 | 9.72 | 24 | 7.39 | 1085.1
2008-07-03 | 10:46 | 10.72 | 94.6 | 9.81 | 24 | 7.39 | 71.7
2008-07-03 | 11:01 | 10.7 | 94.7 | 9.96 | 24 | 7.4 | 369
2008-07-03 | 11:16 | 10.66 | 94.7 | 10.14 | 24 | 7.4 | 323.1
2008-07-03 | 11:31 | 10.62 | 94.8 | 10.28 | 24 | 7.41 | 32.4
2008-07-03 | 11:46 | 10.58 | 94.8 | 10.5 | 24 | 7.42 | 116.8
2008-07-03 | 12:01 | 10.53 | 94.9 | 10.73 | 24 | 7.42 | 24.9
2008-07-03 | 12:16 | 10.47 | 94.8 | 10.93 | 24 | 7.42 | 365.1
2008-07-03 | 12:31 | 10.42 | 94.9 | 11.17 | 24 | 7.42 | 205.8
2008-07-03 | 12:46 | 10.37 | 95 | 11.44 | 24 | 7.44 | -1.1
2008-07-03 | 13:00 | 10.3 | 95 | 11.71 | 24 | 7.45 | 108.5
2008-07-03 | 13:15 | 10.24 | 95 | 11.96 | 24 | 7.44 | 26.1
2008-07-03 | 13:30 | 10.19 | 95 | 12.19 | 24 | 7.43 | 48.4
2008-07-03 | 13:45 | 10.13 | 94.6 | 12.3 | 24 | 7.44 | -1.3
2008-07-03 | 14:00 | 10.1 | 94.6 | 12.4 | 24 | 7.44 | -1.1
2008-07-03 | 14:15 | 10.07 | 94.5 | 12.47 | 24 | 7.43 | 697.5
2008-07-03 | 14:30 | 10.06 | 94.7 | 12.63 | 24 | 7.43 | 36.7
2008-07-03 | 14:45 | 10.03 | 94.4 | 12.63 | 24 | 7.43 | 501.3
2008-07-03 | 15:00 | 10.01 | 94.4 | 12.68 | 24 | 7.43 | 15.6
2008-07-03 | 15:15 | 9.96 | 94.6 | 13.04 | 24 | 7.44 | 0.4
2008-07-03 | 15:30 | 9.88 | 94.6 | 13.43 | 24 | 7.44 | -0.4
2008-07-03 | 15:45 | 9.83 | 94.6 | 13.64 | 24 | 7.45 | 23.6
2008-07-03 | 16:00 | 9.79 | 94.7 | 13.87 | 24 | 7.45 | 56.4
2008-07-03 | 16:15 | 9.74 | 94.6 | 14.06 | 25 | 7.45 | 1121.4
2008-07-03 | 16:30 | 9.7 | 94.6 | 14.23 | 25 | 7.45 | 5.1
2008-07-03 | 16:45 | 9.68 | 94.5 | 14.27 | 25 | 7.46 | 135.2
2008-07-03 | 17:00 | 9.67 | 94.5 | 14.33 | 25 | 7.45 | -0.5
2008-07-03 | 17:15 | 9.64 | 94.4 | 14.45 | 25 | 7.45 | 257.1
2008-07-03 | 17:30 | 9.61 | 94.3 | 14.55 | 25 | 7.46 | 7.2
2008-07-03 | 17:45 | 9.59 | 94.3 | 14.6 | 25 | 7.46 | -0.2
2008-07-03 | 18:00 | 9.58 | 94.2 | 14.64 | 25 | 7.46 | 1126.2
2008-07-03 | 18:15 | 9.57 | 94.2 | 14.66 | 25 | 7.46 | 1112.7
2008-07-03 | 18:30 | 9.56 | 94.2 | 14.69 | 25 | 7.45 | 15.4
2008-07-03 | 18:45 | 9.56 | 94.2 | 14.7 | 25 | 7.46 | 1
2008-07-03 | 19:00 | 9.57 | 94.2 | 14.69 | 25 | 7.46 | 11.7
2008-07-03 | 19:15 | 9.57 | 94.1 | 14.64 | 25 | 7.46 | 1126.2
2008-07-03 | 19:30 | 9.58 | 94.1 | 14.58 | 25 | 7.45 | -1
2008-07-03 | 19:45 | 9.59 | 94.1 | 14.55 | 25 | 7.45 | 13.7
2008-07-03 | 20:00 | 9.59 | 94.1 | 14.52 | 25 | 7.44 | 1101.9
2008-07-03 | 20:15 | 9.58 | 94 | 14.48 | 25 | 7.44 | 28.3
2008-07-03 | 20:30 | 9.6 | 94 | 14.43 | 25 | 7.44 | 5.1
2008-07-03 | 20:45 | 9.61 | 94 | 14.37 | 25 | 7.43 | 5.5
2008-07-03 | 21:00 | 9.62 | 94 | 14.32 | 25 | 7.42 | 44.8
2008-07-03 | 21:15 | 9.64 | 94 | 14.24 | 25 | 7.41 | 518.6
2008-07-03 | 21:30 | 9.67 | 94.1 | 14.11 | 25 | 7.4 | 40.6
2008-07-03 | 21:45 | 9.7 | 94.1 | 13.97 | 25 | 7.41 | 1030.1
2008-07-03 | 22:00 | 9.73 | 94.1 | 13.84 | 25 | 7.4 | 28.6
2008-07-03 | 22:15 | 9.76 | 94.1 | 13.7 | 25 | 7.4 | 1108.2
2008-07-03 | 22:30 | 9.79 | 94.1 | 13.56 | 25 | 7.4 | 934.7
2008-07-03 | 22:45 | 9.81 | 94.1 | 13.45 | 25 | 7.41 | 1
2008-07-03 | 23:00 | 9.84 | 94.2 | 13.35 | 24 | 7.4 | 103.6
2008-07-03 | 23:15 | 9.86 | 94.1 | 13.25 | 24 | 7.41 | 1093.6
2008-07-03 | 23:30 | 9.88 | 94.1 | 13.15 | 24 | 7.4 | 287.7
2008-07-03 | 23:45 | 9.9 | 94.2 | 13.07 | 24 | 7.4 | 5.8
2008-07-04 | 0:00 | 9.92 | 94.2 | 13.01 | 24 | 7.39 | 1111.1
2008-07-04 | 0:15 | 9.94 | 94.2 | 12.95 | 24 | 7.39 | 1112
2008-07-04 | 0:30 | 9.95 | 94.2 | 12.86 | 24 | 7.38 | 1.2
2008-07-04 | 0:45 | 9.98 | 94.2 | 12.76 | 24 | 7.37 | 1110.5
2008-07-04 | 1:01 | 9.99 | 94.1 | 12.67 | 24 | 7.38 | 275.5
2008-07-04 | 1:16 | 10.02 | 94.2 | 12.59 | 24 | 7.37 | 1029.3
2008-07-04 | 1:31 | 10.03 | 94.1 | 12.5 | 24 | 7.36 | 38.5
2008-07-04 | 1:46 | 10.05 | 94.1 | 12.41 | 23 | 7.36 | 178.2
2008-07-04 | 2:01 | 10.07 | 94.1 | 12.33 | 23 | 7.36 | 1039.3
2008-07-04 | 2:16 | 10.09 | 94.1 | 12.25 | 23 | 7.36 | 555.8
2008-07-04 | 2:31 | 10.12 | 94.2 | 12.15 | 23 | 7.35 | 270.5
2008-07-04 | 2:46 | 10.13 | 94.1 | 12.04 | 23 | 7.35 | 115.5
2008-07-04 | 3:01 | 10.16 | 94.1 | 11.93 | 23 | 7.34 | 1103.5
2008-07-04 | 3:16 | 10.2 | 94.2 | 11.8 | 23 | 7.34 | 1102.4
2008-07-04 | 3:31 | 10.23 | 94.2 | 11.66 | 23 | 7.33 | 197.8
2008-07-04 | 3:46 | 10.26 | 94.2 | 11.52 | 23 | 7.33 | 72
2008-07-04 | 4:01 | 10.3 | 94.2 | 11.38 | 23 | 7.33 | 594.6
2008-07-04 | 4:16 | 10.32 | 94.1 | 11.24 | 23 | 7.33 | 142.9
2008-07-04 | 4:31 | 10.35 | 94.1 | 11.11 | 23 | 7.33 | 7.2
2008-07-04 | 4:46 | 10.39 | 94.2 | 10.97 | 23 | 7.32 | 0.9
2008-07-04 | 5:01 | 10.42 | 94.1 | 10.84 | 23 | 7.32 | 20.4
2008-07-04 | 5:16 | 10.46 | 94.2 | 10.71 | 23 | 7.32 | 676.1
2008-07-04 | 5:31 | 10.49 | 94.2 | 10.59 | 23 | 7.31 | 680.1
2008-07-04 | 5:46 | 10.52 | 94.2 | 10.47 | 23 | 7.32 | 583
2008-07-04 | 6:01 | 10.55 | 94.2 | 10.35 | 23 | 7.32 | 513
2008-07-04 | 6:16 | 10.58 | 94.2 | 10.23 | 23 | 7.31 | 1089.2
2008-07-04 | 6:31 | 10.61 | 94.2 | 10.11 | 23 | 7.32 | 514.8
2008-07-04 | 6:46 | 10.63 | 94.2 | 9.99 | 23 | 7.31 | 255.8
2008-07-04 | 7:01 | 10.66 | 94.2 | 9.89 | 23 | 7.32 | 572.6
2008-07-04 | 7:16 | 10.69 | 94.3 | 9.8 | 23 | 7.32 | 592.6
2008-07-04 | 7:31 | 10.72 | 94.3 | 9.7 | 23 | 7.32 | 10.6
2008-07-04 | 7:46 | 10.74 | 94.3 | 9.62 | 23 | 7.32 | 439.6
2008-07-04 | 8:01 | 10.76 | 94.3 | 9.55 | 23 | 7.33 | 421.6
2008-07-04 | 8:16 | 10.78 | 94.4 | 9.49 | 23 | 7.34 | 226.8
2008-07-04 | 8:31 | 10.79 | 94.4 | 9.46 | 23 | 7.35 | 567
2008-07-04 | 8:46 | 10.8 | 94.4 | 9.43 | 23 | 7.35 | 114.9
2008-07-04 | 9:01 | 10.81 | 94.5 | 9.42 | 23 | 7.36 | 90
2008-07-04 | 9:16 | 10.82 | 94.5 | 9.43 | 23 | 7.37 | 200.7
2008-07-04 | 9:31 | 10.82 | 94.6 | 9.44 | 23 | 7.37 | 222
2008-07-04 | 9:46 | 10.81 | 94.5 | 9.46 | 23 | 7.37 | 151.7
2008-07-04 | 10:01 | 10.81 | 94.5 | 9.45 | 23 | 7.37 | 361.3
2008-07-04 | 10:16 | 10.81 | 94.5 | 9.48 | 23 | 7.37 | 477.6
2008-07-04 | 10:31 | 10.8 | 94.6 | 9.54 | 23 | 7.38 | 506.4
2008-07-04 | 10:46 | 10.79 | 94.7 | 9.62 | 23 | 7.38 | 850.9
2008-07-04 | 11:01 | 10.77 | 94.8 | 9.72 | 23 | 7.39 | 504.1
2008-07-04 | 11:16 | 10.74 | 94.9 | 9.92 | 24 | 7.4 | 211.2
2008-07-04 | 11:31 | 10.7 | 95 | 10.11 | 24 | 7.4 | 238.4
2008-07-04 | 11:46 | 10.66 | 95.1 | 10.27 | 24 | 7.4 | 993.7
2008-07-04 | 12:01 | 10.62 | 95.1 | 10.47 | 24 | 7.4 | 613
2008-07-04 | 12:16 | 10.57 | 95.2 | 10.68 | 24 | 7.41 | 606.2
2008-07-04 | 12:31 | 10.52 | 95.2 | 10.91 | 24 | 7.41 | 393.6
2008-07-04 | 12:46 | 10.48 | 95.1 | 11.07 | 24 | 7.41 | -0.2
2008-07-04 | 13:00 | 10.43 | 95.3 | 11.31 | 24 | 7.42 | 56.4
2008-07-04 | 13:15 | 10.37 | 95.4 | 11.63 | 24 | 7.42 | 17.8
2008-07-04 | 13:30 | 10.32 | 95.5 | 11.87 | 24 | 7.42 | 7.1
2008-07-04 | 13:45 | 10.26 | 95.3 | 12.09 | 24 | 7.44 | 1103.2
2008-07-04 | 14:00 | 10.21 | 95.3 | 12.25 | 24 | 7.43 | 0.7
2008-07-04 | 14:15 | 10.16 | 95.2 | 12.44 | 24 | 7.45 | 0.4
2008-07-04 | 14:30 | 10.12 | 95.3 | 12.69 | 24 | 7.44 | 455.2
2008-07-04 | 14:45 | 10.09 | 95.4 | 12.86 | 24 | 7.44 | -1.9
2008-07-04 | 15:00 | 10.01 | 95 | 13.03 | 24 | 7.44 | 4
2008-07-04 | 15:15 | 9.92 | 94.6 | 13.21 | 24 | 7.46 | -2.1
2008-07-04 | 15:30 | 9.95 | 95.2 | 13.37 | 24 | 7.46 | 117.7
2008-07-04 | 15:45 | 9.94 | 95.3 | 13.45 | 24 | 7.46 | -2.1
2008-07-04 | 16:00 | 9.9 | 95.2 | 13.57 | 24 | 7.47 | -1.9
2008-07-04 | 16:15 | 9.86 | 95.1 | 13.73 | 24 | 7.45 | -0.2
2008-07-04 | 16:30 | 9.83 | 95 | 13.82 | 24 | 7.46 | 28.6
2008-07-04 | 16:45 | 9.81 | 94.9 | 13.87 | 24 | 7.45 | 2.5
2008-07-04 | 17:00 | 9.79 | 94.9 | 13.95 | 24 | 7.46 | 37.2
2008-07-04 | 17:15 | 9.73 | 94.6 | 14.06 | 24 | 7.45 | -0.4
2008-07-04 | 17:30 | 9.74 | 94.8 | 14.16 | 24 | 7.47 | -1.9
2008-07-04 | 17:45 | 9.71 | 94.7 | 14.22 | 24 | 7.46 | -2
2008-07-04 | 18:00 | 9.71 | 94.7 | 14.25 | 24 | 7.44 | -1.5
2008-07-04 | 18:15 | 9.7 | 94.7 | 14.25 | 19 | 7.43 | -2.2
2008-07-04 | 18:30 | 9.71 | 94.6 | 14.19 | 21 | 7.41 | -2.1
2008-07-04 | 18:45 | 9.72 | 94.6 | 14.16 | 21 | 7.44 | -2.1
2008-07-04 | 19:00 | 9.71 | 94.5 | 14.11 | 21 | 7.43 | -1.4
2008-07-04 | 19:15 | 9.74 | 94.5 | 14.01 | 21 | 7.42 | 0
2008-07-04 | 19:30 | 9.74 | 94.4 | 13.94 | 21 | 7.43 | -2.2
2008-07-04 | 19:45 | 9.77 | 94.5 | 13.89 | 21 | 7.42 | -2.1
2008-07-04 | 20:00 | 9.77 | 94.4 | 13.82 | 21 | 7.41 | 181.3
2008-07-04 | 20:15 | 9.79 | 94.4 | 13.73 | 21 | 7.4 | 7.2
2008-07-04 | 20:30 | 9.82 | 94.5 | 13.63 | 21 | 7.39 | -2.1
2008-07-04 | 20:45 | 9.83 | 94.5 | 13.55 | 21 | 7.39 | -2
2008-07-04 | 21:00 | 9.84 | 94.5 | 13.5 | 21 | 7.38 | -1.8
2008-07-04 | 21:15 | 9.86 | 94.5 | 13.45 | 21 | 7.4 | -1.8
2008-07-04 | 21:30 | 9.87 | 94.5 | 13.36 | 21 | 7.39 | 8.4
2008-07-04 | 21:45 | 9.88 | 94.3 | 13.26 | 22 | 7.38 | -1.4
2008-07-04 | 22:00 | 9.91 | 94.4 | 13.17 | 22 | 7.39 | 4.5
2008-07-04 | 22:15 | 9.92 | 94.4 | 13.09 | 22 | 7.36 | -0.5
2008-07-04 | 22:30 | 9.96 | 94.5 | 12.98 | 22 | 7.37 | 1.4
2008-07-04 | 22:45 | 9.93 | 93.9 | 12.87 | 21 | 7.38 | -0.3
2008-07-04 | 23:00 | 10.01 | 94.5 | 12.76 | 21 | 7.37 | -0.4
2008-07-04 | 23:15 | 9.84 | 92.6 | 12.65 | 21 | 7.39 | -0.8
2008-07-04 | 23:30 | 10.05 | 94.4 | 12.55 | 21 | 7.37 | -1.7
2008-07-04 | 23:45 | 10.07 | 94.4 | 12.44 | 21 | 7.38 | 4.3
2008-07-05 | 0:00 | 9.95 | 93 | 12.34 | 21 | 7.37 | -1.3
2008-07-05 | 0:15 | 10.12 | 94.5 | 12.25 | 21 | 7.38 | -1.7
2008-07-05 | 0:30 | 10.09 | 94 | 12.16 | 21 | 7.37 | -1.3
2008-07-05 | 0:45 | 10.18 | 94.6 | 12.08 | 21 | 7.36 | -1.3
2008-07-05 | 1:01 | 10.2 | 94.6 | 11.98 | 21 | 7.36 | -0.8
2008-07-05 | 1:16 | 10.1 | 93.4 | 11.85 | 23 | 7.36 | -1.6
2008-07-05 | 1:31 | 10.26 | 94.5 | 11.69 | 23 | 7.37 | -1.6
2008-07-05 | 1:46 | 10.26 | 94.1 | 11.52 | 23 | 7.34 | 0.9
2008-07-05 | 2:01 | 10.34 | 94.5 | 11.33 | 23 | 7.34 | -0.2
2008-07-05 | 2:16 | 10.39 | 94.5 | 11.14 | 23 | 7.35 | -1.1
2008-07-05 | 2:31 | 10.4 | 94.2 | 10.97 | 23 | 7.33 | -1.2
2008-07-05 | 2:46 | 10.46 | 94.5 | 10.82 | 23 | 7.35 | -1.6
2008-07-05 | 3:01 | 10.5 | 94.5 | 10.71 | 23 | 7.35 | -1.6
2008-07-05 | 3:16 | 10.51 | 94.5 | 10.63 | 18 | 7.33 | -1.6
2008-07-05 | 3:31 | 10.44 | 93.8 | 10.57 | 17 | 7.33 | -1.5
2008-07-05 | 3:46 | 10.41 | 93.3 | 10.51 | 17 | 7.32 | -1.4
2008-07-05 | 4:01 | 10.41 | 93.1 | 10.44 | 17 | 7.31 | -1.4
2008-07-05 | 4:16 | 10.57 | 94.5 | 10.36 | 17 | 7.32 | -1.5
2008-07-05 | 4:31 | 10.59 | 94.4 | 10.27 | 17 | 7.31 | -1
2008-07-05 | 4:46 | 10.53 | 93.6 | 10.17 | 17 | 7.3 | 0.1
2008-07-05 | 5:01 | 10.64 | 94.4 | 10.06 | 17 | 7.31 | 0
2008-07-05 | 5:16 | 10.66 | 94.4 | 9.95 | 22 | 7.31 | 0.1
2008-07-05 | 5:31 | 10.71 | 94.5 | 9.83 | 21 | 7.32 | -0.4
2008-07-05 | 5:46 | 10.74 | 94.5 | 9.71 | 22 | 7.31 | -0.6
2008-07-05 | 6:01 | 10.77 | 94.5 | 9.58 | 20 | 7.32 | -0.8
2008-07-05 | 6:16 | 10.8 | 94.4 | 9.46 | 20 | 7.31 | -0.5
2008-07-05 | 6:31 | 10.73 | 93.6 | 9.35 | 21 | 7.3 | -0.3
2008-07-05 | 6:46 | 10.85 | 94.4 | 9.24 | 21 | 7.3 | -0.2
2008-07-05 | 7:01 | 10.88 | 94.5 | 9.14 | 17 | 7.31 | -0.2
2008-07-05 | 7:16 | 10.93 | 94.7 | 9.06 | 16 | 7.3 | -1.2
2008-07-05 | 7:31 | 10.93 | 94.5 | 9 | 22 | 7.32 | 0
2008-07-05 | 7:46 | 10.95 | 94.5 | 8.94 | 22 | 7.31 | -0.1
2008-07-05 | 8:01 | 10.97 | 94.6 | 8.9 | 22 | 7.31 | -0.2
2008-07-05 | 8:16 | 10.97 | 94.6 | 8.88 | 16 | 7.33 | -0.3
2008-07-05 | 8:31 | 10.98 | 94.7 | 8.86 | 16 | 7.32 | -0.4
2008-07-05 | 8:46 | 10.99 | 94.7 | 8.85 | 16 | 7.32 | -0.5
2008-07-05 | 9:01 | 10.99 | 94.8 | 8.87 | 16 | 7.31 | -0.5
2008-07-05 | 9:16 | 10.98 | 94.7 | 8.9 | 16 | 7.32 | -0.5
2008-07-05 | 9:31 | 10.97 | 94.8 | 8.96 | 16 | 7.35 | -0.5
2008-07-05 | 9:46 | 10.96 | 94.9 | 9.05 | 16 | 7.34 | -0.5
2008-07-05 | 10:01 | 10.94 | 94.9 | 9.1 | 16 | 7.34 | -1
2008-07-05 | 10:16 | 10.9 | 94.7 | 9.17 | 16 | 7.35 | -1.1
2008-07-05 | 10:31 | 10.9 | 95 | 9.32 | 16 | 7.35 | -1.1
2008-07-05 | 10:46 | 10.83 | 94.8 | 9.5 | 16 | 7.36 | -1.1
2008-07-05 | 11:01 | 10.82 | 95 | 9.62 | 16 | 7.36 | -1.1
2008-07-05 | 11:16 | 10.7 | 94.5 | 9.87 | 16 | 7.37 | -1.1
2008-07-05 | 11:31 | 10.69 | 94.9 | 10.08 | 16 | 7.37 | -1
2008-07-05 | 11:46 | 10.67 | 94.6 | 10.03 | 17 | 7.37 | -1.1
2008-07-05 | 12:01 | 10.65 | 95.1 | 10.36 | 16 | 7.37 | -1.1
2008-07-05 | 12:16 | 10.43 | 94.3 | 10.85 | 16 | 7.36 | -1.1
2008-07-05 | 12:31 | 10.44 | 94.9 | 11.11 | 16 | 7.37 | -1.1
2008-07-05 | 12:46 | 10.45 | 94.9 | 11.07 | 16 | 7.37 | -1
2008-07-05 | 13:00 | 10.37 | 94.6 | 11.25 | 16 | 7.38 | -0.9
2008-07-05 | 13:15 | 10.38 | 95 | 11.39 | 16 | 7.37 | -0.8
2008-07-05 | 13:30 | 10.31 | 95.2 | 11.77 | 20 | 7.4 | -0.8
2008-07-05 | 13:45 | 10.17 | 94.7 | 12.15 | 17 | 7.39 | -0.8
2008-07-05 | 14:00 | 10.17 | 95.1 | 12.35 | 17 | 7.39 | -0.5
2008-07-05 | 14:15 | 10.14 | 95 | 12.39 | 16 | 7.37 | -1.2
2008-07-05 | 14:30 | 10.1 | 95.1 | 12.64 | 16 | 7.37 | -1.1
2008-07-05 | 14:45 | 10.08 | 95.1 | 12.78 | 16 | 7.37 | -1.1
2008-07-05 | 15:00 | 9.98 | 94.5 | 12.93 | 24 | 7.36 | -1.1
2008-07-05 | 15:15 | 9.96 | 94.7 | 13.09 | 21 | 7.34 | -1.3
2008-07-05 | 15:30 | 9.84 | 93.8 | 13.21 | 21 | 7.36 | -1.4
2008-07-05 | 15:45 | 9.87 | 94.3 | 13.31 | 21 | 7.36 | -1.5
2008-07-05 | 16:00 | 9.85 | 94.1 | 13.32 | 21 | 7.36 | -1.5
2008-07-05 | 16:15 | 9.82 | 94.1 | 13.45 | 21 | 7.37 | -1.4
2008-07-05 | 16:30 | 9.86 | 94.9 | 13.62 | 20 | 7.36 | -1.3
2008-07-05 | 16:45 | 9.85 | 95 | 13.73 | 20 | 7.37 | -1.3
2008-07-05 | 17:00 | 9.79 | 94.9 | 13.94 | 20 | 7.39 | -1.4
2008-07-05 | 17:15 | 9.68 | 94.2 | 14.15 | 16 | 7.39 | -1.5
2008-07-05 | 17:30 | 9.62 | 93.8 | 14.21 | 17 | 7.39 | -1.6
2008-07-05 | 17:45 | 9.68 | 94.4 | 14.21 | 16 | 7.4 | -1.5
2008-07-05 | 18:00 | 9.67 | 94.2 | 14.21 | 16 | 7.41 | -1.6
2008-07-05 | 18:15 | 9.69 | 94.4 | 14.17 | 16 | 7.42 | -1.7
2008-07-05 | 18:30 | 9.7 | 94.4 | 14.16 | 17 | 7.41 | -1.8
2008-07-05 | 18:45 | 9.67 | 94.1 | 14.14 | 17 | 7.39 | -1.7
2008-07-05 | 19:00 | 9.66 | 94.1 | 14.15 | 17 | 7.4 | -1.7
2008-07-05 | 19:15 | 9.62 | 93.7 | 14.21 | 17 | 7.39 | -1.3
2008-07-05 | 19:30 | 9.54 | 93.2 | 14.32 | 17 | 7.39 | -1.1
2008-07-05 | 19:45 | 9.66 | 94.6 | 14.41 | 16 | 7.37 | -1.7
2008-07-05 | 20:00 | 9.64 | 94.3 | 14.4 | 16 | 7.37 | -1.8
2008-07-05 | 20:15 | 9.6 | 93.8 | 14.31 | 16 | 7.38 | -1.7
2008-07-05 | 20:30 | 9.57 | 93.3 | 14.21 | 17 | 7.39 | -1.1
2008-07-05 | 20:45 | 9.69 | 94.2 | 14.12 | 17 | 7.34 | -1.2
2008-07-05 | 21:00 | 9.72 | 94.4 | 14.01 | 17 | 7.37 | -1.3
2008-07-05 | 21:15 | 9.67 | 93.5 | 13.86 | 17 | 7.38 | -1.2
2008-07-05 | 21:30 | 9.74 | 93.9 | 13.73 | 16 | 7.37 | -1.2
2008-07-05 | 21:45 | 9.83 | 94.6 | 13.61 | 16 | 7.37 | -1.3
2008-07-05 | 22:00 | 9.81 | 94.1 | 13.48 | 16 | 7.36 | -1.3
2008-07-05 | 22:15 | 9.85 | 94.2 | 13.36 | 16 | 7.39 | -1.1
2008-07-05 | 22:30 | 9.87 | 94.2 | 13.25 | 16 | 7.34 | -1.1
2008-07-05 | 22:45 | 9.8 | 93.3 | 13.15 | 16 | 7.36 | -1.1
2008-07-05 | 23:00 | 9.96 | 94.6 | 13.06 | 16 | 7.35 | -1.1
2008-07-05 | 23:15 | 9.91 | 93.9 | 12.95 | 16 | 7.35 | -1.1
2008-07-05 | 23:30 | 9.99 | 94.5 | 12.87 | 16 | 7.35 | -0.8
2008-07-05 | 23:45 | 10.01 | 94.6 | 12.81 | 16 | 7.36 | -1
2008-07-06 | 0:00 | 10.03 | 94.6 | 12.75 | 16 | 7.36 | -1
2008-07-06 | 0:15 | 10.04 | 94.6 | 12.7 | 16 | 7.36 | -1.1
2008-07-06 | 0:30 | 10.05 | 94.6 | 12.64 | 16 | 7.36 | -1
2008-07-06 | 0:45 | 10.06 | 94.6 | 12.59 | 16 | 7.33 | -1
2008-07-06 | 1:01 | 10.08 | 94.6 | 12.52 | 16 | 7.32 | -1.1
2008-07-06 | 1:16 | 10.1 | 94.6 | 12.44 | 16 | 7.33 | -1.1
2008-07-06 | 1:31 | 10.12 | 94.6 | 12.32 | 16 | 7.34 | -1.2
2008-07-06 | 1:46 | 10.15 | 94.5 | 12.19 | 16 | 7.33 | -1.6
2008-07-06 | 2:01 | 10.17 | 94.5 | 12.07 | 16 | 7.34 | -1.5
2008-07-06 | 2:16 | 10.21 | 94.6 | 11.91 | 16 | 7.35 | -1.7
2008-07-06 | 2:31 | 10.24 | 94.5 | 11.77 | 16 | 7.33 | -1.6
2008-07-06 | 2:46 | 10.29 | 94.6 | 11.61 | 16 | 7.31 | -1.6
2008-07-06 | 3:01 | 10.32 | 94.6 | 11.45 | 15 | 7.33 | -1.6
2008-07-06 | 3:16 | 10.35 | 94.5 | 11.29 | 15 | 7.32 | -1.6
2008-07-06 | 3:31 | 10.4 | 94.5 | 11.13 | 15 | 7.31 | -1.6
2008-07-06 | 3:46 | 10.44 | 94.5 | 10.96 | 15 | 7.31 | -1.5
2008-07-06 | 4:01 | 10.47 | 94.5 | 10.8 | 15 | 7.32 | -1.6
2008-07-06 | 4:16 | 10.51 | 94.5 | 10.64 | 15 | 7.32 | -1.5
2008-07-06 | 4:31 | 10.55 | 94.5 | 10.48 | 15 | 7.33 | -1.5
2008-07-06 | 4:46 | 10.58 | 94.5 | 10.34 | 15 | 7.33 | -1.4
2008-07-06 | 5:01 | 10.62 | 94.5 | 10.19 | 15 | 7.32 | -1.5
2008-07-06 | 5:16 | 10.64 | 94.3 | 10.04 | 15 | 7.33 | -1.3
2008-07-06 | 5:31 | 10.68 | 94.4 | 9.91 | 15 | 7.32 | -1.5
2008-07-06 | 5:46 | 10.72 | 94.4 | 9.78 | 15 | 7.32 | -1.3
2008-07-06 | 6:01 | 10.75 | 94.5 | 9.66 | 15 | 7.32 | -1.5
2008-07-06 | 6:16 | 10.76 | 94.3 | 9.54 | 16 | 7.33 | -1.5
2008-07-06 | 6:31 | 10.81 | 94.5 | 9.43 | 16 | 7.32 | -1.4
2008-07-06 | 6:46 | 10.82 | 94.4 | 9.32 | 16 | 7.32 | -1.4
2008-07-06 | 7:01 | 10.85 | 94.4 | 9.22 | 16 | 7.32 | -1.1
2008-07-06 | 7:16 | 10.86 | 94.2 | 9.13 | 16 | 7.33 | -1.5
2008-07-06 | 7:31 | 10.88 | 94.2 | 9.05 | 16 | 7.31 | -1.5
2008-07-06 | 7:46 | 10.91 | 94.3 | 8.98 | 16 | 7.32 | -1.6
2008-07-06 | 8:01 | 10.93 | 94.4 | 8.92 | 15 | 7.33 | -1.5
2008-07-06 | 8:16 | 10.97 | 94.6 | 8.9 | 16 | 7.33 | -1.6
2008-07-06 | 8:31 | 10.98 | 94.8 | 8.91 | 16 | 7.33 | -1.5
2008-07-06 | 8:46 | 10.98 | 94.8 | 8.94 | 16 | 7.33 | -1.5
2008-07-06 | 9:01 | 10.96 | 94.8 | 9 | 16 | 7.34 | -1.6
2008-07-06 | 9:16 | 10.95 | 94.8 | 9.04 | 16 | 7.34 | -1.3
2008-07-06 | 9:31 | 10.93 | 94.7 | 9.08 | 16 | 7.34 | -1
2008-07-06 | 9:46 | 10.92 | 94.8 | 9.12 | 16 | 7.34 | -0.5
2008-07-06 | 10:01 | 10.93 | 94.8 | 9.14 | 16 | 7.36 | 0.6
2008-07-06 | 10:16 | 10.9 | 94.8 | 9.22 | 16 | 7.36 | 1
2008-07-06 | 10:31 | 10.87 | 94.8 | 9.32 | 16 | 7.36 | 1.7
2008-07-06 | 10:46 | 10.88 | 95.2 | 9.47 | 16 | 7.36 | 1.8
2008-07-06 | 11:01 | 10.84 | 95.2 | 9.65 | 16 | 7.34 | 2.5
2008-07-06 | 11:16 | 10.74 | 94.8 | 9.82 | 16 | 7.33 | 2.1
2008-07-06 | 11:31 | 10.67 | 94.7 | 10.05 | 16 | 7.35 | 1.7
2008-07-06 | 11:46 | 10.6 | 94.4 | 10.24 | 16 | 7.35 | 0.9
2008-07-06 | 12:01 | 10.57 | 94.6 | 10.43 | 16 | 7.35 | 0.9
2008-07-06 | 12:16 | 10.45 | 94 | 10.66 | 16 | 7.35 | 0.6
2008-07-06 | 12:31 | 10.47 | 94.7 | 10.9 | 16 | 7.36 | 0.5
2008-07-06 | 12:46 | 10.49 | 95.4 | 11.12 | 17 | 7.35 | 0.5
2008-07-06 | 13:00 | 10.32 | 94.4 | 11.38 | 17 | 7.36 | 0.1
2008-07-06 | 13:15 | 10.25 | 94.3 | 11.63 | 17 | 7.37 | 0.1
2008-07-06 | 13:30 | 10.15 | 93.8 | 11.84 | 17 | 7.37 | -0.2
2008-07-06 | 13:45 | 10.14 | 94.2 | 12.09 | 17 | 7.38 | 0.3
2008-07-06 | 14:00 | 9.99 | 93.4 | 12.33 | 17 | 7.38 | 0.3
2008-07-06 | 14:15 | 9.97 | 93.7 | 12.55 | 17 | 7.39 | 0.2
2008-07-06 | 14:30 | 9.98 | 94.2 | 12.78 | 17 | 7.39 | -0.1
2008-07-06 | 14:45 | 9.92 | 94.2 | 13 | 17 | 7.39 | 0.6
2008-07-06 | 15:00 | 9.9 | 94.3 | 13.19 | 17 | 7.39 | 0.8
2008-07-06 | 15:15 | 9.92 | 94.9 | 13.36 | 17 | 7.39 | 0.8
2008-07-06 | 15:30 | 9.92 | 95.2 | 13.53 | 17 | 7.39 | 0.7
2008-07-06 | 15:45 | 9.89 | 95.3 | 13.7 | 17 | 7.39 | 0.6
2008-07-06 | 16:00 | 9.82 | 95 | 13.84 | 17 | 7.39 | 0.4
2008-07-06 | 16:15 | 9.74 | 94.4 | 13.98 | 17 | 7.39 | 0
2008-07-06 | 16:30 | 9.76 | 94.9 | 14.08 | 17 | 7.39 | 0.6
2008-07-06 | 16:45 | 9.75 | 94.9 | 14.16 | 17 | 7.41 | -0.9
2008-07-06 | 17:00 | 9.72 | 94.8 | 14.22 | 17 | 7.41 | -0.8
2008-07-06 | 17:15 | 9.71 | 94.8 | 14.27 | 17 | 7.41 | -0.9
2008-07-06 | 17:30 | 9.66 | 94.4 | 14.32 | 17 | 7.41 | -0.7
2008-07-06 | 17:45 | 9.62 | 94.1 | 14.36 | 17 | 7.41 | -0.7
2008-07-06 | 18:00 | 9.65 | 94.4 | 14.38 | 17 | 7.42 | -1
2008-07-06 | 18:15 | 9.61 | 94 | 14.39 | 17 | 7.42 | -1.1
2008-07-06 | 18:30 | 9.56 | 93.6 | 14.38 | 17 | 7.41 | -0.7
2008-07-06 | 18:45 | 9.63 | 94.2 | 14.37 | 17 | 7.41 | -0.2
2008-07-06 | 19:00 | 9.64 | 94.2 | 14.35 | 17 | 7.41 | 0
2008-07-06 | 19:15 | 9.59 | 93.7 | 14.31 | 17 | 7.41 | -0.4
2008-07-06 | 19:30 | 9.61 | 93.8 | 14.26 | 17 | 7.41 | -0.7
2008-07-06 | 19:45 | 9.69 | 94.4 | 14.21 | 17 | 7.41 | -0.2
2008-07-06 | 20:00 | 9.64 | 93.8 | 14.15 | 17 | 7.41 | 0
2008-07-06 | 20:15 | 9.66 | 93.9 | 14.08 | 17 | 7.41 | -0.4
2008-07-06 | 20:30 | 9.65 | 93.7 | 14.03 | 17 | 7.41 | -0.9
2008-07-06 | 20:45 | 9.7 | 94.1 | 13.97 | 17 | 7.4 | -0.9
2008-07-06 | 21:00 | 9.65 | 93.5 | 13.9 | 17 | 7.4 | -0.7
2008-07-06 | 21:15 | 9.67 | 93.5 | 13.83 | 17 | 7.41 | -0.6
2008-07-06 | 21:30 | 9.7 | 93.6 | 13.76 | 17 | 7.4 | -0.5
2008-07-06 | 21:45 | 9.73 | 93.8 | 13.69 | 17 | 7.4 | -0.7
2008-07-06 | 22:00 | 9.74 | 93.7 | 13.61 | 17 | 7.4 | -0.5
2008-07-06 | 22:15 | 9.76 | 93.7 | 13.52 | 17 | 7.39 | 0.3
2008-07-06 | 22:30 | 9.78 | 93.7 | 13.42 | 17 | 7.39 | 0.1
2008-07-06 | 22:45 | 9.81 | 93.8 | 13.33 | 17 | 7.38 | -0.1
2008-07-06 | 23:00 | 9.85 | 94 | 13.23 | 17 | 7.39 | -0.4
2008-07-06 | 23:15 | 9.87 | 94 | 13.13 | 17 | 7.38 | -0.8
2008-07-06 | 23:30 | 9.89 | 93.9 | 13.04 | 17 | 7.37 | -1.3
2008-07-06 | 23:45 | 9.89 | 93.7 | 12.94 | 17 | 7.37 | -1.1
2008-07-07 | 0:00 | 9.94 | 94 | 12.84 | 17 | 7.37 | -1.2
2008-07-07 | 0:15 | 9.93 | 93.7 | 12.75 | 17 | 7.37 | -1.3
2008-07-07 | 0:30 | 9.93 | 93.6 | 12.67 | 17 | 7.37 | -1
2008-07-07 | 0:45 | 9.9 | 93.1 | 12.58 | 17 | 7.37 | -1
2008-07-07 | 1:01 | 10.04 | 94.2 | 12.49 | 17 | 7.38 | -1
2008-07-07 | 1:16 | 10.03 | 93.9 | 12.39 | 17 | 7.36 | -1
2008-07-07 | 1:31 | 9.9 | 92.4 | 12.29 | 17 | 7.36 | -1.1
2008-07-07 | 1:46 | 10.04 | 93.6 | 12.18 | 16 | 7.36 | -1.3
2008-07-07 | 2:01 | 10.13 | 94.1 | 12.06 | 16 | 7.35 | -1.3
2008-07-07 | 2:16 | 10.1 | 93.6 | 11.95 | 16 | 7.37 | -1.3
2008-07-07 | 2:31 | 10.11 | 93.4 | 11.84 | 16 | 7.36 | -1.4
2008-07-07 | 2:46 | 10.1 | 93.2 | 11.72 | 16 | 7.36 | -1.1
2008-07-07 | 3:01 | 10.15 | 93.3 | 11.59 | 17 | 7.36 | -1
2008-07-07 | 3:16 | 10.11 | 92.7 | 11.47 | 17 | 7.35 | -0.7
2008-07-07 | 3:31 | 10.27 | 93.8 | 11.3 | 17 | 7.36 | -0.4
2008-07-07 | 3:46 | 10.32 | 93.9 | 11.15 | 17 | 7.36 | -0.7
2008-07-07 | 4:01 | 10.33 | 93.6 | 11 | 16 | 7.36 | -0.6
2008-07-07 | 4:16 | 10.28 | 92.9 | 10.85 | 16 | 7.37 | -0.3
2008-07-07 | 4:31 | 10.46 | 94.2 | 10.71 | 16 | 7.37 | -0.5
2008-07-07 | 4:46 | 10.37 | 93.1 | 10.56 | 16 | 7.37 | -0.6
2008-07-07 | 5:01 | 10.43 | 93.3 | 10.42 | 16 | 7.36 | -0.5
2008-07-07 | 5:16 | 10.46 | 93.2 | 10.28 | 16 | 7.36 | -0.8
2008-07-07 | 5:31 | 10.57 | 93.9 | 10.13 | 16 | 7.36 | -0.9
2008-07-07 | 5:46 | 10.55 | 93.5 | 10 | 16 | 7.36 | -0.9
2008-07-07 | 6:01 | 10.63 | 93.9 | 9.87 | 16 | 7.36 | -0.9
2008-07-07 | 6:16 | 10.66 | 93.8 | 9.74 | 16 | 7.36 | -0.7
2008-07-07 | 6:31 | 10.69 | 93.8 | 9.62 | 16 | 7.36 | -1
2008-07-07 | 6:46 | 10.69 | 93.6 | 9.5 | 16 | 7.36 | -1.2
2008-07-07 | 7:01 | 10.69 | 93.4 | 9.41 | 16 | 7.36 | -1.1
2008-07-07 | 7:16 | 10.73 | 93.5 | 9.32 | 16 | 7.36 | -1.1
2008-07-07 | 7:31 | 10.76 | 93.6 | 9.23 | 17 | 7.36 | -1.1
2008-07-07 | 7:46 | 10.72 | 93.1 | 9.16 | 16 | 7.36 | -1.1
2008-07-07 | 8:01 | 10.91 | 94.6 | 9.08 | 16 | 7.37 | -1.2
2008-07-07 | 8:16 | 10.93 | 94.6 | 9.02 | 16 | 7.36 | -1.1
2008-07-07 | 8:31 | 10.87 | 94 | 8.97 | 16 | 7.37 | -1.1
2008-07-07 | 8:46 | 10.89 | 94.1 | 8.93 | 16 | 7.37 | -1.1
2008-07-07 | 9:01 | 10.89 | 94 | 8.92 | 16 | 7.37 | -1.2
2008-07-07 | 9:16 | 10.86 | 93.8 | 8.93 | 16 | 7.37 | -1.1
2008-07-07 | 9:31 | 10.84 | 93.7 | 8.97 | 16 | 7.37 | -1.1
2008-07-07 | 9:46 | 10.87 | 94.1 | 9.03 | 16 | 7.37 | -1.1
2008-07-07 | 10:01 | 10.83 | 94 | 9.12 | 17 | 7.37 | -1.2
2008-07-07 | 10:16 | 10.82 | 94.1 | 9.23 | 17 | 7.38 | -1.3
2008-07-07 | 10:31 | 10.78 | 94 | 9.36 | 17 | 7.38 | -1.3
2008-07-07 | 10:46 | 10.74 | 94.1 | 9.52 | 17 | 7.38 | -1.3
2008-07-07 | 11:01 | 10.8 | 94.9 | 9.68 | 17 | 7.38 | -1.2
2008-07-07 | 11:16 | 10.62 | 93.8 | 9.87 | 17 | 7.38 | -1.4
2008-07-07 | 11:31 | 10.59 | 94 | 10.07 | 16 | 7.38 | -1.2
2008-07-07 | 11:46 | 10.5 | 93.6 | 10.3 | 16 | 7.38 | -1
2008-07-07 | 12:01 | 10.46 | 93.8 | 10.54 | 17 | 7.39 | -1
2008-07-07 | 12:16 | 10.36 | 93.5 | 10.79 | 17 | 7.39 | -1.1
2008-07-07 | 12:31 | 10.37 | 94.1 | 11.06 | 17 | 7.39 | -1.3
2008-07-07 | 12:46 | 10.3 | 94.1 | 11.33 | 17 | 7.39 | -1.3
2008-07-07 | 13:00 | 10.23 | 94.1 | 11.61 | 17 | 7.39 | -1.5
2008-07-07 | 13:15 | 10.09 | 93.3 | 11.89 | 17 | 7.39 | -2.6
2008-07-07 | 13:30 | 10.03 | 93.4 | 12.16 | 17 | 7.39 | -2.6
2008-07-07 | 13:45 | 9.94 | 93.1 | 12.42 | 17 | 7.4 | -2.6
2008-07-07 | 14:00 | 9.81 | 92.5 | 12.69 | 17 | 7.39 | -2.2
2008-07-07 | 14:15 | 9.72 | 92.1 | 12.94 | 17 | 7.39 | -2.6
2008-07-07 | 14:30 | 9.65 | 91.9 | 13.19 | 17 | 7.39 | -2.7
2008-07-07 | 14:45 | 9.6 | 92 | 13.42 | 18 | 7.39 | -2.6
2008-07-07 | 15:00 | 9.92 | 95.6 | 13.65 | 19 | 7.44 | 13.7
2008-07-07 | 15:15 | 9.88 | 95.5 | 13.85 | 30 | 7.45 | 592.8
2008-07-07 | 15:30 | 9.85 | 95.6 | 14.02 | 30 | 7.46 | 593
2008-07-07 | 15:45 | 9.8 | 95.5 | 14.17 | 30 | 7.48 | 647.5
2008-07-07 | 16:00 | 9.77 | 95.5 | 14.31 | 30 | 7.49 | 655
2008-07-07 | 16:15 | 9.74 | 95.4 | 14.43 | 30 | 7.49 | 709.7
2008-07-07 | 16:30 | 9.71 | 95.3 | 14.53 | 30 | 7.49 | 526
2008-07-07 | 16:45 | 9.69 | 95.3 | 14.62 | 30 | 7.5 | 597.4
2008-07-07 | 17:00 | 9.66 | 95.1 | 14.68 | 30 | 7.48 | 599.9
2008-07-07 | 17:15 | 9.64 | 95.1 | 14.74 | 30 | 7.49 | 597.5
2008-07-07 | 17:30 | 9.62 | 95 | 14.78 | 30 | 7.48 | 601.6
2008-07-07 | 17:45 | 9.61 | 94.9 | 14.82 | 30 | 7.49 | 601.7
2008-07-07 | 18:00 | 9.6 | 94.9 | 14.84 | 30 | 7.48 | 598
2008-07-07 | 18:15 | 9.59 | 94.8 | 14.86 | 30 | 7.48 | 596.1
2008-07-07 | 18:30 | 9.58 | 94.8 | 14.87 | 30 | 7.47 | 597.9
2008-07-07 | 18:45 | 9.58 | 94.8 | 14.87 | 30 | 7.48 | 596.5
2008-07-07 | 19:00 | 9.58 | 94.7 | 14.86 | 30 | 7.47 | 601.3
2008-07-07 | 19:15 | 9.58 | 94.7 | 14.84 | 30 | 7.46 | 597.8
2008-07-07 | 19:30 | 9.59 | 94.7 | 14.8 | 30 | 7.46 | 598.3
2008-07-07 | 19:45 | 9.59 | 94.6 | 14.76 | 30 | 7.46 | 595.8
2008-07-07 | 20:00 | 9.61 | 94.7 | 14.71 | 30 | 7.46 | 597.7
2008-07-07 | 20:15 | 9.62 | 94.7 | 14.64 | 30 | 7.46 | 597.1
2008-07-07 | 20:30 | 9.63 | 94.7 | 14.58 | 31 | 7.46 | 604.1
2008-07-07 | 20:45 | 9.65 | 94.7 | 14.52 | 31 | 7.45 | 597.9
2008-07-07 | 21:00 | 9.66 | 94.7 | 14.46 | 31 | 7.44 | 571.5
2008-07-07 | 21:15 | 9.68 | 94.7 | 14.39 | 31 | 7.46 | 655.5
2008-07-07 | 21:30 | 9.69 | 94.7 | 14.32 | 31 | 7.44 | 613.5
2008-07-07 | 21:45 | 9.7 | 94.7 | 14.26 | 31 | 7.44 | 590.3
2008-07-07 | 22:00 | 9.72 | 94.7 | 14.18 | 31 | 7.44 | 587
2008-07-07 | 22:15 | 9.74 | 94.7 | 14.1 | 31 | 7.44 | 587.1
2008-07-07 | 22:30 | 9.76 | 94.7 | 14.01 | 31 | 7.43 | 575.9
2008-07-07 | 22:45 | 9.78 | 94.7 | 13.91 | 31 | 7.43 | 547.6
2008-07-07 | 23:00 | 9.8 | 94.7 | 13.8 | 31 | 7.43 | 560.5
2008-07-07 | 23:15 | 9.83 | 94.7 | 13.68 | 31 | 7.43 | 581.3
2008-07-07 | 23:30 | 9.86 | 94.7 | 13.56 | 31 | 7.43 | 557.6
2008-07-07 | 23:45 | 9.88 | 94.7 | 13.43 | 31 | 7.42 | 551.6
2008-07-08 | 0:00 | 9.92 | 94.7 | 13.29 | 31 | 7.43 | 571.5
2008-07-08 | 0:15 | 9.95 | 94.7 | 13.16 | 31 | 7.42 | 582.7
2008-07-08 | 0:30 | 9.97 | 94.7 | 13.03 | 31 | 7.42 | 533.5
2008-07-08 | 0:45 | 9.99 | 94.7 | 12.92 | 31 | 7.41 | 551.6
2008-07-08 | 1:01 | 10.02 | 94.7 | 12.8 | 31 | 7.41 | 603.4
2008-07-08 | 1:16 | 10.05 | 94.7 | 12.69 | 31 | 7.41 | 354
2008-07-08 | 1:31 | 10.07 | 94.7 | 12.58 | 31 | 7.4 | 281.1
2008-07-08 | 1:46 | 10.1 | 94.7 | 12.48 | 31 | 7.4 | 139.7
2008-07-08 | 2:01 | 10.12 | 94.7 | 12.38 | 31 | 7.4 | 66.7
2008-07-08 | 2:16 | 10.15 | 94.7 | 12.28 | 30 | 7.39 | 24.8
2008-07-08 | 2:31 | 10.16 | 94.7 | 12.17 | 30 | 7.39 | 32.2
2008-07-08 | 2:46 | 10.19 | 94.7 | 12.07 | 30 | 7.39 | 35.1
2008-07-08 | 3:01 | 10.22 | 94.7 | 11.96 | 30 | 7.38 | 232.3
2008-07-08 | 3:16 | 10.24 | 94.7 | 11.85 | 30 | 7.39 | 613
2008-07-08 | 3:31 | 10.27 | 94.7 | 11.73 | 30 | 7.38 | 569.7
2008-07-08 | 3:46 | 10.3 | 94.7 | 11.6 | 30 | 7.39 | 832.6
2008-07-08 | 4:01 | 10.33 | 94.7 | 11.47 | 30 | 7.39 | 976
2008-07-08 | 4:16 | 10.36 | 94.7 | 11.34 | 30 | 7.38 | 1152.1
2008-07-08 | 4:31 | 10.4 | 94.7 | 11.21 | 30 | 7.38 | 1105.1
2008-07-08 | 4:46 | 10.43 | 94.7 | 11.08 | 30 | 7.39 | 1382.6
2008-07-08 | 5:01 | 10.46 | 94.7 | 10.94 | 30 | 7.39 | 1376.7
2008-07-08 | 5:16 | 10.49 | 94.7 | 10.81 | 29 | 7.38 | 1044.1
2008-07-08 | 5:31 | 10.52 | 94.7 | 10.67 | 29 | 7.38 | 1377.3
2008-07-08 | 5:46 | 10.56 | 94.7 | 10.54 | 29 | 7.37 | 1374.3
2008-07-08 | 6:01 | 10.59 | 94.7 | 10.41 | 29 | 7.39 | 1372.6
2008-07-08 | 6:16 | 10.63 | 94.8 | 10.28 | 29 | 7.39 | 1371
2008-07-08 | 6:31 | 10.66 | 94.8 | 10.16 | 29 | 7.38 | 1371.8
2008-07-08 | 6:46 | 10.69 | 94.8 | 10.04 | 29 | 7.38 | 1364.6
2008-07-08 | 7:01 | 10.72 | 94.8 | 9.94 | 29 | 7.4 | 1369.7
2008-07-08 | 7:16 | 10.75 | 94.9 | 9.85 | 29 | 7.38 | 1368.2
2008-07-08 | 7:31 | 10.77 | 94.9 | 9.76 | 29 | 7.39 | 1365.2
2008-07-08 | 7:46 | 10.8 | 95 | 9.69 | 29 | 7.38 | 822.3
2008-07-08 | 8:01 | 10.82 | 95 | 9.61 | 29 | 7.37 | 683.7
2008-07-08 | 8:16 | 10.84 | 95 | 9.55 | 30 | 7.4 | 1365.5
2008-07-08 | 8:31 | 10.86 | 95.1 | 9.49 | 30 | 7.37 | 800.9
2008-07-08 | 8:46 | 10.88 | 95.1 | 9.45 | 30 | 7.39 | 1366.3
2008-07-08 | 9:01 | 10.88 | 95.1 | 9.44 | 30 | 7.4 | 1363.4
2008-07-08 | 9:16 | 10.88 | 95.2 | 9.47 | 30 | 7.4 | 1365.1
2008-07-08 | 9:31 | 10.88 | 95.2 | 9.51 | 30 | 7.42 | 1362.6
2008-07-08 | 9:46 | 10.87 | 95.3 | 9.6 | 30 | 7.42 | 899.3
2008-07-08 | 10:01 | 10.84 | 95.3 | 9.71 | 30 | 7.44 | 1367.6
2008-07-08 | 10:16 | 10.82 | 95.5 | 9.84 | 30 | 7.42 | 564
2008-07-08 | 10:31 | 10.79 | 95.5 | 10 | 30 | 7.43 | 989.7
2008-07-08 | 10:46 | 10.75 | 95.6 | 10.17 | 30 | 7.45 | 1373.1
2008-07-08 | 11:01 | 10.71 | 95.7 | 10.38 | 30 | 7.43 | 576.2
2008-07-08 | 11:16 | 10.66 | 95.8 | 10.6 | 30 | 7.46 | 1366.2
2008-07-08 | 11:31 | 10.61 | 95.8 | 10.85 | 30 | 7.47 | 1375.4
2008-07-08 | 11:46 | 10.54 | 95.9 | 11.12 | 30 | 7.46 | 1382.7
2008-07-08 | 12:01 | 10.48 | 96 | 11.42 | 30 | 7.47 | 1384.5
2008-07-08 | 12:16 | 10.41 | 96 | 11.73 | 30 | 7.47 | 1383.8
2008-07-08 | 12:31 | 10.34 | 96 | 12.04 | 30 | 7.47 | 1391.6
2008-07-08 | 12:46 | 10.26 | 96 | 12.37 | 30 | 7.48 | 1393.6
2008-07-08 | 13:00 | 10.19 | 96.1 | 12.69 | 30 | 7.48 | 239.6
2008-07-08 | 13:15 | 10.12 | 96.1 | 13.01 | 31 | 7.49 | 1403.1
2008-07-08 | 13:30 | 10.06 | 96.2 | 13.32 | 31 | 7.48 | 584.4
2008-07-08 | 13:45 | 9.99 | 96.1 | 13.63 | 31 | 7.49 | 595.3
2008-07-08 | 14:00 | 9.92 | 96.1 | 13.93 | 31 | 7.5 | 1321.9
2008-07-08 | 14:15 | 9.86 | 96.1 | 14.2 | 31 | 7.52 | 1415.9
2008-07-08 | 14:30 | 9.81 | 96.1 | 14.44 | 31 | 7.51 | 100.8
2008-07-08 | 14:45 | 9.75 | 96 | 14.68 | 31 | 7.52 | 1422.7
2008-07-08 | 15:00 | 9.71 | 96 | 14.88 | 32 | 7.52 | 1390.6
2008-07-08 | 15:15 | 9.66 | 95.9 | 15.07 | 32 | 7.51 | 271.6
2008-07-08 | 15:30 | 9.61 | 95.8 | 15.24 | 32 | 7.53 | 709.6
2008-07-08 | 15:45 | 9.57 | 95.8 | 15.39 | 32 | 7.54 | 1425.2
2008-07-08 | 16:00 | 9.54 | 95.7 | 15.52 | 32 | 7.52 | 1234.2
2008-07-08 | 16:15 | 9.51 | 95.6 | 15.63 | 32 | 7.54 | 316.8
2008-07-08 | 16:30 | 9.48 | 95.5 | 15.71 | 32 | 7.52 | 596.4
2008-07-08 | 16:45 | 9.46 | 95.4 | 15.78 | 32 | 7.52 | 598.9
2008-07-08 | 17:00 | 9.44 | 95.3 | 15.82 | 32 | 7.52 | 1137.2
2008-07-08 | 17:15 | 9.43 | 95.2 | 15.85 | 32 | 7.53 | 1435.6
2008-07-08 | 17:30 | 9.41 | 95.1 | 15.88 | 32 | 7.52 | 1431.4
2008-07-08 | 17:45 | 9.41 | 95.1 | 15.89 | 32 | 7.52 | 1435.6
2008-07-08 | 18:00 | 9.39 | 94.9 | 15.89 | 32 | 7.51 | 1430.8
2008-07-08 | 18:15 | 9.39 | 94.9 | 15.89 | 32 | 7.51 | 381
2008-07-08 | 18:30 | 9.38 | 94.8 | 15.88 | 32 | 7.5 | 599.9
2008-07-08 | 18:45 | 9.38 | 94.7 | 15.87 | 32 | 7.49 | 599.6
2008-07-08 | 19:00 | 9.38 | 94.7 | 15.86 | 32 | 7.48 | 361.7
2008-07-08 | 19:15 | 9.38 | 94.7 | 15.82 | 32 | 7.49 | 600.4
2008-07-08 | 19:30 | 9.39 | 94.7 | 15.77 | 32 | 7.48 | 598.4
2008-07-08 | 19:45 | 9.4 | 94.7 | 15.71 | 32 | 7.48 | 604.3
2008-07-08 | 20:00 | 9.42 | 94.7 | 15.64 | 32 | 7.48 | 609.4
2008-07-08 | 20:15 | 9.43 | 94.6 | 15.57 | 32 | 7.48 | 602.1
2008-07-08 | 20:30 | 9.45 | 94.7 | 15.48 | 32 | 7.47 | 600.4
2008-07-08 | 20:45 | 9.46 | 94.6 | 15.4 | 32 | 7.49 | 1431.5
2008-07-08 | 21:00 | 9.47 | 94.6 | 15.32 | 32 | 7.48 | 928.9
2008-07-08 | 21:15 | 9.49 | 94.6 | 15.24 | 32 | 7.45 | 580.9
2008-07-08 | 21:30 | 9.51 | 94.6 | 15.15 | 32 | 7.45 | 600.5
2008-07-08 | 21:45 | 9.53 | 94.6 | 15.07 | 32 | 7.45 | 601.8
2008-07-08 | 22:00 | 9.54 | 94.6 | 14.98 | 32 | 7.45 | 605.1
2008-07-08 | 22:15 | 9.56 | 94.6 | 14.88 | 33 | 7.45 | 281
2008-07-08 | 22:30 | 9.58 | 94.6 | 14.78 | 33 | 7.46 | 84.4
2008-07-08 | 22:45 | 9.61 | 94.6 | 14.67 | 33 | 7.45 | 367.9
2008-07-08 | 23:00 | 9.63 | 94.6 | 14.55 | 33 | 7.46 | 23.4
2008-07-08 | 23:15 | 9.66 | 94.6 | 14.42 | 33 | 7.46 | 1042.5
2008-07-08 | 23:30 | 9.7 | 94.7 | 14.28 | 33 | 7.46 | 326.9
2008-07-08 | 23:45 | 9.72 | 94.6 | 14.15 | 32 | 7.46 | 282.4
2008-07-09 | 0:00 | 9.75 | 94.6 | 14.01 | 32 | 7.44 | 258.5
2008-07-09 | 0:15 | 9.78 | 94.6 | 13.87 | 32 | 7.45 | 432.1
2008-07-09 | 0:30 | 9.81 | 94.7 | 13.73 | 32 | 7.45 | 592.1
2008-07-09 | 0:45 | 9.84 | 94.6 | 13.59 | 32 | 7.45 | 552.3
2008-07-09 | 1:01 | 9.87 | 94.6 | 13.46 | 32 | 7.45 | 393.7
2008-07-09 | 1:16 | 9.89 | 94.6 | 13.35 | 32 | 7.44 | 313.3
2008-07-09 | 1:31 | 9.92 | 94.6 | 13.23 | 32 | 7.44 | 304.7
2008-07-09 | 1:46 | 9.94 | 94.6 | 13.12 | 32 | 7.44 | 303.2
2008-07-09 | 2:01 | 9.97 | 94.6 | 13.02 | 31 | 7.44 | 216.9
2008-07-09 | 2:16 | 9.99 | 94.6 | 12.91 | 31 | 7.44 | 339.6
2008-07-09 | 2:31 | 10.01 | 94.6 | 12.81 | 31 | 7.44 | 270
2008-07-09 | 2:46 | 10.03 | 94.5 | 12.7 | 31 | 7.44 | 526.9
2008-07-09 | 3:01 | 10.05 | 94.5 | 12.6 | 31 | 7.43 | 868.1
2008-07-09 | 3:16 | 10.08 | 94.5 | 12.49 | 31 | 7.44 | 1096.8
2008-07-09 | 3:31 | 10.1 | 94.5 | 12.38 | 31 | 7.43 | 523.4
2008-07-09 | 3:46 | 10.13 | 94.5 | 12.26 | 31 | 7.43 | 841.7
2008-07-09 | 4:01 | 10.15 | 94.5 | 12.14 | 31 | 7.42 | 580.8
2008-07-09 | 4:16 | 10.19 | 94.6 | 12.02 | 31 | 7.42 | 41.6
2008-07-09 | 4:31 | 10.21 | 94.6 | 11.9 | 31 | 7.43 | 541.2
2008-07-09 | 4:46 | 10.24 | 94.5 | 11.77 | 31 | 7.42 | 872.3
2008-07-09 | 5:01 | 10.27 | 94.5 | 11.64 | 31 | 7.41 | 326
2008-07-09 | 5:16 | 10.3 | 94.5 | 11.5 | 30 | 7.42 | 850.3
2008-07-09 | 5:31 | 10.33 | 94.5 | 11.37 | 30 | 7.41 | 797.8
2008-07-09 | 5:46 | 10.37 | 94.5 | 11.23 | 30 | 7.4 | 675.5
2008-07-09 | 6:01 | 10.4 | 94.5 | 11.09 | 30 | 7.41 | 414.4
2008-07-09 | 6:16 | 10.42 | 94.4 | 10.96 | 30 | 7.4 | 314.7
2008-07-09 | 6:31 | 10.46 | 94.5 | 10.84 | 30 | 7.4 | -0.6
2008-07-09 | 6:46 | 10.5 | 94.6 | 10.72 | 30 | 7.41 | 858.7
2008-07-09 | 7:01 | 10.52 | 94.6 | 10.62 | 30 | 7.41 | 540.9
2008-07-09 | 7:16 | 10.55 | 94.6 | 10.51 | 30 | 7.4 | 121.2
2008-07-09 | 7:31 | 10.58 | 94.6 | 10.42 | 30 | 7.41 | 216.6
2008-07-09 | 7:46 | 10.6 | 94.7 | 10.34 | 30 | 7.4 | 76.1
2008-07-09 | 8:01 | 10.63 | 94.7 | 10.26 | 30 | 7.41 | 133.3
2008-07-09 | 8:16 | 10.65 | 94.7 | 10.19 | 30 | 7.41 | -12.7
2008-07-09 | 8:31 | 10.67 | 94.8 | 10.13 | 30 | 7.42 | 555.5
2008-07-09 | 8:46 | 10.69 | 94.9 | 10.08 | 30 | 7.42 | -13.8
2008-07-09 | 9:01 | 10.69 | 94.8 | 10.06 | 31 | 7.42 | 135.3
2008-07-09 | 9:16 | 10.69 | 94.9 | 10.08 | 31 | 7.43 | 393.4
2008-07-09 | 9:31 | 10.69 | 95 | 10.13 | 31 | 7.44 | 488.4
2008-07-09 | 9:46 | 10.67 | 95 | 10.21 | 31 | 7.44 | -21.8
2008-07-09 | 10:01 | 10.65 | 95.1 | 10.33 | 31 | 7.44 | 243.5
2008-07-09 | 10:16 | 10.63 | 95.2 | 10.47 | 31 | 7.46 | 783
2008-07-09 | 10:31 | 10.6 | 95.3 | 10.64 | 31 | 7.46 | 167.3
2008-07-09 | 10:46 | 10.56 | 95.4 | 10.83 | 31 | 7.47 | 454.1
2008-07-09 | 11:01 | 10.51 | 95.4 | 11.04 | 31 | 7.47 | 144.3
2008-07-09 | 11:16 | 10.46 | 95.5 | 11.27 | 31 | 7.48 | 553.4
2008-07-09 | 11:31 | 10.41 | 95.5 | 11.52 | 31 | 7.48 | -14.3
2008-07-09 | 11:46 | 10.35 | 95.5 | 11.79 | 31 | 7.49 | 288.6
2008-07-09 | 12:01 | 10.28 | 95.5 | 12.06 | 31 | 7.49 | 1010.4
2008-07-09 | 12:16 | 10.23 | 95.7 | 12.35 | 31 | 7.5 | 487.1
2008-07-09 | 12:31 | 10.16 | 95.7 | 12.65 | 31 | 7.49 | 175
2008-07-09 | 12:46 | 10.1 | 95.7 | 12.96 | 31 | 7.49 | 455.3
2008-07-09 | 13:00 | 10.03 | 95.8 | 13.28 | 31 | 7.5 | 695.5
2008-07-09 | 13:15 | 9.96 | 95.8 | 13.59 | 31 | 7.51 | 170.6
2008-07-09 | 13:30 | 9.89 | 95.7 | 13.89 | 31 | 7.51 | 454.6
2008-07-09 | 13:45 | 9.83 | 95.8 | 14.19 | 31 | 7.51 | 18.6
2008-07-09 | 14:00 | 9.76 | 95.7 | 14.47 | 31 | 7.51 | 40.8
2008-07-09 | 14:15 | 9.7 | 95.7 | 14.74 | 31 | 7.52 | 477.9
2008-07-09 | 14:30 | 9.65 | 95.7 | 15 | 31 | 7.52 | 327.2
2008-07-09 | 14:45 | 9.59 | 95.6 | 15.24 | 32 | 7.53 | 287.2
2008-07-09 | 15:00 | 9.54 | 95.6 | 15.46 | 32 | 7.53 | 406
2008-07-09 | 15:15 | 9.5 | 95.5 | 15.67 | 32 | 7.53 | 287.5
2008-07-09 | 15:30 | 9.45 | 95.4 | 15.86 | 32 | 7.53 | 664.5
2008-07-09 | 15:45 | 9.41 | 95.4 | 16.02 | 32 | 7.53 | 551.8
2008-07-09 | 16:00 | 9.37 | 95.3 | 16.17 | 32 | 7.53 | 324.2
2008-07-09 | 16:15 | 9.34 | 95.2 | 16.29 | 32 | 7.53 | 299.4
2008-07-09 | 16:30 | 9.31 | 95.1 | 16.4 | 32 | 7.52 | 597.9
2008-07-09 | 16:45 | 9.28 | 95 | 16.48 | 32 | 7.53 | 254.5
2008-07-09 | 17:00 | 9.26 | 94.9 | 16.55 | 32 | 7.53 | 261.1
2008-07-09 | 17:15 | 9.24 | 94.8 | 16.59 | 32 | 7.53 | 105.5
2008-07-09 | 17:30 | 9.22 | 94.7 | 16.62 | 32 | 7.52 | 225.7
2008-07-09 | 17:45 | 9.22 | 94.7 | 16.65 | 32 | 7.53 | 205.8
2008-07-09 | 18:00 | 9.2 | 94.6 | 16.67 | 32 | 7.51 | 229.1
2008-07-09 | 18:15 | 9.19 | 94.5 | 16.68 | 32 | 7.51 | 16.6
2008-07-09 | 18:30 | 9.19 | 94.4 | 16.69 | 32 | 7.51 | 282.3
2008-07-09 | 18:45 | 9.19 | 94.4 | 16.69 | 32 | 7.51 | 17.1
2008-07-09 | 19:00 | 9.18 | 94.3 | 16.69 | 32 | 7.51 | 748.6
2008-07-09 | 19:15 | 9.18 | 94.3 | 16.66 | 32 | 7.5 | 8.4
2008-07-09 | 19:30 | 9.2 | 94.4 | 16.61 | 32 | 7.5 | 177.6
2008-07-09 | 19:45 | 9.2 | 94.3 | 16.55 | 32 | 7.51 | 215.1
2008-07-09 | 20:00 | 9.22 | 94.3 | 16.48 | 33 | 7.5 | 279.7
2008-07-09 | 20:15 | 9.24 | 94.4 | 16.4 | 33 | 7.5 | 190.6
2008-07-09 | 20:30 | 9.25 | 94.3 | 16.3 | 33 | 7.49 | 264.2
2008-07-09 | 20:45 | 9.27 | 94.3 | 16.21 | 33 | 7.49 | 216.4
2008-07-09 | 21:00 | 9.28 | 94.3 | 16.13 | 33 | 7.48 | 301.6
2008-07-09 | 21:15 | 9.3 | 94.2 | 16.03 | 33 | 7.48 | 442.2
2008-07-09 | 21:30 | 9.32 | 94.3 | 15.94 | 33 | 7.48 | 236
2008-07-09 | 21:45 | 9.33 | 94.2 | 15.85 | 33 | 7.47 | 281.8
2008-07-09 | 22:00 | 9.35 | 94.2 | 15.76 | 33 | 7.47 | 200.5
2008-07-09 | 22:15 | 9.37 | 94.3 | 15.66 | 33 | 7.47 | 361.2
2008-07-09 | 22:30 | 9.39 | 94.2 | 15.56 | 33 | 7.46 | 384
2008-07-09 | 22:45 | 9.41 | 94.2 | 15.44 | 33 | 7.46 | 411.9
2008-07-09 | 23:00 | 9.44 | 94.3 | 15.33 | 33 | 7.46 | 380.5
2008-07-09 | 23:15 | 9.47 | 94.3 | 15.2 | 33 | 7.45 | 376.2
2008-07-09 | 23:30 | 9.49 | 94.3 | 15.07 | 33 | 7.45 | 354.7
2008-07-09 | 23:45 | 9.52 | 94.3 | 14.94 | 33 | 7.45 | 320.4
2008-07-10 | 0:00 | 9.55 | 94.3 | 14.8 | 33 | 7.46 | 354.6
2008-07-10 | 0:15 | 9.58 | 94.3 | 14.65 | 33 | 7.45 | 237.6
2008-07-10 | 0:30 | 9.61 | 94.3 | 14.51 | 32 | 7.45 | 8
2008-07-10 | 0:45 | 9.64 | 94.3 | 14.37 | 32 | 7.45 | 1.2
2008-07-10 | 1:01 | 9.67 | 94.3 | 14.23 | 32 | 7.45 | 129
2008-07-10 | 1:16 | 9.7 | 94.3 | 14.09 | 32 | 7.44 | 136.4
2008-07-10 | 1:31 | 9.73 | 94.3 | 13.95 | 32 | 7.44 | 85.3
2008-07-10 | 1:46 | 9.76 | 94.3 | 13.82 | 32 | 7.44 | 160.8
2008-07-10 | 2:01 | 9.78 | 94.3 | 13.7 | 31 | 7.43 | 125.3
2008-07-10 | 2:16 | 9.8 | 94.2 | 13.58 | 31 | 7.42 | 87.8
2008-07-10 | 2:31 | 9.83 | 94.2 | 13.48 | 31 | 7.42 | 150.5
2008-07-10 | 2:46 | 9.85 | 94.3 | 13.37 | 31 | 7.42 | 283.6
2008-07-10 | 3:01 | 9.88 | 94.3 | 13.26 | 31 | 7.41 | 182.4
2008-07-10 | 3:16 | 9.9 | 94.3 | 13.15 | 31 | 7.41 | 150
2008-07-10 | 3:31 | 9.93 | 94.3 | 13.03 | 31 | 7.41 | 152.1
2008-07-10 | 3:46 | 9.95 | 94.2 | 12.9 | 30 | 7.41 | 280.3
2008-07-10 | 4:01 | 9.98 | 94.2 | 12.76 | 30 | 7.4 | 291.8
2008-07-10 | 4:16 | 10.01 | 94.2 | 12.62 | 30 | 7.4 | 371.7
2008-07-10 | 4:31 | 10.05 | 94.2 | 12.48 | 30 | 7.39 | 263
2008-07-10 | 4:46 | 10.08 | 94.2 | 12.34 | 30 | 7.39 | 192.6
2008-07-10 | 5:01 | 10.11 | 94.2 | 12.2 | 30 | 7.38 | 136.9
2008-07-10 | 5:16 | 10.15 | 94.3 | 12.06 | 30 | 7.39 | 145.1
2008-07-10 | 5:31 | 10.17 | 94.2 | 11.92 | 30 | 7.39 | 53.8
2008-07-10 | 5:46 | 10.21 | 94.3 | 11.77 | 30 | 7.39 | 53.6
2008-07-10 | 6:01 | 10.25 | 94.3 | 11.62 | 30 | 7.38 | 7.6
2008-07-10 | 6:16 | 10.28 | 94.2 | 11.48 | 30 | 7.37 | 7.5
2008-07-10 | 6:31 | 10.32 | 94.3 | 11.35 | 29 | 7.38 | 123.5
2008-07-10 | 6:46 | 10.34 | 94.3 | 11.22 | 29 | 7.37 | 224.1
2008-07-10 | 7:01 | 10.38 | 94.4 | 11.1 | 29 | 7.38 | 215.5
2008-07-10 | 7:16 | 10.4 | 94.3 | 10.99 | 29 | 7.38 | 235.7
2008-07-10 | 7:31 | 10.43 | 94.4 | 10.89 | 29 | 7.38 | 219.7
2008-07-10 | 7:46 | 10.47 | 94.4 | 10.79 | 29 | 7.38 | 211.9
2008-07-10 | 8:01 | 10.49 | 94.5 | 10.7 | 29 | 7.38 | 280.8
2008-07-10 | 8:16 | 10.52 | 94.5 | 10.61 | 29 | 7.38 | 287.3
2008-07-10 | 8:31 | 10.54 | 94.6 | 10.54 | 29 | 7.39 | 250.9
2008-07-10 | 8:46 | 10.56 | 94.6 | 10.49 | 29 | 7.39 | 260.8
2008-07-10 | 9:01 | 10.57 | 94.7 | 10.47 | 29 | 7.4 | 247.5
2008-07-10 | 9:16 | 10.58 | 94.8 | 10.48 | 29 | 7.4 | 250.5
2008-07-10 | 9:31 | 10.57 | 94.8 | 10.52 | 30 | 7.4 | 249.5
2008-07-10 | 9:46 | 10.55 | 94.8 | 10.6 | 30 | 7.4 | 232.7
2008-07-10 | 10:01 | 10.54 | 95 | 10.71 | 30 | 7.42 | 240.1
2008-07-10 | 10:16 | 10.51 | 95 | 10.84 | 30 | 7.42 | 252.1
2008-07-10 | 10:31 | 10.49 | 95.1 | 10.99 | 30 | 7.43 | 240.8
2008-07-10 | 10:46 | 10.46 | 95.2 | 11.16 | 30 | 7.43 | 233
2008-07-10 | 11:01 | 10.42 | 95.3 | 11.35 | 30 | 7.44 | 256.8
2008-07-10 | 11:16 | 10.38 | 95.4 | 11.54 | 30 | 7.45 | 41.5
2008-07-10 | 11:31 | 10.33 | 95.3 | 11.77 | 30 | 7.46 | 46.3
2008-07-10 | 11:46 | 10.29 | 95.5 | 12 | 30 | 7.46 | 271
2008-07-10 | 12:01 | 10.23 | 95.4 | 12.24 | 30 | 7.46 | 251.5
2008-07-10 | 12:16 | 10.18 | 95.5 | 12.49 | 30 | 7.47 | 271.8
2008-07-10 | 12:31 | 10.12 | 95.6 | 12.75 | 30 | 7.47 | 129.2
2008-07-10 | 12:46 | 10.08 | 95.7 | 13.02 | 30 | 7.47 | 272.8
2008-07-10 | 13:00 | 10.01 | 95.7 | 13.29 | 30 | 7.49 | 288.1
2008-07-10 | 13:15 | 9.95 | 95.6 | 13.56 | 30 | 7.49 | 428.5
2008-07-10 | 13:30 | 9.9 | 95.7 | 13.84 | 30 | 7.49 | 310.4
2008-07-10 | 13:45 | 9.84 | 95.7 | 14.1 | 30 | 7.49 | 295.8
2008-07-10 | 14:00 | 9.78 | 95.7 | 14.36 | 31 | 7.5 | -1.2
2008-07-10 | 14:15 | 9.73 | 95.7 | 14.61 | 31 | 7.52 | 30.8
2008-07-10 | 14:30 | 9.68 | 95.7 | 14.84 | 31 | 7.51 | -1.6
2008-07-10 | 14:45 | 9.63 | 95.6 | 15.06 | 31 | 7.51 | -1.8
2008-07-10 | 15:00 | 9.59 | 95.6 | 15.25 | 31 | 7.51 | -1.8
2008-07-10 | 15:15 | 9.55 | 95.5 | 15.42 | 31 | 7.51 | 38.8
2008-07-10 | 15:30 | 9.51 | 95.5 | 15.57 | 31 | 7.51 | 295.1
2008-07-10 | 15:45 | 9.48 | 95.4 | 15.7 | 31 | 7.52 | 595.9
2008-07-10 | 16:00 | 9.45 | 95.4 | 15.82 | 31 | 7.52 | 562.1
2008-07-10 | 16:15 | 9.43 | 95.3 | 15.91 | 31 | 7.53 | 367.4
2008-07-10 | 16:30 | 9.4 | 95.2 | 15.97 | 31 | 7.5 | 609.7