## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AB (Audited In Situ against AA, 556, & NIST)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-18 |  | 2008-08-18 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp | VWR NIST traceable surface water |  | 13.4 |  | 13.57 |  | A
DO | In Situ |  | 9.58 |  | 9.56 |  | A
SpC | In Situ |  | 37 |  | 37 |  | A
Ph | In Situ |  | 6.65 |  | 6.48 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2008-08-19 13:20 | 16.39 | 7.18 | 9.36 | 95.5 | 52 | 37
Sonde | 2008-08-19 13:20 | 16.43 | 7.36 | 9.42 | 96.3 | 51 | 33
Grade |  | A | A | A |  | A
Audit 2 | 2008-08-23 14:17 | 17.89 | 7.54 | 9.28 | 97.8 | 60 | 42
Sonde | 2008-08-23 14:15 | 17.86 | 7.72 | 9.58 | 100.9 | 58 | 38
Grade |  | A | A | A |  | A
Audit 4 | 2008-08-28 12:00 | 12.75 | 7.48 | 10.76 | 101.6 | 62 | 43
Sonde | 2008-08-28 | 12.77 | 7.55 | 10.85 | 102.5 | 61 | 39
Grade |  | A | A | A |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AB (Audited In Situ against AA, 556, & NIST)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-28 |  | 2008-08-28 |  | A
Time | Computer set to atomic time |  | 12:00 |  | 12:00 |  | A
Temp | VWR NIST traceable surface water |  | 12.6 |  | 12.77 |  | A
DO | In Situ |  | 10.76 |  | 10.85 |  | A
SpC | In Situ |  | 62 |  | 61 |  | A
Ph | In Situ |  | 7.48 |  | 7.55 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Sonde File: 081908AB |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-18 |  | 2008-08-18 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp | VWR NIST traceable surface water |  | 13.4 |  | 13.57 |  | A
DO | In Situ |  | 9.58 |  | 9.56 |  | A
SpC | In Situ |  | 37 |  | 37 |  | A
Ph | In Situ |  | 6.65 |  | 6.48 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-08-19 13:20 | 16.39 | 7.18 | 9.36 | 95.5 | 52 | 37
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-19 |  | 2008-08-19 |  | A
Time | Computer set to atomic time |  | 15:01 |  | 15:01 |  | A
Temp | VWR NIST traceable surface water |  | 14.8 |  | 14.96 |  | A
DO | Winkler titration of surface water |  | 9.6 |  | 9.35 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 149 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.03 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
This audit was taken in descrete sample mode while the unattended sample was logging |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-20 |  | 2008-08-20 |  | A
Time | Computer set to atomic time |  | 18:15 |  | 18:15 |  | A
Temp | VWR NIST traceable surface water |  | 17.8 |  | 18.11 |  | A
DO | Winkler titration of surface water |  | 8.8 |  | 8.68 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 146 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.03 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-08-23 14:17 | 17.89 | 7.54 | 9.28 | 97.8 | 60 | 42
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-25 |  | 2008-08-25 |  | A
Time | Computer set to atomic time |  | 15:33 |  | 15:33 |  | A
Temp | VWR NIST traceable surface water |  | 12.8 |  | 12.97 |  | A
DO | Winkler titration of surface water |  | 9.8 |  | 9.81 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.04 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3 (not Macro'ed')
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-25 |  | 2008-08-25 |  | A
Time | Computer set to atomic time |  | 15:33 |  | 15:33 |  | A
Temp | VWR NIST traceable surface water |  | 12.8 |  | 12.97 |  | A
DO | Winkler titration of surface water |  | 9.8 |  | 9.81 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.04 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-08-26 13:45 | 15.21 | 7.45 | 9.9 | 98.6 | 57 | 40
Sonde | 2008-08-26 13:45 | 15.23 | 7.67 | 10.16 | 101.2 | 56 | 36
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-28 |  | 2008-08-28 |  | A
Time | Computer set to atomic time |  | 12:00 |  | 12:00 |  | A
Temp | VWR NIST traceable surface water |  | 12.6 |  | 12.75 |  | A
DO | In Situ |  | 10.85 |  | 10.76 |  | A
SpC | In Situ |  | 61 |  | 62 |  | A
Ph | In Situ |  | 7.55 |  | 7.48 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Low on Na THIO SO3-2 so next 556 audit was 8-28-08 |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 4
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-25 |  | 2008-08-25 |  | A
Time | Computer set to atomic time |  | 15:33 |  | 15:33 |  | A
Temp | VWR NIST traceable surface water |  | 12.8 |  | 12.97 |  | A
DO | Winkler titration of surface water |  | 9.8 |  | 9.81 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.04 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-08-28 12:00 | 12.75 | 7.48 | 10.76 | 101.6 | 62 | 43
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-28 |  | 2008-08-28 |  | A
Time | Computer set to atomic time |  | 12:00 |  | 12:00 |  | A
Temp | VWR NIST traceable surface water |  | 12.6 |  | 12.75 |  | A
DO | In Situ |  | 10.85 |  | 10.76 |  | A
SpC | In Situ |  | 61 |  | 62 |  | A
Ph | In Situ |  | 7.55 |  | 7.48 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | TDS | DO% | Temp | SpCond | DO Conc | pH
M/D/Y |  | mg/L | % | C | uS/cm | mg/L
2008-08-19 | 13:15 | 33 | 99.9 | 16.29 | 51 | 9.81 | 7.34
2008-08-19 | 13:30 | 33 | 99.1 | 16.64 | 51 | 9.66 | 7.36
2008-08-19 | 13:45 | 33 | 99.7 | 17.25 | 51 | 9.59 | 7.34
2008-08-19 | 14:00 | 33 | 99.7 | 17.66 | 51 | 9.5 | 7.37
2008-08-19 | 14:15 | 33 | 99.2 | 17.93 | 51 | 9.4 | 7.4
2008-08-19 | 14:30 | 33 | 99.2 | 18.11 | 51 | 9.37 | 7.43
2008-08-19 | 14:45 | 33 | 99.1 | 18.14 | 51 | 9.36 | 7.45
2008-08-19 | 15:00 | 33 | 99 | 18.15 | 51 | 9.34 | 7.46
2008-08-19 | 15:15 | 33 | 99 | 18.26 | 51 | 9.32 | 7.49
2008-08-19 | 15:30 | 33 | 98.7 | 18.27 | 51 | 9.3 | 7.5
2008-08-19 | 15:45 | 33 | 98.6 | 18.44 | 51 | 9.25 | 7.51
2008-08-19 | 16:00 | 33 | 98.1 | 18.48 | 51 | 9.2 | 7.52
2008-08-19 | 16:15 | 33 | 97.8 | 18.58 | 51 | 9.15 | 7.53
2008-08-19 | 16:30 | 33 | 97.1 | 18.56 | 51 | 9.09 | 7.52
2008-08-19 | 16:45 | 33 | 96.6 | 18.54 | 52 | 9.04 | 7.5
2008-08-19 | 17:00 | 34 | 96.4 | 18.53 | 52 | 9.02 | 7.49
2008-08-19 | 17:15 | 34 | 95.8 | 18.51 | 52 | 8.97 | 7.46
2008-08-19 | 17:30 | 34 | 95.7 | 18.5 | 52 | 8.96 | 7.45
2008-08-19 | 17:45 | 34 | 95.3 | 18.5 | 52 | 8.93 | 7.43
2008-08-19 | 18:00 | 34 | 95 | 18.49 | 52 | 8.91 | 7.42
2008-08-19 | 18:15 | 34 | 95.2 | 18.45 | 52 | 8.93 | 7.4
2008-08-19 | 18:30 | 34 | 95.3 | 18.37 | 52 | 8.95 | 7.4
2008-08-19 | 18:45 | 34 | 95.1 | 18.29 | 52 | 8.95 | 7.39
2008-08-19 | 19:00 | 34 | 95.2 | 18.21 | 52 | 8.97 | 7.39
2008-08-19 | 19:15 | 34 | 95.1 | 18.11 | 52 | 8.98 | 7.38
2008-08-19 | 19:30 | 34 | 94.7 | 18 | 52 | 8.96 | 7.37
2008-08-19 | 19:45 | 34 | 94.5 | 17.9 | 52 | 8.96 | 7.36
2008-08-19 | 20:00 | 34 | 94.6 | 17.8 | 52 | 8.99 | 7.34
2008-08-19 | 20:15 | 34 | 94.6 | 17.71 | 52 | 9.01 | 7.32
2008-08-19 | 20:30 | 34 | 94.2 | 17.63 | 52 | 8.99 | 7.32
2008-08-19 | 20:45 | 34 | 94.4 | 17.54 | 52 | 9.02 | 7.31
2008-08-19 | 21:00 | 34 | 94.2 | 17.45 | 52 | 9.02 | 7.3
2008-08-19 | 21:15 | 34 | 94.4 | 17.38 | 52 | 9.05 | 7.3
2008-08-19 | 21:30 | 34 | 94.5 | 17.31 | 52 | 9.07 | 7.3
2008-08-19 | 21:45 | 34 | 94.2 | 17.25 | 52 | 9.06 | 7.3
2008-08-19 | 22:00 | 34 | 94.5 | 17.2 | 52 | 9.09 | 7.3
2008-08-19 | 22:15 | 34 | 94.5 | 17.14 | 52 | 9.1 | 7.29
2008-08-19 | 22:30 | 34 | 94.2 | 17.08 | 52 | 9.09 | 7.29
2008-08-19 | 22:45 | 34 | 94.4 | 17.01 | 52 | 9.12 | 7.29
2008-08-19 | 23:00 | 34 | 94.5 | 16.95 | 53 | 9.14 | 7.29
2008-08-19 | 23:15 | 34 | 94.5 | 16.88 | 53 | 9.16 | 7.29
2008-08-19 | 23:30 | 34 | 94.6 | 16.82 | 53 | 9.18 | 7.29
2008-08-19 | 23:45 | 34 | 94.3 | 16.75 | 53 | 9.16 | 7.29
2008-08-20 | 0:00 | 34 | 94.2 | 16.69 | 53 | 9.16 | 7.29
2008-08-20 | 0:15 | 34 | 94.1 | 16.63 | 53 | 9.17 | 7.29
2008-08-20 | 0:30 | 34 | 94.4 | 16.57 | 53 | 9.21 | 7.29
2008-08-20 | 0:45 | 34 | 94.2 | 16.53 | 53 | 9.2 | 7.29
2008-08-20 | 1:01 | 35 | 94.5 | 16.49 | 53 | 9.23 | 7.29
2008-08-20 | 1:16 | 35 | 94.5 | 16.45 | 53 | 9.24 | 7.29
2008-08-20 | 1:31 | 35 | 94.3 | 16.42 | 53 | 9.22 | 7.29
2008-08-20 | 1:46 | 35 | 94.5 | 16.39 | 53 | 9.25 | 7.3
2008-08-20 | 2:01 | 35 | 94.5 | 16.37 | 54 | 9.26 | 7.29
2008-08-20 | 2:16 | 35 | 94.6 | 16.35 | 54 | 9.27 | 7.29
2008-08-20 | 2:31 | 35 | 94.3 | 16.32 | 54 | 9.24 | 7.29
2008-08-20 | 2:46 | 35 | 94.5 | 16.3 | 54 | 9.27 | 7.29
2008-08-20 | 3:01 | 35 | 94.2 | 16.27 | 54 | 9.25 | 7.3
2008-08-20 | 3:16 | 35 | 94.5 | 16.23 | 54 | 9.28 | 7.29
2008-08-20 | 3:31 | 35 | 94.2 | 16.19 | 54 | 9.26 | 7.31
2008-08-20 | 3:46 | 35 | 94.5 | 16.15 | 55 | 9.3 | 7.3
2008-08-20 | 4:01 | 36 | 94.6 | 16.1 | 55 | 9.32 | 7.3
2008-08-20 | 4:16 | 36 | 94.7 | 16.05 | 55 | 9.33 | 7.31
2008-08-20 | 4:31 | 36 | 94.4 | 16 | 55 | 9.32 | 7.31
2008-08-20 | 4:46 | 36 | 94.6 | 15.96 | 55 | 9.35 | 7.3
2008-08-20 | 5:01 | 36 | 94.7 | 15.91 | 55 | 9.36 | 7.3
2008-08-20 | 5:16 | 36 | 94.8 | 15.86 | 56 | 9.38 | 7.31
2008-08-20 | 5:31 | 36 | 94.8 | 15.82 | 56 | 9.39 | 7.32
2008-08-20 | 5:46 | 36 | 94.5 | 15.77 | 56 | 9.37 | 7.31
2008-08-20 | 6:01 | 36 | 94.3 | 15.73 | 56 | 9.37 | 7.32
2008-08-20 | 6:16 | 37 | 94.7 | 15.69 | 56 | 9.41 | 7.31
2008-08-20 | 6:31 | 37 | 94.4 | 15.64 | 56 | 9.39 | 7.31
2008-08-20 | 6:46 | 37 | 94.4 | 15.6 | 56 | 9.4 | 7.32
2008-08-20 | 7:01 | 37 | 94.5 | 15.57 | 57 | 9.41 | 7.32
2008-08-20 | 7:16 | 37 | 94.6 | 15.53 | 57 | 9.43 | 7.33
2008-08-20 | 7:31 | 37 | 95.1 | 15.5 | 57 | 9.48 | 7.34
2008-08-20 | 7:46 | 37 | 95.5 | 15.49 | 57 | 9.53 | 7.35
2008-08-20 | 8:01 | 37 | 95.6 | 15.49 | 57 | 9.54 | 7.37
2008-08-20 | 8:16 | 37 | 96.3 | 15.5 | 57 | 9.61 | 7.39
2008-08-20 | 8:31 | 37 | 96.6 | 15.53 | 57 | 9.63 | 7.42
2008-08-20 | 8:46 | 37 | 97.1 | 15.56 | 57 | 9.67 | 7.45
2008-08-20 | 9:01 | 37 | 97.6 | 15.61 | 57 | 9.71 | 7.47
2008-08-20 | 9:16 | 37 | 97.7 | 15.68 | 57 | 9.7 | 7.5
2008-08-20 | 9:31 | 37 | 97.8 | 15.76 | 57 | 9.7 | 7.53
2008-08-20 | 9:46 | 37 | 97.9 | 15.77 | 57 | 9.71 | 7.54
2008-08-20 | 10:01 | 37 | 97.9 | 15.77 | 57 | 9.71 | 7.54
2008-08-20 | 10:16 | 37 | 97.6 | 15.79 | 57 | 9.68 | 7.54
2008-08-20 | 10:31 | 37 | 97.6 | 15.77 | 57 | 9.68 | 7.54
2008-08-20 | 10:46 | 37 | 97.9 | 15.79 | 57 | 9.71 | 7.55
2008-08-20 | 11:01 | 37 | 98.6 | 15.89 | 57 | 9.75 | 7.56
2008-08-20 | 11:16 | 37 | 97.9 | 15.83 | 57 | 9.7 | 7.56
2008-08-20 | 11:31 | 37 | 97.7 | 15.8 | 57 | 9.68 | 7.56
2008-08-20 | 11:46 | 37 | 97.5 | 15.79 | 57 | 9.67 | 7.55
2008-08-20 | 12:01 | 37 | 97.8 | 15.78 | 57 | 9.69 | 7.55
2008-08-20 | 12:16 | 37 | 97.9 | 15.77 | 57 | 9.71 | 7.54
2008-08-20 | 12:31 | 37 | 98.1 | 15.77 | 57 | 9.73 | 7.55
2008-08-20 | 12:46 | 37 | 98 | 15.75 | 57 | 9.73 | 7.55
2008-08-20 | 13:00 | 37 | 98.1 | 15.76 | 57 | 9.73 | 7.54
2008-08-20 | 13:15 | 37 | 98 | 15.82 | 57 | 9.71 | 7.55
2008-08-20 | 13:30 | 37 | 97.9 | 15.84 | 57 | 9.7 | 7.56
2008-08-20 | 13:45 | 37 | 98.1 | 15.84 | 57 | 9.71 | 7.56
2008-08-20 | 14:00 | 37 | 98.1 | 15.87 | 57 | 9.71 | 7.55
2008-08-20 | 14:15 | 37 | 97.6 | 15.85 | 57 | 9.66 | 7.55
2008-08-20 | 14:30 | 37 | 97.1 | 15.87 | 57 | 9.61 | 7.52
2008-08-20 | 14:45 | 37 | 97.1 | 15.84 | 57 | 9.62 | 7.53
2008-08-20 | 15:00 | 37 | 97.5 | 15.89 | 57 | 9.64 | 7.53
2008-08-20 | 15:15 | 37 | 97.9 | 16.01 | 57 | 9.66 | 7.54
2008-08-20 | 15:30 | 37 | 98.1 | 16.22 | 57 | 9.63 | 7.55
2008-08-20 | 15:45 | 37 | 98.5 | 16.48 | 57 | 9.62 | 7.57
2008-08-20 | 16:00 | 37 | 98 | 16.67 | 56 | 9.53 | 7.57
2008-08-20 | 16:15 | 36 | 97.7 | 16.77 | 56 | 9.49 | 7.57
2008-08-20 | 16:30 | 36 | 97.6 | 16.83 | 56 | 9.46 | 7.58
2008-08-20 | 16:45 | 36 | 97.5 | 16.81 | 55 | 9.46 | 7.58
2008-08-20 | 17:00 | 35 | 97.2 | 16.84 | 55 | 9.43 | 7.55
2008-08-20 | 17:15 | 35 | 96.9 | 16.94 | 54 | 9.37 | 7.53
2008-08-20 | 17:30 | 34 | 97 | 17.12 | 53 | 9.35 | 7.52
2008-08-20 | 17:45 | 34 | 96.8 | 17.39 | 52 | 9.28 | 7.5
2008-08-20 | 18:00 | 33 | 96.5 | 17.72 | 51 | 9.19 | 7.5
2008-08-20 | 18:15 | 32 | 96.3 | 18.04 | 50 | 9.11 | 7.49
2008-08-20 | 18:30 | 32 | 95.6 | 18.3 | 49 | 8.99 | 7.47
2008-08-20 | 18:45 | 31 | 95.4 | 18.43 | 48 | 8.95 | 7.45
2008-08-20 | 19:00 | 31 | 95.2 | 18.46 | 47 | 8.93 | 7.41
2008-08-20 | 19:15 | 30 | 94.9 | 18.49 | 47 | 8.9 | 7.4
2008-08-20 | 19:30 | 30 | 94.5 | 18.45 | 46 | 8.87 | 7.38
2008-08-20 | 19:45 | 30 | 94.4 | 18.39 | 46 | 8.86 | 7.38
2008-08-20 | 20:00 | 30 | 94.5 | 18.28 | 45 | 8.89 | 7.36
2008-08-20 | 20:15 | 29 | 94.5 | 18.16 | 45 | 8.92 | 7.31
2008-08-20 | 20:30 | 29 | 94.2 | 18 | 45 | 8.92 | 7.31
2008-08-20 | 20:45 | 29 | 94.1 | 17.83 | 44 | 8.94 | 7.28
2008-08-20 | 21:00 | 29 | 94.4 | 17.68 | 44 | 9 | 7.28
2008-08-20 | 21:15 | 28 | 94.5 | 17.53 | 44 | 9.04 | 7.26
2008-08-20 | 21:30 | 28 | 94.7 | 17.36 | 44 | 9.08 | 7.25
2008-08-20 | 21:45 | 28 | 94.8 | 17.17 | 43 | 9.13 | 7.24
2008-08-20 | 22:00 | 28 | 94.5 | 17 | 43 | 9.14 | 7.23
2008-08-20 | 22:15 | 28 | 94.8 | 16.83 | 43 | 9.19 | 7.22
2008-08-20 | 22:30 | 28 | 94.9 | 16.68 | 42 | 9.24 | 7.21
2008-08-20 | 22:45 | 27 | 94.7 | 16.52 | 42 | 9.24 | 7.21
2008-08-20 | 23:00 | 27 | 95.1 | 16.36 | 42 | 9.32 | 7.19
2008-08-20 | 23:15 | 27 | 95.1 | 16.22 | 42 | 9.34 | 7.19
2008-08-20 | 23:30 | 27 | 94.8 | 16.06 | 41 | 9.35 | 7.21
2008-08-20 | 23:45 | 27 | 94.8 | 15.92 | 41 | 9.38 | 7.18
2008-08-21 | 0:00 | 27 | 94.8 | 15.76 | 41 | 9.4 | 7.16
2008-08-21 | 0:15 | 27 | 95.1 | 15.61 | 41 | 9.47 | 7.16
2008-08-21 | 0:30 | 27 | 95.3 | 15.47 | 41 | 9.52 | 7.15
2008-08-21 | 0:45 | 26 | 95.4 | 15.33 | 41 | 9.56 | 7.15
2008-08-21 | 1:01 | 26 | 95.2 | 15.2 | 40 | 9.56 | 7.14
2008-08-21 | 1:16 | 26 | 95.4 | 15.05 | 40 | 9.61 | 7.14
2008-08-21 | 1:31 | 26 | 95.6 | 14.9 | 40 | 9.66 | 7.12
2008-08-21 | 1:46 | 26 | 95.3 | 14.74 | 40 | 9.66 | 7.12
2008-08-21 | 2:01 | 26 | 95.3 | 14.58 | 40 | 9.69 | 7.12
2008-08-21 | 2:16 | 26 | 95.6 | 14.43 | 40 | 9.76 | 7.12
2008-08-21 | 2:31 | 26 | 95.4 | 14.3 | 40 | 9.77 | 7.12
2008-08-21 | 2:46 | 26 | 95.8 | 14.17 | 40 | 9.83 | 7.11
2008-08-21 | 3:01 | 26 | 96 | 14.06 | 40 | 9.88 | 7.08
2008-08-21 | 3:16 | 26 | 96 | 13.95 | 39 | 9.91 | 7.09
2008-08-21 | 3:31 | 25 | 96.1 | 13.84 | 39 | 9.94 | 7.09
2008-08-21 | 3:46 | 25 | 96.1 | 13.75 | 39 | 9.96 | 7.09
2008-08-21 | 4:01 | 25 | 96.2 | 13.66 | 39 | 9.99 | 7.07
2008-08-21 | 4:16 | 25 | 96.2 | 13.59 | 38 | 10.01 | 7.08
2008-08-21 | 4:31 | 25 | 96.3 | 13.53 | 38 | 10.02 | 7.07
2008-08-21 | 4:46 | 24 | 96.3 | 13.47 | 38 | 10.04 | 7.06
2008-08-21 | 5:01 | 24 | 96.3 | 13.42 | 37 | 10.05 | 7.05
2008-08-21 | 5:16 | 24 | 96 | 13.37 | 37 | 10.03 | 7.06
2008-08-21 | 5:31 | 24 | 96.2 | 13.33 | 37 | 10.06 | 7.06
2008-08-21 | 5:46 | 24 | 96 | 13.29 | 37 | 10.05 | 7.07
2008-08-21 | 6:01 | 24 | 96.3 | 13.25 | 37 | 10.09 | 7.06
2008-08-21 | 6:16 | 24 | 96 | 13.21 | 37 | 10.07 | 7.06
2008-08-21 | 6:31 | 24 | 96 | 13.15 | 36 | 10.08 | 7.05
2008-08-21 | 6:46 | 23 | 96.3 | 13.09 | 36 | 10.13 | 7.05
2008-08-21 | 7:01 | 23 | 96.5 | 13.01 | 36 | 10.17 | 7.05
2008-08-21 | 7:16 | 23 | 96.7 | 12.93 | 36 | 10.2 | 7.05
2008-08-21 | 7:31 | 23 | 96.8 | 12.85 | 36 | 10.23 | 7.05
2008-08-21 | 7:46 | 23 | 96.8 | 12.77 | 36 | 10.25 | 7.06
2008-08-21 | 8:01 | 23 | 96.9 | 12.69 | 35 | 10.28 | 7.06
2008-08-21 | 8:16 | 23 | 97 | 12.6 | 35 | 10.31 | 7.06
2008-08-21 | 8:31 | 23 | 96.8 | 12.52 | 35 | 10.31 | 7.09
2008-08-21 | 8:46 | 23 | 96.7 | 12.45 | 35 | 10.32 | 7.04
2008-08-21 | 9:01 | 23 | 96.8 | 12.4 | 35 | 10.34 | 7.06
2008-08-21 | 9:16 | 23 | 97 | 12.33 | 35 | 10.37 | 7.06
2008-08-21 | 9:31 | 23 | 97.1 | 12.28 | 35 | 10.4 | 7.07
2008-08-21 | 9:46 | 23 | 97.4 | 12.24 | 35 | 10.44 | 7.08
2008-08-21 | 10:01 | 23 | 97.6 | 12.24 | 35 | 10.46 | 7.08
2008-08-21 | 10:16 | 23 | 97.8 | 12.3 | 35 | 10.47 | 7.1
2008-08-21 | 10:31 | 23 | 98.3 | 12.33 | 35 | 10.52 | 7.11
2008-08-21 | 10:46 | 23 | 98.6 | 12.39 | 35 | 10.54 | 7.12
2008-08-21 | 11:01 | 23 | 99 | 12.51 | 35 | 10.54 | 7.13
2008-08-21 | 11:16 | 23 | 99 | 12.62 | 35 | 10.52 | 7.14
2008-08-21 | 11:31 | 23 | 99.2 | 12.75 | 35 | 10.51 | 7.15
2008-08-21 | 11:46 | 23 | 99.3 | 12.78 | 35 | 10.51 | 7.16
2008-08-21 | 12:01 | 23 | 98.8 | 12.86 | 35 | 10.44 | 7.16
2008-08-21 | 12:16 | 23 | 98.9 | 13.09 | 35 | 10.4 | 7.17
2008-08-21 | 12:31 | 23 | 99.1 | 13.39 | 35 | 10.36 | 7.19
2008-08-21 | 12:46 | 23 | 99.1 | 13.78 | 36 | 10.27 | 7.21
2008-08-21 | 13:00 | 23 | 99.1 | 14.1 | 36 | 10.19 | 7.22
2008-08-21 | 13:15 | 23 | 99.5 | 14.35 | 36 | 10.17 | 7.24
2008-08-21 | 13:30 | 23 | 99.5 | 14.71 | 36 | 10.1 | 7.26
2008-08-21 | 13:45 | 23 | 99.2 | 15.08 | 36 | 9.98 | 7.26
2008-08-21 | 14:00 | 23 | 99 | 15.43 | 36 | 9.89 | 7.29
2008-08-21 | 14:15 | 23 | 99.3 | 15.75 | 36 | 9.85 | 7.31
2008-08-21 | 14:30 | 24 | 98.9 | 16.03 | 36 | 9.76 | 7.33
2008-08-21 | 14:45 | 24 | 99 | 16.24 | 36 | 9.72 | 7.34
2008-08-21 | 15:00 | 24 | 98.4 | 16.34 | 36 | 9.64 | 7.34
2008-08-21 | 15:15 | 24 | 97.9 | 16.55 | 37 | 9.55 | 7.32
2008-08-21 | 15:30 | 24 | 97.9 | 16.76 | 36 | 9.51 | 7.34
2008-08-21 | 15:45 | 24 | 98.1 | 16.92 | 37 | 9.5 | 7.35
2008-08-21 | 16:00 | 24 | 98.1 | 17.02 | 37 | 9.48 | 7.35
2008-08-21 | 16:15 | 24 | 97.6 | 17.08 | 37 | 9.42 | 7.35
2008-08-21 | 16:30 | 24 | 97.6 | 17.12 | 37 | 9.41 | 7.35
2008-08-21 | 16:45 | 24 | 97.5 | 17.1 | 37 | 9.4 | 7.34
2008-08-21 | 17:00 | 24 | 97.2 | 17.03 | 37 | 9.39 | 7.34
2008-08-21 | 17:15 | 24 | 97 | 16.93 | 37 | 9.39 | 7.32
2008-08-21 | 17:30 | 24 | 96.9 | 16.78 | 38 | 9.41 | 7.3
2008-08-21 | 17:45 | 25 | 96.8 | 16.67 | 38 | 9.42 | 7.29
2008-08-21 | 18:00 | 25 | 96.6 | 16.59 | 38 | 9.42 | 7.27
2008-08-21 | 18:15 | 25 | 96.2 | 16.46 | 38 | 9.4 | 7.26
2008-08-21 | 18:30 | 25 | 96.3 | 16.35 | 38 | 9.44 | 7.25
2008-08-21 | 18:45 | 25 | 95.9 | 16.33 | 38 | 9.4 | 7.24
2008-08-21 | 19:00 | 25 | 95.7 | 16.32 | 38 | 9.38 | 7.23
2008-08-21 | 19:15 | 25 | 96 | 16.29 | 39 | 9.42 | 7.24
2008-08-21 | 19:30 | 25 | 95.7 | 16.23 | 39 | 9.4 | 7.23
2008-08-21 | 19:45 | 25 | 96 | 16.16 | 39 | 9.44 | 7.22
2008-08-21 | 20:00 | 25 | 95.6 | 16.07 | 39 | 9.42 | 7.21
2008-08-21 | 20:15 | 25 | 95.5 | 15.94 | 39 | 9.44 | 7.2
2008-08-21 | 20:30 | 25 | 95.8 | 15.79 | 39 | 9.5 | 7.2
2008-08-21 | 20:45 | 26 | 95.6 | 15.62 | 39 | 9.52 | 7.2
2008-08-21 | 21:00 | 26 | 95.9 | 15.44 | 39 | 9.58 | 7.19
2008-08-21 | 21:15 | 26 | 95.8 | 15.26 | 40 | 9.6 | 7.19
2008-08-21 | 21:30 | 26 | 95.7 | 15.1 | 40 | 9.63 | 7.18
2008-08-21 | 21:45 | 26 | 96.2 | 14.94 | 40 | 9.71 | 7.19
2008-08-21 | 22:00 | 26 | 96.2 | 14.78 | 40 | 9.75 | 7.18
2008-08-21 | 22:15 | 26 | 96.3 | 14.63 | 40 | 9.79 | 7.18
2008-08-21 | 22:30 | 26 | 96.1 | 14.5 | 40 | 9.8 | 7.17
2008-08-21 | 22:45 | 26 | 96.3 | 14.37 | 40 | 9.84 | 7.17
2008-08-21 | 23:00 | 26 | 96.4 | 14.25 | 41 | 9.88 | 7.15
2008-08-21 | 23:15 | 26 | 96.2 | 14.13 | 41 | 9.88 | 7.16
2008-08-21 | 23:30 | 27 | 96.4 | 14 | 41 | 9.94 | 7.16
2008-08-21 | 23:45 | 27 | 96.6 | 13.88 | 41 | 9.98 | 7.15
2008-08-22 | 0:00 | 27 | 96.3 | 13.76 | 41 | 9.98 | 7.16
2008-08-22 | 0:15 | 27 | 96.6 | 13.64 | 41 | 10.03 | 7.16
2008-08-22 | 0:30 | 27 | 96.5 | 13.51 | 41 | 10.05 | 7.16
2008-08-22 | 0:45 | 27 | 96.6 | 13.38 | 41 | 10.1 | 7.15
2008-08-22 | 1:01 | 27 | 96.5 | 13.25 | 42 | 10.11 | 7.16
2008-08-22 | 1:16 | 27 | 96.8 | 13.11 | 42 | 10.17 | 7.15
2008-08-22 | 1:31 | 27 | 96.9 | 12.97 | 42 | 10.22 | 7.15
2008-08-22 | 1:46 | 28 | 97 | 12.83 | 42 | 10.26 | 7.15
2008-08-22 | 2:01 | 28 | 97.1 | 12.68 | 42 | 10.3 | 7.15
2008-08-22 | 2:16 | 28 | 97.2 | 12.54 | 43 | 10.34 | 7.15
2008-08-22 | 2:31 | 28 | 96.8 | 12.39 | 43 | 10.34 | 7.15
2008-08-22 | 2:46 | 28 | 97.2 | 12.25 | 43 | 10.41 | 7.15
2008-08-22 | 3:01 | 28 | 97.2 | 12.12 | 43 | 10.45 | 7.15
2008-08-22 | 3:16 | 28 | 97.3 | 11.98 | 43 | 10.49 | 7.15
2008-08-22 | 3:31 | 28 | 97.4 | 11.84 | 43 | 10.53 | 7.15
2008-08-22 | 3:46 | 28 | 97.1 | 11.71 | 43 | 10.53 | 7.15
2008-08-22 | 4:01 | 28 | 97.4 | 11.59 | 44 | 10.59 | 7.15
2008-08-22 | 4:16 | 28 | 97.1 | 11.46 | 44 | 10.59 | 7.16
2008-08-22 | 4:31 | 29 | 97.4 | 11.33 | 44 | 10.66 | 7.15
2008-08-22 | 4:46 | 29 | 97.2 | 11.22 | 44 | 10.67 | 7.15
2008-08-22 | 5:01 | 29 | 97.5 | 11.1 | 44 | 10.73 | 7.15
2008-08-22 | 5:16 | 29 | 97.3 | 10.99 | 44 | 10.73 | 7.16
2008-08-22 | 5:31 | 29 | 97.6 | 10.88 | 45 | 10.79 | 7.15
2008-08-22 | 5:46 | 29 | 97.7 | 10.78 | 45 | 10.83 | 7.15
2008-08-22 | 6:01 | 29 | 97.8 | 10.67 | 45 | 10.87 | 7.15
2008-08-22 | 6:16 | 29 | 97.9 | 10.57 | 45 | 10.9 | 7.15
2008-08-22 | 6:31 | 29 | 97.6 | 10.48 | 45 | 10.9 | 7.15
2008-08-22 | 6:46 | 29 | 98.1 | 10.39 | 45 | 10.97 | 7.16
2008-08-22 | 7:01 | 29 | 98 | 10.31 | 45 | 10.98 | 7.16
2008-08-22 | 7:16 | 29 | 98.4 | 10.23 | 45 | 11.05 | 7.17
2008-08-22 | 7:31 | 29 | 98.6 | 10.15 | 45 | 11.1 | 7.17
2008-08-22 | 7:46 | 29 | 98.9 | 10.07 | 45 | 11.14 | 7.17
2008-08-22 | 8:01 | 30 | 98.7 | 10.02 | 45 | 11.14 | 7.18
2008-08-22 | 8:16 | 29 | 99.3 | 9.98 | 45 | 11.21 | 7.19
2008-08-22 | 8:31 | 29 | 99.1 | 9.94 | 45 | 11.2 | 7.2
2008-08-22 | 8:46 | 30 | 99.1 | 9.92 | 46 | 11.21 | 7.2
2008-08-22 | 9:01 | 30 | 99.5 | 9.96 | 46 | 11.24 | 7.21
2008-08-22 | 9:16 | 30 | 99.7 | 9.99 | 46 | 11.26 | 7.21
2008-08-22 | 9:31 | 30 | 99.9 | 10.05 | 46 | 11.26 | 7.23
2008-08-22 | 9:46 | 30 | 100 | 10.1 | 46 | 11.27 | 7.24
2008-08-22 | 10:01 | 30 | 100.7 | 10.18 | 46 | 11.32 | 7.26
2008-08-22 | 10:16 | 30 | 101 | 10.27 | 46 | 11.33 | 7.27
2008-08-22 | 10:31 | 30 | 100.9 | 10.36 | 46 | 11.29 | 7.29
2008-08-22 | 10:46 | 30 | 101.5 | 10.51 | 46 | 11.33 | 7.3
2008-08-22 | 11:01 | 30 | 101.8 | 10.61 | 46 | 11.33 | 7.31
2008-08-22 | 11:16 | 30 | 102.2 | 10.83 | 47 | 11.31 | 7.32
2008-08-22 | 11:31 | 30 | 102.4 | 11.13 | 47 | 11.26 | 7.34
2008-08-22 | 11:46 | 30 | 102.4 | 11.46 | 47 | 11.17 | 7.35
2008-08-22 | 12:01 | 30 | 102 | 11.81 | 47 | 11.04 | 7.36
2008-08-22 | 12:16 | 31 | 102.2 | 12.22 | 47 | 10.96 | 7.38
2008-08-22 | 12:31 | 31 | 102.2 | 12.7 | 47 | 10.84 | 7.4
2008-08-22 | 12:46 | 31 | 101.8 | 13.18 | 47 | 10.68 | 7.42
2008-08-22 | 13:00 | 31 | 101.5 | 13.66 | 48 | 10.54 | 7.43
2008-08-22 | 13:15 | 31 | 101.7 | 14.1 | 47 | 10.46 | 7.47
2008-08-22 | 13:30 | 31 | 101.7 | 14.57 | 47 | 10.35 | 7.45
2008-08-22 | 13:45 | 31 | 101.2 | 15.05 | 47 | 10.19 | 7.48
2008-08-22 | 14:00 | 31 | 100.9 | 15.57 | 48 | 10.06 | 7.49
2008-08-22 | 14:15 | 31 | 101.1 | 16.04 | 48 | 9.97 | 7.51
2008-08-22 | 14:30 | 31 | 101 | 16.45 | 48 | 9.87 | 7.55
2008-08-22 | 14:45 | 31 | 100.8 | 16.85 | 48 | 9.78 | 7.58
2008-08-22 | 15:00 | 31 | 100.5 | 17.22 | 48 | 9.67 | 7.6
2008-08-22 | 15:15 | 31 | 99.9 | 17.53 | 48 | 9.54 | 7.62
2008-08-22 | 15:30 | 31 | 99.4 | 17.8 | 48 | 9.45 | 7.64
2008-08-22 | 15:45 | 32 | 99.4 | 18 | 49 | 9.41 | 7.64
2008-08-22 | 16:00 | 32 | 99.2 | 18.15 | 49 | 9.36 | 7.63
2008-08-22 | 16:15 | 32 | 99 | 18.21 | 49 | 9.33 | 7.62
2008-08-22 | 16:30 | 32 | 98.7 | 18.21 | 49 | 9.3 | 7.59
2008-08-22 | 16:45 | 32 | 98.4 | 18.18 | 49 | 9.28 | 7.56
2008-08-22 | 17:00 | 32 | 98.1 | 18.12 | 49 | 9.27 | 7.55
2008-08-22 | 17:15 | 32 | 97.8 | 18.01 | 49 | 9.26 | 7.52
2008-08-22 | 17:30 | 32 | 97.5 | 17.92 | 49 | 9.25 | 7.49
2008-08-22 | 17:45 | 32 | 96.9 | 17.85 | 49 | 9.21 | 7.47
2008-08-22 | 18:00 | 32 | 96.6 | 17.76 | 49 | 9.19 | 7.44
2008-08-22 | 18:15 | 32 | 96.3 | 17.69 | 50 | 9.17 | 7.43
2008-08-22 | 18:30 | 32 | 96.1 | 17.64 | 50 | 9.16 | 7.42
2008-08-22 | 18:45 | 32 | 96.2 | 17.59 | 50 | 9.18 | 7.39
2008-08-22 | 19:00 | 32 | 95.9 | 17.55 | 50 | 9.16 | 7.38
2008-08-22 | 19:15 | 32 | 96 | 17.53 | 50 | 9.18 | 7.37
2008-08-22 | 19:30 | 32 | 95.6 | 17.51 | 50 | 9.14 | 7.36
2008-08-22 | 19:45 | 32 | 95.6 | 17.48 | 50 | 9.15 | 7.35
2008-08-22 | 20:00 | 33 | 95.2 | 17.42 | 50 | 9.12 | 7.34
2008-08-22 | 20:15 | 33 | 95.3 | 17.35 | 50 | 9.14 | 7.33
2008-08-22 | 20:30 | 33 | 95.3 | 17.25 | 50 | 9.16 | 7.32
2008-08-22 | 20:45 | 33 | 95 | 17.15 | 50 | 9.15 | 7.3
2008-08-22 | 21:00 | 33 | 94.9 | 17.02 | 50 | 9.17 | 7.32
2008-08-22 | 21:15 | 33 | 94.9 | 16.87 | 51 | 9.2 | 7.29
2008-08-22 | 21:30 | 33 | 95.2 | 16.7 | 51 | 9.26 | 7.28
2008-08-22 | 21:45 | 33 | 95.4 | 16.54 | 51 | 9.31 | 7.29
2008-08-22 | 22:00 | 33 | 95.5 | 16.37 | 51 | 9.35 | 7.29
2008-08-22 | 22:15 | 33 | 95.3 | 16.2 | 51 | 9.36 | 7.28
2008-08-22 | 22:30 | 33 | 95.2 | 16.03 | 51 | 9.39 | 7.27
2008-08-22 | 22:45 | 33 | 95.5 | 15.86 | 51 | 9.46 | 7.27
2008-08-22 | 23:00 | 33 | 95.7 | 15.7 | 51 | 9.5 | 7.27
2008-08-22 | 23:15 | 33 | 95.8 | 15.55 | 51 | 9.55 | 7.27
2008-08-22 | 23:30 | 33 | 95.8 | 15.4 | 51 | 9.58 | 7.26
2008-08-22 | 23:45 | 33 | 95.6 | 15.26 | 51 | 9.58 | 7.27
2008-08-23 | 0:00 | 33 | 95.9 | 15.11 | 52 | 9.65 | 7.27
2008-08-23 | 0:15 | 34 | 96 | 14.95 | 52 | 9.69 | 7.27
2008-08-23 | 0:30 | 34 | 95.7 | 14.8 | 52 | 9.7 | 7.29
2008-08-23 | 0:45 | 34 | 96 | 14.65 | 52 | 9.76 | 7.27
2008-08-23 | 1:01 | 34 | 96.2 | 14.51 | 52 | 9.8 | 7.27
2008-08-23 | 1:16 | 34 | 95.9 | 14.38 | 52 | 9.8 | 7.27
2008-08-23 | 1:31 | 34 | 95.8 | 14.25 | 53 | 9.82 | 7.28
2008-08-23 | 1:46 | 34 | 96.1 | 14.13 | 53 | 9.88 | 7.27
2008-08-23 | 2:01 | 34 | 96 | 14.01 | 53 | 9.89 | 7.27
2008-08-23 | 2:16 | 35 | 95.9 | 13.89 | 53 | 9.9 | 7.27
2008-08-23 | 2:31 | 35 | 96.2 | 13.77 | 53 | 9.96 | 7.28
2008-08-23 | 2:46 | 35 | 96.4 | 13.66 | 53 | 10.01 | 7.27
2008-08-23 | 3:01 | 35 | 96.1 | 13.55 | 53 | 10 | 7.27
2008-08-23 | 3:16 | 35 | 96.4 | 13.44 | 54 | 10.05 | 7.26
2008-08-23 | 3:31 | 35 | 96.1 | 13.33 | 54 | 10.05 | 7.26
2008-08-23 | 3:46 | 35 | 96.5 | 13.23 | 54 | 10.11 | 7.27
2008-08-23 | 4:01 | 35 | 96.2 | 13.12 | 54 | 10.11 | 7.27
2008-08-23 | 4:16 | 35 | 96.5 | 13.01 | 54 | 10.16 | 7.27
2008-08-23 | 4:31 | 35 | 96.3 | 12.9 | 54 | 10.17 | 7.26
2008-08-23 | 4:46 | 35 | 96.5 | 12.79 | 54 | 10.22 | 7.27
2008-08-23 | 5:01 | 35 | 96.7 | 12.68 | 54 | 10.26 | 7.26
2008-08-23 | 5:16 | 35 | 96.5 | 12.57 | 54 | 10.26 | 7.27
2008-08-23 | 5:31 | 35 | 96.7 | 12.45 | 55 | 10.32 | 7.26
2008-08-23 | 5:46 | 36 | 96.9 | 12.34 | 55 | 10.36 | 7.27
2008-08-23 | 6:01 | 36 | 97 | 12.22 | 55 | 10.4 | 7.27
2008-08-23 | 6:16 | 36 | 96.8 | 12.1 | 55 | 10.4 | 7.26
2008-08-23 | 6:31 | 36 | 96.7 | 11.99 | 55 | 10.42 | 7.26
2008-08-23 | 6:46 | 36 | 97.1 | 11.87 | 55 | 10.5 | 7.27
2008-08-23 | 7:01 | 36 | 97.5 | 11.76 | 55 | 10.56 | 7.27
2008-08-23 | 7:16 | 36 | 97.8 | 11.65 | 55 | 10.62 | 7.28
2008-08-23 | 7:31 | 36 | 98 | 11.56 | 55 | 10.67 | 7.28
2008-08-23 | 7:46 | 36 | 97.9 | 11.49 | 56 | 10.67 | 7.28
2008-08-23 | 8:01 | 36 | 97.9 | 11.42 | 56 | 10.7 | 7.3
2008-08-23 | 8:16 | 36 | 98.4 | 11.37 | 56 | 10.76 | 7.3
2008-08-23 | 8:31 | 36 | 98.7 | 11.34 | 56 | 10.8 | 7.3
2008-08-23 | 8:46 | 36 | 99 | 11.35 | 56 | 10.83 | 7.31
2008-08-23 | 9:01 | 36 | 99.3 | 11.4 | 56 | 10.85 | 7.32
2008-08-23 | 9:16 | 37 | 99.6 | 11.44 | 56 | 10.87 | 7.33
2008-08-23 | 9:31 | 37 | 99.4 | 11.49 | 56 | 10.84 | 7.34
2008-08-23 | 9:46 | 37 | 99.5 | 11.55 | 57 | 10.84 | 7.36
2008-08-23 | 10:01 | 37 | 100 | 11.64 | 57 | 10.87 | 7.38
2008-08-23 | 10:16 | 37 | 100.3 | 11.72 | 57 | 10.88 | 7.39
2008-08-23 | 10:31 | 37 | 100.6 | 11.82 | 57 | 10.88 | 7.4
2008-08-23 | 10:46 | 37 | 100.6 | 11.95 | 57 | 10.85 | 7.42
2008-08-23 | 11:01 | 37 | 101.3 | 12.1 | 57 | 10.89 | 7.43
2008-08-23 | 11:16 | 37 | 101.3 | 12.28 | 57 | 10.85 | 7.45
2008-08-23 | 11:31 | 37 | 101.3 | 12.53 | 57 | 10.79 | 7.46
2008-08-23 | 11:46 | 37 | 101.8 | 12.84 | 57 | 10.76 | 7.48
2008-08-23 | 12:01 | 37 | 101.9 | 13.21 | 57 | 10.69 | 7.5
2008-08-23 | 12:16 | 37 | 101.5 | 13.65 | 57 | 10.55 | 7.52
2008-08-23 | 12:31 | 37 | 101.3 | 14.13 | 58 | 10.41 | 7.54
2008-08-23 | 12:46 | 37 | 101.2 | 14.62 | 58 | 10.29 | 7.56
2008-08-23 | 13:00 | 38 | 101.5 | 15.15 | 58 | 10.2 | 7.59
2008-08-23 | 13:15 | 38 | 101.5 | 15.7 | 58 | 10.08 | 7.61
2008-08-23 | 13:30 | 38 | 101 | 16.23 | 58 | 9.92 | 7.64
2008-08-23 | 13:45 | 38 | 101.1 | 16.78 | 58 | 9.82 | 7.67
2008-08-23 | 14:00 | 38 | 101 | 17.32 | 58 | 9.7 | 7.7
2008-08-23 | 14:15 | 38 | 100.9 | 17.86 | 58 | 9.58 | 7.72
2008-08-23 | 14:30 | 38 | 100.3 | 18.36 | 58 | 9.43 | 7.8
2008-08-23 | 14:45 | 38 | 100.3 | 18.81 | 58 | 9.34 | 7.85
2008-08-23 | 15:00 | 38 | 99.7 | 19.2 | 58 | 9.21 | 7.86
2008-08-23 | 15:15 | 38 | 99.3 | 19.48 | 58 | 9.12 | 7.85
2008-08-23 | 15:30 | 38 | 99.3 | 19.7 | 58 | 9.08 | 7.85
2008-08-23 | 15:45 | 38 | 99.2 | 19.93 | 58 | 9.03 | 7.86
2008-08-23 | 16:00 | 38 | 98.6 | 19.98 | 58 | 8.97 | 7.86
2008-08-23 | 16:15 | 38 | 98.7 | 20.1 | 58 | 8.95 | 7.83
2008-08-23 | 16:30 | 38 | 98.4 | 20.14 | 58 | 8.92 | 7.81
2008-08-23 | 16:45 | 38 | 97.8 | 20.09 | 58 | 8.87 | 7.78
2008-08-23 | 17:00 | 38 | 97.7 | 20.05 | 58 | 8.88 | 7.75
2008-08-23 | 17:15 | 38 | 97.5 | 19.9 | 58 | 8.88 | 7.72
2008-08-23 | 17:30 | 38 | 97.3 | 19.78 | 58 | 8.88 | 7.7
2008-08-23 | 17:45 | 38 | 97.2 | 19.7 | 58 | 8.89 | 7.67
2008-08-23 | 18:00 | 38 | 97 | 19.64 | 58 | 8.89 | 7.64
2008-08-23 | 18:15 | 38 | 96.9 | 19.57 | 58 | 8.89 | 7.63
2008-08-23 | 18:30 | 38 | 96.4 | 19.5 | 58 | 8.85 | 7.61
2008-08-23 | 18:45 | 38 | 96.1 | 19.44 | 58 | 8.84 | 7.58
2008-08-23 | 19:00 | 38 | 95.7 | 19.43 | 58 | 8.8 | 7.55
2008-08-23 | 19:15 | 38 | 95 | 19.39 | 58 | 8.74 | 7.52
2008-08-23 | 19:30 | 38 | 94.9 | 19.33 | 58 | 8.75 | 7.5
2008-08-23 | 19:45 | 38 | 94.5 | 19.28 | 58 | 8.71 | 7.48
2008-08-23 | 20:00 | 38 | 94.1 | 19.21 | 59 | 8.69 | 7.46
2008-08-23 | 20:15 | 38 | 93.9 | 19.12 | 59 | 8.69 | 7.43
2008-08-23 | 20:30 | 38 | 94.1 | 19 | 59 | 8.73 | 7.42
2008-08-23 | 20:45 | 38 | 94.1 | 18.87 | 59 | 8.76 | 7.41
2008-08-23 | 21:00 | 39 | 93.9 | 18.73 | 59 | 8.76 | 7.4
2008-08-23 | 21:15 | 38 | 94.2 | 18.6 | 59 | 8.81 | 7.39
2008-08-23 | 21:30 | 38 | 94.4 | 18.46 | 59 | 8.85 | 7.37
2008-08-23 | 21:45 | 38 | 94.5 | 18.3 | 59 | 8.89 | 7.36
2008-08-23 | 22:00 | 38 | 94.6 | 18.13 | 59 | 8.93 | 7.36
2008-08-23 | 22:15 | 38 | 94.4 | 17.97 | 59 | 8.94 | 7.36
2008-08-23 | 22:30 | 38 | 94.6 | 17.82 | 59 | 8.99 | 7.37
2008-08-23 | 22:45 | 38 | 94.8 | 17.67 | 59 | 9.03 | 7.36
2008-08-23 | 23:00 | 39 | 94.8 | 17.52 | 60 | 9.07 | 7.36
2008-08-23 | 23:15 | 39 | 94.9 | 17.38 | 59 | 9.1 | 7.36
2008-08-23 | 23:30 | 39 | 95 | 17.24 | 60 | 9.13 | 7.35
2008-08-23 | 23:45 | 39 | 95 | 17.1 | 60 | 9.16 | 7.35
2008-08-24 | 0:00 | 39 | 95 | 16.97 | 60 | 9.19 | 7.35
2008-08-24 | 0:15 | 39 | 95 | 16.83 | 60 | 9.21 | 7.35
2008-08-24 | 0:30 | 39 | 95.2 | 16.7 | 60 | 9.25 | 7.35
2008-08-24 | 0:45 | 39 | 94.8 | 16.57 | 59 | 9.24 | 7.35
2008-08-24 | 1:01 | 39 | 95 | 16.45 | 60 | 9.29 | 7.35
2008-08-24 | 1:16 | 39 | 94.8 | 16.34 | 59 | 9.29 | 7.35
2008-08-24 | 1:31 | 39 | 94.8 | 16.23 | 59 | 9.31 | 7.35
2008-08-24 | 1:46 | 39 | 95.1 | 16.12 | 59 | 9.37 | 7.35
2008-08-24 | 2:01 | 39 | 95.3 | 16.01 | 60 | 9.4 | 7.35
2008-08-24 | 2:16 | 39 | 95.4 | 15.91 | 60 | 9.43 | 7.35
2008-08-24 | 2:31 | 39 | 95.5 | 15.82 | 60 | 9.46 | 7.35
2008-08-24 | 2:46 | 39 | 95.1 | 15.72 | 60 | 9.45 | 7.35
2008-08-24 | 3:01 | 39 | 95 | 15.62 | 60 | 9.45 | 7.35
2008-08-24 | 3:16 | 39 | 95 | 15.52 | 60 | 9.47 | 7.35
2008-08-24 | 3:31 | 39 | 95.4 | 15.42 | 60 | 9.53 | 7.34
2008-08-24 | 3:46 | 39 | 95.5 | 15.33 | 60 | 9.57 | 7.35
2008-08-24 | 4:01 | 39 | 95.6 | 15.23 | 60 | 9.59 | 7.34
2008-08-24 | 4:16 | 39 | 95.3 | 15.13 | 60 | 9.59 | 7.34
2008-08-24 | 4:31 | 39 | 95.6 | 15.03 | 60 | 9.63 | 7.33
2008-08-24 | 4:46 | 38 | 95.7 | 14.92 | 59 | 9.67 | 7.33
2008-08-24 | 5:01 | 38 | 95.4 | 14.82 | 59 | 9.66 | 7.33
2008-08-24 | 5:16 | 38 | 95.7 | 14.71 | 58 | 9.71 | 7.32
2008-08-24 | 5:31 | 38 | 95.4 | 14.6 | 58 | 9.71 | 7.31
2008-08-24 | 5:46 | 38 | 95.8 | 14.49 | 58 | 9.77 | 7.32
2008-08-24 | 6:01 | 38 | 96 | 14.38 | 58 | 9.81 | 7.31
2008-08-24 | 6:16 | 38 | 95.7 | 14.28 | 58 | 9.81 | 7.3
2008-08-24 | 6:31 | 38 | 95.7 | 14.17 | 58 | 9.83 | 7.31
2008-08-24 | 6:46 | 38 | 96.2 | 14.05 | 58 | 9.9 | 7.31
2008-08-24 | 7:01 | 38 | 96.7 | 13.95 | 58 | 9.97 | 7.32
2008-08-24 | 7:16 | 38 | 97 | 13.85 | 58 | 10.03 | 7.33
2008-08-24 | 7:31 | 38 | 97 | 13.77 | 58 | 10.04 | 7.34
2008-08-24 | 7:46 | 38 | 97.6 | 13.7 | 58 | 10.13 | 7.35
2008-08-24 | 8:01 | 38 | 98 | 13.65 | 58 | 10.17 | 7.36
2008-08-24 | 8:16 | 38 | 97.7 | 13.6 | 58 | 10.16 | 7.37
2008-08-24 | 8:31 | 38 | 97.7 | 13.56 | 58 | 10.17 | 7.37
2008-08-24 | 8:46 | 38 | 97.8 | 13.53 | 58 | 10.19 | 7.38
2008-08-24 | 9:01 | 38 | 98.5 | 13.56 | 58 | 10.25 | 7.4
2008-08-24 | 9:16 | 38 | 98.9 | 13.59 | 58 | 10.28 | 7.41
2008-08-24 | 9:31 | 38 | 98.7 | 13.62 | 58 | 10.26 | 7.42
2008-08-24 | 9:46 | 38 | 99.1 | 13.69 | 58 | 10.29 | 7.43
2008-08-24 | 10:01 | 38 | 99.4 | 13.77 | 58 | 10.29 | 7.44
2008-08-24 | 10:16 | 38 | 99.6 | 13.85 | 59 | 10.29 | 7.45
2008-08-24 | 10:31 | 38 | 99.8 | 13.95 | 59 | 10.29 | 7.46
2008-08-24 | 10:46 | 38 | 99.8 | 14.07 | 59 | 10.27 | 7.48
2008-08-24 | 11:01 | 38 | 100.1 | 14.21 | 59 | 10.27 | 7.5
2008-08-24 | 11:16 | 38 | 100.4 | 14.37 | 59 | 10.26 | 7.51
2008-08-24 | 11:31 | 38 | 100.7 | 14.62 | 59 | 10.24 | 7.54
2008-08-24 | 11:46 | 38 | 100.9 | 14.93 | 59 | 10.19 | 7.56
2008-08-24 | 12:01 | 38 | 101 | 15.28 | 59 | 10.12 | 7.59
2008-08-24 | 12:16 | 38 | 101.1 | 15.66 | 59 | 10.05 | 7.64
2008-08-24 | 12:31 | 38 | 101.5 | 16.13 | 58 | 9.99 | 7.68
2008-08-24 | 12:46 | 38 | 101.6 | 16.51 | 58 | 9.92 | 7.69
2008-08-24 | 13:00 | 38 | 101.5 | 16.97 | 58 | 9.82 | 7.73
2008-08-24 | 13:15 | 38 | 101.4 | 17.46 | 58 | 9.71 | 7.75
2008-08-24 | 13:30 | 38 | 101.3 | 17.83 | 58 | 9.63 | 7.78
2008-08-24 | 13:45 | 38 | 101.2 | 18.29 | 58 | 9.53 | 7.8
2008-08-24 | 14:00 | 38 | 101 | 18.68 | 58 | 9.43 | 7.83
2008-08-24 | 14:15 | 38 | 101 | 19.09 | 58 | 9.35 | 7.86
2008-08-24 | 14:30 | 38 | 100.3 | 19.5 | 58 | 9.21 | 7.88
2008-08-24 | 14:45 | 38 | 99.8 | 19.78 | 58 | 9.12 | 7.88
2008-08-24 | 15:00 | 38 | 99.9 | 20.06 | 58 | 9.08 | 7.89
2008-08-24 | 15:15 | 38 | 99.5 | 20.29 | 58 | 8.99 | 7.91
2008-08-24 | 15:30 | 38 | 99.1 | 20.46 | 58 | 8.93 | 7.92
2008-08-24 | 15:45 | 38 | 98.9 | 20.53 | 58 | 8.89 | 7.92
2008-08-24 | 16:00 | 38 | 99 | 20.61 | 58 | 8.89 | 7.93
2008-08-24 | 16:15 | 37 | 98.9 | 20.64 | 57 | 8.88 | 7.91
2008-08-24 | 16:30 | 37 | 98.7 | 20.65 | 57 | 8.86 | 7.9
2008-08-24 | 16:45 | 37 | 98.1 | 20.57 | 57 | 8.82 | 7.89
2008-08-24 | 17:00 | 37 | 97.7 | 20.46 | 57 | 8.81 | 7.87
2008-08-24 | 17:15 | 37 | 97.6 | 20.35 | 57 | 8.82 | 7.83
2008-08-24 | 17:30 | 37 | 97.4 | 20.18 | 57 | 8.82 | 7.81
2008-08-24 | 17:45 | 37 | 96.8 | 20.08 | 57 | 8.79 | 7.76
2008-08-24 | 18:00 | 37 | 96.7 | 19.98 | 57 | 8.8 | 7.74
2008-08-24 | 18:15 | 37 | 96.4 | 19.88 | 57 | 8.79 | 7.7
2008-08-24 | 18:30 | 37 | 95.8 | 19.77 | 57 | 8.75 | 7.66
2008-08-24 | 18:45 | 37 | 95.5 | 19.66 | 57 | 8.74 | 7.62
2008-08-24 | 19:00 | 37 | 94.8 | 19.59 | 57 | 8.69 | 7.58
2008-08-24 | 19:15 | 37 | 94.6 | 19.51 | 57 | 8.69 | 7.54
2008-08-24 | 19:30 | 37 | 94.2 | 19.39 | 57 | 8.66 | 7.51
2008-08-24 | 19:45 | 37 | 93.9 | 19.27 | 57 | 8.66 | 7.48
2008-08-24 | 20:00 | 37 | 94 | 19.13 | 57 | 8.7 | 7.44
2008-08-24 | 20:15 | 37 | 93.7 | 19 | 57 | 8.69 | 7.43
2008-08-24 | 20:30 | 37 | 93.5 | 18.87 | 57 | 8.7 | 7.42
2008-08-24 | 20:45 | 37 | 93.8 | 18.7 | 57 | 8.75 | 7.41
2008-08-24 | 21:00 | 37 | 93.9 | 18.52 | 57 | 8.8 | 7.4
2008-08-24 | 21:15 | 37 | 94.1 | 18.36 | 57 | 8.84 | 7.38
2008-08-24 | 21:30 | 37 | 94.2 | 18.22 | 57 | 8.88 | 7.37
2008-08-24 | 21:45 | 37 | 93.9 | 18.1 | 57 | 8.87 | 7.37
2008-08-24 | 22:00 | 37 | 93.9 | 17.97 | 57 | 8.9 | 7.36
2008-08-24 | 22:15 | 37 | 94.2 | 17.86 | 57 | 8.94 | 7.36
2008-08-24 | 22:30 | 37 | 94.4 | 17.76 | 57 | 8.98 | 7.34
2008-08-24 | 22:45 | 37 | 94.1 | 17.66 | 57 | 8.98 | 7.34
2008-08-24 | 23:00 | 37 | 94 | 17.53 | 57 | 8.99 | 7.34
2008-08-24 | 23:15 | 37 | 93.8 | 17.43 | 57 | 8.98 | 7.34
2008-08-24 | 23:30 | 37 | 94 | 17.34 | 57 | 9.02 | 7.34
2008-08-24 | 23:45 | 37 | 94.1 | 17.26 | 57 | 9.04 | 7.33
2008-08-25 | 0:00 | 37 | 94.4 | 17.17 | 57 | 9.09 | 7.34
2008-08-25 | 0:15 | 37 | 94.2 | 17.09 | 57 | 9.09 | 7.35
2008-08-25 | 0:30 | 37 | 94.5 | 17.01 | 57 | 9.13 | 7.34
2008-08-25 | 0:45 | 37 | 94.6 | 16.93 | 57 | 9.15 | 7.34
2008-08-25 | 1:01 | 37 | 94.3 | 16.85 | 57 | 9.14 | 7.34
2008-08-25 | 1:16 | 37 | 94.3 | 16.78 | 57 | 9.15 | 7.34
2008-08-25 | 1:31 | 37 | 94.6 | 16.7 | 57 | 9.2 | 7.34
2008-08-25 | 1:46 | 37 | 94.7 | 16.64 | 57 | 9.22 | 7.34
2008-08-25 | 2:01 | 37 | 94.4 | 16.59 | 57 | 9.2 | 7.34
2008-08-25 | 2:16 | 37 | 94.7 | 16.54 | 57 | 9.24 | 7.33
2008-08-25 | 2:31 | 37 | 94.8 | 16.49 | 57 | 9.26 | 7.34
2008-08-25 | 2:46 | 37 | 94.9 | 16.44 | 57 | 9.28 | 7.34
2008-08-25 | 3:01 | 37 | 95 | 16.39 | 57 | 9.3 | 7.33
2008-08-25 | 3:16 | 37 | 94.9 | 16.33 | 57 | 9.31 | 7.33
2008-08-25 | 3:31 | 37 | 95 | 16.27 | 57 | 9.32 | 7.33
2008-08-25 | 3:46 | 37 | 94.6 | 16.2 | 57 | 9.3 | 7.34
2008-08-25 | 4:01 | 37 | 94.9 | 16.13 | 57 | 9.34 | 7.34
2008-08-25 | 4:16 | 37 | 94.7 | 16.06 | 57 | 9.33 | 7.33
2008-08-25 | 4:31 | 37 | 95 | 16 | 57 | 9.37 | 7.33
2008-08-25 | 4:46 | 37 | 94.6 | 15.94 | 57 | 9.35 | 7.33
2008-08-25 | 5:01 | 37 | 94.8 | 15.88 | 57 | 9.38 | 7.33
2008-08-25 | 5:16 | 37 | 94.6 | 15.83 | 57 | 9.37 | 7.33
2008-08-25 | 5:31 | 37 | 94.9 | 15.77 | 57 | 9.41 | 7.33
2008-08-25 | 5:46 | 37 | 94.7 | 15.71 | 57 | 9.41 | 7.33
2008-08-25 | 6:01 | 37 | 95 | 15.66 | 57 | 9.45 | 7.33
2008-08-25 | 6:16 | 37 | 94.8 | 15.61 | 57 | 9.44 | 7.33
2008-08-25 | 6:31 | 37 | 94.8 | 15.56 | 57 | 9.45 | 7.33
2008-08-25 | 6:46 | 37 | 95.1 | 15.51 | 57 | 9.49 | 7.33
2008-08-25 | 7:01 | 37 | 95.4 | 15.46 | 58 | 9.52 | 7.33
2008-08-25 | 7:16 | 37 | 95.6 | 15.41 | 58 | 9.55 | 7.35
2008-08-25 | 7:31 | 38 | 95.7 | 15.36 | 58 | 9.57 | 7.35
2008-08-25 | 7:46 | 38 | 95.8 | 15.31 | 58 | 9.6 | 7.35
2008-08-25 | 8:01 | 38 | 95.7 | 15.27 | 58 | 9.59 | 7.36
2008-08-25 | 8:16 | 38 | 96.3 | 15.23 | 58 | 9.66 | 7.37
2008-08-25 | 8:31 | 38 | 96.7 | 15.2 | 58 | 9.71 | 7.39
2008-08-25 | 8:46 | 38 | 97 | 15.17 | 58 | 9.75 | 7.41
2008-08-25 | 9:01 | 38 | 96.9 | 15.14 | 58 | 9.74 | 7.43
2008-08-25 | 9:16 | 38 | 97 | 15.12 | 58 | 9.75 | 7.44
2008-08-25 | 9:31 | 38 | 97.9 | 15.14 | 58 | 9.84 | 7.46
2008-08-25 | 9:46 | 37 | 98.4 | 15.19 | 58 | 9.88 | 7.5
2008-08-25 | 10:01 | 37 | 98.6 | 15.21 | 58 | 9.9 | 7.54
2008-08-25 | 10:16 | 37 | 99.1 | 15.23 | 57 | 9.94 | 7.56
2008-08-25 | 10:31 | 37 | 99.2 | 15.29 | 57 | 9.94 | 7.59
2008-08-25 | 10:46 | 37 | 99.9 | 15.37 | 57 | 10 | 7.62
2008-08-25 | 11:01 | 37 | 100.4 | 15.45 | 57 | 10.03 | 7.65
2008-08-25 | 11:16 | 37 | 100.2 | 15.52 | 57 | 9.99 | 7.68
2008-08-25 | 11:31 | 37 | 100.1 | 15.54 | 57 | 9.97 | 7.69
2008-08-25 | 11:46 | 37 | 100.1 | 15.66 | 57 | 9.95 | 7.7
2008-08-25 | 12:01 | 37 | 101 | 15.96 | 57 | 9.98 | 7.74
2008-08-25 | 12:16 | 37 | 101.6 | 16.19 | 56 | 9.98 | 7.76
2008-08-25 | 12:31 | 37 | 100.9 | 16.3 | 56 | 9.89 | 7.78
2008-08-25 | 12:46 | 37 | 101.4 | 16.65 | 56 | 9.87 | 7.8
2008-08-25 | 13:00 | 36 | 101.5 | 16.75 | 56 | 9.86 | 7.81
2008-08-25 | 13:15 | 36 | 101 | 16.97 | 56 | 9.77 | 7.84
2008-08-25 | 13:30 | 36 | 100.8 | 17.37 | 56 | 9.66 | 7.85
2008-08-25 | 13:45 | 36 | 100.8 | 17.58 | 56 | 9.62 | 7.84
2008-08-25 | 14:00 | 36 | 100.9 | 17.93 | 56 | 9.56 | 7.85
2008-08-25 | 14:15 | 36 | 100.6 | 18.29 | 56 | 9.47 | 7.86
2008-08-25 | 14:30 | 36 | 100.1 | 18.52 | 56 | 9.38 | 7.85
2008-08-25 | 14:45 | 36 | 99.6 | 18.8 | 56 | 9.28 | 7.87
2008-08-25 | 15:00 | 36 | 99.2 | 18.99 | 56 | 9.21 | 7.86
2008-08-25 | 15:15 | 36 | 98.9 | 19.13 | 56 | 9.15 | 7.86
2008-08-25 | 15:30 | 36 | 99 | 19.24 | 56 | 9.14 | 7.85
2008-08-25 | 15:45 | 36 | 98.5 | 19.27 | 55 | 9.09 | 7.85
2008-08-25 | 16:00 | 36 | 98.4 | 19.23 | 55 | 9.09 | 7.84
2008-08-25 | 16:15 | 36 | 98.1 | 19.17 | 55 | 9.07 | 7.81
2008-08-25 | 16:30 | 36 | 98 | 19.05 | 55 | 9.08 | 7.79
2008-08-25 | 16:45 | 36 | 97.8 | 18.94 | 55 | 9.08 | 7.76
2008-08-25 | 17:00 | 36 | 97.2 | 18.75 | 55 | 9.06 | 7.73
2008-08-25 | 17:15 | 36 | 96.8 | 18.56 | 55 | 9.06 | 7.69
2008-08-25 | 17:30 | 36 | 96.8 | 18.37 | 55 | 9.09 | 7.65
2008-08-25 | 17:45 | 36 | 96.3 | 18.19 | 55 | 9.08 | 7.62
2008-08-25 | 18:00 | 36 | 96.1 | 18 | 55 | 9.09 | 7.58
2008-08-25 | 18:15 | 36 | 96.1 | 17.85 | 55 | 9.12 | 7.55
2008-08-25 | 18:30 | 36 | 95.7 | 17.73 | 55 | 9.11 | 7.53
2008-08-25 | 18:45 | 36 | 95.5 | 17.62 | 55 | 9.12 | 7.5
2008-08-25 | 19:00 | 36 | 95.7 | 17.46 | 55 | 9.16 | 7.48
2008-08-25 | 19:15 | 36 | 95.7 | 17.32 | 55 | 9.18 | 7.47
2008-08-25 | 19:30 | 36 | 95.2 | 17.19 | 55 | 9.17 | 7.45
2008-08-25 | 19:45 | 36 | 95.2 | 17.07 | 55 | 9.19 | 7.43
2008-08-25 | 20:00 | 36 | 95.2 | 16.97 | 55 | 9.21 | 7.42
2008-08-25 | 20:15 | 36 | 95.1 | 16.85 | 55 | 9.22 | 7.39
2008-08-25 | 20:30 | 36 | 94.8 | 16.71 | 55 | 9.22 | 7.39
2008-08-25 | 20:45 | 36 | 94.6 | 16.55 | 55 | 9.23 | 7.37
2008-08-25 | 21:00 | 36 | 95 | 16.38 | 55 | 9.31 | 7.37
2008-08-25 | 21:15 | 36 | 94.8 | 16.2 | 55 | 9.32 | 7.36
2008-08-25 | 21:30 | 36 | 95.1 | 16.01 | 55 | 9.39 | 7.25
2008-08-25 | 21:45 | 36 | 95.4 | 15.82 | 55 | 9.45 | 7.31
2008-08-25 | 22:00 | 36 | 95.5 | 15.64 | 55 | 9.5 | 7.33
2008-08-25 | 22:15 | 36 | 95.6 | 15.46 | 55 | 9.54 | 7.33
2008-08-25 | 22:30 | 36 | 95.7 | 15.27 | 55 | 9.59 | 7.33
2008-08-25 | 22:45 | 36 | 95.4 | 15.1 | 55 | 9.6 | 7.32
2008-08-25 | 23:00 | 36 | 95.3 | 14.94 | 55 | 9.63 | 7.31
2008-08-25 | 23:15 | 36 | 95.4 | 14.78 | 55 | 9.66 | 7.31
2008-08-25 | 23:30 | 36 | 95.4 | 14.62 | 55 | 9.7 | 7.31
2008-08-25 | 23:45 | 36 | 95.5 | 14.47 | 55 | 9.74 | 7.31
2008-08-26 | 0:00 | 36 | 95.5 | 14.31 | 55 | 9.77 | 7.31
2008-08-26 | 0:15 | 36 | 95.5 | 14.16 | 55 | 9.81 | 7.31
2008-08-26 | 0:30 | 36 | 95.7 | 14.02 | 55 | 9.86 | 7.3
2008-08-26 | 0:45 | 36 | 95.7 | 13.87 | 55 | 9.89 | 7.31
2008-08-26 | 1:01 | 36 | 96.1 | 13.72 | 55 | 9.97 | 7.3
2008-08-26 | 1:16 | 36 | 96 | 13.58 | 55 | 9.98 | 7.3
2008-08-26 | 1:31 | 36 | 96 | 13.43 | 56 | 10.01 | 7.3
2008-08-26 | 1:46 | 36 | 96 | 13.29 | 56 | 10.05 | 7.3
2008-08-26 | 2:01 | 36 | 96.4 | 13.15 | 56 | 10.12 | 7.29
2008-08-26 | 2:16 | 36 | 96.6 | 13.03 | 56 | 10.17 | 7.29
2008-08-26 | 2:31 | 36 | 96.8 | 12.91 | 56 | 10.22 | 7.29
2008-08-26 | 2:46 | 36 | 96.4 | 12.8 | 56 | 10.2 | 7.29
2008-08-26 | 3:01 | 36 | 96.4 | 12.69 | 56 | 10.22 | 7.29
2008-08-26 | 3:16 | 36 | 96.3 | 12.59 | 56 | 10.24 | 7.29
2008-08-26 | 3:31 | 36 | 96.3 | 12.5 | 56 | 10.27 | 7.28
2008-08-26 | 3:46 | 36 | 96.7 | 12.4 | 56 | 10.32 | 7.29
2008-08-26 | 4:01 | 37 | 96.5 | 12.3 | 56 | 10.33 | 7.29
2008-08-26 | 4:16 | 37 | 96.6 | 12.19 | 56 | 10.36 | 7.29
2008-08-26 | 4:31 | 37 | 96.5 | 12.08 | 56 | 10.38 | 7.29
2008-08-26 | 4:46 | 37 | 96.9 | 11.96 | 56 | 10.46 | 7.29
2008-08-26 | 5:01 | 37 | 96.8 | 11.83 | 57 | 10.47 | 7.29
2008-08-26 | 5:16 | 37 | 97.1 | 11.7 | 57 | 10.54 | 7.29
2008-08-26 | 5:31 | 37 | 97 | 11.57 | 57 | 10.56 | 7.29
2008-08-26 | 5:46 | 37 | 97.3 | 11.42 | 57 | 10.62 | 7.29
2008-08-26 | 6:01 | 37 | 97.1 | 11.28 | 57 | 10.64 | 7.28
2008-08-26 | 6:16 | 37 | 97.1 | 11.14 | 57 | 10.67 | 7.29
2008-08-26 | 6:31 | 37 | 97.1 | 11 | 57 | 10.71 | 7.28
2008-08-26 | 6:46 | 37 | 97.8 | 10.87 | 57 | 10.81 | 7.29
2008-08-26 | 7:01 | 37 | 98.1 | 10.78 | 57 | 10.88 | 7.3
2008-08-26 | 7:16 | 37 | 98.2 | 10.7 | 57 | 10.9 | 7.3
2008-08-26 | 7:31 | 37 | 98.4 | 10.61 | 57 | 10.95 | 7.31
2008-08-26 | 7:46 | 37 | 98.7 | 10.54 | 57 | 11 | 7.33
2008-08-26 | 8:01 | 37 | 98.7 | 10.45 | 57 | 11.02 | 7.33
2008-08-26 | 8:16 | 37 | 98.8 | 10.38 | 57 | 11.05 | 7.34
2008-08-26 | 8:31 | 37 | 99.4 | 10.34 | 57 | 11.13 | 7.34
2008-08-26 | 8:46 | 37 | 99.4 | 10.32 | 57 | 11.13 | 7.35
2008-08-26 | 9:01 | 37 | 100 | 10.34 | 57 | 11.2 | 7.36
2008-08-26 | 9:16 | 37 | 99.9 | 10.38 | 57 | 11.18 | 7.37
2008-08-26 | 9:31 | 37 | 100 | 10.42 | 57 | 11.18 | 7.39
2008-08-26 | 9:46 | 37 | 100.6 | 10.48 | 57 | 11.23 | 7.4
2008-08-26 | 10:01 | 37 | 100.6 | 10.57 | 57 | 11.2 | 7.41
2008-08-26 | 10:16 | 37 | 101 | 10.64 | 57 | 11.23 | 7.42
2008-08-26 | 10:31 | 37 | 101.3 | 10.74 | 57 | 11.24 | 7.42
2008-08-26 | 10:46 | 37 | 101.7 | 10.85 | 57 | 11.25 | 7.44
2008-08-26 | 11:01 | 37 | 102 | 10.98 | 57 | 11.25 | 7.45
2008-08-26 | 11:16 | 37 | 101.8 | 11.16 | 57 | 11.19 | 7.46
2008-08-26 | 11:31 | 37 | 101.8 | 11.39 | 57 | 11.13 | 7.47
2008-08-26 | 11:46 | 37 | 102.2 | 11.7 | 57 | 11.09 | 7.49
2008-08-26 | 12:01 | 37 | 102 | 12.04 | 56 | 10.98 | 7.51
2008-08-26 | 12:16 | 37 | 101.8 | 12.44 | 56 | 10.86 | 7.53
2008-08-26 | 12:31 | 37 | 102.1 | 12.88 | 56 | 10.79 | 7.55
2008-08-26 | 12:46 | 37 | 101.8 | 13.33 | 56 | 10.65 | 7.57
2008-08-26 | 13:00 | 37 | 101.7 | 13.8 | 56 | 10.52 | 7.59
2008-08-26 | 13:15 | 37 | 101.9 | 14.28 | 56 | 10.44 | 7.61
2008-08-26 | 13:30 | 37 | 101.6 | 14.76 | 56 | 10.3 | 7.63
2008-08-26 | 13:45 | 36 | 101.2 | 15.23 | 56 | 10.16 | 7.67
2008-08-26 | 14:00 | 36 | 101 | 15.72 | 56 | 10.02 | 7.68
2008-08-26 | 14:15 | 36 | 100.8 | 16.19 | 56 | 9.91 | 7.7
2008-08-26 | 14:30 | 36 | 100.9 | 16.64 | 56 | 9.83 | 7.73
2008-08-26 | 14:45 | 36 | 100.8 | 16.99 | 56 | 9.74 | 7.75
2008-08-26 | 15:00 | 36 | 100.2 | 17.29 | 56 | 9.63 | 7.76
2008-08-26 | 15:15 | 36 | 99.8 | 17.57 | 56 | 9.54 | 7.78
2008-08-26 | 15:30 | 36 | 99.8 | 17.82 | 56 | 9.48 | 7.79
2008-08-26 | 15:45 | 36 | 99.7 | 18.03 | 56 | 9.43 | 7.8
2008-08-26 | 16:00 | 36 | 99.5 | 18.13 | 56 | 9.39 | 7.79
2008-08-26 | 16:15 | 36 | 98.8 | 18.2 | 56 | 9.32 | 7.79
2008-08-26 | 16:30 | 36 | 98.8 | 18.16 | 56 | 9.32 | 7.77
2008-08-26 | 16:45 | 36 | 98.1 | 18.09 | 56 | 9.27 | 7.74
2008-08-26 | 17:00 | 36 | 98 | 18 | 56 | 9.28 | 7.72
2008-08-26 | 17:15 | 36 | 97.5 | 17.84 | 56 | 9.26 | 7.69
2008-08-26 | 17:30 | 36 | 97 | 17.74 | 56 | 9.24 | 7.65
2008-08-26 | 17:45 | 36 | 96.8 | 17.63 | 56 | 9.23 | 7.62
2008-08-26 | 18:00 | 36 | 96.8 | 17.52 | 56 | 9.26 | 7.6
2008-08-26 | 18:15 | 36 | 96.7 | 17.45 | 56 | 9.25 | 7.57
2008-08-26 | 18:30 | 36 | 96.5 | 17.38 | 56 | 9.25 | 7.54
2008-08-26 | 18:45 | 36 | 96 | 17.29 | 56 | 9.22 | 7.52
2008-08-26 | 19:00 | 36 | 96 | 17.25 | 56 | 9.23 | 7.5
2008-08-26 | 19:15 | 36 | 95.5 | 17.22 | 56 | 9.19 | 7.48
2008-08-26 | 19:30 | 37 | 95.5 | 17.16 | 56 | 9.2 | 7.47
2008-08-26 | 19:45 | 37 | 95.1 | 17.11 | 56 | 9.17 | 7.45
2008-08-26 | 20:00 | 37 | 94.8 | 17.06 | 56 | 9.15 | 7.43
2008-08-26 | 20:15 | 37 | 94.9 | 16.99 | 56 | 9.18 | 7.43
2008-08-26 | 20:30 | 37 | 94.7 | 16.87 | 56 | 9.17 | 7.4
2008-08-26 | 20:45 | 37 | 94.6 | 16.74 | 56 | 9.19 | 7.39
2008-08-26 | 21:00 | 37 | 94.8 | 16.61 | 56 | 9.24 | 7.39
2008-08-26 | 21:15 | 37 | 95 | 16.48 | 56 | 9.28 | 7.37
2008-08-26 | 21:30 | 37 | 94.8 | 16.34 | 56 | 9.29 | 7.37
2008-08-26 | 21:45 | 37 | 95.1 | 16.17 | 56 | 9.35 | 7.37
2008-08-26 | 22:00 | 45 | 95.3 | 16.01 | 69 | 9.4 | 7.36
2008-08-26 | 22:15 | 45 | 95.3 | 15.86 | 69 | 9.44 | 7.36
2008-08-26 | 22:30 | 45 | 94.9 | 15.7 | 69 | 9.43 | 7.36
2008-08-26 | 22:45 | 45 | 95.4 | 15.54 | 69 | 9.51 | 7.35
2008-08-26 | 23:00 | 45 | 95.6 | 15.39 | 69 | 9.56 | 7.36
2008-08-26 | 23:15 | 43 | 95.5 | 15.24 | 66 | 9.58 | 7.34
2008-08-26 | 23:30 | 43 | 95.6 | 15.1 | 66 | 9.62 | 7.35
2008-08-26 | 23:45 | 43 | 95.7 | 14.96 | 66 | 9.66 | 7.34
2008-08-27 | 0:00 | 44 | 95.8 | 14.83 | 68 | 9.7 | 7.34
2008-08-27 | 0:15 | 43 | 95.9 | 14.73 | 67 | 9.73 | 7.34
2008-08-27 | 0:30 | 44 | 96 | 14.62 | 68 | 9.76 | 7.34
2008-08-27 | 0:45 | 37 | 96 | 14.53 | 57 | 9.78 | 7.33
2008-08-27 | 1:01 | 37 | 95.7 | 14.43 | 57 | 9.77 | 7.33
2008-08-27 | 1:16 | 37 | 95.9 | 14.33 | 57 | 9.81 | 7.33
2008-08-27 | 1:31 | 37 | 96.1 | 14.24 | 57 | 9.85 | 7.33
2008-08-27 | 1:46 | 37 | 95.8 | 14.13 | 57 | 9.84 | 7.33
2008-08-27 | 2:01 | 37 | 96 | 14.03 | 57 | 9.89 | 7.33
2008-08-27 | 2:16 | 37 | 95.8 | 13.94 | 57 | 9.88 | 7.33
2008-08-27 | 2:31 | 37 | 95.7 | 13.88 | 57 | 9.89 | 7.33
2008-08-27 | 2:46 | 37 | 95.7 | 13.79 | 57 | 9.91 | 7.34
2008-08-27 | 3:01 | 37 | 96.1 | 13.71 | 57 | 9.96 | 7.33
2008-08-27 | 3:16 | 37 | 95.9 | 13.66 | 57 | 9.95 | 7.33
2008-08-27 | 3:31 | 37 | 95.8 | 13.59 | 57 | 9.96 | 7.33
2008-08-27 | 3:46 | 44 | 96.2 | 13.53 | 67 | 10.01 | 7.33
2008-08-27 | 4:01 | 44 | 96.4 | 13.46 | 67 | 10.05 | 7.33
2008-08-27 | 4:16 | 44 | 96.5 | 13.39 | 67 | 10.07 | 7.33
2008-08-27 | 4:31 | 44 | 96.2 | 13.32 | 68 | 10.06 | 7.33
2008-08-27 | 4:46 | 45 | 96.1 | 13.24 | 69 | 10.07 | 7.33
2008-08-27 | 5:01 | 45 | 96.5 | 13.18 | 70 | 10.13 | 7.33
2008-08-27 | 5:16 | 45 | 96.3 | 13.13 | 69 | 10.11 | 7.34
2008-08-27 | 5:31 | 45 | 96.6 | 13.06 | 70 | 10.16 | 7.32
2008-08-27 | 5:46 | 48 | 96.3 | 12.99 | 74 | 10.15 | 7.33
2008-08-27 | 6:01 | 48 | 96.3 | 12.92 | 73 | 10.16 | 7.33
2008-08-27 | 6:16 | 48 | 96.3 | 12.85 | 73 | 10.18 | 7.33
2008-08-27 | 6:31 | 48 | 96.7 | 12.78 | 73 | 10.24 | 7.33
2008-08-27 | 6:46 | 48 | 96.6 | 12.74 | 73 | 10.24 | 7.33
2008-08-27 | 7:01 | 48 | 97.2 | 12.69 | 74 | 10.31 | 7.34
2008-08-27 | 7:16 | 48 | 97.7 | 12.65 | 74 | 10.37 | 7.35
2008-08-27 | 7:31 | 48 | 98.1 | 12.6 | 74 | 10.43 | 7.36
2008-08-27 | 7:46 | 48 | 98.1 | 12.57 | 74 | 10.43 | 7.39
2008-08-27 | 8:01 | 48 | 98.7 | 12.54 | 74 | 10.51 | 7.4
2008-08-27 | 8:16 | 48 | 99.2 | 12.53 | 74 | 10.56 | 7.42
2008-08-27 | 8:31 | 48 | 99.6 | 12.53 | 74 | 10.61 | 7.45
2008-08-27 | 8:46 | 48 | 100.1 | 12.53 | 74 | 10.66 | 7.48
2008-08-27 | 9:01 | 49 | 100.1 | 12.57 | 75 | 10.65 | 7.51
2008-08-27 | 9:16 | 49 | 100.3 | 12.59 | 75 | 10.66 | 7.52
2008-08-27 | 9:31 | 49 | 100.7 | 12.6 | 75 | 10.71 | 7.54
2008-08-27 | 9:46 | 49 | 100.7 | 12.65 | 75 | 10.7 | 7.56
2008-08-27 | 10:01 | 49 | 101.2 | 12.71 | 75 | 10.74 | 7.58
2008-08-27 | 10:16 | 49 | 100.9 | 12.76 | 75 | 10.69 | 7.6
2008-08-27 | 10:31 | 49 | 101.1 | 12.82 | 75 | 10.69 | 7.61
2008-08-27 | 10:46 | 49 | 100.9 | 12.88 | 75 | 10.66 | 7.6
2008-08-27 | 11:01 | 48 | 100.8 | 12.97 | 75 | 10.62 | 7.63
2008-08-27 | 11:16 | 49 | 100.8 | 13.05 | 75 | 10.61 | 7.63
2008-08-27 | 11:31 | 50 | 101.1 | 13.22 | 76 | 10.6 | 7.65
2008-08-27 | 11:46 | 50 | 101.7 | 13.38 | 76 | 10.62 | 7.68
2008-08-27 | 12:01 | 49 | 101.5 | 13.54 | 76 | 10.57 | 7.7
2008-08-27 | 12:16 | 49 | 101.6 | 13.85 | 76 | 10.51 | 7.72
2008-08-27 | 12:31 | 49 | 101.4 | 14.07 | 76 | 10.43 | 7.75
2008-08-27 | 12:46 | 50 | 102.2 | 14.56 | 77 | 10.4 | 7.77
2008-08-27 | 13:00 | 50 | 101.9 | 15.01 | 77 | 10.28 | 7.79
2008-08-27 | 13:15 | 52 | 101.7 | 15.42 | 80 | 10.16 | 7.8
2008-08-27 | 13:30 | 46 | 101.9 | 15.94 | 70 | 10.07 | 7.81
2008-08-27 | 13:45 | 46 | 101.9 | 16.47 | 70 | 9.95 | 7.84
2008-08-27 | 14:00 | 46 | 101.7 | 16.98 | 70 | 9.84 | 7.85
2008-08-27 | 14:15 | 45 | 101.5 | 17.6 | 70 | 9.68 | 7.87
2008-08-27 | 14:30 | 45 | 101.1 | 18.17 | 70 | 9.54 | 7.91
2008-08-27 | 14:45 | 46 | 100.4 | 18.63 | 71 | 9.38 | 7.93
2008-08-27 | 15:00 | 46 | 100 | 18.95 | 71 | 9.29 | 8.03
2008-08-27 | 15:15 | 45 | 100 | 19.24 | 70 | 9.23 | 8.04
2008-08-27 | 15:30 | 50 | 99.8 | 19.42 | 77 | 9.18 | 8.03
2008-08-27 | 15:45 | 49 | 99.2 | 19.59 | 76 | 9.09 | 8.01
2008-08-27 | 16:00 | 49 | 99.2 | 19.68 | 76 | 9.08 | 7.99
2008-08-27 | 16:15 | 49 | 98.7 | 19.66 | 76 | 9.04 | 7.96
2008-08-27 | 16:30 | 49 | 98.6 | 19.49 | 76 | 9.05 | 7.9
2008-08-27 | 16:45 | 49 | 98.4 | 19.34 | 75 | 9.06 | 7.85
2008-08-27 | 17:00 | 49 | 98 | 19.12 | 75 | 9.07 | 7.81
2008-08-27 | 17:15 | 49 | 97.6 | 18.89 | 75 | 9.07 | 7.75
2008-08-27 | 17:30 | 48 | 97.3 | 18.71 | 74 | 9.08 | 7.7
2008-08-27 | 17:45 | 48 | 97.2 | 18.5 | 74 | 9.11 | 7.66
2008-08-27 | 18:00 | 48 | 96.9 | 18.3 | 74 | 9.11 | 7.63
2008-08-27 | 18:15 | 48 | 96.5 | 18.12 | 74 | 9.12 | 7.58
2008-08-27 | 18:30 | 49 | 96 | 17.93 | 75 | 9.1 | 7.55
2008-08-27 | 18:45 | 48 | 96.2 | 17.79 | 74 | 9.15 | 7.53
2008-08-27 | 19:00 | 48 | 96 | 17.66 | 75 | 9.15 | 7.51
2008-08-27 | 19:15 | 49 | 95.6 | 17.58 | 75 | 9.13 | 7.5
2008-08-27 | 19:30 | 49 | 95.4 | 17.51 | 75 | 9.12 | 7.49
2008-08-27 | 19:45 | 49 | 95.5 | 17.43 | 75 | 9.14 | 7.47
2008-08-27 | 20:00 | 49 | 95.4 | 17.33 | 75 | 9.16 | 7.45
2008-08-27 | 20:15 | 44 | 95 | 17.22 | 68 | 9.14 | 7.43
2008-08-27 | 20:30 | 40 | 94.9 | 17.1 | 62 | 9.15 | 7.42
2008-08-27 | 20:45 | 39 | 94.8 | 16.97 | 60 | 9.17 | 7.41
2008-08-27 | 21:00 | 39 | 95.1 | 16.82 | 60 | 9.23 | 7.4
2008-08-27 | 21:15 | 39 | 95 | 16.66 | 60 | 9.25 | 7.4
2008-08-27 | 21:30 | 39 | 95.1 | 16.47 | 60 | 9.29 | 7.39
2008-08-27 | 21:45 | 39 | 95.1 | 16.26 | 60 | 9.34 | 7.38
2008-08-27 | 22:00 | 39 | 95.5 | 16.05 | 60 | 9.42 | 7.38
2008-08-27 | 22:15 | 43 | 95.7 | 15.84 | 65 | 9.48 | 7.38
2008-08-27 | 22:30 | 42 | 95.6 | 15.64 | 65 | 9.51 | 7.37
2008-08-27 | 22:45 | 42 | 95.9 | 15.45 | 65 | 9.57 | 7.37
2008-08-27 | 23:00 | 42 | 95.7 | 15.25 | 65 | 9.6 | 7.36
2008-08-27 | 23:15 | 39 | 96.1 | 15.05 | 60 | 9.68 | 7.36
2008-08-27 | 23:30 | 39 | 96.2 | 14.85 | 60 | 9.74 | 7.36
2008-08-27 | 23:45 | 39 | 96 | 14.67 | 60 | 9.75 | 7.35
2008-08-28 | 0:00 | 39 | 96.3 | 14.49 | 60 | 9.82 | 7.35
2008-08-28 | 0:15 | 39 | 96.5 | 14.34 | 60 | 9.87 | 7.33
2008-08-28 | 0:30 | 39 | 96.3 | 14.19 | 59 | 9.88 | 7.34
2008-08-28 | 0:45 | 39 | 96.5 | 14.05 | 59 | 9.94 | 7.34
2008-08-28 | 1:01 | 39 | 96.3 | 13.91 | 59 | 9.95 | 7.34
2008-08-28 | 1:16 | 39 | 96.3 | 13.79 | 60 | 9.97 | 7.33
2008-08-28 | 1:31 | 39 | 96.6 | 13.66 | 60 | 10.03 | 7.34
2008-08-28 | 1:46 | 39 | 96.8 | 13.54 | 60 | 10.07 | 7.33
2008-08-28 | 2:01 | 39 | 96.5 | 13.44 | 60 | 10.07 | 7.34
2008-08-28 | 2:16 | 39 | 96.4 | 13.33 | 60 | 10.09 | 7.34
2008-08-28 | 2:31 | 39 | 96.8 | 13.22 | 60 | 10.15 | 7.33
2008-08-28 | 2:46 | 39 | 96.9 | 13.11 | 60 | 10.19 | 7.33
2008-08-28 | 3:01 | 39 | 96.7 | 13.01 | 60 | 10.18 | 7.32
2008-08-28 | 3:16 | 39 | 96.6 | 12.9 | 60 | 10.2 | 7.33
2008-08-28 | 3:31 | 39 | 96.6 | 12.79 | 60 | 10.22 | 7.33
2008-08-28 | 3:46 | 39 | 96.6 | 12.7 | 60 | 10.25 | 7.32
2008-08-28 | 4:01 | 39 | 97 | 12.59 | 61 | 10.31 | 7.33
2008-08-28 | 4:16 | 39 | 97.2 | 12.48 | 61 | 10.36 | 7.33
2008-08-28 | 4:31 | 39 | 96.9 | 12.38 | 61 | 10.36 | 7.32
2008-08-28 | 4:46 | 39 | 96.8 | 12.28 | 61 | 10.37 | 7.32
2008-08-28 | 5:01 | 39 | 97.3 | 12.18 | 61 | 10.44 | 7.33
2008-08-28 | 5:16 | 40 | 97.4 | 12.08 | 61 | 10.48 | 7.32
2008-08-28 | 5:31 | 40 | 97.5 | 11.98 | 61 | 10.51 | 7.32
2008-08-28 | 5:46 | 40 | 97.2 | 11.89 | 61 | 10.51 | 7.33
2008-08-28 | 6:01 | 40 | 97.5 | 11.78 | 61 | 10.56 | 7.32
2008-08-28 | 6:16 | 40 | 97.3 | 11.68 | 61 | 10.56 | 7.32
2008-08-28 | 6:31 | 40 | 97.7 | 11.57 | 61 | 10.63 | 7.32
2008-08-28 | 6:46 | 40 | 97.9 | 11.46 | 61 | 10.69 | 7.32
2008-08-28 | 7:01 | 40 | 97.8 | 11.36 | 61 | 10.7 | 7.33
2008-08-28 | 7:16 | 40 | 98.3 | 11.27 | 61 | 10.77 | 7.34
2008-08-28 | 7:31 | 40 | 98.6 | 11.19 | 61 | 10.83 | 7.33
2008-08-28 | 7:46 | 40 | 98.5 | 11.12 | 61 | 10.83 | 7.33
2008-08-28 | 8:01 | 40 | 98.6 | 11.06 | 61 | 10.85 | 7.34
2008-08-28 | 8:16 | 40 | 98.7 | 11.02 | 61 | 10.87 | 7.35
2008-08-28 | 8:31 | 40 | 99.2 | 10.99 | 61 | 10.94 | 7.37
2008-08-28 | 8:46 | 40 | 99.1 | 10.99 | 61 | 10.93 | 7.37
2008-08-28 | 9:01 | 40 | 99.7 | 11.02 | 61 | 10.99 | 7.38
2008-08-28 | 9:16 | 40 | 100.1 | 11.05 | 61 | 11.03 | 7.4
2008-08-28 | 9:31 | 40 | 100.4 | 11.09 | 62 | 11.05 | 7.41
2008-08-28 | 9:46 | 40 | 100.7 | 11.16 | 61 | 11.07 | 7.42
2008-08-28 | 10:01 | 40 | 101 | 11.25 | 61 | 11.07 | 7.42
2008-08-28 | 10:16 | 40 | 101.2 | 11.33 | 61 | 11.07 | 7.44
2008-08-28 | 10:31 | 40 | 101 | 11.43 | 61 | 11.03 | 7.45
2008-08-28 | 10:46 | 40 | 101.1 | 11.55 | 61 | 11.01 | 7.47
2008-08-28 | 11:01 | 40 | 101.4 | 11.7 | 61 | 11 | 7.47
2008-08-28 | 11:16 | 40 | 102 | 11.85 | 61 | 11.03 | 7.48
2008-08-28 | 11:31 | 40 | 102 | 12.08 | 61 | 10.97 | 7.5
2008-08-28 | 11:46 | 40 | 102.4 | 12.39 | 61 | 10.94 | 7.53
2008-08-28 | 12:01 | 39 | 102.5 | 12.77 | 61 | 10.85 | 7.55