## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AB (Aduited Against Sonde AA, 556, Titration & NIST)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-10-03 |  | 2008-10-03 |  | A
Time | Computer set to atomic time |  | 14:45 |  | 14:45 |  | A
Temp | VWR NIST traceable surface water |  | 8.2 |  | 8.45 |  | A
DO | Winkler titration of surface water |  | 10.7 |  | 10.79 |  | A
SpC | In Situ |  | 43 |  | 42 |  | A
Ph | In Situ |  | 7.35 |  | 6.69 |  | Fail
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2008-10-07 15:15 | 13.54 | 7.63 | 10.37 | 99.6 | 65 | 45
Sonde | 2008-10-07 15:15 | 13.57 | 7.45 | 10.4 | 100 | 63 | 41
Grade |  | A | A | A |  | A
Audit 2 | 2008-10-14 15:05 | 9.87 | 7.12 | 11.86 | 104.8 | 54 | 38
Sonde | 2008-10-14 15:00 | 9.86 | 7.49 | 11.4 | 100.7 | 53 | 34
Grade |  | A | B | B |  | A
Audit 3 | 2008-10-19 14:15 | 8.17 | 6.75 | 12.26 | 104 | 45 | 31
Sonde | 2008-10-19 14:15 | 8.17 | 7.24 | 12.13 | 102.9 | 44 | 29
Grade |  | A | B | A |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AB Audited In Situ against AA, 556, NIST, & Titration
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-10-19 |  | 2008-10-19 |  | A
Time | Computer set to atomic time |  | 15:30 |  | 15:30 |  | A
Temp | VWR NIST traceable surface water |  | 4.4 |  | 4.26 |  | A
DO | Winkler titration of surface water |  | 12.2 |  | 12.63 |  | B
SpC | In Situ |  | 43 |  | 42 |  | A
Ph | In Situ |  | 7.41 |  | 7.4 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-10-06 |  | 2008-10-06 |  | A
Time | Computer set to atomic time |  | 14:42 |  | 14:42 |  | A
Temp | VWR NIST traceable surface water |  | 8 |  | 8.16 |  | A
DO | Winkler titration of surface water |  | 11.2 |  | 11.65 |  | B
SpC | 148 µS/mL |  | 153.1 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.05 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-10-07 15:15 | 13.54 | 7.63 | 10.37 | 99.6 | 65 | 45
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-10-07 |  | 2008-10-07 |  | A
Time | Computer set to atomic time |  | 16:47 |  | 16:47 |  | A
Temp | VWR NIST traceable surface water |  | 14.16 |  | 14.32 |  | A
DO | Winkler titration of surface water |  | 11.5 |  | 11.79 |  | A
SpC | 148 µS/mL |  | 153.1 |  | 146 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.17 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-10-13 |  | 2008-10-13 |  | A
Time | Computer set to atomic time |  | 15:26 |  | 15:26 |  | A
Temp | VWR NIST traceable surface water |  | 10.13 |  | 9.97 |  | A
DO | Winkler titration of surface water |  | 12.5 |  | 12.7 |  | A
SpC | 148 µS/mL |  | 147 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.16 |  | A
Turb | Gelex secondary 0-5 NTU |  | 5 |  | 4.85 |  | A
Turb | Gelex secondary 0-500 NTU |  | 50 |  | 50.3 |  | A
Turb | Gelex secondary 0-1000 NTU |  | 478 |  | 477 |  | B
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-10-14 15:05 | 9.87 | 7.12 | 11.86 | 104.8 | 54 | 38
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-10-14 |  | 2008-10-14 |  | A
Time | Computer set to atomic time |  | 16:15 |  | 16:15 |  | A
Temp | VWR NIST traceable surface water |  | 9.2 |  | 9.11 |  | A
DO | Winkler titration of surface water |  | 12.5 |  | 13.42 |  | B
SpC | 148 µS/mL |  | 147 |  | 146 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.19 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-10-18 |  | 2008-10-18 |  | A
Time | Computer set to atomic time |  | 12:27 |  | 12:27 |  | A
Temp | VWR NIST traceable surface water |  | 10.11 |  | 9.91 |  | A
DO | Winkler titration of surface water |  | 11.2 |  | 11.15 |  | A
SpC | 148 µS/mL |  | 147 |  | 146 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.17 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-10-19 14:15 | 8.17 | 6.75 | 12.26 | 104 | 45 | 31
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-10-21 |  | 2008-10-21 |  | A
Time | Computer set to atomic time |  | 15:23 |  | 15:30 |  | A
Temp | VWR NIST traceable surface water |  | 4.4 |  | 4.22 |  | A
DO | 556 Against Sonde AA in Situ |  | 12.79 |  | 13.75 |  | B
SpC | 148 µS/mL |  | 147 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.13 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Note: Tritrition was 12.2, but I have doubts that the 556 was that far off for the samples before this audid, it didn’t have time to warm up before this audit due to an error in the intended DO titration sample; take DO data with this in mind |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | TDS | DO% | Temp | SpCond | DO Conc | pH
M/D/Y |  | mg/L | % | ˚C | uS/cm | mg/L
2008-10-07 | 15:15 | 41 | 100 | 13.57 | 63 | 10.4 | 7.45
2008-10-07 | 15:30 | 41 | 99.9 | 13.57 | 63 | 10.39 | 7.48
2008-10-07 | 15:45 | 41 | 99.7 | 13.57 | 63 | 10.38 | 7.47
2008-10-07 | 16:00 | 41 | 99.2 | 13.5 | 63 | 10.33 | 7.45
2008-10-07 | 16:15 | 41 | 99.1 | 13.38 | 63 | 10.35 | 7.44
2008-10-07 | 16:30 | 41 | 99 | 13.21 | 63 | 10.38 | 7.42
2008-10-07 | 16:45 | 41 | 98.5 | 13.05 | 63 | 10.37 | 7.4
2008-10-07 | 17:00 | 41 | 98.3 | 12.84 | 63 | 10.39 | 7.39
2008-10-07 | 17:15 | 41 | 97.8 | 12.65 | 63 | 10.38 | 7.38
2008-10-07 | 17:30 | 41 | 98 | 12.51 | 63 | 10.44 | 7.37
2008-10-07 | 17:45 | 41 | 97.6 | 12.41 | 63 | 10.42 | 7.35
2008-10-07 | 18:00 | 41 | 97.5 | 12.28 | 63 | 10.44 | 7.35
2008-10-07 | 18:15 | 41 | 97.6 | 12.17 | 63 | 10.48 | 7.34
2008-10-07 | 18:30 | 41 | 97.9 | 12.08 | 63 | 10.53 | 7.33
2008-10-07 | 18:45 | 41 | 97.8 | 11.99 | 63 | 10.54 | 7.32
2008-10-07 | 19:00 | 41 | 97.4 | 11.89 | 63 | 10.53 | 7.32
2008-10-07 | 19:15 | 41 | 97.6 | 11.79 | 63 | 10.56 | 7.3
2008-10-07 | 19:30 | 41 | 97.5 | 11.67 | 63 | 10.58 | 7.3
2008-10-07 | 19:45 | 41 | 97.2 | 11.53 | 63 | 10.58 | 7.31
2008-10-07 | 20:00 | 41 | 97.5 | 11.38 | 63 | 10.65 | 7.29
2008-10-07 | 20:15 | 41 | 97.7 | 11.23 | 62 | 10.72 | 7.3
2008-10-07 | 20:30 | 41 | 97.4 | 11.08 | 62 | 10.72 | 7.3
2008-10-07 | 20:45 | 41 | 97.9 | 10.92 | 62 | 10.82 | 7.3
2008-10-07 | 21:00 | 41 | 98 | 10.74 | 62 | 10.87 | 7.29
2008-10-07 | 21:15 | 40 | 97.6 | 10.57 | 62 | 10.87 | 7.29
2008-10-07 | 21:30 | 40 | 98 | 10.4 | 62 | 10.95 | 7.3
2008-10-07 | 21:45 | 40 | 97.7 | 10.23 | 62 | 10.97 | 7.3
2008-10-07 | 22:00 | 40 | 97.8 | 10.06 | 62 | 11.03 | 7.3
2008-10-07 | 22:15 | 40 | 98.2 | 9.91 | 62 | 11.1 | 7.29
2008-10-07 | 22:30 | 41 | 98.1 | 9.76 | 62 | 11.14 | 7.3
2008-10-07 | 22:45 | 41 | 98.4 | 9.61 | 62 | 11.2 | 7.29
2008-10-07 | 23:00 | 41 | 98.6 | 9.46 | 62 | 11.27 | 7.3
2008-10-07 | 23:15 | 41 | 98.7 | 9.32 | 62 | 11.32 | 7.29
2008-10-07 | 23:30 | 41 | 98.7 | 9.19 | 62 | 11.36 | 7.29
2008-10-07 | 23:45 | 41 | 98.3 | 9.06 | 63 | 11.35 | 7.3
2008-10-08 | 0:00 | 41 | 98.2 | 8.94 | 63 | 11.36 | 7.3
2008-10-08 | 0:15 | 41 | 98.4 | 8.81 | 63 | 11.42 | 7.3
2008-10-08 | 0:30 | 41 | 98.3 | 8.7 | 63 | 11.44 | 7.3
2008-10-08 | 0:45 | 41 | 98.7 | 8.59 | 63 | 11.52 | 7.3
2008-10-08 | 1:01 | 41 | 98.8 | 8.48 | 63 | 11.56 | 7.3
2008-10-08 | 1:16 | 41 | 98.9 | 8.36 | 63 | 11.61 | 7.3
2008-10-08 | 1:31 | 41 | 98.7 | 8.25 | 63 | 11.62 | 7.3
2008-10-08 | 1:46 | 41 | 98.7 | 8.13 | 63 | 11.65 | 7.31
2008-10-08 | 2:01 | 41 | 99 | 8.02 | 63 | 11.71 | 7.31
2008-10-08 | 2:16 | 41 | 99.2 | 7.91 | 63 | 11.77 | 7.3
2008-10-08 | 2:31 | 41 | 98.9 | 7.8 | 63 | 11.76 | 7.31
2008-10-08 | 2:46 | 41 | 99.1 | 7.71 | 63 | 11.82 | 7.3
2008-10-08 | 3:01 | 41 | 99.2 | 7.6 | 62 | 11.87 | 7.3
2008-10-08 | 3:16 | 40 | 98.9 | 7.49 | 62 | 11.86 | 7.31
2008-10-08 | 3:31 | 40 | 98.9 | 7.38 | 62 | 11.89 | 7.3
2008-10-08 | 3:46 | 40 | 99.3 | 7.28 | 62 | 11.97 | 7.3
2008-10-08 | 4:01 | 40 | 99.2 | 7.16 | 62 | 11.99 | 7.31
2008-10-08 | 4:16 | 40 | 99 | 7.05 | 62 | 12.01 | 7.3
2008-10-08 | 4:31 | 40 | 98.9 | 6.94 | 62 | 12.03 | 7.31
2008-10-08 | 4:46 | 40 | 99.4 | 6.82 | 62 | 12.12 | 7.3
2008-10-08 | 5:01 | 40 | 99.6 | 6.71 | 62 | 12.18 | 7.3
2008-10-08 | 5:16 | 40 | 99.3 | 6.6 | 62 | 12.18 | 7.3
2008-10-08 | 5:31 | 40 | 99.7 | 6.48 | 62 | 12.26 | 7.29
2008-10-08 | 5:46 | 40 | 99.4 | 6.36 | 62 | 12.26 | 7.29
2008-10-08 | 6:01 | 40 | 99.5 | 6.26 | 62 | 12.3 | 7.28
2008-10-08 | 6:16 | 40 | 99.3 | 6.15 | 62 | 12.31 | 7.29
2008-10-08 | 6:31 | 40 | 99.7 | 6.03 | 62 | 12.4 | 7.29
2008-10-08 | 6:46 | 40 | 100 | 5.92 | 62 | 12.47 | 7.29
2008-10-08 | 7:01 | 40 | 100.1 | 5.82 | 62 | 12.51 | 7.28
2008-10-08 | 7:16 | 40 | 99.8 | 5.71 | 62 | 12.52 | 7.3
2008-10-08 | 7:31 | 40 | 100.1 | 5.62 | 62 | 12.58 | 7.29
2008-10-08 | 7:46 | 40 | 100.4 | 5.52 | 62 | 12.65 | 7.3
2008-10-08 | 8:01 | 40 | 100.5 | 5.44 | 62 | 12.69 | 7.3
2008-10-08 | 8:16 | 40 | 100.3 | 5.36 | 62 | 12.69 | 7.31
2008-10-08 | 8:31 | 40 | 100.2 | 5.29 | 62 | 12.7 | 7.32
2008-10-08 | 8:46 | 40 | 100.8 | 5.23 | 62 | 12.79 | 7.31
2008-10-08 | 9:01 | 40 | 100.5 | 5.18 | 62 | 12.78 | 7.32
2008-10-08 | 9:16 | 40 | 100.6 | 5.14 | 62 | 12.8 | 7.32
2008-10-08 | 9:31 | 40 | 101.1 | 5.12 | 62 | 12.87 | 7.33
2008-10-08 | 9:46 | 40 | 101.3 | 5.12 | 62 | 12.89 | 7.34
2008-10-08 | 10:01 | 40 | 101.6 | 5.13 | 62 | 12.93 | 7.34
2008-10-08 | 10:16 | 40 | 101.8 | 5.17 | 62 | 12.94 | 7.35
2008-10-08 | 10:31 | 40 | 101.9 | 5.22 | 62 | 12.94 | 7.35
2008-10-08 | 10:46 | 40 | 102.2 | 5.26 | 62 | 12.96 | 7.36
2008-10-08 | 11:01 | 40 | 102.2 | 5.3 | 62 | 12.96 | 7.37
2008-10-08 | 11:16 | 40 | 102.2 | 5.36 | 62 | 12.93 | 7.37
2008-10-08 | 11:31 | 40 | 102.2 | 5.44 | 62 | 12.91 | 7.39
2008-10-08 | 11:46 | 40 | 101.9 | 5.53 | 62 | 12.83 | 7.39
2008-10-08 | 12:01 | 40 | 101.6 | 5.66 | 62 | 12.76 | 7.4
2008-10-08 | 12:16 | 40 | 102 | 5.85 | 62 | 12.74 | 7.42
2008-10-08 | 12:31 | 40 | 101.6 | 6.1 | 62 | 12.61 | 7.43
2008-10-08 | 12:46 | 40 | 101.8 | 6.4 | 62 | 12.54 | 7.44
2008-10-08 | 13:00 | 40 | 101.3 | 6.7 | 62 | 12.39 | 7.44
2008-10-08 | 13:15 | 40 | 101 | 7.03 | 62 | 12.26 | 7.46
2008-10-08 | 13:30 | 40 | 100.9 | 7.37 | 62 | 12.13 | 7.46
2008-10-08 | 13:45 | 40 | 100.6 | 7.77 | 62 | 11.99 | 7.48
2008-10-08 | 14:00 | 40 | 101 | 8.18 | 62 | 11.91 | 7.49
2008-10-08 | 14:15 | 40 | 100.5 | 8.56 | 62 | 11.74 | 7.49
2008-10-08 | 14:30 | 40 | 100.6 | 8.9 | 62 | 11.66 | 7.5
2008-10-08 | 14:45 | 40 | 100.1 | 9.18 | 62 | 11.52 | 7.51
2008-10-08 | 15:00 | 40 | 100.2 | 9.38 | 62 | 11.48 | 7.51
2008-10-08 | 15:15 | 40 | 99.8 | 9.51 | 62 | 11.39 | 7.51
2008-10-08 | 15:30 | 40 | 99.5 | 9.58 | 62 | 11.34 | 7.51
2008-10-08 | 15:45 | 40 | 99.3 | 9.63 | 62 | 11.31 | 7.51
2008-10-08 | 16:00 | 40 | 99.2 | 9.65 | 62 | 11.29 | 7.52
2008-10-08 | 16:15 | 40 | 99 | 9.6 | 62 | 11.28 | 7.52
2008-10-08 | 16:30 | 40 | 99.3 | 9.51 | 62 | 11.34 | 7.51
2008-10-08 | 16:45 | 40 | 99.4 | 9.43 | 62 | 11.37 | 7.5
2008-10-08 | 17:00 | 40 | 98.9 | 9.31 | 62 | 11.34 | 7.49
2008-10-08 | 17:15 | 40 | 98.6 | 9.21 | 62 | 11.34 | 7.48
2008-10-08 | 17:30 | 40 | 98.8 | 9.14 | 62 | 11.38 | 7.47
2008-10-08 | 17:45 | 40 | 98.4 | 9.09 | 62 | 11.35 | 7.47
2008-10-08 | 18:00 | 40 | 98.8 | 9.03 | 62 | 11.41 | 7.46
2008-10-08 | 18:15 | 40 | 98.7 | 8.97 | 62 | 11.41 | 7.46
2008-10-08 | 18:30 | 40 | 98.6 | 8.91 | 62 | 11.42 | 7.45
2008-10-08 | 18:45 | 40 | 98.5 | 8.84 | 62 | 11.43 | 7.44
2008-10-08 | 19:00 | 40 | 98.4 | 8.76 | 62 | 11.44 | 7.43
2008-10-08 | 19:15 | 40 | 98 | 8.67 | 62 | 11.42 | 7.42
2008-10-08 | 19:30 | 40 | 98.2 | 8.58 | 62 | 11.47 | 7.41
2008-10-08 | 19:45 | 40 | 98.4 | 8.47 | 62 | 11.52 | 7.41
2008-10-08 | 20:00 | 40 | 98.4 | 8.34 | 62 | 11.56 | 7.4
2008-10-08 | 20:15 | 40 | 98.5 | 8.2 | 62 | 11.61 | 7.4
2008-10-08 | 20:30 | 40 | 98.1 | 8.07 | 62 | 11.6 | 7.4
2008-10-08 | 20:45 | 40 | 98 | 7.93 | 62 | 11.63 | 7.39
2008-10-08 | 21:00 | 40 | 98.4 | 7.79 | 62 | 11.71 | 7.39
2008-10-08 | 21:15 | 40 | 98.6 | 7.65 | 62 | 11.77 | 7.39
2008-10-08 | 21:30 | 40 | 98.7 | 7.51 | 62 | 11.83 | 7.39
2008-10-08 | 21:45 | 40 | 98.9 | 7.38 | 62 | 11.9 | 7.38
2008-10-08 | 22:00 | 40 | 98.5 | 7.26 | 62 | 11.88 | 7.38
2008-10-08 | 22:15 | 40 | 98.8 | 7.14 | 62 | 11.95 | 7.38
2008-10-08 | 22:30 | 40 | 98.9 | 7.03 | 62 | 12 | 7.37
2008-10-08 | 22:45 | 40 | 98.6 | 6.92 | 62 | 11.99 | 7.38
2008-10-08 | 23:00 | 41 | 98.5 | 6.82 | 62 | 12.01 | 7.38
2008-10-08 | 23:15 | 41 | 98.9 | 6.72 | 62 | 12.09 | 7.37
2008-10-08 | 23:30 | 41 | 99.1 | 6.63 | 62 | 12.14 | 7.37
2008-10-08 | 23:45 | 41 | 98.7 | 6.55 | 63 | 12.12 | 7.38
2008-10-09 | 0:00 | 41 | 98.6 | 6.47 | 63 | 12.13 | 7.38
2008-10-09 | 0:15 | 41 | 98.9 | 6.4 | 63 | 12.19 | 7.37
2008-10-09 | 0:30 | 41 | 99.1 | 6.33 | 63 | 12.24 | 7.38
2008-10-09 | 0:45 | 41 | 99.3 | 6.27 | 63 | 12.27 | 7.37
2008-10-09 | 1:01 | 41 | 99.4 | 6.23 | 63 | 12.3 | 7.37
2008-10-09 | 1:16 | 41 | 99.4 | 6.18 | 63 | 12.31 | 7.37
2008-10-09 | 1:31 | 41 | 99.4 | 6.13 | 63 | 12.33 | 7.37
2008-10-09 | 1:46 | 41 | 99.4 | 6.09 | 63 | 12.34 | 7.37
2008-10-09 | 2:01 | 41 | 99 | 6.05 | 63 | 12.3 | 7.38
2008-10-09 | 2:16 | 41 | 99.3 | 6.01 | 63 | 12.35 | 7.38
2008-10-09 | 2:31 | 41 | 99.3 | 5.99 | 63 | 12.37 | 7.37
2008-10-09 | 2:46 | 41 | 99.4 | 5.97 | 63 | 12.39 | 7.37
2008-10-09 | 3:01 | 41 | 99.1 | 5.94 | 64 | 12.35 | 7.38
2008-10-09 | 3:16 | 41 | 99.4 | 5.91 | 64 | 12.39 | 7.37
2008-10-09 | 3:31 | 41 | 99.4 | 5.88 | 64 | 12.42 | 7.38
2008-10-09 | 3:46 | 41 | 99.1 | 5.85 | 64 | 12.38 | 7.38
2008-10-09 | 4:01 | 41 | 99.4 | 5.81 | 64 | 12.43 | 7.38
2008-10-09 | 4:16 | 42 | 99.5 | 5.75 | 64 | 12.46 | 7.38
2008-10-09 | 4:31 | 42 | 99.5 | 5.69 | 64 | 12.49 | 7.38
2008-10-09 | 4:46 | 42 | 99.2 | 5.61 | 64 | 12.47 | 7.38
2008-10-09 | 5:01 | 42 | 99 | 5.52 | 64 | 12.48 | 7.38
2008-10-09 | 5:16 | 42 | 99.4 | 5.42 | 64 | 12.56 | 7.38
2008-10-09 | 5:31 | 42 | 99.7 | 5.32 | 64 | 12.62 | 7.38
2008-10-09 | 5:46 | 42 | 99.7 | 5.23 | 64 | 12.66 | 7.38
2008-10-09 | 6:01 | 42 | 99.4 | 5.13 | 64 | 12.64 | 7.38
2008-10-09 | 6:16 | 42 | 99.2 | 5.04 | 65 | 12.65 | 7.38
2008-10-09 | 6:31 | 42 | 99.2 | 4.96 | 65 | 12.68 | 7.38
2008-10-09 | 6:46 | 42 | 99.6 | 4.88 | 65 | 12.76 | 7.38
2008-10-09 | 7:01 | 42 | 99.4 | 4.8 | 65 | 12.76 | 7.38
2008-10-09 | 7:16 | 42 | 99.8 | 4.71 | 65 | 12.84 | 7.38
2008-10-09 | 7:31 | 42 | 99.6 | 4.62 | 65 | 12.84 | 7.38
2008-10-09 | 7:46 | 42 | 99.6 | 4.54 | 65 | 12.86 | 7.39
2008-10-09 | 8:01 | 42 | 99.7 | 4.47 | 65 | 12.9 | 7.39
2008-10-09 | 8:16 | 42 | 100.2 | 4.4 | 64 | 12.99 | 7.4
2008-10-09 | 8:31 | 42 | 100.1 | 4.34 | 64 | 13 | 7.4
2008-10-09 | 8:46 | 42 | 100.2 | 4.3 | 64 | 13.02 | 7.41
2008-10-09 | 9:01 | 42 | 100.7 | 4.28 | 64 | 13.11 | 7.41
2008-10-09 | 9:16 | 42 | 100.6 | 4.28 | 64 | 13.09 | 7.42
2008-10-09 | 9:31 | 42 | 101 | 4.29 | 64 | 13.14 | 7.42
2008-10-09 | 9:46 | 42 | 100.8 | 4.32 | 64 | 13.11 | 7.43
2008-10-09 | 10:01 | 42 | 100.8 | 4.37 | 64 | 13.09 | 7.44
2008-10-09 | 10:16 | 41 | 100.8 | 4.41 | 64 | 13.07 | 7.44
2008-10-09 | 10:31 | 41 | 101 | 4.48 | 64 | 13.07 | 7.45
2008-10-09 | 10:46 | 41 | 101.4 | 4.53 | 64 | 13.11 | 7.45
2008-10-09 | 11:01 | 41 | 101.7 | 4.6 | 64 | 13.12 | 7.46
2008-10-09 | 11:16 | 41 | 101.8 | 4.66 | 63 | 13.12 | 7.45
2008-10-09 | 11:31 | 41 | 101.4 | 4.71 | 63 | 13.05 | 7.47
2008-10-09 | 11:46 | 41 | 101.8 | 4.83 | 63 | 13.05 | 7.46
2008-10-09 | 12:01 | 41 | 101.3 | 4.91 | 63 | 12.97 | 7.48
2008-10-09 | 12:16 | 41 | 101.6 | 5.08 | 63 | 12.95 | 7.47
2008-10-09 | 12:31 | 41 | 101.8 | 5.34 | 63 | 12.89 | 7.47
2008-10-09 | 12:46 | 41 | 102 | 5.63 | 63 | 12.81 | 7.47
2008-10-09 | 13:00 | 41 | 101.4 | 5.92 | 63 | 12.65 | 7.49
2008-10-09 | 13:15 | 41 | 101.4 | 6.22 | 63 | 12.56 | 7.5
2008-10-09 | 13:30 | 41 | 101 | 6.53 | 63 | 12.41 | 7.51
2008-10-09 | 13:45 | 41 | 100.6 | 6.82 | 63 | 12.27 | 7.51
2008-10-09 | 14:00 | 41 | 100.3 | 7.08 | 63 | 12.14 | 7.52
2008-10-09 | 14:15 | 41 | 100.4 | 7.32 | 63 | 12.09 | 7.53
2008-10-09 | 14:30 | 41 | 100.4 | 7.48 | 63 | 12.05 | 7.52
2008-10-09 | 14:45 | 41 | 99.8 | 7.59 | 63 | 11.94 | 7.52
2008-10-09 | 15:00 | 41 | 99.6 | 7.68 | 63 | 11.89 | 7.53
2008-10-09 | 15:15 | 41 | 99.9 | 7.7 | 63 | 11.92 | 7.53
2008-10-09 | 15:30 | 41 | 99.8 | 7.74 | 63 | 11.89 | 7.53
2008-10-09 | 15:45 | 41 | 100.2 | 7.79 | 63 | 11.93 | 7.53
2008-10-09 | 16:00 | 41 | 100.3 | 7.81 | 63 | 11.94 | 7.51
2008-10-09 | 16:15 | 41 | 100.4 | 7.86 | 63 | 11.93 | 7.54
2008-10-09 | 16:30 | 41 | 100.2 | 7.8 | 63 | 11.93 | 7.54
2008-10-09 | 16:45 | 41 | 100.1 | 7.74 | 63 | 11.93 | 7.53
2008-10-09 | 17:00 | 41 | 99.6 | 7.68 | 63 | 11.89 | 7.53
2008-10-09 | 17:15 | 41 | 99.7 | 7.55 | 63 | 11.94 | 7.52
2008-10-09 | 17:30 | 41 | 99.8 | 7.46 | 63 | 11.98 | 7.51
2008-10-09 | 17:45 | 41 | 99.4 | 7.38 | 63 | 11.95 | 7.5
2008-10-09 | 18:00 | 41 | 99.5 | 7.29 | 63 | 11.99 | 7.49
2008-10-09 | 18:15 | 41 | 99.5 | 7.22 | 63 | 12.01 | 7.48
2008-10-09 | 18:30 | 41 | 98.9 | 7.18 | 63 | 11.95 | 7.46
2008-10-09 | 18:45 | 41 | 98.6 | 7.17 | 63 | 11.92 | 7.47
2008-10-09 | 19:00 | 41 | 98.5 | 7.17 | 63 | 11.9 | 7.45
2008-10-09 | 19:15 | 41 | 98.5 | 7.18 | 63 | 11.9 | 7.45
2008-10-09 | 19:30 | 41 | 98.7 | 7.18 | 63 | 11.93 | 7.44
2008-10-09 | 19:45 | 41 | 98.4 | 7.17 | 63 | 11.9 | 7.44
2008-10-09 | 20:00 | 41 | 98.9 | 7.16 | 63 | 11.96 | 7.44
2008-10-09 | 20:15 | 41 | 98.9 | 7.13 | 63 | 11.97 | 7.43
2008-10-09 | 20:30 | 41 | 99 | 7.06 | 63 | 12 | 7.43
2008-10-09 | 20:45 | 41 | 99.1 | 6.98 | 63 | 12.03 | 7.43
2008-10-09 | 21:00 | 41 | 99.1 | 6.9 | 63 | 12.06 | 7.43
2008-10-09 | 21:15 | 41 | 98.8 | 6.85 | 63 | 12.04 | 7.43
2008-10-09 | 21:30 | 41 | 99 | 6.8 | 63 | 12.08 | 7.41
2008-10-09 | 21:45 | 41 | 99.1 | 6.75 | 63 | 12.11 | 7.42
2008-10-09 | 22:00 | 41 | 98.7 | 6.71 | 63 | 12.07 | 7.42
2008-10-09 | 22:15 | 41 | 99 | 6.68 | 63 | 12.12 | 7.42
2008-10-09 | 22:30 | 41 | 99.1 | 6.65 | 63 | 12.13 | 7.41
2008-10-09 | 22:45 | 41 | 98.7 | 6.62 | 63 | 12.09 | 7.41
2008-10-09 | 23:00 | 41 | 99 | 6.58 | 63 | 12.14 | 7.42
2008-10-09 | 23:15 | 41 | 99.1 | 6.53 | 63 | 12.17 | 7.41
2008-10-09 | 23:30 | 41 | 99.2 | 6.49 | 63 | 12.19 | 7.41
2008-10-09 | 23:45 | 41 | 98.7 | 6.44 | 63 | 12.15 | 7.42
2008-10-10 | 0:00 | 41 | 99 | 6.4 | 63 | 12.2 | 7.42
2008-10-10 | 0:15 | 41 | 99.1 | 6.37 | 63 | 12.22 | 7.41
2008-10-10 | 0:30 | 41 | 99.3 | 6.33 | 63 | 12.25 | 7.41
2008-10-10 | 0:45 | 41 | 99.3 | 6.29 | 64 | 12.27 | 7.42
2008-10-10 | 1:01 | 41 | 98.9 | 6.25 | 64 | 12.23 | 7.41
2008-10-10 | 1:16 | 41 | 98.7 | 6.21 | 64 | 12.22 | 7.41
2008-10-10 | 1:31 | 42 | 99.1 | 6.18 | 64 | 12.28 | 7.42
2008-10-10 | 1:46 | 42 | 99.3 | 6.14 | 64 | 12.32 | 7.41
2008-10-10 | 2:01 | 42 | 98.9 | 6.09 | 64 | 12.28 | 7.42
2008-10-10 | 2:16 | 42 | 98.8 | 6.03 | 64 | 12.28 | 7.42
2008-10-10 | 2:31 | 42 | 98.7 | 5.97 | 64 | 12.3 | 7.42
2008-10-10 | 2:46 | 42 | 99.1 | 5.92 | 64 | 12.36 | 7.41
2008-10-10 | 3:01 | 42 | 99.4 | 5.88 | 64 | 12.41 | 7.42
2008-10-10 | 3:16 | 42 | 99.5 | 5.84 | 64 | 12.43 | 7.41
2008-10-10 | 3:31 | 42 | 99.5 | 5.8 | 64 | 12.45 | 7.42
2008-10-10 | 3:46 | 42 | 99.6 | 5.75 | 65 | 12.47 | 7.42
2008-10-10 | 4:01 | 42 | 99.6 | 5.7 | 65 | 12.5 | 7.41
2008-10-10 | 4:16 | 42 | 99.7 | 5.64 | 65 | 12.52 | 7.41
2008-10-10 | 4:31 | 42 | 99.2 | 5.59 | 64 | 12.48 | 7.42
2008-10-10 | 4:46 | 42 | 99.1 | 5.54 | 64 | 12.48 | 7.42
2008-10-10 | 5:01 | 42 | 99.6 | 5.49 | 64 | 12.55 | 7.41
2008-10-10 | 5:16 | 42 | 99.3 | 5.44 | 64 | 12.54 | 7.41
2008-10-10 | 5:31 | 42 | 99.7 | 5.4 | 64 | 12.6 | 7.41
2008-10-10 | 5:46 | 42 | 99.8 | 5.35 | 64 | 12.63 | 7.41
2008-10-10 | 6:01 | 42 | 99.9 | 5.31 | 64 | 12.66 | 7.41
2008-10-10 | 6:16 | 42 | 100.1 | 5.27 | 64 | 12.69 | 7.41
2008-10-10 | 6:31 | 42 | 100.1 | 5.24 | 64 | 12.7 | 7.4
2008-10-10 | 6:46 | 42 | 99.6 | 5.2 | 64 | 12.66 | 7.41
2008-10-10 | 7:01 | 42 | 99.4 | 5.14 | 64 | 12.65 | 7.41
2008-10-10 | 7:16 | 42 | 99.8 | 5.09 | 64 | 12.72 | 7.41
2008-10-10 | 7:31 | 42 | 100 | 5.04 | 64 | 12.76 | 7.41
2008-10-10 | 7:46 | 42 | 100.3 | 4.98 | 64 | 12.82 | 7.41
2008-10-10 | 8:01 | 42 | 100 | 4.93 | 64 | 12.79 | 7.42
2008-10-10 | 8:16 | 42 | 99.8 | 4.89 | 64 | 12.78 | 7.42
2008-10-10 | 8:31 | 42 | 99.9 | 4.85 | 64 | 12.81 | 7.42
2008-10-10 | 8:46 | 42 | 99.9 | 4.81 | 64 | 12.82 | 7.42
2008-10-10 | 9:01 | 42 | 100.4 | 4.78 | 64 | 12.9 | 7.42
2008-10-10 | 9:16 | 42 | 100.7 | 4.76 | 64 | 12.95 | 7.43
2008-10-10 | 9:31 | 42 | 100.5 | 4.75 | 64 | 12.92 | 7.44
2008-10-10 | 9:46 | 42 | 101 | 4.76 | 64 | 12.97 | 7.44
2008-10-10 | 10:01 | 41 | 101.3 | 4.8 | 64 | 13.01 | 7.45
2008-10-10 | 10:16 | 41 | 101.5 | 4.83 | 64 | 13.02 | 7.45
2008-10-10 | 10:31 | 41 | 101.8 | 4.88 | 64 | 13.04 | 7.45
2008-10-10 | 10:46 | 41 | 102 | 4.95 | 63 | 13.04 | 7.46
2008-10-10 | 11:01 | 41 | 102.1 | 5.02 | 63 | 13.03 | 7.46
2008-10-10 | 11:16 | 41 | 101.7 | 5.08 | 63 | 12.96 | 7.48
2008-10-10 | 11:31 | 41 | 102.2 | 5.18 | 63 | 12.98 | 7.48
2008-10-10 | 11:46 | 41 | 102.2 | 5.26 | 63 | 12.97 | 7.49
2008-10-10 | 12:01 | 41 | 102 | 5.45 | 63 | 12.87 | 7.5
2008-10-10 | 12:16 | 41 | 102.2 | 5.65 | 63 | 12.84 | 7.48
2008-10-10 | 12:31 | 41 | 101.6 | 5.82 | 63 | 12.7 | 7.49
2008-10-10 | 12:46 | 41 | 101.3 | 6.04 | 63 | 12.6 | 7.5
2008-10-10 | 13:00 | 41 | 101.1 | 6.22 | 63 | 12.51 | 7.51
2008-10-10 | 13:15 | 41 | 100.9 | 6.29 | 63 | 12.47 | 7.52
2008-10-10 | 13:30 | 41 | 101.2 | 6.52 | 63 | 12.43 | 7.51
2008-10-10 | 13:45 | 41 | 101.3 | 6.73 | 63 | 12.38 | 7.52
2008-10-10 | 14:00 | 41 | 101.3 | 6.93 | 63 | 12.32 | 7.53
2008-10-10 | 14:15 | 41 | 100.6 | 7.07 | 63 | 12.18 | 7.52
2008-10-10 | 14:30 | 41 | 100.4 | 7.26 | 63 | 12.1 | 7.53
2008-10-10 | 14:45 | 41 | 100.6 | 7.32 | 63 | 12.11 | 7.53
2008-10-10 | 15:00 | 41 | 100.4 | 7.34 | 63 | 12.08 | 7.53
2008-10-10 | 15:15 | 41 | 100.6 | 7.39 | 63 | 12.09 | 7.53
2008-10-10 | 15:30 | 41 | 100.1 | 7.37 | 63 | 12.04 | 7.53
2008-10-10 | 15:45 | 41 | 100.2 | 7.29 | 63 | 12.08 | 7.53
2008-10-10 | 16:00 | 41 | 99.9 | 7.27 | 63 | 12.04 | 7.52
2008-10-10 | 16:15 | 41 | 99.8 | 7.25 | 63 | 12.04 | 7.52
2008-10-10 | 16:30 | 41 | 100.3 | 7.25 | 63 | 12.1 | 7.52
2008-10-10 | 16:45 | 41 | 100.3 | 7.23 | 62 | 12.11 | 7.51
2008-10-10 | 17:00 | 41 | 100.3 | 7.18 | 62 | 12.12 | 7.5
2008-10-10 | 17:15 | 41 | 100.2 | 7.12 | 63 | 12.13 | 7.5
2008-10-10 | 17:30 | 41 | 100.2 | 7.06 | 63 | 12.14 | 7.49
2008-10-10 | 17:45 | 41 | 100.2 | 7 | 63 | 12.16 | 7.48
2008-10-10 | 18:00 | 41 | 99.7 | 6.93 | 63 | 12.12 | 7.48
2008-10-10 | 18:15 | 41 | 99.8 | 6.88 | 63 | 12.15 | 7.47
2008-10-10 | 18:30 | 41 | 99.8 | 6.8 | 63 | 12.18 | 7.46
2008-10-10 | 18:45 | 41 | 99.8 | 6.75 | 63 | 12.19 | 7.45
2008-10-10 | 19:00 | 41 | 99.8 | 6.69 | 63 | 12.2 | 7.44
2008-10-10 | 19:15 | 41 | 99.7 | 6.63 | 63 | 12.22 | 7.43
2008-10-10 | 19:30 | 41 | 99.8 | 6.57 | 63 | 12.24 | 7.43
2008-10-10 | 19:45 | 41 | 99.3 | 6.5 | 63 | 12.2 | 7.42
2008-10-10 | 20:00 | 41 | 99.5 | 6.42 | 63 | 12.26 | 7.42
2008-10-10 | 20:15 | 41 | 99.6 | 6.33 | 63 | 12.3 | 7.42
2008-10-10 | 20:30 | 41 | 99.7 | 6.24 | 63 | 12.34 | 7.42
2008-10-10 | 20:45 | 41 | 99.3 | 6.13 | 63 | 12.32 | 7.41
2008-10-10 | 21:00 | 41 | 99.2 | 6.02 | 63 | 12.34 | 7.42
2008-10-10 | 21:15 | 41 | 99.6 | 5.91 | 63 | 12.42 | 7.41
2008-10-10 | 21:30 | 41 | 99.3 | 5.81 | 63 | 12.42 | 7.41
2008-10-10 | 21:45 | 41 | 99.2 | 5.71 | 63 | 12.44 | 7.42
2008-10-10 | 22:00 | 41 | 99.6 | 5.62 | 63 | 12.52 | 7.41
2008-10-10 | 22:15 | 41 | 99.4 | 5.52 | 63 | 12.53 | 7.41
2008-10-10 | 22:30 | 41 | 99.8 | 5.42 | 63 | 12.61 | 7.41
2008-10-10 | 22:45 | 41 | 100 | 5.32 | 63 | 12.67 | 7.41
2008-10-10 | 23:00 | 41 | 100.1 | 5.22 | 63 | 12.71 | 7.41
2008-10-10 | 23:15 | 41 | 100.2 | 5.13 | 63 | 12.76 | 7.41
2008-10-10 | 23:30 | 41 | 100.3 | 5.03 | 63 | 12.79 | 7.41
2008-10-10 | 23:45 | 41 | 99.9 | 4.94 | 63 | 12.77 | 7.41
2008-10-11 | 0:00 | 41 | 100.1 | 4.85 | 63 | 12.84 | 7.41
2008-10-11 | 0:15 | 41 | 100.3 | 4.78 | 63 | 12.89 | 7.41
2008-10-11 | 0:30 | 41 | 100 | 4.7 | 63 | 12.87 | 7.41
2008-10-11 | 0:45 | 41 | 100.3 | 4.61 | 64 | 12.93 | 7.41
2008-10-11 | 1:01 | 41 | 99.9 | 4.53 | 64 | 12.92 | 7.41
2008-10-11 | 1:16 | 41 | 99.8 | 4.44 | 64 | 12.93 | 7.42
2008-10-11 | 1:31 | 42 | 100.3 | 4.36 | 64 | 13.02 | 7.42
2008-10-11 | 1:46 | 42 | 100.4 | 4.27 | 64 | 13.07 | 7.42
2008-10-11 | 2:01 | 42 | 100.6 | 4.18 | 64 | 13.13 | 7.41
2008-10-11 | 2:16 | 42 | 100.7 | 4.1 | 64 | 13.16 | 7.42
2008-10-11 | 2:31 | 42 | 100.8 | 4.01 | 64 | 13.2 | 7.41
2008-10-11 | 2:46 | 42 | 100.9 | 3.92 | 64 | 13.25 | 7.41
2008-10-11 | 3:01 | 42 | 100.4 | 3.83 | 64 | 13.22 | 7.42
2008-10-11 | 3:16 | 42 | 100.3 | 3.75 | 64 | 13.23 | 7.41
2008-10-11 | 3:31 | 42 | 100.7 | 3.65 | 64 | 13.32 | 7.42
2008-10-11 | 3:46 | 42 | 100.9 | 3.58 | 64 | 13.38 | 7.41
2008-10-11 | 4:01 | 42 | 100.7 | 3.49 | 64 | 13.37 | 7.41
2008-10-11 | 4:16 | 42 | 100.4 | 3.4 | 64 | 13.37 | 7.42
2008-10-11 | 4:31 | 42 | 100.9 | 3.31 | 64 | 13.46 | 7.41
2008-10-11 | 4:46 | 42 | 101 | 3.23 | 64 | 13.52 | 7.41
2008-10-11 | 5:01 | 42 | 101.2 | 3.14 | 64 | 13.57 | 7.41
2008-10-11 | 5:16 | 42 | 101.3 | 3.05 | 64 | 13.61 | 7.41
2008-10-11 | 5:31 | 42 | 101.4 | 2.97 | 64 | 13.65 | 7.41
2008-10-11 | 5:46 | 42 | 101.5 | 2.89 | 64 | 13.7 | 7.41
2008-10-11 | 6:01 | 41 | 101 | 2.82 | 64 | 13.66 | 7.41
2008-10-11 | 6:16 | 41 | 101.4 | 2.74 | 64 | 13.75 | 7.4
2008-10-11 | 6:31 | 41 | 101.1 | 2.67 | 63 | 13.73 | 7.4
2008-10-11 | 6:46 | 41 | 101 | 2.6 | 63 | 13.74 | 7.4
2008-10-11 | 7:01 | 41 | 100.9 | 2.51 | 63 | 13.76 | 7.4
2008-10-11 | 7:16 | 41 | 101 | 2.43 | 63 | 13.8 | 7.4
2008-10-11 | 7:31 | 41 | 101.5 | 2.35 | 63 | 13.9 | 7.39
2008-10-11 | 7:46 | 41 | 101.8 | 2.27 | 63 | 13.98 | 7.39
2008-10-11 | 8:01 | 41 | 101.5 | 2.19 | 62 | 13.97 | 7.4
2008-10-11 | 8:16 | 41 | 101.9 | 2.12 | 62 | 14.05 | 7.39
2008-10-11 | 8:31 | 40 | 102.2 | 2.05 | 62 | 14.12 | 7.39
2008-10-11 | 8:46 | 40 | 101.9 | 1.99 | 62 | 14.1 | 7.4
2008-10-11 | 9:01 | 40 | 101.9 | 1.95 | 62 | 14.12 | 7.39
2008-10-11 | 9:16 | 40 | 102.4 | 1.95 | 62 | 14.18 | 7.4
2008-10-11 | 9:31 | 40 | 102.7 | 1.96 | 62 | 14.21 | 7.39
2008-10-11 | 9:46 | 40 | 102.4 | 1.99 | 62 | 14.17 | 7.4
2008-10-11 | 10:01 | 40 | 102.8 | 2.03 | 62 | 14.21 | 7.41
2008-10-11 | 10:16 | 40 | 103.1 | 2.06 | 62 | 14.23 | 7.41
2008-10-11 | 10:31 | 40 | 102.8 | 2.11 | 62 | 14.18 | 7.42
2008-10-11 | 10:46 | 40 | 102.7 | 2.15 | 61 | 14.15 | 7.41
2008-10-11 | 11:01 | 40 | 103.1 | 2.19 | 61 | 14.19 | 7.41
2008-10-11 | 11:16 | 40 | 102.9 | 2.23 | 61 | 14.14 | 7.42
2008-10-11 | 11:31 | 40 | 103.2 | 2.3 | 61 | 14.16 | 7.43
2008-10-11 | 11:46 | 40 | 102.9 | 2.4 | 61 | 14.08 | 7.43
2008-10-11 | 12:01 | 40 | 103.3 | 2.53 | 61 | 14.08 | 7.43
2008-10-11 | 12:16 | 40 | 103 | 2.73 | 61 | 13.96 | 7.43
2008-10-11 | 12:31 | 40 | 102.7 | 2.97 | 61 | 13.83 | 7.44
2008-10-11 | 12:46 | 40 | 102.9 | 3.27 | 61 | 13.75 | 7.44
2008-10-11 | 13:00 | 40 | 102.5 | 3.57 | 61 | 13.59 | 7.44
2008-10-11 | 13:15 | 40 | 102.2 | 3.89 | 61 | 13.43 | 7.45
2008-10-11 | 13:30 | 40 | 102.5 | 4.23 | 61 | 13.35 | 7.46
2008-10-11 | 13:45 | 40 | 102.5 | 4.63 | 61 | 13.21 | 7.46
2008-10-11 | 14:00 | 40 | 101.9 | 5.03 | 61 | 13 | 7.47
2008-10-11 | 14:15 | 40 | 101.6 | 5.37 | 61 | 12.85 | 7.49
2008-10-11 | 14:30 | 40 | 101.3 | 5.69 | 61 | 12.71 | 7.5
2008-10-11 | 14:45 | 40 | 101.5 | 5.95 | 61 | 12.65 | 7.5
2008-10-11 | 15:00 | 40 | 101.5 | 6.11 | 61 | 12.6 | 7.5
2008-10-11 | 15:15 | 40 | 101 | 6.23 | 61 | 12.49 | 7.5
2008-10-11 | 15:30 | 40 | 100.7 | 6.28 | 61 | 12.44 | 7.51
2008-10-11 | 15:45 | 40 | 100.5 | 6.34 | 61 | 12.4 | 7.51
2008-10-11 | 16:00 | 40 | 100.3 | 6.36 | 61 | 12.37 | 7.5
2008-10-11 | 16:15 | 40 | 100.2 | 6.31 | 61 | 12.37 | 7.48
2008-10-11 | 16:30 | 40 | 100 | 6.24 | 61 | 12.38 | 7.51
2008-10-11 | 16:45 | 40 | 100.4 | 6.16 | 61 | 12.45 | 7.49
2008-10-11 | 17:00 | 40 | 100 | 6.04 | 61 | 12.43 | 7.48
2008-10-11 | 17:15 | 40 | 99.8 | 5.97 | 61 | 12.44 | 7.49
2008-10-11 | 17:30 | 40 | 99.7 | 5.92 | 61 | 12.43 | 7.49
2008-10-11 | 17:45 | 40 | 99.6 | 5.89 | 61 | 12.43 | 7.48
2008-10-11 | 18:00 | 40 | 100 | 5.87 | 61 | 12.48 | 7.47
2008-10-11 | 18:15 | 40 | 100.1 | 5.85 | 61 | 12.51 | 7.46
2008-10-11 | 18:30 | 40 | 100.1 | 5.81 | 61 | 12.52 | 7.45
2008-10-11 | 18:45 | 40 | 100.1 | 5.77 | 61 | 12.53 | 7.44
2008-10-11 | 19:00 | 40 | 100 | 5.71 | 61 | 12.54 | 7.44
2008-10-11 | 19:15 | 40 | 100 | 5.65 | 61 | 12.56 | 7.43
2008-10-11 | 19:30 | 40 | 100.1 | 5.59 | 61 | 12.59 | 7.43
2008-10-11 | 19:45 | 40 | 100.4 | 5.5 | 61 | 12.66 | 7.43
2008-10-11 | 20:00 | 40 | 100 | 5.38 | 61 | 12.64 | 7.43
2008-10-11 | 20:15 | 40 | 100.3 | 5.27 | 61 | 12.72 | 7.43
2008-10-11 | 20:30 | 40 | 100.4 | 5.16 | 61 | 12.76 | 7.42
2008-10-11 | 20:45 | 40 | 100 | 5.04 | 61 | 12.76 | 7.42
2008-10-11 | 21:00 | 40 | 99.9 | 4.93 | 61 | 12.78 | 7.42
2008-10-11 | 21:15 | 40 | 99.9 | 4.82 | 61 | 12.81 | 7.42
2008-10-11 | 21:30 | 40 | 100.2 | 4.7 | 62 | 12.9 | 7.42
2008-10-11 | 21:45 | 40 | 100.5 | 4.59 | 62 | 12.97 | 7.42
2008-10-11 | 22:00 | 40 | 100.2 | 4.49 | 62 | 12.96 | 7.42
2008-10-11 | 22:15 | 40 | 100.5 | 4.4 | 62 | 13.03 | 7.41
2008-10-11 | 22:30 | 40 | 100.7 | 4.32 | 62 | 13.08 | 7.42
2008-10-11 | 22:45 | 40 | 101 | 4.26 | 62 | 13.14 | 7.42
2008-10-11 | 23:00 | 41 | 101 | 4.2 | 62 | 13.16 | 7.41
2008-10-11 | 23:15 | 41 | 100.9 | 4.14 | 63 | 13.18 | 7.42
2008-10-11 | 23:30 | 41 | 100.9 | 4.09 | 63 | 13.2 | 7.42
2008-10-11 | 23:45 | 41 | 100.5 | 4.05 | 63 | 13.16 | 7.42
2008-10-12 | 0:00 | 41 | 100.8 | 3.99 | 63 | 13.21 | 7.42
2008-10-12 | 0:15 | 40 | 100.5 | 3.95 | 62 | 13.18 | 7.43
2008-10-12 | 0:30 | 40 | 100.8 | 3.89 | 61 | 13.25 | 7.42
2008-10-12 | 0:45 | 39 | 100.5 | 3.83 | 60 | 13.23 | 7.41
2008-10-12 | 1:01 | 39 | 100.9 | 3.76 | 60 | 13.31 | 7.4
2008-10-12 | 1:16 | 39 | 100.6 | 3.68 | 59 | 13.29 | 7.39
2008-10-12 | 1:31 | 38 | 101 | 3.59 | 59 | 13.38 | 7.4
2008-10-12 | 1:46 | 38 | 100.8 | 3.49 | 59 | 13.39 | 7.39
2008-10-12 | 2:01 | 38 | 101.1 | 3.39 | 58 | 13.47 | 7.38
2008-10-12 | 2:16 | 38 | 101 | 3.31 | 58 | 13.48 | 7.38
2008-10-12 | 2:31 | 38 | 100.8 | 3.21 | 58 | 13.49 | 7.38
2008-10-12 | 2:46 | 38 | 100.8 | 3.13 | 58 | 13.53 | 7.38
2008-10-12 | 3:01 | 37 | 100.8 | 3.04 | 58 | 13.55 | 7.38
2008-10-12 | 3:16 | 37 | 100.8 | 2.95 | 57 | 13.58 | 7.37
2008-10-12 | 3:31 | 37 | 101.3 | 2.87 | 57 | 13.68 | 7.37
2008-10-12 | 3:46 | 37 | 101.5 | 2.8 | 57 | 13.74 | 7.37
2008-10-12 | 4:01 | 37 | 101.7 | 2.73 | 57 | 13.78 | 7.37
2008-10-12 | 4:16 | 37 | 101.8 | 2.66 | 57 | 13.82 | 7.36
2008-10-12 | 4:31 | 37 | 101.9 | 2.59 | 57 | 13.87 | 7.36
2008-10-12 | 4:46 | 37 | 101.5 | 2.53 | 56 | 13.83 | 7.36
2008-10-12 | 5:01 | 36 | 101.3 | 2.46 | 56 | 13.83 | 7.36
2008-10-12 | 5:16 | 36 | 101.2 | 2.39 | 56 | 13.85 | 7.35
2008-10-12 | 5:31 | 36 | 101.2 | 2.31 | 55 | 13.87 | 7.35
2008-10-12 | 5:46 | 36 | 101.7 | 2.26 | 55 | 13.97 | 7.34
2008-10-12 | 6:01 | 36 | 101.5 | 2.2 | 55 | 13.96 | 7.33
2008-10-12 | 6:16 | 35 | 102 | 2.14 | 54 | 14.06 | 7.34
2008-10-12 | 6:31 | 35 | 101.7 | 2.08 | 54 | 14.03 | 7.34
2008-10-12 | 6:46 | 35 | 102.1 | 2.02 | 54 | 14.11 | 7.33
2008-10-12 | 7:01 | 35 | 102.3 | 1.97 | 54 | 14.15 | 7.33
2008-10-12 | 7:16 | 35 | 102.4 | 1.92 | 54 | 14.19 | 7.32
2008-10-12 | 7:31 | 35 | 102 | 1.88 | 53 | 14.16 | 7.33
2008-10-12 | 7:46 | 35 | 101.9 | 1.85 | 53 | 14.16 | 7.33
2008-10-12 | 8:01 | 35 | 102.5 | 1.82 | 53 | 14.25 | 7.33
2008-10-12 | 8:16 | 34 | 102.7 | 1.8 | 53 | 14.29 | 7.33
2008-10-12 | 8:31 | 34 | 102.5 | 1.8 | 53 | 14.25 | 7.33
2008-10-12 | 8:46 | 34 | 103 | 1.81 | 53 | 14.32 | 7.33
2008-10-12 | 9:01 | 34 | 103.1 | 1.83 | 53 | 14.33 | 7.33
2008-10-12 | 9:16 | 34 | 102.9 | 1.87 | 53 | 14.28 | 7.33
2008-10-12 | 9:31 | 34 | 103.3 | 1.92 | 53 | 14.31 | 7.33
2008-10-12 | 9:46 | 34 | 103 | 1.98 | 53 | 14.25 | 7.34
2008-10-12 | 10:01 | 34 | 102.9 | 2.06 | 53 | 14.21 | 7.34
2008-10-12 | 10:16 | 34 | 103.4 | 2.16 | 53 | 14.23 | 7.35
2008-10-12 | 10:31 | 34 | 103.6 | 2.27 | 53 | 14.22 | 7.36
2008-10-12 | 10:46 | 34 | 103.7 | 2.4 | 53 | 14.18 | 7.36
2008-10-12 | 11:01 | 34 | 103.8 | 2.53 | 53 | 14.15 | 7.36
2008-10-12 | 11:16 | 34 | 103.3 | 2.64 | 53 | 14.03 | 7.37
2008-10-12 | 11:31 | 34 | 103.6 | 2.81 | 53 | 14.01 | 7.37
2008-10-12 | 11:46 | 34 | 103.2 | 2.97 | 53 | 13.9 | 7.38
2008-10-12 | 12:01 | 34 | 103.4 | 3.16 | 53 | 13.86 | 7.38
2008-10-12 | 12:16 | 34 | 102.9 | 3.4 | 53 | 13.71 | 7.39
2008-10-12 | 12:31 | 34 | 103.2 | 3.62 | 53 | 13.66 | 7.39
2008-10-12 | 12:46 | 34 | 102.8 | 3.89 | 53 | 13.51 | 7.4
2008-10-12 | 13:00 | 34 | 102.5 | 4.13 | 53 | 13.39 | 7.41
2008-10-12 | 13:15 | 34 | 102.8 | 4.4 | 53 | 13.33 | 7.41
2008-10-12 | 13:30 | 34 | 102.3 | 4.66 | 53 | 13.18 | 7.42
2008-10-12 | 13:45 | 34 | 102 | 4.92 | 53 | 13.05 | 7.42
2008-10-12 | 14:00 | 34 | 102.4 | 5.17 | 53 | 13.01 | 7.43
2008-10-12 | 14:15 | 34 | 102 | 5.39 | 53 | 12.9 | 7.43
2008-10-12 | 14:30 | 34 | 102.3 | 5.56 | 53 | 12.87 | 7.44
2008-10-12 | 14:45 | 34 | 102.2 | 5.72 | 53 | 12.82 | 7.45
2008-10-12 | 15:00 | 34 | 102.1 | 5.85 | 53 | 12.76 | 7.45
2008-10-12 | 15:15 | 34 | 102 | 5.96 | 53 | 12.71 | 7.46
2008-10-12 | 15:30 | 34 | 101.9 | 6.06 | 53 | 12.67 | 7.46
2008-10-12 | 15:45 | 34 | 101.9 | 6.12 | 53 | 12.65 | 7.46
2008-10-12 | 16:00 | 34 | 101.9 | 6.18 | 53 | 12.63 | 7.45
2008-10-12 | 16:15 | 35 | 101.3 | 6.24 | 53 | 12.53 | 7.45
2008-10-12 | 16:30 | 35 | 101.5 | 6.28 | 53 | 12.54 | 7.44
2008-10-12 | 16:45 | 35 | 101.1 | 6.28 | 53 | 12.49 | 7.45
2008-10-12 | 17:00 | 35 | 101.2 | 6.3 | 53 | 12.51 | 7.44
2008-10-12 | 17:15 | 35 | 100.8 | 6.31 | 53 | 12.44 | 7.43
2008-10-12 | 17:30 | 35 | 100.9 | 6.32 | 54 | 12.45 | 7.43
2008-10-12 | 17:45 | 35 | 100.8 | 6.34 | 54 | 12.44 | 7.42
2008-10-12 | 18:00 | 35 | 100.8 | 6.32 | 54 | 12.44 | 7.42
2008-10-12 | 18:15 | 35 | 100.3 | 6.27 | 54 | 12.4 | 7.41
2008-10-12 | 18:30 | 35 | 100.1 | 6.2 | 54 | 12.39 | 7.41
2008-10-12 | 18:45 | 35 | 100.3 | 6.16 | 54 | 12.44 | 7.4
2008-10-12 | 19:00 | 35 | 100.4 | 6.12 | 54 | 12.47 | 7.38
2008-10-12 | 19:15 | 35 | 100 | 6.07 | 54 | 12.43 | 7.39
2008-10-12 | 19:30 | 35 | 100.4 | 6.01 | 54 | 12.49 | 7.38
2008-10-12 | 19:45 | 35 | 100.6 | 5.94 | 54 | 12.54 | 7.38
2008-10-12 | 20:00 | 35 | 100.1 | 5.86 | 55 | 12.51 | 7.38
2008-10-12 | 20:15 | 36 | 100 | 5.77 | 55 | 12.52 | 7.38
2008-10-12 | 20:30 | 36 | 100.6 | 5.68 | 55 | 12.62 | 7.37
2008-10-12 | 20:45 | 36 | 100.3 | 5.59 | 55 | 12.62 | 7.38
2008-10-12 | 21:00 | 36 | 100.7 | 5.5 | 55 | 12.69 | 7.37
2008-10-12 | 21:15 | 36 | 100.9 | 5.42 | 55 | 12.74 | 7.37
2008-10-12 | 21:30 | 36 | 101 | 5.36 | 55 | 12.77 | 7.37
2008-10-12 | 21:45 | 36 | 100.6 | 5.3 | 55 | 12.75 | 7.38
2008-10-12 | 22:00 | 36 | 101 | 5.25 | 55 | 12.81 | 7.37
2008-10-12 | 22:15 | 36 | 100.5 | 5.2 | 55 | 12.77 | 7.37
2008-10-12 | 22:30 | 36 | 101 | 5.16 | 55 | 12.84 | 7.37
2008-10-12 | 22:45 | 36 | 101 | 5.12 | 55 | 12.87 | 7.37
2008-10-12 | 23:00 | 36 | 101.1 | 5.08 | 55 | 12.89 | 7.37
2008-10-12 | 23:15 | 36 | 101.2 | 5.04 | 55 | 12.91 | 7.37
2008-10-12 | 23:30 | 36 | 101.3 | 4.99 | 55 | 12.94 | 7.37
2008-10-12 | 23:45 | 36 | 100.8 | 4.95 | 55 | 12.89 | 7.36
2008-10-13 | 0:00 | 36 | 100.7 | 4.89 | 55 | 12.9 | 7.37
2008-10-13 | 0:15 | 36 | 101.1 | 4.84 | 55 | 12.97 | 7.37
2008-10-13 | 0:30 | 36 | 101 | 4.79 | 55 | 12.96 | 7.37
2008-10-13 | 0:45 | 36 | 101.2 | 4.76 | 55 | 13.01 | 7.36
2008-10-13 | 1:01 | 36 | 100.9 | 4.72 | 55 | 12.98 | 7.37
2008-10-13 | 1:16 | 36 | 101.2 | 4.68 | 55 | 13.04 | 7.37
2008-10-13 | 1:31 | 36 | 101.4 | 4.63 | 55 | 13.07 | 7.37
2008-10-13 | 1:46 | 36 | 101.5 | 4.59 | 55 | 13.1 | 7.37
2008-10-13 | 2:01 | 35 | 101.6 | 4.54 | 55 | 13.13 | 7.36
2008-10-13 | 2:16 | 35 | 101.7 | 4.48 | 54 | 13.17 | 7.36
2008-10-13 | 2:31 | 35 | 101.8 | 4.42 | 54 | 13.19 | 7.36
2008-10-13 | 2:46 | 35 | 101.8 | 4.36 | 54 | 13.22 | 7.36
2008-10-13 | 3:01 | 35 | 101.8 | 4.31 | 54 | 13.24 | 7.36
2008-10-13 | 3:16 | 35 | 102 | 4.25 | 54 | 13.28 | 7.35
2008-10-13 | 3:31 | 35 | 101.9 | 4.2 | 53 | 13.29 | 7.35
2008-10-13 | 3:46 | 34 | 101.9 | 4.15 | 53 | 13.3 | 7.34
2008-10-13 | 4:01 | 34 | 101.4 | 4.1 | 53 | 13.26 | 7.33
2008-10-13 | 4:16 | 34 | 101.9 | 4.04 | 52 | 13.34 | 7.34
2008-10-13 | 4:31 | 34 | 102 | 3.99 | 52 | 13.38 | 7.33
2008-10-13 | 4:46 | 34 | 101.6 | 3.93 | 52 | 13.34 | 7.33
2008-10-13 | 5:01 | 34 | 101.4 | 3.86 | 52 | 13.33 | 7.32
2008-10-13 | 5:16 | 34 | 101.3 | 3.8 | 52 | 13.34 | 7.33
2008-10-13 | 5:31 | 34 | 101.3 | 3.74 | 52 | 13.36 | 7.33
2008-10-13 | 5:46 | 34 | 101.3 | 3.69 | 52 | 13.39 | 7.33
2008-10-13 | 6:01 | 33 | 101.8 | 3.63 | 52 | 13.48 | 7.32
2008-10-13 | 6:16 | 33 | 101.6 | 3.57 | 51 | 13.47 | 7.32
2008-10-13 | 6:31 | 33 | 102 | 3.52 | 51 | 13.54 | 7.32
2008-10-13 | 6:46 | 33 | 101.7 | 3.47 | 51 | 13.51 | 7.31
2008-10-13 | 7:01 | 33 | 101.6 | 3.42 | 51 | 13.52 | 7.31
2008-10-13 | 7:16 | 33 | 101.5 | 3.37 | 51 | 13.53 | 7.32
2008-10-13 | 7:31 | 33 | 102.1 | 3.32 | 51 | 13.62 | 7.31
2008-10-13 | 7:46 | 33 | 102.4 | 3.27 | 51 | 13.68 | 7.31
2008-10-13 | 8:01 | 33 | 102.6 | 3.23 | 51 | 13.72 | 7.31
2008-10-13 | 8:16 | 33 | 102.7 | 3.19 | 51 | 13.75 | 7.31
2008-10-13 | 8:31 | 33 | 102.4 | 3.16 | 51 | 13.72 | 7.31
2008-10-13 | 8:46 | 33 | 102.7 | 3.14 | 51 | 13.78 | 7.32
2008-10-13 | 9:01 | 33 | 102.9 | 3.13 | 52 | 13.81 | 7.31
2008-10-13 | 9:16 | 34 | 102.5 | 3.12 | 52 | 13.76 | 7.32
2008-10-13 | 9:31 | 34 | 103 | 3.12 | 52 | 13.81 | 7.31
2008-10-13 | 9:46 | 34 | 102.7 | 3.14 | 52 | 13.77 | 7.31
2008-10-13 | 10:01 | 34 | 102.6 | 3.17 | 52 | 13.75 | 7.32
2008-10-13 | 10:16 | 34 | 103.1 | 3.2 | 52 | 13.8 | 7.31
2008-10-13 | 10:31 | 34 | 103.4 | 3.27 | 52 | 13.82 | 7.32
2008-10-13 | 10:46 | 34 | 103.1 | 3.32 | 52 | 13.76 | 7.33
2008-10-13 | 11:01 | 34 | 102.9 | 3.38 | 52 | 13.71 | 7.33
2008-10-13 | 11:16 | 34 | 102.9 | 3.46 | 52 | 13.69 | 7.34
2008-10-13 | 11:31 | 34 | 102.9 | 3.57 | 52 | 13.64 | 7.34
2008-10-13 | 11:46 | 34 | 103.3 | 3.73 | 52 | 13.64 | 7.35
2008-10-13 | 12:01 | 34 | 103.6 | 3.95 | 52 | 13.6 | 7.35
2008-10-13 | 12:16 | 34 | 103.2 | 4.23 | 52 | 13.45 | 7.36
2008-10-13 | 12:31 | 34 | 103.4 | 4.54 | 52 | 13.36 | 7.37
2008-10-13 | 12:46 | 34 | 102.8 | 4.89 | 52 | 13.17 | 7.38
2008-10-13 | 13:00 | 34 | 102.9 | 5.28 | 52 | 13.05 | 7.38
2008-10-13 | 13:15 | 34 | 103 | 5.66 | 52 | 12.93 | 7.39
2008-10-13 | 13:30 | 34 | 102.8 | 6.08 | 52 | 12.77 | 7.4
2008-10-13 | 13:45 | 34 | 102.8 | 6.47 | 52 | 12.64 | 7.4
2008-10-13 | 14:00 | 34 | 102.1 | 6.82 | 52 | 12.44 | 7.42
2008-10-13 | 14:15 | 34 | 102.2 | 7.16 | 52 | 12.35 | 7.43
2008-10-13 | 14:30 | 34 | 101.8 | 7.44 | 52 | 12.22 | 7.43
2008-10-13 | 14:45 | 34 | 101.9 | 7.67 | 52 | 12.16 | 7.43
2008-10-13 | 15:00 | 34 | 101.4 | 7.86 | 52 | 12.05 | 7.43
2008-10-13 | 15:15 | 34 | 100.9 | 8.01 | 52 | 11.94 | 7.43
2008-10-13 | 15:30 | 34 | 101.1 | 8.13 | 52 | 11.93 | 7.43
2008-10-13 | 15:45 | 34 | 100.6 | 8.27 | 52 | 11.84 | 7.43
2008-10-13 | 16:00 | 34 | 100.4 | 8.37 | 52 | 11.78 | 7.44
2008-10-13 | 16:15 | 34 | 100.6 | 8.4 | 52 | 11.8 | 7.43
2008-10-13 | 16:30 | 34 | 100.2 | 8.44 | 52 | 11.74 | 7.43
2008-10-13 | 16:45 | 34 | 100.3 | 8.46 | 52 | 11.75 | 7.43
2008-10-13 | 17:00 | 34 | 99.8 | 8.48 | 52 | 11.69 | 7.43
2008-10-13 | 17:15 | 34 | 99.5 | 8.5 | 52 | 11.64 | 7.42
2008-10-13 | 17:30 | 34 | 99.8 | 8.53 | 52 | 11.67 | 7.41
2008-10-13 | 17:45 | 34 | 99.9 | 8.58 | 52 | 11.66 | 7.41
2008-10-13 | 18:00 | 34 | 99.4 | 8.62 | 52 | 11.6 | 7.41
2008-10-13 | 18:15 | 34 | 99.6 | 8.67 | 52 | 11.6 | 7.4
2008-10-13 | 18:30 | 34 | 99.5 | 8.69 | 52 | 11.59 | 7.39
2008-10-13 | 18:45 | 34 | 99.6 | 8.71 | 52 | 11.59 | 7.39
2008-10-13 | 19:00 | 34 | 99.6 | 8.7 | 52 | 11.59 | 7.38
2008-10-13 | 19:15 | 34 | 99 | 8.68 | 52 | 11.53 | 7.38
2008-10-13 | 19:30 | 34 | 98.9 | 8.67 | 52 | 11.52 | 7.38
2008-10-13 | 19:45 | 34 | 99.2 | 8.65 | 52 | 11.57 | 7.38
2008-10-13 | 20:00 | 34 | 99.3 | 8.62 | 52 | 11.59 | 7.37
2008-10-13 | 20:15 | 34 | 99 | 8.57 | 52 | 11.57 | 7.37
2008-10-13 | 20:30 | 34 | 99.4 | 8.53 | 52 | 11.62 | 7.36
2008-10-13 | 20:45 | 34 | 99.4 | 8.47 | 52 | 11.64 | 7.37
2008-10-13 | 21:00 | 34 | 99.5 | 8.43 | 52 | 11.66 | 7.36
2008-10-13 | 21:15 | 34 | 99.6 | 8.4 | 52 | 11.68 | 7.36
2008-10-13 | 21:30 | 34 | 99.6 | 8.37 | 52 | 11.69 | 7.36
2008-10-13 | 21:45 | 34 | 99.7 | 8.35 | 52 | 11.7 | 7.36
2008-10-13 | 22:00 | 34 | 99.7 | 8.34 | 52 | 11.72 | 7.36
2008-10-13 | 22:15 | 34 | 99.2 | 8.32 | 53 | 11.66 | 7.36
2008-10-13 | 22:30 | 34 | 99 | 8.3 | 53 | 11.64 | 7.36
2008-10-13 | 22:45 | 34 | 99 | 8.27 | 53 | 11.65 | 7.36
2008-10-13 | 23:00 | 35 | 98.9 | 8.23 | 53 | 11.65 | 7.36
2008-10-13 | 23:15 | 35 | 99.4 | 8.18 | 53 | 11.71 | 7.36
2008-10-13 | 23:30 | 34 | 99.1 | 8.14 | 53 | 11.7 | 7.36
2008-10-13 | 23:45 | 35 | 99 | 8.09 | 53 | 11.7 | 7.36
2008-10-14 | 0:00 | 34 | 99.1 | 8.05 | 53 | 11.72 | 7.36
2008-10-14 | 0:15 | 35 | 99.4 | 8 | 53 | 11.77 | 7.36
2008-10-14 | 0:30 | 34 | 99.2 | 7.97 | 53 | 11.75 | 7.36
2008-10-14 | 0:45 | 34 | 99.1 | 7.94 | 53 | 11.75 | 7.36
2008-10-14 | 1:01 | 35 | 99 | 7.91 | 54 | 11.75 | 7.36
2008-10-14 | 1:16 | 34 | 99.5 | 7.86 | 53 | 11.82 | 7.35
2008-10-14 | 1:31 | 34 | 99.2 | 7.81 | 53 | 11.8 | 7.35
2008-10-14 | 1:46 | 34 | 99.5 | 7.77 | 53 | 11.85 | 7.35
2008-10-14 | 2:01 | 35 | 99.2 | 7.73 | 54 | 11.83 | 7.35
2008-10-14 | 2:16 | 35 | 99.1 | 7.68 | 54 | 11.83 | 7.35
2008-10-14 | 2:31 | 35 | 99.6 | 7.63 | 54 | 11.9 | 7.35
2008-10-14 | 2:46 | 35 | 99.3 | 7.59 | 54 | 11.88 | 7.36
2008-10-14 | 3:01 | 35 | 99.2 | 7.54 | 54 | 11.88 | 7.35
2008-10-14 | 3:16 | 35 | 99.2 | 7.49 | 53 | 11.9 | 7.36
2008-10-14 | 3:31 | 35 | 99.1 | 7.44 | 54 | 11.9 | 7.36
2008-10-14 | 3:46 | 35 | 99.6 | 7.41 | 54 | 11.97 | 7.35
2008-10-14 | 4:01 | 35 | 99.9 | 7.38 | 54 | 12.01 | 7.35
2008-10-14 | 4:16 | 35 | 100 | 7.34 | 53 | 12.04 | 7.36
2008-10-14 | 4:31 | 35 | 100.2 | 7.3 | 53 | 12.07 | 7.34
2008-10-14 | 4:46 | 35 | 100.1 | 7.25 | 53 | 12.07 | 7.35
2008-10-14 | 5:01 | 35 | 99.7 | 7.2 | 53 | 12.04 | 7.35
2008-10-14 | 5:16 | 35 | 100 | 7.14 | 53 | 12.1 | 7.35
2008-10-14 | 5:31 | 35 | 100.2 | 7.06 | 54 | 12.15 | 7.35
2008-10-14 | 5:46 | 35 | 100.3 | 6.99 | 53 | 12.18 | 7.35
2008-10-14 | 6:01 | 35 | 100.4 | 6.91 | 53 | 12.21 | 7.35
2008-10-14 | 6:16 | 35 | 99.9 | 6.84 | 53 | 12.17 | 7.35
2008-10-14 | 6:31 | 35 | 99.8 | 6.78 | 53 | 12.18 | 7.34
2008-10-14 | 6:46 | 35 | 99.7 | 6.72 | 53 | 12.19 | 7.34
2008-10-14 | 7:01 | 35 | 99.8 | 6.67 | 53 | 12.21 | 7.35
2008-10-14 | 7:16 | 35 | 99.8 | 6.62 | 53 | 12.23 | 7.35
2008-10-14 | 7:31 | 35 | 99.8 | 6.55 | 53 | 12.26 | 7.34
2008-10-14 | 7:46 | 35 | 99.9 | 6.47 | 53 | 12.28 | 7.35
2008-10-14 | 8:01 | 34 | 100.5 | 6.42 | 53 | 12.38 | 7.35
2008-10-14 | 8:16 | 34 | 100.8 | 6.37 | 53 | 12.43 | 7.35
2008-10-14 | 8:31 | 35 | 101.1 | 6.34 | 53 | 12.47 | 7.36
2008-10-14 | 8:46 | 35 | 101.3 | 6.33 | 53 | 12.51 | 7.36
2008-10-14 | 9:01 | 34 | 101 | 6.33 | 53 | 12.47 | 7.37
2008-10-14 | 9:16 | 34 | 101.4 | 6.36 | 53 | 12.51 | 7.37
2008-10-14 | 9:31 | 35 | 101.2 | 6.4 | 53 | 12.47 | 7.38
2008-10-14 | 9:46 | 35 | 101.7 | 6.46 | 53 | 12.51 | 7.39
2008-10-14 | 10:01 | 34 | 101.8 | 6.5 | 53 | 12.51 | 7.4
2008-10-14 | 10:16 | 34 | 101.8 | 6.53 | 53 | 12.51 | 7.4
2008-10-14 | 10:31 | 34 | 101.9 | 6.62 | 53 | 12.49 | 7.4
2008-10-14 | 10:46 | 34 | 102 | 6.68 | 53 | 12.47 | 7.4
2008-10-14 | 11:01 | 34 | 101.9 | 6.74 | 53 | 12.45 | 7.4
2008-10-14 | 11:16 | 34 | 101.9 | 6.82 | 53 | 12.42 | 7.41
2008-10-14 | 11:31 | 34 | 102 | 6.89 | 53 | 12.41 | 7.41
2008-10-14 | 11:46 | 34 | 102 | 6.99 | 53 | 12.38 | 7.4
2008-10-14 | 12:01 | 34 | 101.4 | 7.07 | 53 | 12.28 | 7.42
2008-10-14 | 12:16 | 34 | 101.6 | 7.33 | 53 | 12.23 | 7.41
2008-10-14 | 12:31 | 34 | 101.3 | 7.56 | 53 | 12.12 | 7.42
2008-10-14 | 12:46 | 34 | 100.9 | 7.76 | 52 | 12.02 | 7.43
2008-10-14 | 13:00 | 34 | 100.7 | 8.11 | 52 | 11.89 | 7.43
2008-10-14 | 13:15 | 34 | 101.1 | 8.45 | 52 | 11.84 | 7.45
2008-10-14 | 13:30 | 34 | 101.2 | 8.74 | 52 | 11.77 | 7.46
2008-10-14 | 13:45 | 34 | 101.1 | 8.96 | 53 | 11.7 | 7.47
2008-10-14 | 14:00 | 34 | 100.8 | 9.15 | 53 | 11.61 | 7.48
2008-10-14 | 14:15 | 34 | 100.9 | 9.42 | 53 | 11.54 | 7.55
2008-10-14 | 14:30 | 34 | 100.8 | 9.59 | 53 | 11.49 | 7.56
2008-10-14 | 14:45 | 34 | 100.9 | 9.78 | 53 | 11.45 | 7.56
2008-10-14 | 15:00 | 34 | 100.7 | 9.86 | 53 | 11.4 | 7.49
2008-10-14 | 15:15 | 34 | 100.3 | 9.94 | 53 | 11.34 | 7.49
2008-10-14 | 15:30 | 35 | 100.6 | 9.94 | 53 | 11.37 | 7.49
2008-10-14 | 15:45 | 34 | 100.9 | 9.89 | 53 | 11.42 | 7.47
2008-10-14 | 16:00 | 34 | 101 | 9.71 | 53 | 11.48 | 7.46
2008-10-14 | 16:15 | 34 | 100.9 | 9.56 | 53 | 11.51 | 7.45
2008-10-14 | 16:30 | 34 | 100.4 | 9.48 | 53 | 11.47 | 7.43
2008-10-14 | 16:45 | 34 | 100.4 | 9.44 | 53 | 11.49 | 7.44
2008-10-14 | 17:00 | 34 | 100.5 | 9.4 | 53 | 11.51 | 7.41
2008-10-14 | 17:15 | 34 | 100.4 | 9.37 | 53 | 11.5 | 7.41
2008-10-14 | 17:30 | 34 | 99.7 | 9.4 | 52 | 11.42 | 7.4
2008-10-14 | 17:45 | 34 | 99.8 | 9.41 | 53 | 11.42 | 7.39
2008-10-14 | 18:00 | 34 | 99.3 | 9.4 | 53 | 11.36 | 7.39
2008-10-14 | 18:15 | 34 | 99.5 | 9.36 | 53 | 11.4 | 7.37
2008-10-14 | 18:30 | 34 | 99 | 9.3 | 53 | 11.36 | 7.37
2008-10-14 | 18:45 | 35 | 98.7 | 9.27 | 53 | 11.34 | 7.36
2008-10-14 | 19:00 | 35 | 99 | 9.24 | 53 | 11.38 | 7.35
2008-10-14 | 19:15 | 35 | 99.2 | 9.18 | 53 | 11.42 | 7.35
2008-10-14 | 19:30 | 34 | 99.3 | 9.1 | 53 | 11.45 | 7.34
2008-10-14 | 19:45 | 34 | 99.3 | 9.03 | 53 | 11.47 | 7.34
2008-10-14 | 20:00 | 34 | 98.9 | 8.96 | 53 | 11.44 | 7.33
2008-10-14 | 20:15 | 34 | 99.2 | 8.89 | 53 | 11.5 | 7.34
2008-10-14 | 20:30 | 34 | 99.3 | 8.83 | 53 | 11.52 | 7.33
2008-10-14 | 20:45 | 34 | 98.9 | 8.76 | 53 | 11.5 | 7.32
2008-10-14 | 21:00 | 34 | 99.2 | 8.69 | 53 | 11.56 | 7.32
2008-10-14 | 21:15 | 34 | 98.9 | 8.63 | 53 | 11.53 | 7.34
2008-10-14 | 21:30 | 35 | 98.7 | 8.56 | 53 | 11.53 | 7.33
2008-10-14 | 21:45 | 35 | 99.1 | 8.49 | 53 | 11.6 | 7.33
2008-10-14 | 22:00 | 35 | 99.3 | 8.41 | 53 | 11.64 | 7.34
2008-10-14 | 22:15 | 35 | 99.5 | 8.32 | 53 | 11.69 | 7.33
2008-10-14 | 22:30 | 35 | 99 | 8.23 | 53 | 11.66 | 7.33
2008-10-14 | 22:45 | 35 | 99.3 | 8.14 | 53 | 11.72 | 7.34
2008-10-14 | 23:00 | 35 | 99 | 8.05 | 53 | 11.71 | 7.33
2008-10-14 | 23:15 | 35 | 98.9 | 7.96 | 53 | 11.72 | 7.33
2008-10-14 | 23:30 | 35 | 99.3 | 7.88 | 53 | 11.79 | 7.33
2008-10-14 | 23:45 | 35 | 99.5 | 7.8 | 54 | 11.84 | 7.33
2008-10-15 | 0:00 | 35 | 99.6 | 7.71 | 54 | 11.87 | 7.33
2008-10-15 | 0:15 | 35 | 99.7 | 7.63 | 54 | 11.91 | 7.34
2008-10-15 | 0:30 | 35 | 99.7 | 7.54 | 54 | 11.94 | 7.34
2008-10-15 | 0:45 | 35 | 99.8 | 7.45 | 54 | 11.98 | 7.33
2008-10-15 | 1:01 | 35 | 99.8 | 7.36 | 54 | 12.01 | 7.34
2008-10-15 | 1:16 | 35 | 99.9 | 7.27 | 54 | 12.04 | 7.34
2008-10-15 | 1:31 | 35 | 99.9 | 7.18 | 54 | 12.08 | 7.33
2008-10-15 | 1:46 | 35 | 100 | 7.08 | 54 | 12.11 | 7.34
2008-10-15 | 2:01 | 35 | 100 | 6.98 | 54 | 12.15 | 7.33
2008-10-15 | 2:16 | 35 | 99.6 | 6.89 | 54 | 12.12 | 7.33
2008-10-15 | 2:31 | 35 | 99.6 | 6.78 | 54 | 12.15 | 7.33
2008-10-15 | 2:46 | 35 | 99.5 | 6.67 | 54 | 12.17 | 7.34
2008-10-15 | 3:01 | 35 | 99.8 | 6.56 | 54 | 12.25 | 7.33
2008-10-15 | 3:16 | 35 | 100.1 | 6.44 | 54 | 12.32 | 7.33
2008-10-15 | 3:31 | 35 | 99.8 | 6.33 | 54 | 12.31 | 7.34
2008-10-15 | 3:46 | 35 | 99.6 | 6.23 | 54 | 12.33 | 7.33
2008-10-15 | 4:01 | 35 | 100 | 6.12 | 54 | 12.42 | 7.33
2008-10-15 | 4:16 | 35 | 100.3 | 6.02 | 54 | 12.48 | 7.33
2008-10-15 | 4:31 | 35 | 100.4 | 5.92 | 54 | 12.52 | 7.33
2008-10-15 | 4:46 | 35 | 100.6 | 5.83 | 54 | 12.57 | 7.32
2008-10-15 | 5:01 | 35 | 100.1 | 5.73 | 54 | 12.55 | 7.33
2008-10-15 | 5:16 | 35 | 100.4 | 5.63 | 54 | 12.62 | 7.33
2008-10-15 | 5:31 | 35 | 100.1 | 5.53 | 54 | 12.61 | 7.32
2008-10-15 | 5:46 | 35 | 100.5 | 5.44 | 54 | 12.69 | 7.32
2008-10-15 | 6:01 | 35 | 100.7 | 5.34 | 54 | 12.75 | 7.32
2008-10-15 | 6:16 | 35 | 100.3 | 5.24 | 54 | 12.73 | 7.32
2008-10-15 | 6:31 | 35 | 100.2 | 5.15 | 54 | 12.75 | 7.32
2008-10-15 | 6:46 | 35 | 100.7 | 5.05 | 54 | 12.84 | 7.3
2008-10-15 | 7:01 | 35 | 100.9 | 4.96 | 54 | 12.9 | 7.31
2008-10-15 | 7:16 | 35 | 101 | 4.86 | 54 | 12.95 | 7.32
2008-10-15 | 7:31 | 35 | 101.2 | 4.76 | 54 | 13 | 7.31
2008-10-15 | 7:46 | 35 | 101.3 | 4.68 | 54 | 13.05 | 7.31
2008-10-15 | 8:01 | 35 | 101.5 | 4.59 | 54 | 13.1 | 7.32
2008-10-15 | 8:16 | 35 | 101.8 | 4.51 | 54 | 13.16 | 7.32
2008-10-15 | 8:31 | 35 | 101.5 | 4.45 | 54 | 13.15 | 7.32
2008-10-15 | 8:46 | 35 | 102.1 | 4.4 | 54 | 13.24 | 7.33
2008-10-15 | 9:01 | 35 | 102.4 | 4.34 | 54 | 13.3 | 7.33
2008-10-15 | 9:16 | 35 | 102 | 4.28 | 54 | 13.27 | 7.33
2008-10-15 | 9:31 | 35 | 102.5 | 4.23 | 54 | 13.35 | 7.33
2008-10-15 | 9:46 | 35 | 102.2 | 4.21 | 54 | 13.33 | 7.34
2008-10-15 | 10:01 | 35 | 102.9 | 4.2 | 54 | 13.42 | 7.34
2008-10-15 | 10:16 | 35 | 103.5 | 4.21 | 54 | 13.49 | 7.34
2008-10-15 | 10:31 | 35 | 103.4 | 4.24 | 54 | 13.47 | 7.35
2008-10-15 | 10:46 | 35 | 104 | 4.28 | 54 | 13.54 | 7.35
2008-10-15 | 11:01 | 35 | 104.4 | 4.32 | 54 | 13.56 | 7.35
2008-10-15 | 11:16 | 35 | 104.5 | 4.36 | 53 | 13.56 | 7.35
2008-10-15 | 11:31 | 35 | 104.1 | 4.44 | 53 | 13.49 | 7.35
2008-10-15 | 11:46 | 35 | 103.8 | 4.55 | 53 | 13.42 | 7.36
2008-10-15 | 12:01 | 35 | 104.2 | 4.72 | 53 | 13.4 | 7.37
2008-10-15 | 12:16 | 35 | 104.2 | 4.94 | 53 | 13.33 | 7.37
2008-10-15 | 12:31 | 35 | 104.2 | 5.18 | 53 | 13.24 | 7.38
2008-10-15 | 12:46 | 35 | 103.6 | 5.47 | 53 | 13.08 | 7.39
2008-10-15 | 13:00 | 35 | 103.2 | 5.78 | 53 | 12.92 | 7.4
2008-10-15 | 13:15 | 35 | 103.4 | 6.11 | 53 | 12.84 | 7.41
2008-10-15 | 13:30 | 35 | 103 | 6.46 | 53 | 12.67 | 7.41
2008-10-15 | 13:45 | 35 | 103.3 | 6.8 | 53 | 12.6 | 7.41
2008-10-15 | 14:00 | 35 | 102.8 | 7.11 | 53 | 12.44 | 7.42
2008-10-15 | 14:15 | 35 | 102.6 | 7.41 | 53 | 12.33 | 7.43
2008-10-15 | 14:30 | 34 | 102.9 | 7.65 | 53 | 12.3 | 7.44
2008-10-15 | 14:45 | 34 | 102.5 | 7.83 | 53 | 12.19 | 7.44
2008-10-15 | 15:00 | 34 | 102.7 | 7.92 | 53 | 12.19 | 7.45
2008-10-15 | 15:15 | 34 | 102.7 | 7.97 | 53 | 12.17 | 7.45
2008-10-15 | 15:30 | 34 | 102.3 | 8.03 | 53 | 12.1 | 7.45
2008-10-15 | 15:45 | 34 | 101.8 | 8 | 53 | 12.05 | 7.45
2008-10-15 | 16:00 | 34 | 102 | 7.98 | 53 | 12.08 | 7.44
2008-10-15 | 16:15 | 34 | 102.1 | 7.96 | 53 | 12.1 | 7.44
2008-10-15 | 16:30 | 34 | 102.1 | 7.96 | 53 | 12.1 | 7.43
2008-10-15 | 16:45 | 34 | 101.6 | 7.98 | 53 | 12.03 | 7.44
2008-10-15 | 17:00 | 34 | 101.7 | 7.99 | 53 | 12.04 | 7.43
2008-10-15 | 17:15 | 34 | 101 | 8 | 53 | 11.97 | 7.44
2008-10-15 | 17:30 | 34 | 101.1 | 7.99 | 53 | 11.98 | 7.41
2008-10-15 | 17:45 | 34 | 101.1 | 7.97 | 53 | 11.98 | 7.41
2008-10-15 | 18:00 | 34 | 100.9 | 7.94 | 53 | 11.97 | 7.4
2008-10-15 | 18:15 | 34 | 100.7 | 7.87 | 53 | 11.97 | 7.38
2008-10-15 | 18:30 | 34 | 100.6 | 7.83 | 53 | 11.96 | 7.38
2008-10-15 | 18:45 | 34 | 99.9 | 7.8 | 53 | 11.9 | 7.38
2008-10-15 | 19:00 | 34 | 100.2 | 7.77 | 53 | 11.93 | 7.37
2008-10-15 | 19:15 | 34 | 100.2 | 7.74 | 53 | 11.94 | 7.36
2008-10-15 | 19:30 | 34 | 100.2 | 7.71 | 53 | 11.95 | 7.35
2008-10-15 | 19:45 | 34 | 100.2 | 7.67 | 53 | 11.96 | 7.34
2008-10-15 | 20:00 | 34 | 99.8 | 7.62 | 53 | 11.92 | 7.34
2008-10-15 | 20:15 | 34 | 99.6 | 7.58 | 53 | 11.91 | 7.34
2008-10-15 | 20:30 | 34 | 99.9 | 7.54 | 53 | 11.96 | 7.34
2008-10-15 | 20:45 | 35 | 99.6 | 7.5 | 53 | 11.94 | 7.33
2008-10-15 | 21:00 | 35 | 99.9 | 7.46 | 53 | 11.99 | 7.34
2008-10-15 | 21:15 | 35 | 99.6 | 7.43 | 53 | 11.96 | 7.34
2008-10-15 | 21:30 | 35 | 100 | 7.4 | 53 | 12.02 | 7.32
2008-10-15 | 21:45 | 35 | 100.1 | 7.37 | 53 | 12.03 | 7.33
2008-10-15 | 22:00 | 35 | 100.1 | 7.35 | 53 | 12.05 | 7.34
2008-10-15 | 22:15 | 35 | 100.2 | 7.35 | 53 | 12.06 | 7.33
2008-10-15 | 22:30 | 35 | 99.7 | 7.35 | 54 | 12 | 7.33
2008-10-15 | 22:45 | 35 | 100 | 7.34 | 54 | 12.04 | 7.34
2008-10-15 | 23:00 | 35 | 100 | 7.32 | 54 | 12.05 | 7.34
2008-10-15 | 23:15 | 35 | 99.7 | 7.29 | 54 | 12.01 | 7.34
2008-10-15 | 23:30 | 35 | 99.9 | 7.26 | 54 | 12.05 | 7.35
2008-10-15 | 23:45 | 35 | 99.6 | 7.22 | 54 | 12.02 | 7.34
2008-10-16 | 0:00 | 35 | 99.5 | 7.19 | 54 | 12.02 | 7.34
2008-10-16 | 0:15 | 35 | 99.9 | 7.14 | 54 | 12.08 | 7.35
2008-10-16 | 0:30 | 35 | 100.1 | 7.1 | 54 | 12.12 | 7.33
2008-10-16 | 0:45 | 35 | 100.2 | 7.05 | 54 | 12.14 | 7.33
2008-10-16 | 1:01 | 35 | 99.8 | 7 | 54 | 12.11 | 7.34
2008-10-16 | 1:16 | 35 | 99.7 | 6.95 | 54 | 12.11 | 7.36
2008-10-16 | 1:31 | 35 | 100.1 | 6.9 | 54 | 12.18 | 7.34
2008-10-16 | 1:46 | 35 | 100.2 | 6.84 | 54 | 12.21 | 7.33
2008-10-16 | 2:01 | 35 | 99.9 | 6.8 | 54 | 12.18 | 7.33
2008-10-16 | 2:16 | 35 | 100.2 | 6.75 | 54 | 12.24 | 7.33
2008-10-16 | 2:31 | 35 | 100.4 | 6.72 | 54 | 12.27 | 7.34
2008-10-16 | 2:46 | 35 | 100.4 | 6.68 | 54 | 12.28 | 7.34
2008-10-16 | 3:01 | 35 | 100.5 | 6.66 | 54 | 12.3 | 7.34
2008-10-16 | 3:16 | 35 | 100.1 | 6.62 | 54 | 12.26 | 7.35
2008-10-16 | 3:31 | 35 | 100.3 | 6.59 | 54 | 12.3 | 7.33
2008-10-16 | 3:46 | 35 | 100 | 6.55 | 54 | 12.28 | 7.33
2008-10-16 | 4:01 | 35 | 100.3 | 6.51 | 54 | 12.32 | 7.34
2008-10-16 | 4:16 | 35 | 99.9 | 6.46 | 54 | 12.29 | 7.34
2008-10-16 | 4:31 | 35 | 100.3 | 6.41 | 54 | 12.35 | 7.34
2008-10-16 | 4:46 | 35 | 100 | 6.37 | 54 | 12.33 | 7.34
2008-10-16 | 5:01 | 35 | 100.3 | 6.32 | 54 | 12.39 | 7.34
2008-10-16 | 5:16 | 35 | 100.5 | 6.28 | 54 | 12.42 | 7.33
2008-10-16 | 5:31 | 35 | 100.6 | 6.24 | 54 | 12.45 | 7.33
2008-10-16 | 5:46 | 35 | 100.7 | 6.2 | 54 | 12.47 | 7.32
2008-10-16 | 6:01 | 35 | 100.8 | 6.16 | 54 | 12.49 | 7.33
2008-10-16 | 6:16 | 35 | 100.8 | 6.12 | 54 | 12.51 | 7.32
2008-10-16 | 6:31 | 35 | 100.4 | 6.08 | 54 | 12.47 | 7.33
2008-10-16 | 6:46 | 35 | 100.7 | 6.04 | 54 | 12.52 | 7.31
2008-10-16 | 7:01 | 35 | 100.3 | 6 | 54 | 12.49 | 7.32
2008-10-16 | 7:16 | 35 | 100.2 | 5.98 | 54 | 12.48 | 7.33
2008-10-16 | 7:31 | 35 | 100.7 | 5.95 | 54 | 12.55 | 7.32
2008-10-16 | 7:46 | 35 | 100.9 | 5.94 | 54 | 12.58 | 7.32
2008-10-16 | 8:01 | 35 | 101.2 | 5.91 | 54 | 12.62 | 7.32
2008-10-16 | 8:16 | 35 | 100.9 | 5.89 | 54 | 12.59 | 7.34
2008-10-16 | 8:31 | 35 | 100.8 | 5.86 | 54 | 12.6 | 7.34
2008-10-16 | 8:46 | 35 | 101.5 | 5.85 | 54 | 12.69 | 7.34
2008-10-16 | 9:01 | 35 | 102 | 5.85 | 54 | 12.75 | 7.34
2008-10-16 | 9:16 | 35 | 101.9 | 5.86 | 54 | 12.73 | 7.35
2008-10-16 | 9:31 | 35 | 102.5 | 5.9 | 54 | 12.79 | 7.36
2008-10-16 | 9:46 | 35 | 102.2 | 5.91 | 54 | 12.75 | 7.37
2008-10-16 | 10:01 | 35 | 102.1 | 5.9 | 54 | 12.74 | 7.37
2008-10-16 | 10:16 | 35 | 102.8 | 5.9 | 54 | 12.83 | 7.37
2008-10-16 | 10:31 | 35 | 103.2 | 5.92 | 54 | 12.88 | 7.38
2008-10-16 | 10:46 | 35 | 103.6 | 5.96 | 54 | 12.91 | 7.38
2008-10-16 | 11:01 | 35 | 103.7 | 6 | 54 | 12.91 | 7.38
2008-10-16 | 11:16 | 35 | 103.9 | 6.05 | 54 | 12.91 | 7.38
2008-10-16 | 11:31 | 35 | 103.4 | 6.15 | 54 | 12.82 | 7.38
2008-10-16 | 11:46 | 35 | 103.6 | 6.28 | 54 | 12.8 | 7.39
2008-10-16 | 12:01 | 35 | 103.6 | 6.45 | 54 | 12.75 | 7.4
2008-10-16 | 12:16 | 35 | 103.1 | 6.68 | 54 | 12.62 | 7.4
2008-10-16 | 12:31 | 35 | 102.8 | 6.92 | 54 | 12.5 | 7.41
2008-10-16 | 12:46 | 35 | 103.1 | 7.19 | 54 | 12.46 | 7.42
2008-10-16 | 13:00 | 35 | 102.8 | 7.54 | 53 | 12.32 | 7.43
2008-10-16 | 13:15 | 35 | 102.5 | 7.9 | 53 | 12.17 | 7.44
2008-10-16 | 13:30 | 35 | 102.2 | 8.25 | 53 | 12.03 | 7.45
2008-10-16 | 13:45 | 35 | 102.6 | 8.61 | 53 | 11.97 | 7.47
2008-10-16 | 14:00 | 35 | 102.7 | 8.9 | 53 | 11.9 | 7.51
2008-10-16 | 14:15 | 35 | 102.7 | 9.2 | 53 | 11.81 | 7.53
2008-10-16 | 14:30 | 35 | 102 | 9.56 | 53 | 11.63 | 7.56
2008-10-16 | 14:45 | 35 | 102.1 | 9.84 | 53 | 11.57 | 7.57
2008-10-16 | 15:00 | 35 | 102 | 10.02 | 53 | 11.5 | 7.55
2008-10-16 | 15:15 | 35 | 102 | 10.12 | 53 | 11.48 | 7.53
2008-10-16 | 15:30 | 35 | 101.2 | 10.21 | 53 | 11.37 | 7.52
2008-10-16 | 15:45 | 35 | 100.8 | 10.25 | 53 | 11.31 | 7.51
2008-10-16 | 16:00 | 35 | 100.8 | 10.26 | 53 | 11.3 | 7.5
2008-10-16 | 16:15 | 35 | 100.5 | 10.25 | 53 | 11.28 | 7.49
2008-10-16 | 16:30 | 35 | 100.3 | 10.23 | 53 | 11.26 | 7.47
2008-10-16 | 16:45 | 35 | 100.5 | 10.2 | 53 | 11.29 | 7.47
2008-10-16 | 17:00 | 35 | 100.5 | 10.16 | 53 | 11.3 | 7.44
2008-10-16 | 17:15 | 35 | 99.9 | 10.15 | 53 | 11.23 | 7.44
2008-10-16 | 17:30 | 35 | 100.1 | 10.17 | 53 | 11.25 | 7.42
2008-10-16 | 17:45 | 35 | 100.1 | 10.17 | 53 | 11.25 | 7.41
2008-10-16 | 18:00 | 35 | 99.4 | 10.17 | 53 | 11.18 | 7.4
2008-10-16 | 18:15 | 35 | 99.6 | 10.16 | 53 | 11.2 | 7.38
2008-10-16 | 18:30 | 35 | 99.5 | 10.15 | 53 | 11.19 | 7.38
2008-10-16 | 18:45 | 35 | 99.4 | 10.12 | 53 | 11.19 | 7.37
2008-10-16 | 19:00 | 35 | 99 | 10.08 | 53 | 11.16 | 7.36
2008-10-16 | 19:15 | 35 | 99.3 | 10 | 54 | 11.2 | 7.36
2008-10-16 | 19:30 | 35 | 98.9 | 9.91 | 54 | 11.19 | 7.35
2008-10-16 | 19:45 | 35 | 98.8 | 9.81 | 53 | 11.2 | 7.35
2008-10-16 | 20:00 | 35 | 99.2 | 9.72 | 54 | 11.27 | 7.34
2008-10-16 | 20:15 | 35 | 99.4 | 9.64 | 54 | 11.31 | 7.33
2008-10-16 | 20:30 | 35 | 99 | 9.57 | 54 | 11.29 | 7.34
2008-10-16 | 20:45 | 35 | 99.3 | 9.5 | 54 | 11.35 | 7.34
2008-10-16 | 21:00 | 35 | 99.4 | 9.44 | 54 | 11.37 | 7.34
2008-10-16 | 21:15 | 35 | 99.5 | 9.38 | 54 | 11.4 | 7.34
2008-10-16 | 21:30 | 35 | 99.1 | 9.32 | 54 | 11.37 | 7.34
2008-10-16 | 21:45 | 35 | 99.4 | 9.27 | 54 | 11.42 | 7.34
2008-10-16 | 22:00 | 35 | 99.6 | 9.21 | 54 | 11.45 | 7.35
2008-10-16 | 22:15 | 35 | 99.2 | 9.15 | 54 | 11.42 | 7.34
2008-10-16 | 22:30 | 35 | 99 | 9.1 | 54 | 11.42 | 7.35
2008-10-16 | 22:45 | 35 | 99.5 | 9.04 | 54 | 11.48 | 7.35
2008-10-16 | 23:00 | 35 | 99.6 | 8.98 | 54 | 11.52 | 7.35
2008-10-16 | 23:15 | 35 | 99.3 | 8.92 | 54 | 11.5 | 7.35
2008-10-16 | 23:30 | 35 | 99.1 | 8.86 | 54 | 11.49 | 7.34
2008-10-16 | 23:45 | 35 | 99.5 | 8.79 | 54 | 11.56 | 7.34
2008-10-17 | 0:00 | 35 | 99.7 | 8.73 | 54 | 11.6 | 7.34
2008-10-17 | 0:15 | 35 | 99.3 | 8.67 | 54 | 11.57 | 7.35
2008-10-17 | 0:30 | 35 | 99.7 | 8.6 | 54 | 11.63 | 7.34
2008-10-17 | 0:45 | 35 | 99.8 | 8.53 | 54 | 11.67 | 7.35
2008-10-17 | 1:01 | 35 | 100 | 8.46 | 54 | 11.71 | 7.34
2008-10-17 | 1:16 | 35 | 100.1 | 8.38 | 54 | 11.74 | 7.35
2008-10-17 | 1:31 | 35 | 99.7 | 8.31 | 54 | 11.71 | 7.27
2008-10-17 | 1:46 | 35 | 99.5 | 8.24 | 54 | 11.71 | 7.33
2008-10-17 | 2:01 | 35 | 99.4 | 8.16 | 54 | 11.73 | 7.33
2008-10-17 | 2:16 | 35 | 99.3 | 8.09 | 53 | 11.74 | 7.32
2008-10-17 | 2:31 | 35 | 99.8 | 8.01 | 53 | 11.82 | 7.34
2008-10-17 | 2:46 | 35 | 100 | 7.94 | 53 | 11.86 | 7.34
2008-10-17 | 3:01 | 35 | 100.1 | 7.88 | 53 | 11.89 | 7.31
2008-10-17 | 3:16 | 35 | 99.8 | 7.81 | 53 | 11.87 | 7.29
2008-10-17 | 3:31 | 35 | 99.6 | 7.75 | 53 | 11.87 | 7.28
2008-10-17 | 3:46 | 35 | 100.1 | 7.67 | 53 | 11.95 | 7.29
2008-10-17 | 4:01 | 34 | 99.8 | 7.58 | 53 | 11.94 | 7.29
2008-10-17 | 4:16 | 34 | 99 | 7.48 | 52 | 11.87 | 7.31
2008-10-17 | 4:31 | 33 | 98.5 | 7.38 | 51 | 11.85 | 7.32
2008-10-17 | 4:46 | 32 | 99.5 | 7.28 | 50 | 11.99 | 7.3
2008-10-17 | 5:01 | 32 | 99.7 | 7.16 | 49 | 12.06 | 7.29
2008-10-17 | 5:16 | 31 | 100.4 | 7.05 | 48 | 12.17 | 7.26
2008-10-17 | 5:31 | 31 | 100.6 | 6.94 | 47 | 12.23 | 7.24
2008-10-17 | 5:46 | 30 | 100.2 | 6.83 | 47 | 12.22 | 7.22
2008-10-17 | 6:01 | 30 | 100.6 | 6.72 | 47 | 12.29 | 7.22
2008-10-17 | 6:16 | 30 | 100.3 | 6.63 | 46 | 12.29 | 7.25
2008-10-17 | 6:31 | 30 | 100.2 | 6.54 | 46 | 12.3 | 7.22
2008-10-17 | 6:46 | 30 | 100.6 | 6.46 | 46 | 12.38 | 7.26
2008-10-17 | 7:01 | 30 | 100.8 | 6.39 | 46 | 12.43 | 7.23
2008-10-17 | 7:16 | 30 | 100.9 | 6.31 | 46 | 12.46 | 7.25
2008-10-17 | 7:31 | 30 | 101 | 6.24 | 46 | 12.5 | 7.23
2008-10-17 | 7:46 | 30 | 100.6 | 6.18 | 45 | 12.47 | 7.23
2008-10-17 | 8:01 | 30 | 101 | 6.12 | 45 | 12.54 | 7.21
2008-10-17 | 8:16 | 29 | 101.2 | 6.06 | 45 | 12.57 | 7.2
2008-10-17 | 8:31 | 29 | 101.4 | 6.02 | 45 | 12.61 | 7.2
2008-10-17 | 8:46 | 29 | 101 | 5.99 | 45 | 12.58 | 7.23
2008-10-17 | 9:01 | 30 | 101 | 5.97 | 45 | 12.58 | 7.24
2008-10-17 | 9:16 | 29 | 101 | 5.97 | 45 | 12.58 | 7.22
2008-10-17 | 9:31 | 29 | 101.6 | 5.97 | 45 | 12.65 | 7.23
2008-10-17 | 9:46 | 29 | 101.8 | 5.98 | 45 | 12.68 | 7.23
2008-10-17 | 10:01 | 29 | 101.6 | 5.99 | 45 | 12.65 | 7.23
2008-10-17 | 10:16 | 29 | 101.7 | 6 | 45 | 12.66 | 7.23
2008-10-17 | 10:31 | 29 | 102.3 | 6.03 | 45 | 12.73 | 7.23
2008-10-17 | 10:46 | 29 | 102.2 | 6.06 | 45 | 12.71 | 7.25
2008-10-17 | 11:01 | 29 | 102.3 | 6.1 | 45 | 12.7 | 7.24
2008-10-17 | 11:16 | 29 | 102.3 | 6.16 | 45 | 12.68 | 7.23
2008-10-17 | 11:31 | 29 | 102.7 | 6.26 | 45 | 12.7 | 7.26
2008-10-17 | 11:46 | 29 | 102.4 | 6.4 | 45 | 12.62 | 7.24
2008-10-17 | 12:01 | 29 | 102.7 | 6.58 | 45 | 12.6 | 7.27
2008-10-17 | 12:16 | 29 | 102.8 | 6.79 | 45 | 12.55 | 7.28
2008-10-17 | 12:31 | 29 | 102.8 | 7.02 | 45 | 12.48 | 7.27
2008-10-17 | 12:46 | 29 | 102.3 | 7.28 | 45 | 12.34 | 7.29
2008-10-17 | 13:00 | 29 | 102.5 | 7.53 | 45 | 12.28 | 7.29
2008-10-17 | 13:15 | 29 | 102.1 | 7.78 | 45 | 12.15 | 7.31
2008-10-17 | 13:30 | 29 | 102.3 | 8.04 | 45 | 12.1 | 7.31
2008-10-17 | 13:45 | 29 | 101.9 | 8.3 | 44 | 11.98 | 7.32
2008-10-17 | 14:00 | 29 | 101.6 | 8.55 | 44 | 11.88 | 7.34
2008-10-17 | 14:15 | 29 | 102 | 8.78 | 44 | 11.86 | 7.35
2008-10-17 | 14:30 | 29 | 102.1 | 8.97 | 44 | 11.81 | 7.34
2008-10-17 | 14:45 | 29 | 101.5 | 9.14 | 44 | 11.7 | 7.36
2008-10-17 | 15:00 | 29 | 101.2 | 9.26 | 44 | 11.63 | 7.37
2008-10-17 | 15:15 | 29 | 101.4 | 9.36 | 44 | 11.63 | 7.36
2008-10-17 | 15:30 | 29 | 101.5 | 9.44 | 44 | 11.6 | 7.35
2008-10-17 | 15:45 | 29 | 100.9 | 9.52 | 44 | 11.53 | 7.37
2008-10-17 | 16:00 | 29 | 101.1 | 9.56 | 44 | 11.53 | 7.36
2008-10-17 | 16:15 | 29 | 101 | 9.59 | 44 | 11.51 | 7.33
2008-10-17 | 16:30 | 29 | 100.9 | 9.62 | 44 | 11.49 | 7.33
2008-10-17 | 16:45 | 29 | 100.2 | 9.67 | 44 | 11.4 | 7.32
2008-10-17 | 17:00 | 29 | 100.3 | 9.71 | 44 | 11.4 | 7.33
2008-10-17 | 17:15 | 29 | 99.8 | 9.76 | 44 | 11.33 | 7.32
2008-10-17 | 17:30 | 29 | 99.5 | 9.8 | 44 | 11.28 | 7.33
2008-10-17 | 17:45 | 29 | 99.3 | 9.83 | 44 | 11.25 | 7.27
2008-10-17 | 18:00 | 29 | 99.6 | 9.87 | 45 | 11.27 | 7.28
2008-10-17 | 18:15 | 29 | 99.1 | 9.9 | 45 | 11.22 | 7.29
2008-10-17 | 18:30 | 29 | 99.3 | 9.89 | 45 | 11.24 | 7.26
2008-10-17 | 18:45 | 29 | 98.9 | 9.85 | 45 | 11.2 | 7.24
2008-10-17 | 19:00 | 29 | 98.7 | 9.8 | 45 | 11.2 | 7.26
2008-10-17 | 19:15 | 29 | 98.6 | 9.75 | 45 | 11.2 | 7.27
2008-10-17 | 19:30 | 29 | 99 | 9.7 | 45 | 11.25 | 7.2
2008-10-17 | 19:45 | 29 | 98.7 | 9.65 | 45 | 11.23 | 7.2
2008-10-17 | 20:00 | 29 | 99 | 9.6 | 45 | 11.28 | 7.22
2008-10-17 | 20:15 | 29 | 99.2 | 9.56 | 45 | 11.31 | 7.23
2008-10-17 | 20:30 | 29 | 98.8 | 9.53 | 45 | 11.28 | 7.19
2008-10-17 | 20:45 | 29 | 99.1 | 9.51 | 45 | 11.31 | 7.27
2008-10-17 | 21:00 | 29 | 99.2 | 9.48 | 45 | 11.33 | 7.25
2008-10-17 | 21:15 | 29 | 99.2 | 9.45 | 45 | 11.35 | 7.26
2008-10-17 | 21:30 | 29 | 99.3 | 9.41 | 45 | 11.37 | 7.24
2008-10-17 | 21:45 | 29 | 99.3 | 9.37 | 45 | 11.38 | 7.23
2008-10-17 | 22:00 | 29 | 98.9 | 9.34 | 45 | 11.34 | 7.27
2008-10-17 | 22:15 | 29 | 99.1 | 9.31 | 45 | 11.38 | 7.26
2008-10-17 | 22:30 | 29 | 98.8 | 9.27 | 45 | 11.35 | 7.26
2008-10-17 | 22:45 | 29 | 98.7 | 9.24 | 45 | 11.34 | 7.26
2008-10-17 | 23:00 | 29 | 99.1 | 9.21 | 45 | 11.4 | 7.26
2008-10-17 | 23:15 | 29 | 99.3 | 9.16 | 45 | 11.43 | 7.26
2008-10-17 | 23:30 | 29 | 99.4 | 9.12 | 45 | 11.45 | 7.25
2008-10-17 | 23:45 | 29 | 99 | 9.06 | 45 | 11.42 | 7.25
2008-10-18 | 0:00 | 29 | 98.8 | 9.01 | 45 | 11.42 | 7.26
2008-10-18 | 0:15 | 29 | 99.2 | 8.96 | 45 | 11.48 | 7.25
2008-10-18 | 0:30 | 29 | 99 | 8.91 | 45 | 11.46 | 7.24
2008-10-18 | 0:45 | 29 | 98.8 | 8.87 | 45 | 11.45 | 7.26
2008-10-18 | 1:01 | 29 | 98.8 | 8.83 | 45 | 11.47 | 7.24
2008-10-18 | 1:16 | 29 | 98.7 | 8.8 | 45 | 11.47 | 7.23
2008-10-18 | 1:31 | 29 | 98.7 | 8.76 | 45 | 11.47 | 7.23
2008-10-18 | 1:46 | 29 | 99.1 | 8.72 | 45 | 11.54 | 7.24
2008-10-18 | 2:01 | 29 | 98.9 | 8.68 | 45 | 11.53 | 7.24
2008-10-18 | 2:16 | 29 | 99.3 | 8.64 | 45 | 11.58 | 7.24
2008-10-18 | 2:31 | 29 | 99.5 | 8.6 | 45 | 11.61 | 7.22
2008-10-18 | 2:46 | 29 | 99.6 | 8.56 | 45 | 11.64 | 7.23
2008-10-18 | 3:01 | 29 | 99.7 | 8.51 | 45 | 11.66 | 7.23
2008-10-18 | 3:16 | 29 | 99.2 | 8.47 | 45 | 11.62 | 7.23
2008-10-18 | 3:31 | 29 | 99.1 | 8.42 | 45 | 11.62 | 7.24
2008-10-18 | 3:46 | 29 | 99 | 8.36 | 45 | 11.62 | 7.25
2008-10-18 | 4:01 | 29 | 99.4 | 8.29 | 45 | 11.69 | 7.24
2008-10-18 | 4:16 | 29 | 99.6 | 8.22 | 45 | 11.73 | 7.22
2008-10-18 | 4:31 | 29 | 100 | 8.16 | 45 | 11.79 | 7.22
2008-10-18 | 4:46 | 29 | 100 | 8.1 | 45 | 11.82 | 7.21
2008-10-18 | 5:01 | 29 | 100 | 8.04 | 45 | 11.84 | 7.22
2008-10-18 | 5:16 | 29 | 100 | 7.99 | 45 | 11.84 | 7.23
2008-10-18 | 5:31 | 29 | 100 | 7.94 | 45 | 11.86 | 7.22
2008-10-18 | 5:46 | 29 | 100.1 | 7.89 | 44 | 11.88 | 7.22
2008-10-18 | 6:01 | 29 | 100.1 | 7.84 | 45 | 11.9 | 7.22
2008-10-18 | 6:16 | 29 | 100.3 | 7.79 | 45 | 11.94 | 7.22
2008-10-18 | 6:31 | 29 | 99.7 | 7.74 | 44 | 11.88 | 7.22
2008-10-18 | 6:46 | 29 | 100 | 7.7 | 44 | 11.93 | 7.22
2008-10-18 | 7:01 | 29 | 100.2 | 7.67 | 44 | 11.96 | 7.21
2008-10-18 | 7:16 | 29 | 100.2 | 7.64 | 44 | 11.97 | 7.21
2008-10-18 | 7:31 | 29 | 99.7 | 7.63 | 44 | 11.92 | 7.21
2008-10-18 | 7:46 | 29 | 99.6 | 7.61 | 44 | 11.91 | 7.21
2008-10-18 | 8:01 | 29 | 100.1 | 7.61 | 44 | 11.96 | 7.22
2008-10-18 | 8:16 | 29 | 99.9 | 7.62 | 44 | 11.95 | 7.22
2008-10-18 | 8:31 | 29 | 100.6 | 7.64 | 44 | 12.02 | 7.23
2008-10-18 | 8:46 | 29 | 100.9 | 7.67 | 44 | 12.05 | 7.24
2008-10-18 | 9:01 | 29 | 100.7 | 7.72 | 44 | 12.01 | 7.26
2008-10-18 | 9:16 | 29 | 101.1 | 7.77 | 44 | 12.05 | 7.26
2008-10-18 | 9:31 | 29 | 101.4 | 7.82 | 44 | 12.06 | 7.26
2008-10-18 | 9:46 | 29 | 101.6 | 7.87 | 44 | 12.07 | 7.26
2008-10-18 | 10:01 | 29 | 101.8 | 7.95 | 44 | 12.06 | 7.28
2008-10-18 | 10:16 | 29 | 101.5 | 8.04 | 44 | 12 | 7.29
2008-10-18 | 10:31 | 29 | 101.4 | 8.13 | 44 | 11.97 | 7.29
2008-10-18 | 10:46 | 28 | 101.9 | 8.23 | 44 | 12 | 7.3
2008-10-18 | 11:01 | 28 | 101.7 | 8.31 | 44 | 11.95 | 7.31
2008-10-18 | 11:16 | 28 | 101.6 | 8.4 | 44 | 11.91 | 7.31
2008-10-18 | 11:31 | 28 | 102.1 | 8.5 | 44 | 11.95 | 7.32
2008-10-18 | 11:46 | 28 | 102.4 | 8.63 | 44 | 11.94 | 7.32
2008-10-18 | 12:01 | 28 | 102.1 | 8.75 | 44 | 11.87 | 7.33
2008-10-18 | 12:16 | 28 | 101.7 | 8.88 | 44 | 11.79 | 7.34
2008-10-18 | 12:31 | 28 | 102 | 9.02 | 44 | 11.79 | 7.33
2008-10-18 | 12:46 | 28 | 102.2 | 9.18 | 44 | 11.76 | 7.33
2008-10-18 | 13:00 | 28 | 101.7 | 9.32 | 43 | 11.67 | 7.34
2008-10-18 | 13:15 | 28 | 102 | 9.44 | 43 | 11.66 | 7.34
2008-10-18 | 13:30 | 28 | 102.1 | 9.55 | 43 | 11.65 | 7.35
2008-10-18 | 13:45 | 28 | 102.1 | 9.63 | 43 | 11.63 | 7.35
2008-10-18 | 14:00 | 28 | 101.7 | 9.68 | 43 | 11.56 | 7.36
2008-10-18 | 14:15 | 28 | 101.5 | 9.76 | 43 | 11.52 | 7.35
2008-10-18 | 14:30 | 28 | 101.7 | 9.83 | 43 | 11.53 | 7.35
2008-10-18 | 14:45 | 28 | 101.2 | 9.94 | 43 | 11.44 | 7.35
2008-10-18 | 15:00 | 28 | 101.1 | 10 | 43 | 11.41 | 7.35
2008-10-18 | 15:15 | 28 | 101.4 | 10.04 | 43 | 11.43 | 7.36
2008-10-18 | 15:30 | 28 | 101.4 | 10.05 | 43 | 11.43 | 7.36
2008-10-18 | 15:45 | 28 | 101.3 | 10.08 | 43 | 11.42 | 7.36
2008-10-18 | 16:00 | 28 | 101.3 | 10.11 | 43 | 11.4 | 7.35
2008-10-18 | 16:15 | 28 | 101.2 | 10.15 | 43 | 11.38 | 7.35
2008-10-18 | 16:30 | 28 | 100.6 | 10.2 | 43 | 11.3 | 7.35
2008-10-18 | 16:45 | 28 | 100.7 | 10.25 | 43 | 11.31 | 7.34
2008-10-18 | 17:00 | 28 | 100.2 | 10.28 | 43 | 11.23 | 7.34
2008-10-18 | 17:15 | 28 | 99.8 | 10.27 | 43 | 11.2 | 7.34
2008-10-18 | 17:30 | 28 | 99.6 | 10.25 | 43 | 11.17 | 7.31
2008-10-18 | 17:45 | 28 | 99.8 | 10.23 | 43 | 11.2 | 7.32
2008-10-18 | 18:00 | 28 | 99.8 | 10.21 | 43 | 11.21 | 7.31
2008-10-18 | 18:15 | 28 | 99.8 | 10.2 | 44 | 11.21 | 7.3
2008-10-18 | 18:30 | 28 | 99.1 | 10.18 | 44 | 11.14 | 7.29
2008-10-18 | 18:45 | 28 | 99.3 | 10.15 | 44 | 11.17 | 7.28
2008-10-18 | 19:00 | 28 | 99.4 | 10.1 | 44 | 11.19 | 7.27
2008-10-18 | 19:15 | 28 | 99.4 | 10.03 | 44 | 11.22 | 7.26
2008-10-18 | 19:30 | 28 | 99.4 | 9.95 | 44 | 11.23 | 7.26
2008-10-18 | 19:45 | 28 | 99.4 | 9.86 | 44 | 11.26 | 7.25
2008-10-18 | 20:00 | 29 | 99.4 | 9.77 | 44 | 11.28 | 7.24
2008-10-18 | 20:15 | 29 | 99 | 9.7 | 44 | 11.26 | 7.24
2008-10-18 | 20:30 | 29 | 98.8 | 9.62 | 44 | 11.25 | 7.23
2008-10-18 | 20:45 | 29 | 99.2 | 9.56 | 44 | 11.31 | 7.25
2008-10-18 | 21:00 | 29 | 98.9 | 9.49 | 44 | 11.3 | 7.25
2008-10-18 | 21:15 | 29 | 98.7 | 9.42 | 44 | 11.3 | 7.24
2008-10-18 | 21:30 | 29 | 99.2 | 9.36 | 44 | 11.37 | 7.24
2008-10-18 | 21:45 | 29 | 98.9 | 9.29 | 44 | 11.35 | 7.24
2008-10-18 | 22:00 | 29 | 99.2 | 9.23 | 44 | 11.41 | 7.24
2008-10-18 | 22:15 | 29 | 99 | 9.16 | 44 | 11.4 | 7.24
2008-10-18 | 22:30 | 29 | 99.3 | 9.09 | 44 | 11.45 | 7.25
2008-10-18 | 22:45 | 29 | 99.5 | 9.02 | 44 | 11.5 | 7.24
2008-10-18 | 23:00 | 29 | 99.6 | 8.95 | 44 | 11.53 | 7.24
2008-10-18 | 23:15 | 29 | 99.7 | 8.88 | 44 | 11.55 | 7.24
2008-10-18 | 23:30 | 29 | 99.3 | 8.8 | 44 | 11.53 | 7.24
2008-10-18 | 23:45 | 29 | 99.6 | 8.72 | 44 | 11.58 | 7.24
2008-10-19 | 0:00 | 29 | 99.2 | 8.65 | 44 | 11.57 | 7.24
2008-10-19 | 0:15 | 29 | 99.6 | 8.58 | 44 | 11.63 | 7.23
2008-10-19 | 0:30 | 29 | 99.3 | 8.5 | 44 | 11.62 | 7.23
2008-10-19 | 0:45 | 29 | 99.2 | 8.42 | 44 | 11.62 | 7.24
2008-10-19 | 1:01 | 29 | 99.6 | 8.35 | 44 | 11.7 | 7.24
2008-10-19 | 1:16 | 29 | 99.3 | 8.26 | 44 | 11.68 | 7.23
2008-10-19 | 1:31 | 29 | 99.7 | 8.19 | 44 | 11.75 | 7.23
2008-10-19 | 1:46 | 29 | 99.4 | 8.11 | 44 | 11.74 | 7.23
2008-10-19 | 2:01 | 29 | 99.8 | 8.04 | 44 | 11.81 | 7.23
2008-10-19 | 2:16 | 29 | 100 | 7.96 | 44 | 11.85 | 7.23
2008-10-19 | 2:31 | 29 | 100.1 | 7.87 | 44 | 11.89 | 7.22
2008-10-19 | 2:46 | 29 | 99.7 | 7.79 | 44 | 11.87 | 7.22
2008-10-19 | 3:01 | 29 | 99.5 | 7.71 | 44 | 11.87 | 7.22
2008-10-19 | 3:16 | 29 | 99.9 | 7.61 | 44 | 11.95 | 7.23
2008-10-19 | 3:31 | 29 | 99.7 | 7.52 | 44 | 11.94 | 7.22
2008-10-19 | 3:46 | 29 | 99.6 | 7.44 | 44 | 11.96 | 7.23
2008-10-19 | 4:01 | 29 | 100 | 7.35 | 44 | 12.04 | 7.22
2008-10-19 | 4:16 | 29 | 100.2 | 7.25 | 44 | 12.09 | 7.22
2008-10-19 | 4:31 | 29 | 100.3 | 7.16 | 44 | 12.13 | 7.21
2008-10-19 | 4:46 | 29 | 100.4 | 7.08 | 44 | 12.16 | 7.21
2008-10-19 | 5:01 | 29 | 100 | 6.99 | 44 | 12.15 | 7.21
2008-10-19 | 5:16 | 29 | 100.3 | 6.9 | 44 | 12.21 | 7.21
2008-10-19 | 5:31 | 29 | 100.5 | 6.81 | 44 | 12.26 | 7.2
2008-10-19 | 5:46 | 29 | 100.6 | 6.73 | 44 | 12.3 | 7.2
2008-10-19 | 6:01 | 29 | 100.7 | 6.65 | 44 | 12.33 | 7.19
2008-10-19 | 6:16 | 29 | 100.3 | 6.57 | 44 | 12.31 | 7.19
2008-10-19 | 6:31 | 29 | 100.1 | 6.5 | 44 | 12.31 | 7.19
2008-10-19 | 6:46 | 29 | 100.1 | 6.42 | 44 | 12.33 | 7.2
2008-10-19 | 7:01 | 29 | 100.6 | 6.33 | 44 | 12.41 | 7.19
2008-10-19 | 7:16 | 29 | 100.3 | 6.25 | 44 | 12.4 | 7.2
2008-10-19 | 7:31 | 29 | 100.7 | 6.18 | 44 | 12.48 | 7.19
2008-10-19 | 7:46 | 29 | 100.9 | 6.1 | 44 | 12.53 | 7.18
2008-10-19 | 8:01 | 29 | 101.1 | 6.03 | 44 | 12.58 | 7.2
2008-10-19 | 8:16 | 29 | 100.9 | 5.96 | 44 | 12.57 | 7.19
2008-10-19 | 8:31 | 29 | 101.3 | 5.9 | 44 | 12.64 | 7.19
2008-10-19 | 8:46 | 29 | 101.6 | 5.87 | 44 | 12.69 | 7.19
2008-10-19 | 9:01 | 29 | 101.5 | 5.86 | 44 | 12.68 | 7.2
2008-10-19 | 9:16 | 29 | 102 | 5.85 | 44 | 12.75 | 7.2
2008-10-19 | 9:31 | 29 | 102.2 | 5.83 | 44 | 12.77 | 7.2
2008-10-19 | 9:46 | 29 | 102.3 | 5.82 | 44 | 12.79 | 7.2
2008-10-19 | 10:01 | 29 | 102 | 5.81 | 44 | 12.75 | 7.21
2008-10-19 | 10:16 | 29 | 102 | 5.81 | 44 | 12.76 | 7.21
2008-10-19 | 10:31 | 29 | 102.2 | 5.82 | 44 | 12.78 | 7.22
2008-10-19 | 10:46 | 29 | 102.8 | 5.85 | 44 | 12.85 | 7.21
2008-10-19 | 11:01 | 29 | 103.1 | 5.88 | 44 | 12.87 | 7.23
2008-10-19 | 11:16 | 28 | 103.2 | 5.93 | 44 | 12.87 | 7.23
2008-10-19 | 11:31 | 29 | 102.9 | 6.02 | 44 | 12.8 | 7.24
2008-10-19 | 11:46 | 28 | 102.7 | 6.15 | 44 | 12.73 | 7.25
2008-10-19 | 12:01 | 28 | 103.1 | 6.33 | 44 | 12.72 | 7.25
2008-10-19 | 12:16 | 28 | 102.7 | 6.54 | 44 | 12.61 | 7.26
2008-10-19 | 12:31 | 28 | 103 | 6.8 | 44 | 12.56 | 7.26
2008-10-19 | 12:46 | 28 | 102.6 | 7.04 | 44 | 12.44 | 7.28
2008-10-19 | 13:00 | 28 | 102.4 | 7.26 | 44 | 12.34 | 7.28
2008-10-19 | 13:15 | 28 | 102.3 | 7.47 | 44 | 12.27 | 7.3
2008-10-19 | 13:30 | 28 | 102.2 | 7.69 | 44 | 12.2 | 7.33
2008-10-19 | 13:45 | 28 | 102.7 | 7.87 | 44 | 12.2 | 7.31
2008-10-19 | 14:00 | 28 | 102.8 | 8.01 | 44 | 12.17 | 7.31
2008-10-19 | 14:15 | 29 | 102.9 | 8.17 | 44 | 12.13 | 7.24