## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AA (Audited against AB, 556, NIST, & Titration)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-29 |  | 2008-08-29 |  | A
Time | Computer set to atomic time |  | 13:25 |  | 13:25 |  | A
Temp | VWR NIST traceable surface water |  | 12.3 |  | 12.5 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 10.17 |  | A
SpC | In Situ |  | 40 |  | 40 |  | A
Ph | In Situ |  | 6.8 |  | 6.9 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2008-09-02 13:18 | 12.32 | 7.58 | 11 | 103 | 63 | 44
Sonde | 2008-09-02 13:19 | 12.34 | 7.11 | 11.12 | 103.9 | 61 | 39
Grade |  | A | B | A |  | A
Audit 2 | 2008-09-09 14:45 | 16.37 | 7.84 | 9.69 | 98.9 | 61 | 43
Sonde | 2008-09-09 14:45 | 16.36 | 7.73 | 9.96 | 101.7 | 59 | 39
Grade |  | A | A | A |  | A
Audit 3 | 2008-09-12 13:15 | 12.34 | 7.14 | 11.02 | 103 | 64 | 45
Sonde
Grade |  | Fail | Fail | 0 |  | Fail
Post-sampling Field Equipment Check
Field equipment ID | Sonde AA (Audited against, 556 &NIST)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-12 |  | 2008-09-12 |  | A
Time | Computer set to atomic time |  | 13:15 |  | 13:15 |  | A
Temp | VWR NIST traceable surface water |  | 12.2 |  | 12.3 |  | A
DO | In Situ |  | 11.02 |  | 11.3 |  | A
SpC | In Situ |  | 64 |  | 62 |  | A
Ph | In Situ |  | 7.14 |  | 7.66 |  | Fail
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-29 |  | 2008-08-29 |  | A
Time | Computer set to atomic time |  | 13:34 |  | 13:34 |  | A
Temp | VWR NIST traceable surface water |  | 12.4 |  | 12.49 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 10.09 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 150 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.06 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-09-02 13:18 | 12.32 | 7.58 | 11 | 103 | 63 | 44
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-02 |  | 2008-09-02 |  | A
Time | Computer set to atomic time |  | 13:18 |  | 13:18 |  | A
Temp | VWR NIST traceable surface water |  | 12.1 |  | 12.32 |  | A
DO | Winkler titration of surface water |  | 10.7 |  | 11 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 150 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.04 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-09 |  | 2008-09-09 |  | A
Time | Computer set to atomic time |  | 9:47 |  | 9:47 |  | A
Temp | VWR NIST traceable surface water |  | 13.7 |  | 13.84 |  | A
DO | Winkler titration of surface water |  | 9.7 |  | 9.64 |  | A
SpC | 148 µS/mL |  | 153.1 |  | 150 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.09 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-09-09 14:45 | 16.37 | 7.84 | 9.69 | 98.9 | 61 | 43
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-09 |  | 2008-09-09 |  | A
Time | Computer set to atomic time |  | 19:16 |  | 19:16 |  | A
Temp | VWR NIST traceable surface water |  | 15.5 |  | 15.68 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 8.99 |  | A
SpC | 148 µS/mL |  | 153.1 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.09 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-09 |  | 2008-09-09 |  | A
Time | Computer set to atomic time |  | 19:16 |  | 19:16 |  | A
Temp | VWR NIST traceable surface water |  | 15.5 |  | 15.68 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 8.99 |  | A
SpC | 148 µS/mL |  | 153.1 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.09 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-09-12 13:15 | 12.34 | 7.14 | 11.02 | 103 | 64 | 45
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-12 |  | 2008-09-12 |  | A
Time | Computer set to atomic time |  | 14:45 |  | 14:45 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17.46 |  | A
DO | Winkler titration of surface water |  | 10.4 |  | 10.76 |  | B
SpC | In Situ |  | 67 |  | 69 |  | A
Ph | In Situ |  | 8.44 |  | 8.23 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | TDS | DO% | Temp | SpCond | DO Conc | pH
M/D/Y |  | mg/L | % | ˚C | mS/cm | mg/L
2008-09-02 | 13:15 | 40 | 111.1 | 12.25 | 61 | 11.91 | 7.17
2008-09-02 | 13:30 | 39 | 106.6 | 12.56 | 61 | 11.34 | 7.19
2008-09-02 | 13:45 | 40 | 106.2 | 12.89 | 61 | 11.22 | 7.4
2008-09-02 | 14:00 | 39 | 105.5 | 13.16 | 61 | 11.07 | 7.58
2008-09-02 | 14:15 | 39 | 104.9 | 13.47 | 61 | 10.94 | 7.68
2008-09-02 | 14:30 | 39 | 104.3 | 13.72 | 61 | 10.82 | 7.7
2008-09-02 | 14:45 | 39 | 103.8 | 14 | 60 | 10.7 | 7.73
2008-09-02 | 15:00 | 39 | 103.3 | 14.24 | 60 | 10.59 | 7.73
2008-09-02 | 15:15 | 39 | 102.7 | 14.38 | 60 | 10.5 | 7.76
2008-09-02 | 15:30 | 39 | 102.4 | 14.68 | 60 | 10.4 | 7.78
2008-09-02 | 15:45 | 39 | 101.7 | 14.95 | 60 | 10.27 | 7.79
2008-09-02 | 16:00 | 39 | 101.1 | 15.16 | 60 | 10.16 | 7.77
2008-09-02 | 16:15 | 39 | 100.7 | 15.33 | 60 | 10.08 | 7.75
2008-09-02 | 16:30 | 39 | 100.7 | 15.4 | 60 | 10.06 | 7.73
2008-09-02 | 16:45 | 39 | 100.1 | 15.42 | 60 | 10 | 7.69
2008-09-02 | 17:00 | 39 | 100 | 15.39 | 60 | 10 | 7.65
2008-09-02 | 17:15 | 39 | 99.8 | 15.33 | 60 | 9.99 | 7.61
2008-09-02 | 17:30 | 39 | 99.2 | 15.25 | 60 | 9.95 | 7.58
2008-09-02 | 17:45 | 39 | 99.1 | 15.15 | 60 | 9.96 | 7.55
2008-09-02 | 18:00 | 39 | 98.6 | 15.07 | 60 | 9.93 | 7.52
2008-09-02 | 18:15 | 39 | 98.3 | 15.01 | 60 | 9.91 | 7.51
2008-09-02 | 18:30 | 39 | 98.1 | 14.98 | 60 | 9.89 | 7.48
2008-09-02 | 18:45 | 39 | 98.2 | 14.96 | 60 | 9.91 | 7.49
2008-09-02 | 19:00 | 39 | 97.8 | 14.94 | 60 | 9.87 | 7.46
2008-09-02 | 19:15 | 39 | 97.5 | 14.92 | 60 | 9.84 | 7.46
2008-09-02 | 19:30 | 39 | 97.3 | 14.89 | 60 | 9.83 | 7.44
2008-09-02 | 19:45 | 39 | 97.4 | 14.83 | 60 | 9.86 | 7.42
2008-09-02 | 20:00 | 39 | 97.4 | 14.76 | 60 | 9.87 | 7.39
2008-09-02 | 20:15 | 39 | 97.3 | 14.67 | 60 | 9.88 | 7.39
2008-09-02 | 20:30 | 39 | 97 | 14.58 | 60 | 9.87 | 7.37
2008-09-02 | 20:45 | 39 | 96.9 | 14.45 | 60 | 9.88 | 7.37
2008-09-02 | 21:00 | 39 | 97.2 | 14.31 | 60 | 9.95 | 7.35
2008-09-02 | 21:15 | 39 | 97.4 | 14.17 | 60 | 10 | 7.34
2008-09-02 | 21:30 | 39 | 97.5 | 14.01 | 60 | 10.04 | 7.34
2008-09-02 | 21:45 | 39 | 97.2 | 13.84 | 60 | 10.05 | 7.33
2008-09-02 | 22:00 | 39 | 97.5 | 13.68 | 60 | 10.12 | 7.33
2008-09-02 | 22:15 | 39 | 97.3 | 13.52 | 60 | 10.13 | 7.33
2008-09-02 | 22:30 | 39 | 97.6 | 13.36 | 60 | 10.2 | 7.33
2008-09-02 | 22:45 | 39 | 97.8 | 13.19 | 60 | 10.26 | 7.33
2008-09-02 | 23:00 | 39 | 97.9 | 13.03 | 60 | 10.31 | 7.33
2008-09-02 | 23:15 | 39 | 97.6 | 12.87 | 60 | 10.32 | 7.33
2008-09-02 | 23:30 | 39 | 98 | 12.71 | 60 | 10.39 | 7.33
2008-09-02 | 23:45 | 39 | 98.2 | 12.55 | 60 | 10.45 | 7.34
2008-09-03 | 0:00 | 39 | 97.9 | 12.37 | 60 | 10.46 | 7.34
2008-09-03 | 0:15 | 39 | 98.2 | 12.21 | 60 | 10.53 | 7.33
2008-09-03 | 0:30 | 39 | 98.4 | 12.05 | 60 | 10.59 | 7.33
2008-09-03 | 0:45 | 39 | 98.5 | 11.88 | 60 | 10.64 | 7.33
2008-09-03 | 1:01 | 39 | 98.2 | 11.72 | 60 | 10.65 | 7.34
2008-09-03 | 1:16 | 39 | 98.2 | 11.57 | 60 | 10.68 | 7.32
2008-09-03 | 1:31 | 39 | 98.6 | 11.42 | 60 | 10.76 | 7.32
2008-09-03 | 1:46 | 39 | 98.4 | 11.29 | 60 | 10.78 | 7.33
2008-09-03 | 2:01 | 39 | 98.7 | 11.16 | 60 | 10.85 | 7.32
2008-09-03 | 2:16 | 39 | 98.5 | 11.03 | 60 | 10.85 | 7.33
2008-09-03 | 2:31 | 39 | 98.3 | 10.91 | 60 | 10.87 | 7.33
2008-09-03 | 2:46 | 39 | 98.8 | 10.79 | 60 | 10.95 | 7.33
2008-09-03 | 3:01 | 39 | 99 | 10.68 | 60 | 11 | 7.32
2008-09-03 | 3:16 | 39 | 98.8 | 10.56 | 60 | 11.01 | 7.33
2008-09-03 | 3:31 | 39 | 99.1 | 10.45 | 60 | 11.07 | 7.32
2008-09-03 | 3:46 | 39 | 99.3 | 10.36 | 60 | 11.11 | 7.32
2008-09-03 | 4:01 | 39 | 99 | 10.26 | 60 | 11.1 | 7.32
2008-09-03 | 4:16 | 39 | 99.3 | 10.16 | 60 | 11.16 | 7.33
2008-09-03 | 4:31 | 39 | 99.5 | 10.07 | 60 | 11.21 | 7.32
2008-09-03 | 4:46 | 39 | 99.5 | 9.97 | 61 | 11.24 | 7.32
2008-09-03 | 5:01 | 39 | 99.7 | 9.88 | 61 | 11.28 | 7.32
2008-09-03 | 5:16 | 39 | 99.2 | 9.77 | 61 | 11.26 | 7.32
2008-09-03 | 5:31 | 39 | 99.1 | 9.67 | 61 | 11.28 | 7.33
2008-09-03 | 5:46 | 40 | 99.1 | 9.58 | 61 | 11.29 | 7.32
2008-09-03 | 6:01 | 40 | 99.5 | 9.47 | 61 | 11.37 | 7.31
2008-09-03 | 6:16 | 40 | 99.8 | 9.36 | 61 | 11.43 | 7.32
2008-09-03 | 6:31 | 40 | 99.9 | 9.26 | 61 | 11.48 | 7.31
2008-09-03 | 6:46 | 40 | 99.7 | 9.16 | 61 | 11.48 | 7.32
2008-09-03 | 7:01 | 40 | 99.7 | 9.07 | 61 | 11.5 | 7.33
2008-09-03 | 7:16 | 40 | 100.2 | 8.97 | 61 | 11.59 | 7.33
2008-09-03 | 7:31 | 40 | 100.2 | 8.88 | 61 | 11.62 | 7.34
2008-09-03 | 7:46 | 40 | 100.4 | 8.8 | 61 | 11.66 | 7.34
2008-09-03 | 8:01 | 40 | 100.5 | 8.73 | 61 | 11.69 | 7.35
2008-09-03 | 8:16 | 40 | 101 | 8.67 | 62 | 11.77 | 7.35
2008-09-03 | 8:31 | 40 | 100.9 | 8.64 | 62 | 11.77 | 7.37
2008-09-03 | 8:46 | 40 | 101.6 | 8.63 | 61 | 11.84 | 7.37
2008-09-03 | 9:01 | 40 | 101.4 | 8.65 | 61 | 11.82 | 7.39
2008-09-03 | 9:16 | 40 | 101.9 | 8.67 | 61 | 11.87 | 7.39
2008-09-03 | 9:31 | 40 | 101.8 | 8.72 | 61 | 11.84 | 7.42
2008-09-03 | 9:46 | 40 | 102 | 8.8 | 61 | 11.84 | 7.43
2008-09-03 | 10:01 | 40 | 102.5 | 8.89 | 61 | 11.88 | 7.44
2008-09-03 | 10:16 | 40 | 102.9 | 8.96 | 61 | 11.9 | 7.44
2008-09-03 | 10:31 | 40 | 102.7 | 9.06 | 61 | 11.85 | 7.47
2008-09-03 | 10:46 | 40 | 102.9 | 9.15 | 61 | 11.85 | 7.46
2008-09-03 | 11:01 | 40 | 103 | 9.27 | 61 | 11.83 | 7.48
2008-09-03 | 11:16 | 40 | 103.2 | 9.43 | 61 | 11.8 | 7.49
2008-09-03 | 11:31 | 40 | 103.7 | 9.6 | 61 | 11.82 | 7.52
2008-09-03 | 11:46 | 40 | 103.9 | 9.86 | 61 | 11.77 | 7.54
2008-09-03 | 12:01 | 40 | 103.6 | 10.2 | 61 | 11.64 | 7.58
2008-09-03 | 12:16 | 40 | 103.8 | 10.6 | 61 | 11.56 | 7.59
2008-09-03 | 12:31 | 40 | 103.3 | 11.06 | 61 | 11.38 | 7.6
2008-09-03 | 12:46 | 40 | 103.5 | 11.53 | 61 | 11.28 | 7.63
2008-09-03 | 13:00 | 40 | 103.1 | 12.04 | 61 | 11.1 | 7.64
2008-09-03 | 13:15 | 40 | 102.8 | 12.57 | 61 | 10.94 | 7.65
2008-09-03 | 13:30 | 40 | 102.9 | 13.09 | 61 | 10.82 | 7.66
2008-09-03 | 13:45 | 40 | 102.5 | 13.62 | 61 | 10.65 | 7.68
2008-09-03 | 14:00 | 40 | 102.2 | 14.14 | 61 | 10.5 | 7.69
2008-09-03 | 14:15 | 40 | 101.9 | 14.67 | 61 | 10.35 | 7.72
2008-09-03 | 14:30 | 40 | 102 | 15.18 | 61 | 10.25 | 7.72
2008-09-03 | 14:45 | 39 | 102 | 15.59 | 61 | 10.15 | 7.74
2008-09-03 | 15:00 | 39 | 101.3 | 15.93 | 61 | 10.02 | 7.73
2008-09-03 | 15:15 | 39 | 101.4 | 16.22 | 61 | 9.96 | 7.74
2008-09-03 | 15:30 | 39 | 100.7 | 16.45 | 61 | 9.85 | 7.74
2008-09-03 | 15:45 | 39 | 100.8 | 16.7 | 60 | 9.8 | 7.71
2008-09-03 | 16:00 | 39 | 100.2 | 16.86 | 60 | 9.71 | 7.71
2008-09-03 | 16:15 | 39 | 100.2 | 16.91 | 60 | 9.7 | 7.7
2008-09-03 | 16:30 | 39 | 100 | 16.88 | 60 | 9.69 | 7.64
2008-09-03 | 16:45 | 39 | 99.8 | 16.79 | 60 | 9.69 | 7.62
2008-09-03 | 17:00 | 39 | 99.5 | 16.69 | 60 | 9.68 | 7.59
2008-09-03 | 17:15 | 39 | 98.9 | 16.55 | 60 | 9.65 | 7.55
2008-09-03 | 17:30 | 39 | 98.9 | 16.42 | 60 | 9.67 | 7.53
2008-09-03 | 17:45 | 39 | 98.4 | 16.29 | 60 | 9.65 | 7.51
2008-09-03 | 18:00 | 39 | 98 | 16.17 | 60 | 9.64 | 7.48
2008-09-03 | 18:15 | 39 | 98.1 | 16.06 | 60 | 9.67 | 7.47
2008-09-03 | 18:30 | 39 | 97.7 | 15.98 | 60 | 9.65 | 7.46
2008-09-03 | 18:45 | 39 | 97.8 | 15.9 | 60 | 9.67 | 7.45
2008-09-03 | 19:00 | 39 | 97.5 | 15.86 | 60 | 9.65 | 7.44
2008-09-03 | 19:15 | 39 | 97.2 | 15.84 | 60 | 9.63 | 7.44
2008-09-03 | 19:30 | 39 | 97 | 15.81 | 60 | 9.61 | 7.43
2008-09-03 | 19:45 | 39 | 96.8 | 15.77 | 60 | 9.6 | 7.42
2008-09-03 | 20:00 | 39 | 96.7 | 15.7 | 60 | 9.61 | 7.4
2008-09-03 | 20:15 | 39 | 96.6 | 15.63 | 60 | 9.61 | 7.4
2008-09-03 | 20:30 | 39 | 96.6 | 15.52 | 60 | 9.63 | 7.38
2008-09-03 | 20:45 | 39 | 96.6 | 15.4 | 60 | 9.65 | 7.36
2008-09-03 | 21:00 | 39 | 97 | 15.28 | 60 | 9.72 | 7.35
2008-09-03 | 21:15 | 39 | 97.2 | 15.14 | 60 | 9.77 | 7.36
2008-09-03 | 21:30 | 39 | 97.3 | 14.98 | 60 | 9.81 | 7.34
2008-09-03 | 21:45 | 39 | 97.4 | 14.81 | 60 | 9.86 | 7.34
2008-09-03 | 22:00 | 39 | 97.1 | 14.63 | 60 | 9.87 | 7.33
2008-09-03 | 22:15 | 39 | 97.1 | 14.47 | 60 | 9.9 | 7.34
2008-09-03 | 22:30 | 39 | 97.5 | 14.3 | 60 | 9.98 | 7.33
2008-09-03 | 22:45 | 39 | 97.3 | 14.12 | 59 | 10 | 7.33
2008-09-03 | 23:00 | 39 | 97.7 | 13.93 | 60 | 10.08 | 7.34
2008-09-03 | 23:15 | 39 | 97.9 | 13.75 | 59 | 10.14 | 7.34
2008-09-03 | 23:30 | 39 | 97.6 | 13.56 | 59 | 10.16 | 7.34
2008-09-03 | 23:45 | 39 | 97.6 | 13.37 | 59 | 10.2 | 7.34
2008-09-04 | 0:00 | 39 | 98 | 13.18 | 59 | 10.29 | 7.34
2008-09-04 | 0:15 | 39 | 98.2 | 13.01 | 59 | 10.34 | 7.34
2008-09-04 | 0:30 | 39 | 98 | 12.83 | 59 | 10.36 | 7.34
2008-09-04 | 0:45 | 38 | 98.3 | 12.67 | 59 | 10.44 | 7.34
2008-09-04 | 1:01 | 38 | 98.5 | 12.51 | 59 | 10.49 | 7.34
2008-09-04 | 1:16 | 38 | 98.2 | 12.37 | 59 | 10.5 | 7.34
2008-09-04 | 1:31 | 39 | 98.2 | 12.24 | 59 | 10.52 | 7.34
2008-09-04 | 1:46 | 39 | 98.6 | 12.11 | 59 | 10.6 | 7.34
2008-09-04 | 2:01 | 39 | 98.3 | 11.99 | 59 | 10.6 | 7.35
2008-09-04 | 2:16 | 39 | 98.3 | 11.87 | 59 | 10.62 | 7.35
2008-09-04 | 2:31 | 39 | 98.7 | 11.75 | 59 | 10.69 | 7.34
2008-09-04 | 2:46 | 39 | 98.9 | 11.63 | 59 | 10.75 | 7.34
2008-09-04 | 3:01 | 39 | 99 | 11.51 | 60 | 10.79 | 7.33
2008-09-04 | 3:16 | 39 | 99.1 | 11.41 | 60 | 10.83 | 7.34
2008-09-04 | 3:31 | 39 | 98.8 | 11.3 | 60 | 10.82 | 7.34
2008-09-04 | 3:46 | 39 | 98.7 | 11.19 | 60 | 10.84 | 7.33
2008-09-04 | 4:01 | 39 | 99.1 | 11.08 | 60 | 10.91 | 7.34
2008-09-04 | 4:16 | 39 | 98.9 | 10.97 | 60 | 10.92 | 7.33
2008-09-04 | 4:31 | 39 | 98.8 | 10.86 | 60 | 10.93 | 7.34
2008-09-04 | 4:46 | 39 | 99.3 | 10.76 | 60 | 11.01 | 7.34
2008-09-04 | 5:01 | 39 | 99.5 | 10.65 | 60 | 11.06 | 7.34
2008-09-04 | 5:16 | 39 | 99.6 | 10.54 | 60 | 11.09 | 7.32
2008-09-04 | 5:31 | 39 | 99.3 | 10.43 | 60 | 11.1 | 7.33
2008-09-04 | 5:46 | 39 | 99.6 | 10.33 | 60 | 11.16 | 7.33
2008-09-04 | 6:01 | 39 | 99.4 | 10.22 | 60 | 11.16 | 7.34
2008-09-04 | 6:16 | 39 | 99.3 | 10.11 | 60 | 11.18 | 7.34
2008-09-04 | 6:31 | 39 | 99.7 | 10.01 | 61 | 11.25 | 7.32
2008-09-04 | 6:46 | 39 | 99.6 | 9.93 | 61 | 11.26 | 7.33
2008-09-04 | 7:01 | 39 | 99.7 | 9.84 | 61 | 11.29 | 7.34
2008-09-04 | 7:16 | 40 | 99.8 | 9.75 | 61 | 11.33 | 7.34
2008-09-04 | 7:31 | 40 | 100.3 | 9.67 | 61 | 11.41 | 7.36
2008-09-04 | 7:46 | 40 | 100.7 | 9.59 | 61 | 11.48 | 7.35
2008-09-04 | 8:01 | 40 | 100.6 | 9.52 | 61 | 11.48 | 7.35
2008-09-04 | 8:16 | 40 | 101.1 | 9.46 | 61 | 11.56 | 7.36
2008-09-04 | 8:31 | 40 | 101.3 | 9.42 | 61 | 11.6 | 7.37
2008-09-04 | 8:46 | 40 | 101.2 | 9.39 | 61 | 11.59 | 7.38
2008-09-04 | 9:01 | 40 | 101.3 | 9.41 | 61 | 11.59 | 7.39
2008-09-04 | 9:16 | 40 | 101.9 | 9.44 | 61 | 11.65 | 7.4
2008-09-04 | 9:31 | 40 | 101.8 | 9.5 | 62 | 11.62 | 7.43
2008-09-04 | 9:46 | 40 | 102.3 | 9.58 | 61 | 11.66 | 7.44
2008-09-04 | 10:01 | 40 | 102.6 | 9.67 | 61 | 11.68 | 7.44
2008-09-04 | 10:16 | 40 | 102.4 | 9.75 | 61 | 11.63 | 7.46
2008-09-04 | 10:31 | 40 | 102.9 | 9.82 | 61 | 11.66 | 7.47
2008-09-04 | 10:46 | 40 | 102.8 | 9.91 | 61 | 11.63 | 7.48
2008-09-04 | 11:01 | 40 | 103.4 | 10.03 | 61 | 11.66 | 7.49
2008-09-04 | 11:16 | 40 | 103.4 | 10.16 | 61 | 11.62 | 7.51
2008-09-04 | 11:31 | 40 | 103.4 | 10.32 | 61 | 11.58 | 7.52
2008-09-04 | 11:46 | 40 | 103.7 | 10.56 | 61 | 11.56 | 7.54
2008-09-04 | 12:01 | 40 | 103.9 | 10.86 | 61 | 11.49 | 7.57
2008-09-04 | 12:16 | 40 | 103.9 | 11.26 | 61 | 11.39 | 7.59
2008-09-04 | 12:31 | 40 | 103.9 | 11.68 | 61 | 11.28 | 7.61
2008-09-04 | 12:46 | 40 | 103.3 | 12.13 | 61 | 11.09 | 7.62
2008-09-04 | 13:00 | 40 | 103.4 | 12.6 | 61 | 11 | 7.63
2008-09-04 | 13:15 | 40 | 103.4 | 13.08 | 61 | 10.88 | 7.65
2008-09-04 | 13:30 | 40 | 102.8 | 13.58 | 61 | 10.69 | 7.66
2008-09-04 | 13:45 | 40 | 102.9 | 14.08 | 61 | 10.58 | 7.67
2008-09-04 | 14:00 | 40 | 102.3 | 14.57 | 61 | 10.42 | 7.69
2008-09-04 | 14:15 | 40 | 102.4 | 15.05 | 61 | 10.31 | 7.71
2008-09-04 | 14:30 | 40 | 101.9 | 15.52 | 61 | 10.15 | 7.73
2008-09-04 | 14:45 | 40 | 102 | 15.93 | 61 | 10.08 | 7.73
2008-09-04 | 15:00 | 39 | 101.8 | 16.26 | 61 | 10 | 7.73
2008-09-04 | 15:15 | 39 | 101.5 | 16.52 | 61 | 9.91 | 7.71
2008-09-04 | 15:30 | 39 | 100.8 | 16.76 | 61 | 9.79 | 7.72
2008-09-04 | 15:45 | 39 | 100.5 | 16.94 | 60 | 9.72 | 7.73
2008-09-04 | 16:00 | 39 | 100.2 | 17.09 | 60 | 9.66 | 7.7
2008-09-04 | 16:15 | 39 | 100.3 | 17.14 | 60 | 9.67 | 7.66
2008-09-04 | 16:30 | 39 | 100.2 | 17.09 | 60 | 9.66 | 7.64
2008-09-04 | 16:45 | 39 | 99.9 | 17.01 | 60 | 9.65 | 7.59
2008-09-04 | 17:00 | 39 | 99.6 | 16.89 | 60 | 9.65 | 7.58
2008-09-04 | 17:15 | 39 | 99.1 | 16.74 | 60 | 9.63 | 7.53
2008-09-04 | 17:30 | 39 | 98.7 | 16.61 | 60 | 9.62 | 7.5
2008-09-04 | 17:45 | 39 | 98.8 | 16.47 | 60 | 9.65 | 7.48
2008-09-04 | 18:00 | 39 | 98.7 | 16.35 | 60 | 9.67 | 7.47
2008-09-04 | 18:15 | 39 | 98.6 | 16.23 | 60 | 9.68 | 7.45
2008-09-04 | 18:30 | 39 | 98.4 | 16.15 | 60 | 9.68 | 7.45
2008-09-04 | 18:45 | 39 | 98.4 | 16.08 | 59 | 9.69 | 7.44
2008-09-04 | 19:00 | 39 | 97.9 | 16.04 | 60 | 9.66 | 7.43
2008-09-04 | 19:15 | 39 | 97.9 | 16 | 59 | 9.66 | 7.43
2008-09-04 | 19:30 | 39 | 97.8 | 15.97 | 59 | 9.66 | 7.41
2008-09-04 | 19:45 | 39 | 97.4 | 15.92 | 59 | 9.63 | 7.42
2008-09-04 | 20:00 | 39 | 97.1 | 15.86 | 59 | 9.61 | 7.39
2008-09-04 | 20:15 | 39 | 97 | 15.77 | 59 | 9.62 | 7.38
2008-09-04 | 20:30 | 39 | 96.9 | 15.66 | 59 | 9.63 | 7.37
2008-09-04 | 20:45 | 38 | 96.9 | 15.54 | 59 | 9.66 | 7.35
2008-09-04 | 21:00 | 38 | 97.3 | 15.4 | 59 | 9.73 | 7.35
2008-09-04 | 21:15 | 39 | 97.5 | 15.24 | 59 | 9.78 | 7.34
2008-09-04 | 21:30 | 39 | 97.3 | 15.06 | 59 | 9.8 | 7.33
2008-09-04 | 21:45 | 39 | 97.2 | 14.86 | 59 | 9.83 | 7.34
2008-09-04 | 22:00 | 39 | 97.3 | 14.68 | 59 | 9.88 | 7.34
2008-09-04 | 22:15 | 38 | 97.8 | 14.5 | 59 | 9.97 | 7.33
2008-09-04 | 22:30 | 38 | 98 | 14.31 | 59 | 10.03 | 7.34
2008-09-04 | 22:45 | 38 | 98.2 | 14.12 | 59 | 10.09 | 7.33
2008-09-04 | 23:00 | 38 | 97.9 | 13.94 | 59 | 10.1 | 7.34
2008-09-04 | 23:15 | 38 | 98.3 | 13.75 | 59 | 10.18 | 7.34
2008-09-04 | 23:30 | 38 | 98.1 | 13.57 | 59 | 10.2 | 7.34
2008-09-04 | 23:45 | 38 | 98.4 | 13.4 | 59 | 10.28 | 7.33
2008-09-05 | 0:00 | 38 | 98.6 | 13.23 | 59 | 10.33 | 7.34
2008-09-05 | 0:15 | 38 | 98.7 | 13.07 | 59 | 10.38 | 7.34
2008-09-05 | 0:30 | 38 | 98.8 | 12.92 | 59 | 10.43 | 7.34
2008-09-05 | 0:45 | 38 | 98.9 | 12.79 | 59 | 10.47 | 7.34
2008-09-05 | 1:01 | 38 | 99 | 12.65 | 59 | 10.51 | 7.34
2008-09-05 | 1:16 | 38 | 98.7 | 12.52 | 59 | 10.51 | 7.33
2008-09-05 | 1:31 | 38 | 98.6 | 12.4 | 59 | 10.53 | 7.34
2008-09-05 | 1:46 | 38 | 99 | 12.28 | 59 | 10.6 | 7.34
2008-09-05 | 2:01 | 38 | 99.3 | 12.16 | 59 | 10.66 | 7.33
2008-09-05 | 2:16 | 38 | 98.9 | 12.05 | 59 | 10.65 | 7.33
2008-09-05 | 2:31 | 38 | 98.8 | 11.94 | 59 | 10.67 | 7.34
2008-09-05 | 2:46 | 39 | 99.1 | 11.82 | 59 | 10.72 | 7.34
2008-09-05 | 3:01 | 39 | 98.9 | 11.71 | 59 | 10.73 | 7.33
2008-09-05 | 3:16 | 39 | 99.3 | 11.59 | 59 | 10.8 | 7.34
2008-09-05 | 3:31 | 39 | 99.5 | 11.48 | 59 | 10.85 | 7.33
2008-09-05 | 3:46 | 39 | 99.5 | 11.37 | 59 | 10.88 | 7.34
2008-09-05 | 4:01 | 39 | 99.3 | 11.26 | 60 | 10.88 | 7.34
2008-09-05 | 4:16 | 39 | 99.6 | 11.16 | 60 | 10.94 | 7.33
2008-09-05 | 4:31 | 39 | 99.7 | 11.06 | 60 | 10.98 | 7.33
2008-09-05 | 4:46 | 39 | 99.8 | 10.96 | 60 | 11.02 | 7.33
2008-09-05 | 5:01 | 39 | 99.9 | 10.85 | 60 | 11.05 | 7.33
2008-09-05 | 5:16 | 39 | 99.6 | 10.75 | 60 | 11.05 | 7.34
2008-09-05 | 5:31 | 39 | 99.5 | 10.64 | 60 | 11.06 | 7.33
2008-09-05 | 5:46 | 39 | 99.8 | 10.53 | 60 | 11.13 | 7.33
2008-09-05 | 6:01 | 39 | 100.1 | 10.41 | 60 | 11.19 | 7.33
2008-09-05 | 6:16 | 39 | 100.2 | 10.3 | 60 | 11.23 | 7.33
2008-09-05 | 6:31 | 39 | 99.9 | 10.2 | 60 | 11.23 | 7.33
2008-09-05 | 6:46 | 39 | 99.9 | 10.08 | 60 | 11.25 | 7.33
2008-09-05 | 7:01 | 39 | 100.4 | 9.97 | 61 | 11.34 | 7.33
2008-09-05 | 7:16 | 39 | 100.7 | 9.86 | 61 | 11.41 | 7.33
2008-09-05 | 7:31 | 40 | 101 | 9.75 | 61 | 11.46 | 7.33
2008-09-05 | 7:46 | 40 | 101.3 | 9.64 | 61 | 11.53 | 7.33
2008-09-05 | 8:01 | 40 | 101.1 | 9.56 | 61 | 11.53 | 7.35
2008-09-05 | 8:16 | 40 | 101.1 | 9.5 | 61 | 11.54 | 7.36
2008-09-05 | 8:31 | 40 | 101.6 | 9.46 | 61 | 11.62 | 7.37
2008-09-05 | 8:46 | 40 | 102 | 9.43 | 61 | 11.67 | 7.37
2008-09-05 | 9:01 | 40 | 102.1 | 9.44 | 61 | 11.68 | 7.39
2008-09-05 | 9:16 | 40 | 102.5 | 9.47 | 61 | 11.71 | 7.4
2008-09-05 | 9:31 | 40 | 102.2 | 9.51 | 61 | 11.67 | 7.41
2008-09-05 | 9:46 | 40 | 102.7 | 9.6 | 61 | 11.7 | 7.43
2008-09-05 | 10:01 | 40 | 103.1 | 9.69 | 61 | 11.73 | 7.44
2008-09-05 | 10:16 | 40 | 102.9 | 9.75 | 61 | 11.68 | 7.45
2008-09-05 | 10:31 | 40 | 102.9 | 9.85 | 61 | 11.65 | 7.46
2008-09-05 | 10:46 | 40 | 103 | 9.94 | 61 | 11.64 | 7.47
2008-09-05 | 11:01 | 40 | 103.7 | 10.08 | 61 | 11.68 | 7.49
2008-09-05 | 11:16 | 40 | 103.6 | 10.22 | 61 | 11.63 | 7.5
2008-09-05 | 11:31 | 40 | 104.1 | 10.4 | 61 | 11.64 | 7.52
2008-09-05 | 11:46 | 40 | 104.2 | 10.64 | 61 | 11.59 | 7.53
2008-09-05 | 12:01 | 40 | 103.7 | 10.99 | 61 | 11.44 | 7.56
2008-09-05 | 12:16 | 40 | 103.5 | 11.4 | 61 | 11.31 | 7.58
2008-09-05 | 12:31 | 40 | 103.4 | 11.85 | 61 | 11.18 | 7.6
2008-09-05 | 12:46 | 40 | 103.5 | 12.34 | 61 | 11.07 | 7.62
2008-09-05 | 13:00 | 40 | 103.2 | 12.86 | 61 | 10.9 | 7.63
2008-09-05 | 13:15 | 39 | 102.8 | 13.38 | 61 | 10.74 | 7.65
2008-09-05 | 13:30 | 39 | 102.5 | 13.94 | 61 | 10.58 | 7.66
2008-09-05 | 13:45 | 39 | 102.7 | 14.47 | 61 | 10.48 | 7.69
2008-09-05 | 14:00 | 39 | 102.7 | 15.02 | 60 | 10.35 | 7.69
2008-09-05 | 14:15 | 39 | 102.1 | 15.52 | 61 | 10.18 | 7.72
2008-09-05 | 14:30 | 39 | 101.8 | 16.01 | 60 | 10.05 | 7.72
2008-09-05 | 14:45 | 39 | 101.8 | 16.38 | 60 | 9.97 | 7.73
2008-09-05 | 15:00 | 39 | 101.6 | 16.7 | 60 | 9.88 | 7.72
2008-09-05 | 15:15 | 39 | 101.1 | 17.01 | 60 | 9.77 | 7.72
2008-09-05 | 15:30 | 39 | 101.1 | 17.23 | 60 | 9.73 | 7.73
2008-09-05 | 15:45 | 39 | 100.6 | 17.42 | 60 | 9.63 | 7.72
2008-09-05 | 16:00 | 39 | 100.7 | 17.58 | 60 | 9.61 | 7.72
2008-09-05 | 16:15 | 39 | 100.4 | 17.64 | 60 | 9.58 | 7.69
2008-09-05 | 16:30 | 39 | 100.3 | 17.62 | 60 | 9.57 | 7.64
2008-09-05 | 16:45 | 39 | 100 | 17.56 | 59 | 9.56 | 7.61
2008-09-05 | 17:00 | 39 | 99.8 | 17.46 | 59 | 9.55 | 7.59
2008-09-05 | 17:15 | 39 | 99.5 | 17.34 | 59 | 9.55 | 7.55
2008-09-05 | 17:30 | 39 | 99.4 | 17.22 | 59 | 9.56 | 7.51
2008-09-05 | 17:45 | 38 | 98.8 | 17.09 | 59 | 9.53 | 7.51
2008-09-05 | 18:00 | 38 | 98.9 | 16.97 | 59 | 9.57 | 7.48
2008-09-05 | 18:15 | 38 | 98.8 | 16.87 | 59 | 9.58 | 7.47
2008-09-05 | 18:30 | 38 | 98.3 | 16.8 | 59 | 9.54 | 7.47
2008-09-05 | 18:45 | 38 | 98.4 | 16.73 | 59 | 9.56 | 7.46
2008-09-05 | 19:00 | 38 | 98 | 16.69 | 59 | 9.53 | 7.45
2008-09-05 | 19:15 | 38 | 97.6 | 16.66 | 59 | 9.5 | 7.44
2008-09-05 | 19:30 | 38 | 97.7 | 16.64 | 59 | 9.51 | 7.43
2008-09-05 | 19:45 | 38 | 97.6 | 16.56 | 59 | 9.52 | 7.4
2008-09-05 | 20:00 | 38 | 97.5 | 16.47 | 59 | 9.52 | 7.39
2008-09-05 | 20:15 | 38 | 97.4 | 16.35 | 59 | 9.54 | 7.38
2008-09-05 | 20:30 | 38 | 97.4 | 16.22 | 59 | 9.57 | 7.35
2008-09-05 | 20:45 | 38 | 97.1 | 16.09 | 59 | 9.56 | 7.35
2008-09-05 | 21:00 | 38 | 97 | 15.92 | 59 | 9.59 | 7.34
2008-09-05 | 21:15 | 38 | 97.3 | 15.74 | 59 | 9.66 | 7.33
2008-09-05 | 21:30 | 38 | 97.2 | 15.56 | 59 | 9.69 | 7.32
2008-09-05 | 21:45 | 38 | 97.6 | 15.37 | 59 | 9.76 | 7.33
2008-09-05 | 22:00 | 38 | 97.4 | 15.2 | 59 | 9.78 | 7.33
2008-09-05 | 22:15 | 38 | 97.4 | 15.04 | 59 | 9.82 | 7.33
2008-09-05 | 22:30 | 38 | 97.8 | 14.88 | 59 | 9.89 | 7.33
2008-09-05 | 22:45 | 38 | 98 | 14.71 | 59 | 9.94 | 7.33
2008-09-05 | 23:00 | 38 | 98.1 | 14.55 | 59 | 9.99 | 7.33
2008-09-05 | 23:15 | 38 | 97.9 | 14.39 | 59 | 10 | 7.33
2008-09-05 | 23:30 | 38 | 97.8 | 14.23 | 59 | 10.03 | 7.33
2008-09-05 | 23:45 | 38 | 97.9 | 14.08 | 59 | 10.07 | 7.33
2008-09-06 | 0:00 | 38 | 98.2 | 13.94 | 59 | 10.13 | 7.34
2008-09-06 | 0:15 | 38 | 98.4 | 13.81 | 59 | 10.19 | 7.34
2008-09-06 | 0:30 | 38 | 98.6 | 13.7 | 59 | 10.23 | 7.34
2008-09-06 | 0:45 | 38 | 98.7 | 13.58 | 59 | 10.26 | 7.34
2008-09-06 | 1:01 | 38 | 98.4 | 13.47 | 59 | 10.26 | 7.33
2008-09-06 | 1:16 | 38 | 98.3 | 13.37 | 59 | 10.27 | 7.34
2008-09-06 | 1:31 | 38 | 98.3 | 13.28 | 59 | 10.29 | 7.35
2008-09-06 | 1:46 | 38 | 98.7 | 13.19 | 59 | 10.35 | 7.35
2008-09-06 | 2:01 | 38 | 98.9 | 13.1 | 59 | 10.39 | 7.35
2008-09-06 | 2:16 | 38 | 99 | 13.01 | 59 | 10.43 | 7.34
2008-09-06 | 2:31 | 38 | 99.1 | 12.92 | 59 | 10.46 | 7.35
2008-09-06 | 2:46 | 38 | 99.1 | 12.82 | 59 | 10.48 | 7.34
2008-09-06 | 3:01 | 38 | 99.1 | 12.72 | 59 | 10.51 | 7.34
2008-09-06 | 3:16 | 38 | 98.8 | 12.63 | 59 | 10.5 | 7.34
2008-09-06 | 3:31 | 38 | 98.7 | 12.53 | 59 | 10.51 | 7.35
2008-09-06 | 3:46 | 39 | 98.7 | 12.43 | 59 | 10.53 | 7.34
2008-09-06 | 4:01 | 39 | 99.1 | 12.34 | 59 | 10.6 | 7.34
2008-09-06 | 4:16 | 39 | 99.3 | 12.25 | 59 | 10.64 | 7.34
2008-09-06 | 4:31 | 39 | 99 | 12.15 | 60 | 10.63 | 7.34
2008-09-06 | 4:46 | 39 | 99 | 12.05 | 60 | 10.66 | 7.33
2008-09-06 | 5:01 | 39 | 99.1 | 11.95 | 60 | 10.7 | 7.34
2008-09-06 | 5:16 | 39 | 99 | 11.85 | 60 | 10.7 | 7.34
2008-09-06 | 5:31 | 39 | 99 | 11.76 | 60 | 10.73 | 7.35
2008-09-06 | 5:46 | 39 | 99.5 | 11.65 | 60 | 10.8 | 7.34
2008-09-06 | 6:01 | 39 | 99.5 | 11.55 | 60 | 10.83 | 7.36
2008-09-06 | 6:16 | 39 | 99.7 | 11.44 | 60 | 10.88 | 7.33
2008-09-06 | 6:31 | 39 | 99.5 | 11.34 | 60 | 10.89 | 7.34
2008-09-06 | 6:46 | 39 | 99.5 | 11.23 | 60 | 10.91 | 7.34
2008-09-06 | 7:01 | 39 | 99.7 | 11.13 | 60 | 10.96 | 7.33
2008-09-06 | 7:16 | 39 | 100.2 | 11.03 | 60 | 11.04 | 7.33
2008-09-06 | 7:31 | 39 | 100.2 | 10.93 | 61 | 11.06 | 7.35
2008-09-06 | 7:46 | 39 | 100.2 | 10.85 | 61 | 11.08 | 7.35
2008-09-06 | 8:01 | 39 | 100.8 | 10.77 | 61 | 11.18 | 7.35
2008-09-06 | 8:16 | 39 | 100.7 | 10.7 | 61 | 11.19 | 7.36
2008-09-06 | 8:31 | 40 | 101.2 | 10.66 | 61 | 11.25 | 7.37
2008-09-06 | 8:46 | 40 | 101.1 | 10.64 | 61 | 11.25 | 7.37
2008-09-06 | 9:01 | 40 | 101.2 | 10.64 | 61 | 11.25 | 7.39
2008-09-06 | 9:16 | 40 | 101.4 | 10.65 | 61 | 11.27 | 7.41
2008-09-06 | 9:31 | 40 | 101.9 | 10.68 | 61 | 11.32 | 7.41
2008-09-06 | 9:46 | 40 | 102.3 | 10.76 | 61 | 11.34 | 7.43
2008-09-06 | 10:01 | 40 | 102.6 | 10.83 | 61 | 11.36 | 7.44
2008-09-06 | 10:16 | 40 | 102.9 | 10.9 | 61 | 11.37 | 7.45
2008-09-06 | 10:31 | 40 | 102.7 | 10.97 | 61 | 11.33 | 7.47
2008-09-06 | 10:46 | 40 | 103.2 | 11.05 | 61 | 11.37 | 7.48
2008-09-06 | 11:01 | 40 | 103.7 | 11.17 | 61 | 11.39 | 7.5
2008-09-06 | 11:16 | 40 | 103.8 | 11.28 | 61 | 11.37 | 7.51
2008-09-06 | 11:31 | 40 | 104 | 11.43 | 61 | 11.35 | 7.53
2008-09-06 | 11:46 | 40 | 103.6 | 11.66 | 61 | 11.25 | 7.55
2008-09-06 | 12:01 | 40 | 103.8 | 11.98 | 61 | 11.2 | 7.57
2008-09-06 | 12:16 | 40 | 103.8 | 12.38 | 61 | 11.09 | 7.59
2008-09-06 | 12:31 | 39 | 103.4 | 12.8 | 61 | 10.94 | 7.62
2008-09-06 | 12:46 | 39 | 103.3 | 13.26 | 61 | 10.82 | 7.64
2008-09-06 | 13:00 | 39 | 103.3 | 13.75 | 61 | 10.7 | 7.65
2008-09-06 | 13:15 | 39 | 102.7 | 14.26 | 60 | 10.52 | 7.67
2008-09-06 | 13:30 | 39 | 102.8 | 14.77 | 60 | 10.42 | 7.68
2008-09-06 | 13:45 | 39 | 102.3 | 15.29 | 60 | 10.25 | 7.69
2008-09-06 | 14:00 | 39 | 102.4 | 15.79 | 60 | 10.15 | 7.71
2008-09-06 | 14:15 | 39 | 102.2 | 16.3 | 60 | 10.03 | 7.71
2008-09-06 | 14:30 | 39 | 101.6 | 16.77 | 60 | 9.86 | 7.72
2008-09-06 | 14:45 | 39 | 101.1 | 17.17 | 60 | 9.74 | 7.73
2008-09-06 | 15:00 | 39 | 100.8 | 17.49 | 60 | 9.64 | 7.73
2008-09-06 | 15:15 | 39 | 100.5 | 17.76 | 60 | 9.56 | 7.71
2008-09-06 | 15:30 | 39 | 100.7 | 17.99 | 60 | 9.53 | 7.72
2008-09-06 | 15:45 | 39 | 100.1 | 18.18 | 60 | 9.44 | 7.71
2008-09-06 | 16:00 | 39 | 99.7 | 18.31 | 59 | 9.38 | 7.7
2008-09-06 | 16:15 | 39 | 99.4 | 18.38 | 59 | 9.34 | 7.66
2008-09-06 | 16:30 | 39 | 99.5 | 18.34 | 59 | 9.36 | 7.62
2008-09-06 | 16:45 | 38 | 99 | 18.26 | 59 | 9.32 | 7.58
2008-09-06 | 17:00 | 38 | 98.7 | 18.15 | 59 | 9.31 | 7.54
2008-09-06 | 17:15 | 38 | 98.7 | 17.99 | 59 | 9.35 | 7.51
2008-09-06 | 17:30 | 38 | 98.6 | 17.84 | 59 | 9.37 | 7.47
2008-09-06 | 17:45 | 38 | 98 | 17.7 | 59 | 9.34 | 7.46
2008-09-06 | 18:00 | 38 | 98.1 | 17.58 | 59 | 9.37 | 7.44
2008-09-06 | 18:15 | 38 | 97.7 | 17.5 | 59 | 9.34 | 7.43
2008-09-06 | 18:30 | 38 | 97.8 | 17.43 | 59 | 9.37 | 7.42
2008-09-06 | 18:45 | 38 | 97.5 | 17.38 | 59 | 9.35 | 7.42
2008-09-06 | 19:00 | 38 | 97.6 | 17.31 | 59 | 9.38 | 7.42
2008-09-06 | 19:15 | 38 | 97.2 | 17.28 | 59 | 9.34 | 7.41
2008-09-06 | 19:30 | 38 | 97 | 17.25 | 59 | 9.33 | 7.41
2008-09-06 | 19:45 | 38 | 96.8 | 17.21 | 59 | 9.32 | 7.4
2008-09-06 | 20:00 | 38 | 96.7 | 17.14 | 59 | 9.32 | 7.38
2008-09-06 | 20:15 | 38 | 97 | 17.06 | 59 | 9.36 | 7.36
2008-09-06 | 20:30 | 38 | 96.8 | 16.98 | 59 | 9.36 | 7.36
2008-09-06 | 20:45 | 38 | 97 | 16.87 | 59 | 9.4 | 7.36
2008-09-06 | 21:00 | 38 | 97.1 | 16.75 | 59 | 9.43 | 7.35
2008-09-06 | 21:15 | 38 | 96.9 | 16.6 | 59 | 9.45 | 7.33
2008-09-06 | 21:30 | 38 | 96.9 | 16.45 | 59 | 9.47 | 7.34
2008-09-06 | 21:45 | 38 | 96.9 | 16.29 | 59 | 9.51 | 7.33
2008-09-06 | 22:00 | 38 | 97 | 16.15 | 59 | 9.54 | 7.33
2008-09-06 | 22:15 | 38 | 96.9 | 15.99 | 59 | 9.57 | 7.33
2008-09-06 | 22:30 | 38 | 97.5 | 15.82 | 59 | 9.66 | 7.32
2008-09-06 | 22:45 | 38 | 97.6 | 15.66 | 59 | 9.7 | 7.33
2008-09-06 | 23:00 | 38 | 97.7 | 15.51 | 59 | 9.75 | 7.32
2008-09-06 | 23:15 | 38 | 97.6 | 15.37 | 59 | 9.76 | 7.32
2008-09-06 | 23:30 | 38 | 97.8 | 15.22 | 59 | 9.82 | 7.31
2008-09-06 | 23:45 | 38 | 97.6 | 15.08 | 59 | 9.82 | 7.32
2008-09-07 | 0:00 | 38 | 97.6 | 14.94 | 59 | 9.86 | 7.32
2008-09-07 | 0:15 | 38 | 97.7 | 14.81 | 59 | 9.89 | 7.32
2008-09-07 | 0:30 | 38 | 98.1 | 14.68 | 59 | 9.96 | 7.32
2008-09-07 | 0:45 | 38 | 98.4 | 14.54 | 59 | 10.02 | 7.33
2008-09-07 | 1:01 | 38 | 98 | 14.41 | 59 | 10.01 | 7.33
2008-09-07 | 1:16 | 38 | 98.4 | 14.27 | 59 | 10.08 | 7.34
2008-09-07 | 1:31 | 38 | 98.6 | 14.14 | 59 | 10.12 | 7.34
2008-09-07 | 1:46 | 38 | 98.3 | 14.01 | 59 | 10.13 | 7.34
2008-09-07 | 2:01 | 38 | 98.5 | 13.88 | 59 | 10.18 | 7.33
2008-09-07 | 2:16 | 38 | 98.8 | 13.76 | 59 | 10.23 | 7.34
2008-09-07 | 2:31 | 38 | 98.9 | 13.63 | 59 | 10.27 | 7.34
2008-09-07 | 2:46 | 38 | 99 | 13.51 | 59 | 10.31 | 7.33
2008-09-07 | 3:01 | 38 | 98.7 | 13.39 | 59 | 10.31 | 7.34
2008-09-07 | 3:16 | 38 | 98.9 | 13.27 | 59 | 10.36 | 7.34
2008-09-07 | 3:31 | 38 | 99.2 | 13.16 | 59 | 10.41 | 7.34
2008-09-07 | 3:46 | 39 | 99.2 | 13.06 | 59 | 10.44 | 7.33
2008-09-07 | 4:01 | 39 | 98.9 | 12.96 | 59 | 10.43 | 7.34
2008-09-07 | 4:16 | 39 | 98.8 | 12.87 | 60 | 10.44 | 7.34
2008-09-07 | 4:31 | 39 | 99.3 | 12.76 | 60 | 10.51 | 7.34
2008-09-07 | 4:46 | 39 | 98.9 | 12.66 | 60 | 10.5 | 7.34
2008-09-07 | 5:01 | 39 | 99.3 | 12.55 | 60 | 10.57 | 7.34
2008-09-07 | 5:16 | 39 | 99.5 | 12.45 | 60 | 10.62 | 7.34
2008-09-07 | 5:31 | 39 | 99.3 | 12.34 | 60 | 10.61 | 7.33
2008-09-07 | 5:46 | 39 | 99.2 | 12.24 | 60 | 10.63 | 7.33
2008-09-07 | 6:01 | 39 | 99.3 | 12.13 | 61 | 10.67 | 7.34
2008-09-07 | 6:16 | 39 | 99.7 | 12.02 | 61 | 10.74 | 7.34
2008-09-07 | 6:31 | 39 | 99.9 | 11.92 | 61 | 10.79 | 7.34
2008-09-07 | 6:46 | 40 | 100.1 | 11.81 | 61 | 10.84 | 7.34
2008-09-07 | 7:01 | 40 | 99.9 | 11.7 | 61 | 10.84 | 7.34
2008-09-07 | 7:16 | 40 | 100 | 11.59 | 61 | 10.87 | 7.35
2008-09-07 | 7:31 | 40 | 100.5 | 11.49 | 61 | 10.96 | 7.35
2008-09-07 | 7:46 | 40 | 100.5 | 11.4 | 61 | 10.98 | 7.35
2008-09-07 | 8:01 | 40 | 101 | 11.32 | 61 | 11.05 | 7.37
2008-09-07 | 8:16 | 40 | 101.3 | 11.25 | 61 | 11.1 | 7.37
2008-09-07 | 8:31 | 40 | 101.6 | 11.19 | 61 | 11.15 | 7.38
2008-09-07 | 8:46 | 40 | 101.3 | 11.18 | 62 | 11.13 | 7.39
2008-09-07 | 9:01 | 40 | 101.9 | 11.19 | 62 | 11.19 | 7.41
2008-09-07 | 9:16 | 40 | 102.1 | 11.2 | 62 | 11.21 | 7.42
2008-09-07 | 9:31 | 40 | 102.4 | 11.23 | 62 | 11.23 | 7.43
2008-09-07 | 9:46 | 40 | 102.2 | 11.31 | 62 | 11.19 | 7.45
2008-09-07 | 10:01 | 40 | 102.3 | 11.38 | 62 | 11.18 | 7.47
2008-09-07 | 10:16 | 40 | 102.4 | 11.45 | 62 | 11.18 | 7.49
2008-09-07 | 10:31 | 40 | 102.9 | 11.53 | 62 | 11.21 | 7.5
2008-09-07 | 10:46 | 40 | 102.9 | 11.62 | 62 | 11.19 | 7.5
2008-09-07 | 11:01 | 40 | 103.2 | 11.71 | 62 | 11.19 | 7.51
2008-09-07 | 11:16 | 40 | 103.3 | 11.83 | 62 | 11.17 | 7.53
2008-09-07 | 11:31 | 40 | 103.8 | 11.97 | 62 | 11.19 | 7.54
2008-09-07 | 11:46 | 40 | 104 | 12.2 | 62 | 11.15 | 7.58
2008-09-07 | 12:01 | 40 | 104.1 | 12.48 | 62 | 11.09 | 7.58
2008-09-07 | 12:16 | 40 | 104 | 12.85 | 62 | 10.99 | 7.62
2008-09-07 | 12:31 | 40 | 103.5 | 13.24 | 62 | 10.85 | 7.63
2008-09-07 | 12:46 | 40 | 103.6 | 13.67 | 62 | 10.75 | 7.66
2008-09-07 | 13:00 | 40 | 103.5 | 14.13 | 61 | 10.63 | 7.66
2008-09-07 | 13:15 | 40 | 102.9 | 14.6 | 61 | 10.47 | 7.67
2008-09-07 | 13:30 | 40 | 103 | 15.07 | 61 | 10.37 | 7.69
2008-09-07 | 13:45 | 40 | 102.5 | 15.55 | 61 | 10.22 | 7.69
2008-09-07 | 14:00 | 40 | 102.2 | 16.03 | 61 | 10.08 | 7.71
2008-09-07 | 14:15 | 40 | 102.3 | 16.49 | 61 | 9.99 | 7.73
2008-09-07 | 14:30 | 40 | 101.8 | 16.95 | 61 | 9.84 | 7.75
2008-09-07 | 14:45 | 40 | 101.8 | 17.35 | 61 | 9.76 | 7.73
2008-09-07 | 15:00 | 40 | 101.6 | 17.64 | 61 | 9.69 | 7.74
2008-09-07 | 15:15 | 39 | 101.3 | 17.88 | 61 | 9.62 | 7.74
2008-09-07 | 15:30 | 39 | 100.8 | 18.11 | 61 | 9.52 | 7.74
2008-09-07 | 15:45 | 39 | 100.7 | 18.26 | 61 | 9.48 | 7.74
2008-09-07 | 16:00 | 39 | 100.2 | 18.38 | 60 | 9.41 | 7.7
2008-09-07 | 16:15 | 39 | 100.1 | 18.43 | 60 | 9.4 | 7.68
2008-09-07 | 16:30 | 39 | 100 | 18.35 | 60 | 9.4 | 7.64
2008-09-07 | 16:45 | 39 | 99.7 | 18.22 | 60 | 9.4 | 7.6
2008-09-07 | 17:00 | 39 | 99.6 | 18.07 | 60 | 9.41 | 7.55
2008-09-07 | 17:15 | 39 | 99.3 | 17.88 | 60 | 9.43 | 7.51
2008-09-07 | 17:30 | 39 | 99.1 | 17.71 | 60 | 9.43 | 7.47
2008-09-07 | 17:45 | 39 | 98.9 | 17.54 | 60 | 9.45 | 7.46
2008-09-07 | 18:00 | 39 | 98.4 | 17.39 | 60 | 9.43 | 7.45
2008-09-07 | 18:15 | 39 | 98 | 17.23 | 60 | 9.43 | 7.44
2008-09-07 | 18:30 | 39 | 98.4 | 17.11 | 60 | 9.49 | 7.44
2008-09-07 | 18:45 | 39 | 97.9 | 17.02 | 60 | 9.46 | 7.42
2008-09-07 | 19:00 | 39 | 98.1 | 16.92 | 59 | 9.5 | 7.43
2008-09-07 | 19:15 | 39 | 97.7 | 16.86 | 59 | 9.47 | 7.42
2008-09-07 | 19:30 | 39 | 97.5 | 16.81 | 59 | 9.46 | 7.42
2008-09-07 | 19:45 | 39 | 97.7 | 16.74 | 59 | 9.49 | 7.41
2008-09-07 | 20:00 | 39 | 97.7 | 16.67 | 59 | 9.51 | 7.39
2008-09-07 | 20:15 | 39 | 97.7 | 16.57 | 59 | 9.53 | 7.38
2008-09-07 | 20:30 | 39 | 97.4 | 16.46 | 59 | 9.52 | 7.38
2008-09-07 | 20:45 | 39 | 97.3 | 16.34 | 59 | 9.54 | 7.37
2008-09-07 | 21:00 | 39 | 97.3 | 16.18 | 59 | 9.57 | 7.35
2008-09-07 | 21:15 | 39 | 97.8 | 16.03 | 59 | 9.64 | 7.35
2008-09-07 | 21:30 | 39 | 97.6 | 15.86 | 59 | 9.66 | 7.34
2008-09-07 | 21:45 | 39 | 97.7 | 15.69 | 59 | 9.7 | 7.35
2008-09-07 | 22:00 | 39 | 97.7 | 15.53 | 59 | 9.74 | 7.35
2008-09-07 | 22:15 | 39 | 97.8 | 15.35 | 59 | 9.79 | 7.34
2008-09-07 | 22:30 | 39 | 98.2 | 15.19 | 59 | 9.86 | 7.34
2008-09-07 | 22:45 | 39 | 98.1 | 15.03 | 59 | 9.88 | 7.34
2008-09-07 | 23:00 | 39 | 98 | 14.86 | 59 | 9.92 | 7.35
2008-09-07 | 23:15 | 39 | 98 | 14.69 | 59 | 9.95 | 7.34
2008-09-07 | 23:30 | 39 | 98.5 | 14.53 | 59 | 10.03 | 7.34
2008-09-07 | 23:45 | 39 | 98.3 | 14.39 | 59 | 10.05 | 7.34
2008-09-08 | 0:00 | 39 | 98.4 | 14.24 | 59 | 10.08 | 7.34
2008-09-08 | 0:15 | 39 | 98.7 | 14.1 | 59 | 10.15 | 7.33
2008-09-08 | 0:30 | 39 | 98.6 | 13.96 | 59 | 10.17 | 7.34
2008-09-08 | 0:45 | 39 | 98.8 | 13.84 | 59 | 10.22 | 7.34
2008-09-08 | 1:01 | 39 | 99.1 | 13.7 | 59 | 10.28 | 7.3
2008-09-08 | 1:16 | 39 | 99.2 | 13.55 | 59 | 10.33 | 7.34
2008-09-08 | 1:31 | 39 | 99 | 13.38 | 59 | 10.34 | 7.35
2008-09-08 | 1:46 | 39 | 98.9 | 13.21 | 59 | 10.37 | 7.34
2008-09-08 | 2:01 | 39 | 99.1 | 13.04 | 59 | 10.43 | 7.34
2008-09-08 | 2:16 | 39 | 99 | 12.89 | 60 | 10.46 | 7.34
2008-09-08 | 2:31 | 39 | 99 | 12.73 | 60 | 10.5 | 7.34
2008-09-08 | 2:46 | 39 | 99.1 | 12.59 | 60 | 10.54 | 7.34
2008-09-08 | 3:01 | 39 | 99.2 | 12.46 | 60 | 10.58 | 7.35
2008-09-08 | 3:16 | 39 | 99.7 | 12.34 | 60 | 10.66 | 7.35
2008-09-08 | 3:31 | 39 | 99.9 | 12.22 | 60 | 10.71 | 7.35
2008-09-08 | 3:46 | 39 | 100 | 12.09 | 60 | 10.76 | 7.34
2008-09-08 | 4:01 | 39 | 100.1 | 11.98 | 60 | 10.79 | 7.34
2008-09-08 | 4:16 | 39 | 100.3 | 11.86 | 60 | 10.84 | 7.34
2008-09-08 | 4:31 | 39 | 100.3 | 11.75 | 60 | 10.87 | 7.35
2008-09-08 | 4:46 | 39 | 100.4 | 11.64 | 60 | 10.9 | 7.34
2008-09-08 | 5:01 | 39 | 100.5 | 11.54 | 60 | 10.94 | 7.35
2008-09-08 | 5:16 | 39 | 100.5 | 11.43 | 60 | 10.97 | 7.34
2008-09-08 | 5:31 | 39 | 100.6 | 11.31 | 60 | 11.02 | 7.34
2008-09-08 | 5:46 | 39 | 100.6 | 11.19 | 60 | 11.04 | 7.34
2008-09-08 | 6:01 | 39 | 100.2 | 11.07 | 61 | 11.03 | 7.34
2008-09-08 | 6:16 | 39 | 100.6 | 10.96 | 61 | 11.1 | 7.33
2008-09-08 | 6:31 | 39 | 100.4 | 10.83 | 61 | 11.11 | 7.34
2008-09-08 | 6:46 | 40 | 100.4 | 10.69 | 61 | 11.15 | 7.33
2008-09-08 | 7:01 | 40 | 100.5 | 10.56 | 61 | 11.2 | 7.34
2008-09-08 | 7:16 | 40 | 100.6 | 10.43 | 61 | 11.24 | 7.34
2008-09-08 | 7:31 | 40 | 100.9 | 10.3 | 61 | 11.31 | 7.35
2008-09-08 | 7:46 | 40 | 101.1 | 10.18 | 61 | 11.36 | 7.35
2008-09-08 | 8:01 | 40 | 101.3 | 10.07 | 61 | 11.41 | 7.36
2008-09-08 | 8:16 | 40 | 101.9 | 9.99 | 61 | 11.5 | 7.36
2008-09-08 | 8:31 | 40 | 102.2 | 9.92 | 61 | 11.56 | 7.37
2008-09-08 | 8:46 | 40 | 102 | 9.88 | 61 | 11.55 | 7.38
2008-09-08 | 9:01 | 40 | 102.4 | 9.86 | 61 | 11.6 | 7.39
2008-09-08 | 9:16 | 40 | 102.8 | 9.87 | 61 | 11.64 | 7.4
2008-09-08 | 9:31 | 40 | 102.6 | 9.91 | 61 | 11.61 | 7.43
2008-09-08 | 9:46 | 40 | 103.2 | 9.99 | 61 | 11.65 | 7.44
2008-09-08 | 10:01 | 40 | 103.2 | 10.05 | 61 | 11.63 | 7.45
2008-09-08 | 10:16 | 40 | 103.2 | 10.12 | 61 | 11.62 | 7.46
2008-09-08 | 10:31 | 40 | 103.4 | 10.18 | 61 | 11.62 | 7.47
2008-09-08 | 10:46 | 40 | 103.9 | 10.25 | 61 | 11.66 | 7.47
2008-09-08 | 11:01 | 40 | 104.4 | 10.35 | 61 | 11.68 | 7.49
2008-09-08 | 11:16 | 40 | 104.2 | 10.48 | 61 | 11.63 | 7.51
2008-09-08 | 11:31 | 40 | 104.6 | 10.63 | 61 | 11.64 | 7.52
2008-09-08 | 11:46 | 39 | 104.3 | 10.82 | 61 | 11.55 | 7.54
2008-09-08 | 12:01 | 39 | 104.5 | 11.11 | 61 | 11.49 | 7.55
2008-09-08 | 12:16 | 39 | 104.5 | 11.49 | 60 | 11.39 | 7.58
2008-09-08 | 12:31 | 39 | 104.4 | 11.89 | 60 | 11.28 | 7.6
2008-09-08 | 12:46 | 39 | 104 | 12.32 | 60 | 11.13 | 7.62
2008-09-08 | 13:00 | 39 | 104 | 12.8 | 60 | 11.01 | 7.63
2008-09-08 | 13:15 | 39 | 103.4 | 13.28 | 60 | 10.83 | 7.65
2008-09-08 | 13:30 | 39 | 103 | 13.78 | 60 | 10.67 | 7.67
2008-09-08 | 13:45 | 39 | 102.7 | 14.29 | 60 | 10.52 | 7.68
2008-09-08 | 14:00 | 39 | 102.4 | 14.79 | 60 | 10.38 | 7.71
2008-09-08 | 14:15 | 39 | 102.7 | 15.29 | 60 | 10.29 | 7.71
2008-09-08 | 14:30 | 39 | 102.1 | 15.77 | 60 | 10.13 | 7.71
2008-09-08 | 14:45 | 39 | 102.1 | 16.19 | 60 | 10.04 | 7.71
2008-09-08 | 15:00 | 39 | 101.5 | 16.52 | 60 | 9.91 | 7.71
2008-09-08 | 15:15 | 39 | 101.5 | 16.78 | 60 | 9.86 | 7.69
2008-09-08 | 15:30 | 39 | 100.9 | 17.03 | 60 | 9.75 | 7.71
2008-09-08 | 15:45 | 39 | 100.5 | 17.24 | 60 | 9.66 | 7.71
2008-09-08 | 16:00 | 39 | 100.2 | 17.38 | 60 | 9.61 | 7.68
2008-09-08 | 16:15 | 39 | 100 | 17.45 | 60 | 9.57 | 7.66
2008-09-08 | 16:30 | 39 | 99.7 | 17.42 | 59 | 9.55 | 7.62
2008-09-08 | 16:45 | 39 | 99.8 | 17.33 | 59 | 9.58 | 7.58
2008-09-08 | 17:00 | 39 | 99.6 | 17.23 | 59 | 9.58 | 7.55
2008-09-08 | 17:15 | 39 | 99.2 | 17.08 | 59 | 9.57 | 7.52
2008-09-08 | 17:30 | 39 | 98.8 | 16.93 | 59 | 9.56 | 7.48
2008-09-08 | 17:45 | 38 | 99 | 16.82 | 59 | 9.6 | 7.48
2008-09-08 | 18:00 | 39 | 98.8 | 16.71 | 59 | 9.61 | 7.46
2008-09-08 | 18:15 | 38 | 98.3 | 16.6 | 59 | 9.58 | 7.45
2008-09-08 | 18:30 | 38 | 98.1 | 16.56 | 59 | 9.57 | 7.45
2008-09-08 | 18:45 | 38 | 97.8 | 16.51 | 59 | 9.55 | 7.44
2008-09-08 | 19:00 | 38 | 98.1 | 16.48 | 59 | 9.59 | 7.44
2008-09-08 | 19:15 | 38 | 98 | 16.44 | 59 | 9.59 | 7.43
2008-09-08 | 19:30 | 38 | 97.9 | 16.39 | 59 | 9.58 | 7.41
2008-09-08 | 19:45 | 38 | 97.4 | 16.32 | 59 | 9.55 | 7.4
2008-09-08 | 20:00 | 38 | 97.5 | 16.24 | 59 | 9.57 | 7.39
2008-09-08 | 20:15 | 38 | 97.2 | 16.14 | 59 | 9.56 | 7.38
2008-09-08 | 20:30 | 38 | 97.5 | 16.02 | 59 | 9.62 | 7.37
2008-09-08 | 20:45 | 38 | 97.2 | 15.9 | 59 | 9.61 | 7.36
2008-09-08 | 21:00 | 38 | 97.5 | 15.76 | 59 | 9.67 | 7.35
2008-09-08 | 21:15 | 38 | 97.3 | 15.61 | 59 | 9.68 | 7.34
2008-09-08 | 21:30 | 38 | 97.3 | 15.44 | 59 | 9.72 | 7.33
2008-09-08 | 21:45 | 39 | 97.3 | 15.26 | 59 | 9.76 | 7.33
2008-09-08 | 22:00 | 39 | 97.8 | 15.08 | 59 | 9.84 | 7.33
2008-09-08 | 22:15 | 39 | 98 | 14.91 | 59 | 9.9 | 7.33
2008-09-08 | 22:30 | 39 | 98.2 | 14.72 | 59 | 9.96 | 7.34
2008-09-08 | 22:45 | 39 | 98 | 14.54 | 59 | 9.98 | 7.33
2008-09-08 | 23:00 | 39 | 97.9 | 14.35 | 59 | 10.01 | 7.34
2008-09-08 | 23:15 | 39 | 98.3 | 14.17 | 59 | 10.09 | 7.33
2008-09-08 | 23:30 | 39 | 98.6 | 13.99 | 59 | 10.16 | 7.34
2008-09-08 | 23:45 | 39 | 98.7 | 13.82 | 59 | 10.21 | 7.34
2008-09-09 | 0:00 | 39 | 98.9 | 13.65 | 59 | 10.27 | 7.34
2008-09-09 | 0:15 | 39 | 99 | 13.5 | 59 | 10.31 | 7.34
2008-09-09 | 0:30 | 39 | 99 | 13.35 | 59 | 10.35 | 7.35
2008-09-09 | 0:45 | 39 | 99 | 13.21 | 59 | 10.38 | 7.34
2008-09-09 | 1:01 | 39 | 99.1 | 13.09 | 59 | 10.42 | 7.34
2008-09-09 | 1:16 | 38 | 98.8 | 12.96 | 59 | 10.42 | 7.34
2008-09-09 | 1:31 | 38 | 99.1 | 12.85 | 59 | 10.48 | 7.34
2008-09-09 | 1:46 | 38 | 99.3 | 12.73 | 59 | 10.52 | 7.34
2008-09-09 | 2:01 | 38 | 99.4 | 12.62 | 59 | 10.56 | 7.34
2008-09-09 | 2:16 | 38 | 99.1 | 12.5 | 58 | 10.56 | 7.33
2008-09-09 | 2:31 | 38 | 99.4 | 12.4 | 58 | 10.61 | 7.34
2008-09-09 | 2:46 | 38 | 99.6 | 12.29 | 59 | 10.66 | 7.33
2008-09-09 | 3:01 | 38 | 99.6 | 12.17 | 59 | 10.69 | 7.34
2008-09-09 | 3:16 | 38 | 99.3 | 12.07 | 59 | 10.69 | 7.33
2008-09-09 | 3:31 | 38 | 99.2 | 11.96 | 59 | 10.7 | 7.33
2008-09-09 | 3:46 | 38 | 99.6 | 11.86 | 59 | 10.77 | 7.33
2008-09-09 | 4:01 | 38 | 99.3 | 11.76 | 59 | 10.77 | 7.33
2008-09-09 | 4:16 | 38 | 99.7 | 11.65 | 59 | 10.83 | 7.34
2008-09-09 | 4:31 | 38 | 99.6 | 11.54 | 59 | 10.84 | 7.33
2008-09-09 | 4:46 | 38 | 99.9 | 11.44 | 59 | 10.91 | 7.34
2008-09-09 | 5:01 | 39 | 100.1 | 11.35 | 59 | 10.95 | 7.33
2008-09-09 | 5:16 | 39 | 99.9 | 11.25 | 59 | 10.95 | 7.33
2008-09-09 | 5:31 | 39 | 100.1 | 11.15 | 60 | 11 | 7.33
2008-09-09 | 5:46 | 39 | 99.8 | 11.05 | 60 | 10.99 | 7.32
2008-09-09 | 6:01 | 39 | 99.9 | 10.96 | 60 | 11.03 | 7.34
2008-09-09 | 6:16 | 39 | 100.3 | 10.87 | 60 | 11.09 | 7.33
2008-09-09 | 6:31 | 39 | 100.5 | 10.78 | 60 | 11.14 | 7.34
2008-09-09 | 6:46 | 39 | 100.2 | 10.69 | 60 | 11.13 | 7.33
2008-09-09 | 7:01 | 39 | 100.2 | 10.6 | 60 | 11.15 | 7.34
2008-09-09 | 7:16 | 39 | 100.8 | 10.51 | 60 | 11.24 | 7.34
2008-09-09 | 7:31 | 39 | 101.2 | 10.43 | 60 | 11.31 | 7.35
2008-09-09 | 7:46 | 39 | 101.3 | 10.38 | 60 | 11.33 | 7.37
2008-09-09 | 8:01 | 39 | 101.3 | 10.33 | 60 | 11.35 | 7.38
2008-09-09 | 8:16 | 39 | 101.3 | 10.27 | 60 | 11.36 | 7.39
2008-09-09 | 8:31 | 39 | 101.4 | 10.22 | 60 | 11.39 | 7.39
2008-09-09 | 8:46 | 39 | 101.9 | 10.17 | 60 | 11.46 | 7.39
2008-09-09 | 9:01 | 39 | 102.2 | 10.15 | 60 | 11.49 | 7.4
2008-09-09 | 9:16 | 39 | 102.2 | 10.16 | 60 | 11.49 | 7.42
2008-09-09 | 9:31 | 39 | 102.3 | 10.21 | 60 | 11.49 | 7.44
2008-09-09 | 9:46 | 39 | 102.9 | 10.28 | 60 | 11.54 | 7.46
2008-09-09 | 10:01 | 39 | 103.3 | 10.35 | 60 | 11.56 | 7.46
2008-09-09 | 10:16 | 39 | 103.2 | 10.43 | 60 | 11.53 | 7.48
2008-09-09 | 10:31 | 39 | 103.4 | 10.51 | 60 | 11.53 | 7.49
2008-09-09 | 10:46 | 39 | 103.4 | 10.61 | 60 | 11.51 | 7.5
2008-09-09 | 11:01 | 39 | 104 | 10.72 | 60 | 11.55 | 7.51
2008-09-09 | 11:16 | 39 | 103.9 | 10.83 | 60 | 11.5 | 7.53
2008-09-09 | 11:31 | 39 | 104 | 10.98 | 60 | 11.47 | 7.54
2008-09-09 | 11:46 | 39 | 104.4 | 11.18 | 60 | 11.46 | 7.56
2008-09-09 | 12:01 | 39 | 104.5 | 11.47 | 60 | 11.4 | 7.58
2008-09-09 | 12:16 | 39 | 104.1 | 11.82 | 60 | 11.26 | 7.62
2008-09-09 | 12:31 | 39 | 104.2 | 12.21 | 60 | 11.17 | 7.63
2008-09-09 | 12:46 | 39 | 104.1 | 12.64 | 60 | 11.06 | 7.65
2008-09-09 | 13:00 | 39 | 104 | 13.09 | 60 | 10.93 | 7.66
2008-09-09 | 13:15 | 39 | 103.3 | 13.56 | 60 | 10.75 | 7.67
2008-09-09 | 13:30 | 39 | 103.3 | 14.05 | 59 | 10.63 | 7.69
2008-09-09 | 13:45 | 39 | 103.2 | 14.55 | 59 | 10.51 | 7.7
2008-09-09 | 14:00 | 39 | 103 | 15.03 | 59 | 10.38 | 7.73
2008-09-09 | 14:15 | 39 | 102.8 | 15.51 | 59 | 10.25 | 7.73
2008-09-09 | 14:30 | 39 | 102.1 | 15.96 | 59 | 10.08 | 7.74
2008-09-09 | 14:45 | 39 | 101.7 | 16.36 | 59 | 9.96 | 7.73
2008-09-09 | 15:00 | 38 | 101.4 | 16.69 | 59 | 9.87 | 7.69
2008-09-09 | 15:15 | 38 | 101.5 | 16.94 | 59 | 9.82 | 7.69
2008-09-09 | 15:30 | 38 | 101.4 | 17.18 | 59 | 9.76 | 7.68
2008-09-09 | 15:45 | 38 | 101.2 | 17.38 | 59 | 9.7 | 7.67
2008-09-09 | 16:00 | 38 | 100.6 | 17.51 | 59 | 9.62 | 7.65
2008-09-09 | 16:15 | 38 | 100.2 | 17.56 | 59 | 9.57 | 7.63
2008-09-09 | 16:30 | 38 | 99.9 | 17.54 | 59 | 9.54 | 7.6
2008-09-09 | 16:45 | 38 | 99.6 | 17.47 | 59 | 9.53 | 7.57
2008-09-09 | 17:00 | 38 | 99.3 | 17.37 | 59 | 9.52 | 7.54
2008-09-09 | 17:15 | 38 | 99.4 | 17.23 | 59 | 9.56 | 7.49
2008-09-09 | 17:30 | 38 | 99 | 17.09 | 59 | 9.55 | 7.47
2008-09-09 | 17:45 | 38 | 98.8 | 16.98 | 59 | 9.55 | 7.46
2008-09-09 | 18:00 | 38 | 98.7 | 16.86 | 59 | 9.57 | 7.44
2008-09-09 | 18:15 | 38 | 98.8 | 16.76 | 59 | 9.6 | 7.42
2008-09-09 | 18:30 | 38 | 98.7 | 16.7 | 59 | 9.6 | 7.41
2008-09-09 | 18:45 | 38 | 98.6 | 16.64 | 59 | 9.6 | 7.41
2008-09-09 | 19:00 | 38 | 98.5 | 16.59 | 59 | 9.6 | 7.4
2008-09-09 | 19:15 | 38 | 98.3 | 16.56 | 59 | 9.59 | 7.4
2008-09-09 | 19:30 | 38 | 98.1 | 16.52 | 59 | 9.58 | 7.39
2008-09-09 | 19:45 | 38 | 98.1 | 16.45 | 59 | 9.59 | 7.37
2008-09-09 | 20:00 | 38 | 97.9 | 16.39 | 59 | 9.58 | 7.36
2008-09-09 | 20:15 | 38 | 97.4 | 16.29 | 59 | 9.56 | 7.35
2008-09-09 | 20:30 | 38 | 97.4 | 16.18 | 59 | 9.58 | 7.34
2008-09-09 | 20:45 | 38 | 97.4 | 16.05 | 59 | 9.6 | 7.34
2008-09-09 | 21:00 | 38 | 97.4 | 15.89 | 59 | 9.64 | 7.33
2008-09-09 | 21:15 | 38 | 97.5 | 15.71 | 59 | 9.68 | 7.33
2008-09-09 | 21:30 | 38 | 98 | 15.53 | 59 | 9.77 | 7.32
2008-09-09 | 21:45 | 38 | 98.1 | 15.36 | 59 | 9.82 | 7.32
2008-09-09 | 22:00 | 38 | 98 | 15.18 | 59 | 9.84 | 7.32
2008-09-09 | 22:15 | 38 | 98.3 | 14.99 | 59 | 9.92 | 7.32
2008-09-09 | 22:30 | 38 | 98.2 | 14.8 | 59 | 9.94 | 7.33
2008-09-09 | 22:45 | 38 | 98.2 | 14.61 | 59 | 9.98 | 7.33
2008-09-09 | 23:00 | 38 | 98.5 | 14.43 | 59 | 10.06 | 7.33
2008-09-09 | 23:15 | 38 | 98.9 | 14.24 | 59 | 10.14 | 7.33
2008-09-09 | 23:30 | 38 | 98.6 | 14.07 | 59 | 10.14 | 7.33
2008-09-09 | 23:45 | 38 | 98.9 | 13.89 | 59 | 10.22 | 7.33
2008-09-10 | 0:00 | 38 | 99.1 | 13.72 | 59 | 10.28 | 7.33
2008-09-10 | 0:15 | 38 | 98.9 | 13.56 | 59 | 10.29 | 7.33
2008-09-10 | 0:30 | 38 | 99.2 | 13.41 | 59 | 10.36 | 7.33
2008-09-10 | 0:45 | 38 | 99.5 | 13.26 | 59 | 10.42 | 7.34
2008-09-10 | 1:01 | 38 | 99.7 | 13.11 | 59 | 10.47 | 7.33
2008-09-10 | 1:16 | 38 | 99.3 | 12.96 | 59 | 10.48 | 7.33
2008-09-10 | 1:31 | 38 | 99.3 | 12.81 | 59 | 10.5 | 7.33
2008-09-10 | 1:46 | 38 | 99.4 | 12.66 | 59 | 10.56 | 7.33
2008-09-10 | 2:01 | 38 | 99.8 | 12.51 | 59 | 10.63 | 7.32
2008-09-10 | 2:16 | 38 | 100.1 | 12.37 | 59 | 10.7 | 7.32
2008-09-10 | 2:31 | 38 | 100.2 | 12.23 | 59 | 10.74 | 7.33
2008-09-10 | 2:46 | 39 | 100.4 | 12.1 | 59 | 10.79 | 7.33
2008-09-10 | 3:01 | 38 | 100 | 11.96 | 59 | 10.79 | 7.33
2008-09-10 | 3:16 | 39 | 99.9 | 11.84 | 59 | 10.81 | 7.34
2008-09-10 | 3:31 | 39 | 100.4 | 11.72 | 59 | 10.89 | 7.32
2008-09-10 | 3:46 | 39 | 100.3 | 11.59 | 60 | 10.91 | 7.32
2008-09-10 | 4:01 | 39 | 100.7 | 11.47 | 60 | 10.98 | 7.33
2008-09-10 | 4:16 | 39 | 100.9 | 11.35 | 60 | 11.03 | 7.32
2008-09-10 | 4:31 | 39 | 100.6 | 11.24 | 60 | 11.03 | 7.32
2008-09-10 | 4:46 | 39 | 100.6 | 11.13 | 60 | 11.06 | 7.33
2008-09-10 | 5:01 | 39 | 100.9 | 11.03 | 60 | 11.12 | 7.31
2008-09-10 | 5:16 | 39 | 100.8 | 10.92 | 60 | 11.14 | 7.32
2008-09-10 | 5:31 | 39 | 101.2 | 10.81 | 60 | 11.21 | 7.32
2008-09-10 | 5:46 | 39 | 101 | 10.71 | 60 | 11.21 | 7.32
2008-09-10 | 6:01 | 39 | 101.4 | 10.6 | 60 | 11.28 | 7.32
2008-09-10 | 6:16 | 39 | 101.3 | 10.49 | 60 | 11.3 | 7.31
2008-09-10 | 6:31 | 39 | 101.2 | 10.37 | 60 | 11.32 | 7.32
2008-09-10 | 6:46 | 39 | 101.7 | 10.25 | 60 | 11.41 | 7.32
2008-09-10 | 7:01 | 39 | 101.5 | 10.14 | 60 | 11.42 | 7.32
2008-09-10 | 7:16 | 39 | 102.1 | 10.03 | 61 | 11.51 | 7.32
2008-09-10 | 7:31 | 39 | 102.5 | 9.93 | 61 | 11.59 | 7.33
2008-09-10 | 7:46 | 39 | 102.8 | 9.84 | 61 | 11.65 | 7.34
2008-09-10 | 8:01 | 39 | 102.6 | 9.76 | 61 | 11.65 | 7.34
2008-09-10 | 8:16 | 39 | 102.7 | 9.69 | 61 | 11.68 | 7.35
2008-09-10 | 8:31 | 40 | 102.8 | 9.64 | 61 | 11.71 | 7.36
2008-09-10 | 8:46 | 40 | 103.4 | 9.61 | 61 | 11.78 | 7.37
2008-09-10 | 9:01 | 40 | 103.8 | 9.61 | 61 | 11.83 | 7.37
2008-09-10 | 9:16 | 40 | 104.1 | 9.63 | 61 | 11.85 | 7.38
2008-09-10 | 9:31 | 40 | 103.9 | 9.68 | 61 | 11.82 | 7.4
2008-09-10 | 9:46 | 40 | 104.5 | 9.75 | 61 | 11.87 | 7.41
2008-09-10 | 10:01 | 40 | 104.3 | 9.83 | 61 | 11.82 | 7.42
2008-09-10 | 10:16 | 40 | 104.8 | 9.92 | 61 | 11.86 | 7.44
2008-09-10 | 10:31 | 40 | 105.4 | 9.99 | 61 | 11.9 | 7.45
2008-09-10 | 10:46 | 40 | 105.7 | 10.08 | 61 | 11.91 | 7.45
2008-09-10 | 11:01 | 40 | 105.6 | 10.17 | 61 | 11.87 | 7.46
2008-09-10 | 11:16 | 40 | 106 | 10.3 | 61 | 11.89 | 7.48
2008-09-10 | 11:31 | 39 | 106.3 | 10.44 | 61 | 11.88 | 7.48
2008-09-10 | 11:46 | 39 | 105.9 | 10.62 | 61 | 11.78 | 7.5
2008-09-10 | 12:01 | 39 | 105.7 | 10.91 | 61 | 11.67 | 7.52
2008-09-10 | 12:16 | 39 | 105.6 | 11.25 | 61 | 11.57 | 7.54
2008-09-10 | 12:31 | 39 | 105.8 | 11.66 | 60 | 11.49 | 7.57
2008-09-10 | 12:46 | 39 | 105.7 | 12.08 | 60 | 11.37 | 7.6
2008-09-10 | 13:00 | 39 | 105.6 | 12.52 | 60 | 11.25 | 7.6
2008-09-10 | 13:15 | 39 | 105.1 | 12.98 | 60 | 11.08 | 7.63
2008-09-10 | 13:30 | 39 | 105.1 | 13.45 | 60 | 10.96 | 7.63
2008-09-10 | 13:45 | 39 | 104.6 | 13.93 | 60 | 10.79 | 7.65
2008-09-10 | 14:00 | 39 | 104.6 | 14.4 | 60 | 10.69 | 7.66
2008-09-10 | 14:15 | 39 | 104.5 | 14.88 | 60 | 10.57 | 7.68
2008-09-10 | 14:30 | 39 | 103.9 | 15.33 | 60 | 10.4 | 7.68
2008-09-10 | 14:45 | 39 | 103.9 | 15.72 | 60 | 10.32 | 7.69
2008-09-10 | 15:00 | 39 | 103.7 | 16.03 | 60 | 10.22 | 7.67
2008-09-10 | 15:15 | 39 | 103.4 | 16.25 | 60 | 10.15 | 7.67
2008-09-10 | 15:30 | 39 | 103.1 | 16.46 | 60 | 10.08 | 7.67
2008-09-10 | 15:45 | 39 | 102.9 | 16.63 | 60 | 10.02 | 7.67
2008-09-10 | 16:00 | 39 | 102.2 | 16.72 | 60 | 9.94 | 7.65
2008-09-10 | 16:15 | 39 | 101.8 | 16.75 | 60 | 9.89 | 7.62
2008-09-10 | 16:30 | 39 | 101.6 | 16.71 | 60 | 9.88 | 7.6
2008-09-10 | 16:45 | 39 | 101.6 | 16.6 | 60 | 9.9 | 7.55
2008-09-10 | 17:00 | 39 | 101.1 | 16.44 | 60 | 9.89 | 7.52
2008-09-10 | 17:15 | 39 | 101.1 | 16.28 | 60 | 9.92 | 7.48
2008-09-10 | 17:30 | 39 | 100.6 | 16.11 | 60 | 9.91 | 7.46
2008-09-10 | 17:45 | 39 | 100.7 | 16 | 60 | 9.94 | 7.45
2008-09-10 | 18:00 | 39 | 100.7 | 15.85 | 60 | 9.97 | 7.43
2008-09-10 | 18:15 | 39 | 100.2 | 15.74 | 60 | 9.94 | 7.42
2008-09-10 | 18:30 | 39 | 100 | 15.64 | 60 | 9.94 | 7.42
2008-09-10 | 18:45 | 39 | 100.1 | 15.56 | 60 | 9.97 | 7.4
2008-09-10 | 19:00 | 39 | 100.1 | 15.5 | 60 | 9.98 | 7.4
2008-09-10 | 19:15 | 39 | 99.7 | 15.43 | 60 | 9.96 | 7.4
2008-09-10 | 19:30 | 39 | 99.8 | 15.36 | 60 | 9.98 | 7.38
2008-09-10 | 19:45 | 39 | 99.7 | 15.27 | 60 | 9.99 | 7.36
2008-09-10 | 20:00 | 39 | 99.3 | 15.16 | 60 | 9.97 | 7.37
2008-09-10 | 20:15 | 39 | 99.5 | 15.05 | 60 | 10.02 | 7.36
2008-09-10 | 20:30 | 39 | 99.6 | 14.93 | 60 | 10.06 | 7.34
2008-09-10 | 20:45 | 39 | 99.4 | 14.78 | 60 | 10.07 | 7.34
2008-09-10 | 21:00 | 39 | 99.7 | 14.61 | 60 | 10.14 | 7.34
2008-09-10 | 21:15 | 39 | 99.9 | 14.42 | 60 | 10.2 | 7.33
2008-09-10 | 21:30 | 39 | 99.6 | 14.22 | 60 | 10.22 | 7.33
2008-09-10 | 21:45 | 39 | 100.1 | 14.01 | 60 | 10.31 | 7.33
2008-09-10 | 22:00 | 39 | 100.3 | 13.81 | 60 | 10.38 | 7.33
2008-09-10 | 22:15 | 39 | 100.4 | 13.6 | 60 | 10.44 | 7.33
2008-09-10 | 22:30 | 39 | 100.6 | 13.38 | 60 | 10.51 | 7.33
2008-09-10 | 22:45 | 39 | 100.7 | 13.18 | 60 | 10.57 | 7.34
2008-09-10 | 23:00 | 39 | 100.9 | 12.96 | 60 | 10.64 | 7.33
2008-09-10 | 23:15 | 39 | 101 | 12.76 | 60 | 10.7 | 7.34
2008-09-10 | 23:30 | 39 | 101.1 | 12.55 | 60 | 10.76 | 7.33
2008-09-10 | 23:45 | 39 | 101.2 | 12.35 | 60 | 10.82 | 7.34
2008-09-11 | 0:00 | 39 | 101.3 | 12.18 | 60 | 10.87 | 7.34
2008-09-11 | 0:15 | 39 | 101.4 | 12 | 60 | 10.93 | 7.33
2008-09-11 | 0:30 | 39 | 101.1 | 11.83 | 60 | 10.94 | 7.33
2008-09-11 | 0:45 | 39 | 101 | 11.68 | 60 | 10.96 | 7.33
2008-09-11 | 1:01 | 39 | 101.4 | 11.54 | 60 | 11.05 | 7.32
2008-09-11 | 1:16 | 39 | 101.4 | 11.4 | 60 | 11.07 | 7.33
2008-09-11 | 1:31 | 39 | 101.6 | 11.26 | 60 | 11.14 | 7.33
2008-09-11 | 1:46 | 39 | 101.9 | 11.12 | 60 | 11.2 | 7.32
2008-09-11 | 2:01 | 39 | 101.6 | 10.99 | 60 | 11.21 | 7.32
2008-09-11 | 2:16 | 39 | 101.9 | 10.87 | 60 | 11.28 | 7.32
2008-09-11 | 2:31 | 39 | 102.1 | 10.74 | 60 | 11.33 | 7.33
2008-09-11 | 2:46 | 39 | 102.2 | 10.62 | 60 | 11.37 | 7.32
2008-09-11 | 3:01 | 39 | 101.9 | 10.5 | 60 | 11.37 | 7.32
2008-09-11 | 3:16 | 39 | 101.9 | 10.39 | 60 | 11.4 | 7.32
2008-09-11 | 3:31 | 39 | 102.2 | 10.27 | 60 | 11.47 | 7.32
2008-09-11 | 3:46 | 39 | 102.1 | 10.16 | 60 | 11.48 | 7.32
2008-09-11 | 4:01 | 39 | 101.9 | 10.05 | 60 | 11.49 | 7.32
2008-09-11 | 4:16 | 39 | 102.5 | 9.94 | 60 | 11.58 | 7.31
2008-09-11 | 4:31 | 39 | 102.7 | 9.83 | 60 | 11.64 | 7.32
2008-09-11 | 4:46 | 39 | 102.8 | 9.72 | 61 | 11.68 | 7.31
2008-09-11 | 5:01 | 39 | 102.9 | 9.63 | 61 | 11.72 | 7.32
2008-09-11 | 5:16 | 39 | 103 | 9.52 | 61 | 11.76 | 7.31
2008-09-11 | 5:31 | 40 | 103.1 | 9.41 | 61 | 11.8 | 7.3
2008-09-11 | 5:46 | 40 | 102.8 | 9.29 | 61 | 11.8 | 7.31
2008-09-11 | 6:01 | 40 | 103.1 | 9.17 | 61 | 11.87 | 7.31
2008-09-11 | 6:16 | 40 | 103.2 | 9.05 | 61 | 11.92 | 7.31
2008-09-11 | 6:31 | 40 | 103 | 8.93 | 61 | 11.93 | 7.3
2008-09-11 | 6:46 | 40 | 103.4 | 8.81 | 61 | 12.01 | 7.31
2008-09-11 | 7:01 | 40 | 103.7 | 8.69 | 61 | 12.08 | 7.31
2008-09-11 | 7:16 | 40 | 104 | 8.59 | 61 | 12.14 | 7.32
2008-09-11 | 7:31 | 40 | 104.2 | 8.48 | 61 | 12.19 | 7.32
2008-09-11 | 7:46 | 40 | 104 | 8.37 | 61 | 12.21 | 7.32
2008-09-11 | 8:01 | 40 | 104 | 8.28 | 61 | 12.23 | 7.32
2008-09-11 | 8:16 | 40 | 104.5 | 8.21 | 62 | 12.32 | 7.34
2008-09-11 | 8:31 | 40 | 105 | 8.14 | 62 | 12.39 | 7.33
2008-09-11 | 8:46 | 40 | 104.8 | 8.1 | 62 | 12.38 | 7.34
2008-09-11 | 9:01 | 40 | 105.3 | 8.08 | 62 | 12.45 | 7.35
2008-09-11 | 9:16 | 40 | 105.6 | 8.09 | 62 | 12.48 | 7.37
2008-09-11 | 9:31 | 40 | 105.4 | 8.14 | 62 | 12.44 | 7.38
2008-09-11 | 9:46 | 40 | 105.5 | 8.21 | 62 | 12.43 | 7.39
2008-09-11 | 10:01 | 40 | 106.1 | 8.3 | 62 | 12.47 | 7.39
2008-09-11 | 10:16 | 40 | 106 | 8.39 | 62 | 12.43 | 7.42
2008-09-11 | 10:31 | 40 | 106.6 | 8.46 | 62 | 12.49 | 7.42
2008-09-11 | 10:46 | 40 | 107.1 | 8.55 | 62 | 12.51 | 7.42
2008-09-11 | 11:01 | 40 | 107.4 | 8.66 | 62 | 12.51 | 7.43
2008-09-11 | 11:16 | 40 | 107.5 | 8.77 | 61 | 12.5 | 7.45
2008-09-11 | 11:31 | 40 | 107.3 | 8.92 | 61 | 12.42 | 7.45
2008-09-11 | 11:46 | 40 | 107.1 | 9.11 | 61 | 12.35 | 7.47
2008-09-11 | 12:01 | 40 | 106.9 | 9.37 | 61 | 12.25 | 7.49
2008-09-11 | 12:16 | 40 | 106.7 | 9.74 | 61 | 12.12 | 7.53
2008-09-11 | 12:31 | 40 | 107 | 10.17 | 61 | 12.03 | 7.54
2008-09-11 | 12:46 | 40 | 106.9 | 10.63 | 61 | 11.89 | 7.57
2008-09-11 | 13:00 | 40 | 106.8 | 11.11 | 61 | 11.75 | 7.58
2008-09-11 | 13:15 | 40 | 106.1 | 11.61 | 61 | 11.53 | 7.6
2008-09-11 | 13:30 | 40 | 106.1 | 12.13 | 61 | 11.4 | 7.62
2008-09-11 | 13:45 | 40 | 106 | 12.66 | 61 | 11.25 | 7.63
2008-09-11 | 14:00 | 40 | 105.7 | 13.19 | 61 | 11.09 | 7.63
2008-09-11 | 14:15 | 40 | 105.1 | 13.71 | 61 | 10.9 | 7.65
2008-09-11 | 14:30 | 40 | 104.6 | 14.24 | 61 | 10.72 | 7.67
2008-09-11 | 14:45 | 40 | 104.2 | 14.68 | 61 | 10.58 | 7.67
2008-09-11 | 15:00 | 40 | 104.2 | 15.03 | 61 | 10.5 | 7.68
2008-09-11 | 15:15 | 40 | 104.1 | 15.32 | 61 | 10.42 | 7.66
2008-09-11 | 15:30 | 40 | 103.8 | 15.56 | 61 | 10.34 | 7.67
2008-09-11 | 15:45 | 40 | 103.6 | 15.77 | 61 | 10.27 | 7.65
2008-09-11 | 16:00 | 40 | 103 | 15.92 | 61 | 10.18 | 7.64
2008-09-11 | 16:15 | 40 | 102.5 | 16 | 61 | 10.12 | 7.62
2008-09-11 | 16:30 | 40 | 102.6 | 15.97 | 61 | 10.13 | 7.6
2008-09-11 | 16:45 | 40 | 102.4 | 15.89 | 61 | 10.13 | 7.56
2008-09-11 | 17:00 | 40 | 101.8 | 15.78 | 61 | 10.1 | 7.53
2008-09-11 | 17:15 | 40 | 101.8 | 15.64 | 61 | 10.12 | 7.49
2008-09-11 | 17:30 | 39 | 101.2 | 15.51 | 61 | 10.09 | 7.47
2008-09-11 | 17:45 | 39 | 100.9 | 15.4 | 61 | 10.09 | 7.47
2008-09-11 | 18:00 | 39 | 101 | 15.29 | 61 | 10.12 | 7.43
2008-09-11 | 18:15 | 39 | 100.6 | 15.19 | 61 | 10.1 | 7.42
2008-09-11 | 18:30 | 39 | 100.8 | 15.1 | 61 | 10.14 | 7.42
2008-09-11 | 18:45 | 39 | 100.7 | 15.03 | 61 | 10.15 | 7.41
2008-09-11 | 19:00 | 39 | 100.2 | 14.96 | 61 | 10.11 | 7.4
2008-09-11 | 19:15 | 39 | 100.3 | 14.91 | 61 | 10.13 | 7.39
2008-09-11 | 19:30 | 39 | 99.7 | 14.85 | 61 | 10.09 | 7.39
2008-09-11 | 19:45 | 39 | 99.9 | 14.79 | 61 | 10.12 | 7.38
2008-09-11 | 20:00 | 39 | 99.9 | 14.71 | 61 | 10.14 | 7.36
2008-09-11 | 20:15 | 39 | 99.9 | 14.61 | 61 | 10.16 | 7.35
2008-09-11 | 20:30 | 40 | 100 | 14.51 | 61 | 10.19 | 7.36
2008-09-11 | 20:45 | 39 | 99.6 | 14.4 | 61 | 10.18 | 7.34
2008-09-11 | 21:00 | 40 | 99.6 | 14.27 | 61 | 10.2 | 7.34
2008-09-11 | 21:15 | 40 | 99.9 | 14.11 | 61 | 10.27 | 7.35
2008-09-11 | 21:30 | 40 | 100.1 | 13.94 | 61 | 10.33 | 7.33
2008-09-11 | 21:45 | 40 | 99.9 | 13.77 | 61 | 10.35 | 7.34
2008-09-11 | 22:00 | 40 | 99.8 | 13.58 | 61 | 10.38 | 7.34
2008-09-11 | 22:15 | 40 | 99.9 | 13.4 | 61 | 10.43 | 7.34
2008-09-11 | 22:30 | 40 | 100.4 | 13.23 | 61 | 10.52 | 7.34
2008-09-11 | 22:45 | 40 | 100.2 | 13.04 | 61 | 10.55 | 7.33
2008-09-11 | 23:00 | 40 | 100.2 | 12.86 | 61 | 10.59 | 7.34
2008-09-11 | 23:15 | 40 | 100.3 | 12.68 | 61 | 10.64 | 7.34
2008-09-11 | 23:30 | 40 | 100.3 | 12.51 | 61 | 10.69 | 7.34
2008-09-11 | 23:45 | 40 | 100.4 | 12.34 | 61 | 10.74 | 7.34
2008-09-12 | 0:00 | 40 | 100.9 | 12.18 | 61 | 10.83 | 7.33
2008-09-12 | 0:15 | 40 | 100.8 | 12.04 | 61 | 10.85 | 7.34
2008-09-12 | 0:30 | 40 | 100.8 | 11.9 | 61 | 10.88 | 7.34
2008-09-12 | 0:45 | 40 | 100.8 | 11.76 | 61 | 10.92 | 7.33
2008-09-12 | 1:01 | 39 | 101.3 | 11.63 | 61 | 11.01 | 7.33
2008-09-12 | 1:16 | 39 | 101.1 | 11.5 | 61 | 11.02 | 7.33
2008-09-12 | 1:31 | 40 | 101.5 | 11.38 | 61 | 11.09 | 7.33
2008-09-12 | 1:46 | 40 | 101.2 | 11.27 | 61 | 11.09 | 7.33
2008-09-12 | 2:01 | 40 | 101.3 | 11.15 | 61 | 11.13 | 7.33
2008-09-12 | 2:16 | 40 | 101.3 | 11.04 | 61 | 11.15 | 7.33
2008-09-12 | 2:31 | 40 | 101.7 | 10.94 | 61 | 11.23 | 7.33
2008-09-12 | 2:46 | 40 | 101.5 | 10.83 | 61 | 11.24 | 7.33
2008-09-12 | 3:01 | 40 | 101.4 | 10.72 | 61 | 11.26 | 7.33
2008-09-12 | 3:16 | 40 | 101.9 | 10.62 | 61 | 11.33 | 7.33
2008-09-12 | 3:31 | 40 | 102.1 | 10.52 | 61 | 11.38 | 7.32
2008-09-12 | 3:46 | 40 | 102.2 | 10.42 | 61 | 11.43 | 7.32
2008-09-12 | 4:01 | 40 | 101.9 | 10.31 | 61 | 11.42 | 7.32
2008-09-12 | 4:16 | 40 | 102.2 | 10.21 | 61 | 11.48 | 7.32
2008-09-12 | 4:31 | 40 | 102 | 10.11 | 61 | 11.48 | 7.33
2008-09-12 | 4:46 | 40 | 102.3 | 10.01 | 61 | 11.55 | 7.32
2008-09-12 | 5:01 | 40 | 102.5 | 9.92 | 61 | 11.6 | 7.32
2008-09-12 | 5:16 | 40 | 102.7 | 9.82 | 62 | 11.64 | 7.32
2008-09-12 | 5:31 | 40 | 102.8 | 9.73 | 62 | 11.68 | 7.32
2008-09-12 | 5:46 | 40 | 102.4 | 9.63 | 62 | 11.66 | 7.31
2008-09-12 | 6:01 | 40 | 102.7 | 9.54 | 62 | 11.72 | 7.32
2008-09-12 | 6:16 | 40 | 102.5 | 9.43 | 62 | 11.73 | 7.32
2008-09-12 | 6:31 | 40 | 102.9 | 9.33 | 62 | 11.8 | 7.31
2008-09-12 | 6:46 | 40 | 103.1 | 9.22 | 62 | 11.86 | 7.31
2008-09-12 | 7:01 | 41 | 103.4 | 9.13 | 62 | 11.91 | 7.31
2008-09-12 | 7:16 | 41 | 103.6 | 9.03 | 62 | 11.96 | 7.33
2008-09-12 | 7:31 | 41 | 103.4 | 8.93 | 62 | 11.97 | 7.32
2008-09-12 | 7:46 | 41 | 103.4 | 8.84 | 62 | 12 | 7.33
2008-09-12 | 8:01 | 41 | 103.9 | 8.76 | 63 | 12.08 | 7.33
2008-09-12 | 8:16 | 41 | 104.4 | 8.69 | 63 | 12.16 | 7.34
2008-09-12 | 8:31 | 41 | 104.2 | 8.64 | 63 | 12.15 | 7.35
2008-09-12 | 8:46 | 41 | 104.2 | 8.59 | 63 | 12.17 | 7.35
2008-09-12 | 9:01 | 41 | 104.8 | 8.57 | 63 | 12.24 | 7.36
2008-09-12 | 9:16 | 41 | 105.2 | 8.6 | 63 | 12.27 | 7.37
2008-09-12 | 9:31 | 41 | 104.9 | 8.65 | 63 | 12.23 | 7.38
2008-09-12 | 9:46 | 41 | 105 | 8.73 | 63 | 12.22 | 7.39
2008-09-12 | 10:01 | 41 | 105.7 | 8.83 | 63 | 12.27 | 7.41
2008-09-12 | 10:16 | 41 | 105.7 | 8.92 | 63 | 12.24 | 7.42
2008-09-12 | 10:31 | 41 | 105.8 | 9.02 | 63 | 12.23 | 7.44
2008-09-12 | 10:46 | 41 | 106.1 | 9.12 | 63 | 12.22 | 7.46
2008-09-12 | 11:01 | 41 | 106.3 | 9.23 | 63 | 12.22 | 7.45
2008-09-12 | 11:16 | 41 | 106.5 | 9.37 | 63 | 12.2 | 7.47
2008-09-12 | 11:31 | 41 | 107.1 | 9.52 | 63 | 12.22 | 7.48
2008-09-12 | 11:46 | 41 | 106.8 | 9.71 | 63 | 12.14 | 7.49
2008-09-12 | 12:01 | 41 | 107 | 10 | 63 | 12.08 | 7.52
2008-09-12 | 12:16 | 41 | 106.6 | 10.36 | 63 | 11.93 | 7.55
2008-09-12 | 12:31 | 41 | 106.8 | 10.8 | 62 | 11.83 | 7.57
2008-09-12 | 12:46 | 41 | 106.7 | 11.27 | 62 | 11.69 | 7.59
2008-09-12 | 13:00 | 41 | 106.1 | 11.78 | 62 | 11.49 | 7.61
2008-09-12 | 13:15 | 41 | 105.6 | 12.3 | 62 | 11.3 | 7.66
2008-09-12 | 13:30 | 40 | 105.7 | 12.84 | 62 | 11.17 | 7.66