## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AB (Audited In Situ against NIST, 556, AA & Titration)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-29 |  | 2008-08-29 |  | A
Time | Computer set to atomic time |  | 13:34 |  | 13:34 |  | A
Temp | VWR NIST traceable surface water |  | 12.4 |  | 12.67 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 9.88 |  | A
SpC | In Situ |  | 40 |  | 40 |  | A
Ph | In Situ |  | 6.89 |  | 7.36 |  | B
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2008-09-02 12:20 | 14.79 | 7.7 | 11.25 | 111.1 | 68 | 47
Sonde | 2008-09-02 12:22 | 14.82 | 7.89 | 11.5 | 113.6 | 66 | 43
Grade |  | A | A | A |  | A
Audit 2 | 2008-09-09 13:45 | 18.12 | 7.99 | 10.45 | 110.7 | 68 | 48
Sonde | 2008-09-09 13:45 | 18.15 | 8.31 | 10.7 | 113.3 | 66 | 43
Grade |  | A | B | A |  | A
Audit 3 | 2008-09-12 14:45 | 17.46 | 8.23 | 10.76 | 112.4 | 69 | 48
Sonde | 2008-09-12 14:45 | 17.5 | 8.44 | 11.22 | 117.3 | 67 | 44
Grade |  | A | A | B |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AB (Audited In Situ against NIST & 556)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-12 |  | 2008-09-12 |  | A
Time | Computer set to atomic time |  | 14:45 |  | 14:45 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17.5 |  | A
DO | In Situ |  | 10.76 |  | 11.22 |  | B
SpC | In Situ |  | 69 |  | 67 |  | A
Ph | In Situ |  | 8.23 |  | 8.44 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-08-29 |  | 2008-08-29 |  | A
Time | Computer set to atomic time |  | 13:34 |  | 13:34 |  | A
Temp | VWR NIST traceable surface water |  | 12.4 |  | 12.49 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 10.09 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 150 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.06 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-09-02 12:20 | 14.79 | 7.7 | 11.25 | 111.1 | 68 | 47
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-02 |  | 2008-09-02 |  | A
Time | Computer set to atomic time |  | 13:18 |  | 13:18 |  | A
Temp | VWR NIST traceable surface water |  | 12.1 |  | 12.32 |  | A
DO | Winkler titration of surface water |  | 10.7 |  | 11 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 150 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.04 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-09 |  | 2008-09-09 |  | A
Time | Computer set to atomic time |  | 9:47 |  | 9:47 |  | A
Temp | VWR NIST traceable surface water |  | 13.7 |  | 13.84 |  | A
DO | Winkler titration of surface water |  | 9.7 |  | 9.64 |  | A
SpC | 148 µS/mL |  | 153.1 |  | 150 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.09 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-09-09 13:45 | 18.12 | 7.99 | 10.45 | 110.7 | 68 | 48
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-09 |  | 2008-09-09 |  | A
Time | Computer set to atomic time |  | 19:16 |  | 19:16 |  | A
Temp | VWR NIST traceable surface water |  | 15.5 |  | 15.68 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 8.99 |  | A
SpC | 148 µS/mL |  | 153.1 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.09 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-09 |  | 2008-09-09 |  | A
Time | Computer set to atomic time |  | 19:16 |  | 19:16 |  | A
Temp | VWR NIST traceable surface water |  | 15.5 |  | 15.68 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 8.99 |  | A
SpC | 148 µS/mL |  | 153.1 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.09 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-09-12 14:45 | 17.46 | 8.23 | 10.76 | 112.4 | 69 | 48
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-09-12 |  | 2008-09-12 |  | A
Time | Computer set to atomic time |  | 14:45 |  | 14:45 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17.46 |  | A
DO | Winkler titration of surface water |  | 10.4 |  | 10.76 |  | B
SpC | In Situ |  | 67 |  | 69 |  | A
Ph | In Situ |  | 8.44 |  | 8.23 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | TDS | DO% | Temp | SpCond | DO Conc | pH
M/D/Y |  | mg/L | % | ˚C | uS/cm | mg/L
2008-09-02 | 12:31 | 43 | 118.7 | 14.93 | 66 | 11.98 | 7.98
2008-09-02 | 12:46 | 43 | 119 | 15.1 | 66 | 11.97 | 8.04
2008-09-02 | 13:00 | 43 | 118.4 | 15.29 | 66 | 11.86 | 8.09
2008-09-02 | 13:15 | 43 | 117.3 | 15.44 | 66 | 11.71 | 8.13
2008-09-02 | 13:30 | 43 | 116.6 | 15.55 | 66 | 11.62 | 8.17
2008-09-02 | 13:45 | 43 | 116.6 | 15.66 | 66 | 11.6 | 8.21
2008-09-02 | 14:00 | 43 | 116.8 | 15.81 | 66 | 11.57 | 8.24
2008-09-02 | 14:15 | 43 | 116.1 | 15.89 | 66 | 11.48 | 8.28
2008-09-02 | 14:30 | 43 | 116.4 | 16.03 | 66 | 11.48 | 8.32
2008-09-02 | 14:45 | 43 | 116.4 | 16.18 | 66 | 11.45 | 8.35
2008-09-02 | 15:00 | 43 | 115.4 | 16.23 | 66 | 11.34 | 8.37
2008-09-02 | 15:15 | 43 | 115.3 | 16.27 | 66 | 11.31 | 8.39
2008-09-02 | 15:30 | 43 | 114.7 | 16.37 | 66 | 11.23 | 8.42
2008-09-02 | 15:45 | 43 | 114.5 | 16.51 | 66 | 11.18 | 8.44
2008-09-02 | 16:00 | 43 | 114.1 | 16.6 | 66 | 11.12 | 8.46
2008-09-02 | 16:15 | 43 | 113.8 | 16.64 | 66 | 11.08 | 8.47
2008-09-02 | 16:30 | 43 | 113.1 | 16.67 | 66 | 11 | 8.48
2008-09-02 | 16:45 | 43 | 112.4 | 16.69 | 66 | 10.93 | 8.5
2008-09-02 | 17:00 | 43 | 112.2 | 16.7 | 66 | 10.91 | 8.5
2008-09-02 | 17:15 | 43 | 111.3 | 16.7 | 65 | 10.82 | 8.51
2008-09-02 | 17:30 | 43 | 110.4 | 16.68 | 65 | 10.74 | 8.5
2008-09-02 | 17:45 | 43 | 109.6 | 16.63 | 66 | 10.67 | 8.5
2008-09-02 | 18:00 | 43 | 108.6 | 16.56 | 65 | 10.59 | 8.49
2008-09-02 | 18:15 | 43 | 107.7 | 16.49 | 65 | 10.52 | 8.48
2008-09-02 | 18:30 | 43 | 106.7 | 16.4 | 65 | 10.44 | 8.46
2008-09-02 | 18:45 | 43 | 105.8 | 16.3 | 66 | 10.38 | 8.43
2008-09-02 | 19:00 | 43 | 104.3 | 16.21 | 66 | 10.25 | 8.4
2008-09-02 | 19:15 | 43 | 103.3 | 16.11 | 66 | 10.17 | 8.36
2008-09-02 | 19:30 | 43 | 102.1 | 16.01 | 66 | 10.07 | 8.3
2008-09-02 | 19:45 | 43 | 100.4 | 15.9 | 66 | 9.93 | 8.25
2008-09-02 | 20:00 | 43 | 99.1 | 15.8 | 66 | 9.82 | 8.18
2008-09-02 | 20:15 | 43 | 98.3 | 15.69 | 66 | 9.76 | 8.11
2008-09-02 | 20:30 | 43 | 97.4 | 15.58 | 66 | 9.69 | 8.03
2008-09-02 | 20:45 | 43 | 96.2 | 15.47 | 66 | 9.6 | 7.94
2008-09-02 | 21:00 | 43 | 95.2 | 15.36 | 66 | 9.53 | 7.85
2008-09-02 | 21:15 | 43 | 94.4 | 15.24 | 66 | 9.47 | 7.76
2008-09-02 | 21:30 | 43 | 93.6 | 15.13 | 67 | 9.42 | 7.68
2008-09-02 | 21:45 | 43 | 93 | 15 | 67 | 9.38 | 7.62
2008-09-02 | 22:00 | 43 | 92.8 | 14.87 | 67 | 9.38 | 7.54
2008-09-02 | 22:15 | 44 | 92.4 | 14.75 | 67 | 9.37 | 7.49
2008-09-02 | 22:30 | 44 | 92.1 | 14.62 | 67 | 9.36 | 7.44
2008-09-02 | 22:45 | 44 | 91.7 | 14.5 | 67 | 9.35 | 7.4
2008-09-02 | 23:00 | 44 | 91.1 | 14.38 | 67 | 9.31 | 7.37
2008-09-02 | 23:15 | 44 | 90.7 | 14.28 | 67 | 9.3 | 7.33
2008-09-02 | 23:30 | 44 | 90.5 | 14.17 | 67 | 9.29 | 7.31
2008-09-02 | 23:45 | 44 | 90.6 | 14.08 | 67 | 9.32 | 7.28
2008-09-03 | 0:00 | 44 | 90.7 | 13.99 | 67 | 9.35 | 7.27
2008-09-03 | 0:15 | 44 | 90.2 | 13.9 | 67 | 9.32 | 7.25
2008-09-03 | 0:30 | 44 | 90.4 | 13.82 | 67 | 9.35 | 7.23
2008-09-03 | 0:45 | 44 | 90.4 | 13.75 | 67 | 9.37 | 7.22
2008-09-03 | 1:01 | 44 | 90.3 | 13.67 | 67 | 9.37 | 7.21
2008-09-03 | 1:16 | 44 | 90 | 13.61 | 67 | 9.36 | 7.2
2008-09-03 | 1:31 | 44 | 89.9 | 13.54 | 67 | 9.36 | 7.19
2008-09-03 | 1:46 | 44 | 89.7 | 13.47 | 67 | 9.35 | 7.19
2008-09-03 | 2:01 | 43 | 90 | 13.41 | 67 | 9.4 | 7.18
2008-09-03 | 2:16 | 43 | 89.8 | 13.36 | 67 | 9.39 | 7.17
2008-09-03 | 2:31 | 43 | 89.7 | 13.31 | 67 | 9.39 | 7.17
2008-09-03 | 2:46 | 43 | 89.9 | 13.26 | 67 | 9.42 | 7.16
2008-09-03 | 3:01 | 43 | 89.7 | 13.22 | 67 | 9.4 | 7.16
2008-09-03 | 3:16 | 43 | 89.9 | 13.18 | 67 | 9.43 | 7.16
2008-09-03 | 3:31 | 43 | 90 | 13.16 | 67 | 9.45 | 7.15
2008-09-03 | 3:46 | 43 | 90 | 13.13 | 67 | 9.46 | 7.15
2008-09-03 | 4:01 | 43 | 90 | 13.1 | 67 | 9.46 | 7.14
2008-09-03 | 4:16 | 43 | 90 | 13.09 | 67 | 9.46 | 7.14
2008-09-03 | 4:31 | 43 | 89.6 | 13.06 | 67 | 9.43 | 7.14
2008-09-03 | 4:46 | 43 | 89.8 | 13.05 | 66 | 9.45 | 7.14
2008-09-03 | 5:01 | 43 | 89.5 | 13.03 | 66 | 9.43 | 7.14
2008-09-03 | 5:16 | 43 | 89.4 | 13.03 | 66 | 9.42 | 7.14
2008-09-03 | 5:31 | 43 | 89.3 | 13.02 | 66 | 9.41 | 7.13
2008-09-03 | 5:46 | 43 | 89.6 | 13.01 | 66 | 9.44 | 7.13
2008-09-03 | 6:01 | 43 | 89.4 | 13.01 | 66 | 9.42 | 7.13
2008-09-03 | 6:16 | 43 | 89.2 | 12.98 | 66 | 9.4 | 7.13
2008-09-03 | 6:31 | 43 | 89.5 | 12.99 | 66 | 9.44 | 7.12
2008-09-03 | 6:46 | 43 | 89.5 | 12.99 | 66 | 9.44 | 7.13
2008-09-03 | 7:01 | 43 | 90.2 | 12.98 | 66 | 9.5 | 7.13
2008-09-03 | 7:16 | 43 | 90.5 | 12.99 | 66 | 9.53 | 7.14
2008-09-03 | 7:31 | 43 | 91.5 | 13.02 | 66 | 9.63 | 7.15
2008-09-03 | 7:46 | 43 | 92.2 | 13.06 | 66 | 9.7 | 7.18
2008-09-03 | 8:01 | 43 | 93.1 | 13.11 | 66 | 9.79 | 7.2
2008-09-03 | 8:16 | 43 | 94.5 | 13.18 | 66 | 9.92 | 7.22
2008-09-03 | 8:31 | 43 | 95.7 | 13.25 | 66 | 10.03 | 7.24
2008-09-03 | 8:46 | 43 | 96.5 | 13.34 | 66 | 10.09 | 7.28
2008-09-03 | 9:01 | 43 | 97.9 | 13.43 | 66 | 10.22 | 7.31
2008-09-03 | 9:16 | 43 | 99 | 13.54 | 66 | 10.31 | 7.36
2008-09-03 | 9:31 | 43 | 100.2 | 13.66 | 66 | 10.4 | 7.4
2008-09-03 | 9:46 | 43 | 101.4 | 13.8 | 66 | 10.5 | 7.45
2008-09-03 | 10:01 | 43 | 103 | 13.96 | 66 | 10.63 | 7.51
2008-09-03 | 10:16 | 43 | 104.5 | 14.13 | 66 | 10.73 | 7.56
2008-09-03 | 10:31 | 43 | 105.4 | 14.31 | 66 | 10.79 | 7.64
2008-09-03 | 10:46 | 43 | 106.9 | 14.48 | 66 | 10.9 | 7.7
2008-09-03 | 11:01 | 43 | 107.9 | 14.67 | 66 | 10.95 | 7.78
2008-09-03 | 11:16 | 43 | 109.4 | 14.86 | 66 | 11.06 | 7.85
2008-09-03 | 11:31 | 43 | 110 | 15.04 | 66 | 11.08 | 7.91
2008-09-03 | 11:46 | 43 | 110.9 | 15.24 | 66 | 11.12 | 7.98
2008-09-03 | 12:01 | 43 | 111.6 | 15.44 | 66 | 11.15 | 8.04
2008-09-03 | 12:16 | 43 | 112.7 | 15.64 | 66 | 11.21 | 8.1
2008-09-03 | 12:31 | 43 | 113.5 | 15.82 | 66 | 11.25 | 8.15
2008-09-03 | 12:46 | 43 | 113.6 | 16.03 | 66 | 11.21 | 8.2
2008-09-03 | 13:00 | 43 | 114.2 | 16.21 | 66 | 11.22 | 8.24
2008-09-03 | 13:15 | 43 | 114.2 | 16.38 | 66 | 11.18 | 8.29
2008-09-03 | 13:30 | 43 | 114.6 | 16.55 | 67 | 11.18 | 8.32
2008-09-03 | 13:45 | 43 | 114.9 | 16.7 | 67 | 11.18 | 8.35
2008-09-03 | 14:00 | 43 | 114.7 | 16.86 | 67 | 11.12 | 8.38
2008-09-03 | 14:15 | 43 | 115.1 | 17.01 | 67 | 11.13 | 8.42
2008-09-03 | 14:30 | 43 | 115.3 | 17.13 | 67 | 11.11 | 8.43
2008-09-03 | 14:45 | 43 | 114.9 | 17.24 | 67 | 11.04 | 8.46
2008-09-03 | 15:00 | 43 | 114.6 | 17.34 | 67 | 11 | 8.48
2008-09-03 | 15:15 | 43 | 114.8 | 17.43 | 67 | 11 | 8.5
2008-09-03 | 15:30 | 43 | 114.3 | 17.52 | 67 | 10.93 | 8.52
2008-09-03 | 15:45 | 43 | 114.4 | 17.58 | 67 | 10.92 | 8.54
2008-09-03 | 16:00 | 43 | 113.7 | 17.64 | 67 | 10.85 | 8.54
2008-09-03 | 16:15 | 43 | 113.2 | 17.66 | 67 | 10.79 | 8.56
2008-09-03 | 16:30 | 43 | 113 | 17.7 | 67 | 10.77 | 8.55
2008-09-03 | 16:45 | 43 | 112.2 | 17.69 | 67 | 10.69 | 8.56
2008-09-03 | 17:00 | 43 | 111.3 | 17.67 | 67 | 10.6 | 8.56
2008-09-03 | 17:15 | 43 | 110.5 | 17.63 | 67 | 10.54 | 8.56
2008-09-03 | 17:30 | 43 | 110.3 | 17.62 | 67 | 10.52 | 8.56
2008-09-03 | 17:45 | 43 | 109.6 | 17.56 | 67 | 10.47 | 8.55
2008-09-03 | 18:00 | 43 | 108.9 | 17.48 | 67 | 10.42 | 8.53
2008-09-03 | 18:15 | 43 | 107.4 | 17.36 | 67 | 10.31 | 8.5
2008-09-03 | 18:30 | 43 | 106.1 | 17.23 | 67 | 10.2 | 8.45
2008-09-03 | 18:45 | 43 | 104.9 | 17.06 | 67 | 10.12 | 8.38
2008-09-03 | 19:00 | 44 | 103.6 | 16.88 | 67 | 10.03 | 8.3
2008-09-03 | 19:15 | 44 | 102 | 16.68 | 67 | 9.92 | 8.2
2008-09-03 | 19:30 | 44 | 100.6 | 16.48 | 67 | 9.83 | 8.09
2008-09-03 | 19:45 | 44 | 99.1 | 16.3 | 67 | 9.72 | 7.95
2008-09-03 | 20:00 | 43 | 97.5 | 16.11 | 67 | 9.6 | 7.85
2008-09-03 | 20:15 | 43 | 96.5 | 15.91 | 66 | 9.55 | 7.73
2008-09-03 | 20:30 | 43 | 95.6 | 15.73 | 66 | 9.49 | 7.63
2008-09-03 | 20:45 | 42 | 94.4 | 15.56 | 65 | 9.41 | 7.54
2008-09-03 | 21:00 | 42 | 93.6 | 15.41 | 65 | 9.36 | 7.47
2008-09-03 | 21:15 | 42 | 93.3 | 15.27 | 64 | 9.35 | 7.42
2008-09-03 | 21:30 | 42 | 92.5 | 15.17 | 64 | 9.3 | 7.38
2008-09-03 | 21:45 | 42 | 92 | 15.09 | 64 | 9.26 | 7.34
2008-09-03 | 22:00 | 42 | 91.8 | 15.03 | 64 | 9.25 | 7.3
2008-09-03 | 22:15 | 41 | 91.5 | 15.01 | 64 | 9.23 | 7.27
2008-09-03 | 22:30 | 41 | 91.3 | 15 | 64 | 9.2 | 7.25
2008-09-03 | 22:45 | 41 | 90.7 | 15 | 64 | 9.14 | 7.25
2008-09-03 | 23:00 | 41 | 90.3 | 15.01 | 64 | 9.1 | 7.23
2008-09-03 | 23:15 | 41 | 90.3 | 15.02 | 64 | 9.1 | 7.21
2008-09-03 | 23:30 | 41 | 90.1 | 15.04 | 64 | 9.08 | 7.22
2008-09-03 | 23:45 | 41 | 90 | 15.06 | 64 | 9.06 | 7.19
2008-09-04 | 0:00 | 41 | 89.5 | 15.06 | 64 | 9.02 | 7.18
2008-09-04 | 0:15 | 41 | 89.3 | 15.06 | 64 | 8.99 | 7.18
2008-09-04 | 0:30 | 41 | 89.4 | 15.05 | 64 | 9 | 7.18
2008-09-04 | 0:45 | 41 | 89.1 | 15.03 | 64 | 8.97 | 7.17
2008-09-04 | 1:01 | 42 | 89.2 | 15.02 | 64 | 8.99 | 7.16
2008-09-04 | 1:16 | 42 | 89.2 | 14.98 | 64 | 9 | 7.15
2008-09-04 | 1:31 | 42 | 88.9 | 14.96 | 64 | 8.98 | 7.15
2008-09-04 | 1:46 | 42 | 89 | 14.92 | 64 | 8.99 | 7.15
2008-09-04 | 2:01 | 42 | 88.8 | 14.88 | 64 | 8.98 | 7.15
2008-09-04 | 2:16 | 42 | 89 | 14.84 | 64 | 9 | 7.14
2008-09-04 | 2:31 | 42 | 88.9 | 14.79 | 64 | 9.01 | 7.14
2008-09-04 | 2:46 | 42 | 88.7 | 14.74 | 64 | 8.99 | 7.13
2008-09-04 | 3:01 | 42 | 88.6 | 14.7 | 64 | 8.99 | 7.13
2008-09-04 | 3:16 | 42 | 88.6 | 14.66 | 65 | 9 | 7.13
2008-09-04 | 3:31 | 42 | 88.8 | 14.62 | 65 | 9.03 | 7.13
2008-09-04 | 3:46 | 42 | 89 | 14.59 | 65 | 9.06 | 7.14
2008-09-04 | 4:01 | 42 | 88.8 | 14.55 | 65 | 9.04 | 7.12
2008-09-04 | 4:16 | 42 | 89 | 14.52 | 65 | 9.07 | 7.12
2008-09-04 | 4:31 | 42 | 89.1 | 14.49 | 65 | 9.09 | 7.11
2008-09-04 | 4:46 | 42 | 89.3 | 14.45 | 65 | 9.11 | 7.12
2008-09-04 | 5:01 | 42 | 89.1 | 14.42 | 65 | 9.09 | 7.12
2008-09-04 | 5:16 | 43 | 89.4 | 14.38 | 66 | 9.13 | 7.12
2008-09-04 | 5:31 | 43 | 89.5 | 14.35 | 65 | 9.15 | 7.12
2008-09-04 | 5:46 | 43 | 89.6 | 14.3 | 65 | 9.18 | 7.12
2008-09-04 | 6:01 | 42 | 89.7 | 14.25 | 65 | 9.2 | 7.11
2008-09-04 | 6:16 | 42 | 89.5 | 14.19 | 65 | 9.18 | 7.11
2008-09-04 | 6:31 | 42 | 89.4 | 14.14 | 65 | 9.19 | 7.12
2008-09-04 | 6:46 | 42 | 89.6 | 14.08 | 65 | 9.22 | 7.11
2008-09-04 | 7:01 | 42 | 90 | 14.01 | 65 | 9.27 | 7.12
2008-09-04 | 7:16 | 42 | 90.4 | 13.96 | 65 | 9.33 | 7.12
2008-09-04 | 7:31 | 42 | 91.1 | 13.91 | 65 | 9.41 | 7.14
2008-09-04 | 7:46 | 42 | 92 | 13.89 | 65 | 9.51 | 7.15
2008-09-04 | 8:01 | 42 | 93 | 13.87 | 64 | 9.61 | 7.17
2008-09-04 | 8:16 | 42 | 94.3 | 13.86 | 64 | 9.75 | 7.19
2008-09-04 | 8:31 | 42 | 95.6 | 13.87 | 64 | 9.88 | 7.22
2008-09-04 | 8:46 | 42 | 96.4 | 13.89 | 64 | 9.96 | 7.25
2008-09-04 | 9:01 | 42 | 97.6 | 13.93 | 64 | 10.07 | 7.29
2008-09-04 | 9:16 | 42 | 99.1 | 14 | 64 | 10.21 | 7.32
2008-09-04 | 9:31 | 42 | 100.6 | 14.08 | 64 | 10.35 | 7.36
2008-09-04 | 9:46 | 42 | 101.6 | 14.18 | 64 | 10.42 | 7.41
2008-09-04 | 10:01 | 42 | 103 | 14.29 | 64 | 10.55 | 7.45
2008-09-04 | 10:16 | 42 | 104.4 | 14.43 | 64 | 10.66 | 7.51
2008-09-04 | 10:31 | 42 | 105.6 | 14.57 | 64 | 10.74 | 7.57
2008-09-04 | 10:46 | 42 | 106.9 | 14.73 | 64 | 10.85 | 7.63
2008-09-04 | 11:01 | 42 | 107.8 | 14.89 | 64 | 10.9 | 7.7
2008-09-04 | 11:16 | 42 | 108.7 | 15.05 | 64 | 10.94 | 7.76
2008-09-04 | 11:31 | 42 | 109.6 | 15.2 | 65 | 11 | 7.83
2008-09-04 | 11:46 | 42 | 110.9 | 15.37 | 65 | 11.09 | 7.9
2008-09-04 | 12:01 | 42 | 111.9 | 15.56 | 65 | 11.15 | 7.96
2008-09-04 | 12:16 | 42 | 112.2 | 15.75 | 65 | 11.14 | 8.03
2008-09-04 | 12:31 | 42 | 112.7 | 15.92 | 65 | 11.14 | 8.08
2008-09-04 | 12:46 | 42 | 113.5 | 16.08 | 65 | 11.18 | 8.13
2008-09-04 | 13:00 | 42 | 113.5 | 16.24 | 65 | 11.15 | 8.17
2008-09-04 | 13:15 | 42 | 114.1 | 16.42 | 65 | 11.16 | 8.22
2008-09-04 | 13:30 | 43 | 114.6 | 16.57 | 65 | 11.17 | 8.26
2008-09-04 | 13:45 | 43 | 114.8 | 16.7 | 66 | 11.16 | 8.29
2008-09-04 | 14:00 | 43 | 115 | 16.85 | 66 | 11.15 | 8.33
2008-09-04 | 14:15 | 43 | 115.1 | 16.99 | 66 | 11.13 | 8.36
2008-09-04 | 14:30 | 43 | 115 | 17.09 | 66 | 11.09 | 8.38
2008-09-04 | 14:45 | 43 | 114.5 | 17.21 | 66 | 11.02 | 8.41
2008-09-04 | 15:00 | 43 | 114.3 | 17.3 | 66 | 10.98 | 8.42
2008-09-04 | 15:15 | 43 | 114 | 17.39 | 66 | 10.93 | 8.44
2008-09-04 | 15:30 | 43 | 114.3 | 17.47 | 66 | 10.94 | 8.46
2008-09-04 | 15:45 | 43 | 114.1 | 17.54 | 66 | 10.9 | 8.47
2008-09-04 | 16:00 | 43 | 113.9 | 17.59 | 66 | 10.87 | 8.49
2008-09-04 | 16:15 | 43 | 113.5 | 17.63 | 67 | 10.83 | 8.5
2008-09-04 | 16:30 | 43 | 113 | 17.67 | 67 | 10.77 | 8.5
2008-09-04 | 16:45 | 43 | 112.1 | 17.69 | 67 | 10.68 | 8.51
2008-09-04 | 17:00 | 43 | 111.3 | 17.7 | 67 | 10.6 | 8.51
2008-09-04 | 17:15 | 43 | 111 | 17.7 | 67 | 10.57 | 8.51
2008-09-04 | 17:30 | 44 | 110.3 | 17.67 | 67 | 10.51 | 8.5
2008-09-04 | 17:45 | 44 | 109 | 17.64 | 67 | 10.4 | 8.49
2008-09-04 | 18:00 | 44 | 108.5 | 17.58 | 67 | 10.36 | 8.49
2008-09-04 | 18:15 | 44 | 106.9 | 17.5 | 67 | 10.23 | 8.46
2008-09-04 | 18:30 | 44 | 106 | 17.42 | 67 | 10.15 | 8.44
2008-09-04 | 18:45 | 44 | 104.7 | 17.32 | 67 | 10.05 | 8.4
2008-09-04 | 19:00 | 44 | 103.4 | 17.21 | 68 | 9.95 | 8.35
2008-09-04 | 19:15 | 44 | 101.6 | 17.11 | 68 | 9.8 | 8.3
2008-09-04 | 19:30 | 44 | 100.3 | 17 | 68 | 9.69 | 8.24
2008-09-04 | 19:45 | 44 | 98.9 | 16.88 | 68 | 9.58 | 8.18
2008-09-04 | 20:00 | 44 | 97.6 | 16.77 | 68 | 9.47 | 8.1
2008-09-04 | 20:15 | 44 | 96.9 | 16.65 | 68 | 9.44 | 8.01
2008-09-04 | 20:30 | 45 | 95.9 | 16.54 | 69 | 9.36 | 7.92
2008-09-04 | 20:45 | 45 | 94.8 | 16.42 | 69 | 9.28 | 7.83
2008-09-04 | 21:00 | 45 | 94 | 16.31 | 69 | 9.21 | 7.74
2008-09-04 | 21:15 | 45 | 93.6 | 16.2 | 69 | 9.19 | 7.67
2008-09-04 | 21:30 | 45 | 93 | 16.08 | 69 | 9.16 | 7.59
2008-09-04 | 21:45 | 45 | 92.1 | 15.95 | 69 | 9.1 | 7.54
2008-09-04 | 22:00 | 45 | 91.9 | 15.83 | 69 | 9.1 | 7.49
2008-09-04 | 22:15 | 45 | 91.4 | 15.72 | 70 | 9.08 | 7.44
2008-09-04 | 22:30 | 45 | 91.2 | 15.6 | 70 | 9.08 | 7.4
2008-09-04 | 22:45 | 45 | 90.6 | 15.49 | 70 | 9.04 | 7.37
2008-09-04 | 23:00 | 45 | 90.5 | 15.38 | 70 | 9.05 | 7.34
2008-09-04 | 23:15 | 45 | 89.9 | 15.28 | 70 | 9.01 | 7.32
2008-09-04 | 23:30 | 45 | 90 | 15.19 | 70 | 9.04 | 7.3
2008-09-04 | 23:45 | 45 | 89.6 | 15.1 | 70 | 9.01 | 7.28
2008-09-05 | 0:00 | 45 | 89.4 | 15.02 | 70 | 9.01 | 7.26
2008-09-05 | 0:15 | 45 | 89.5 | 14.94 | 70 | 9.03 | 7.25
2008-09-05 | 0:30 | 45 | 89.6 | 14.87 | 70 | 9.06 | 7.24
2008-09-05 | 0:45 | 45 | 89.2 | 14.8 | 70 | 9.04 | 7.23
2008-09-05 | 1:01 | 45 | 89.1 | 14.74 | 70 | 9.04 | 7.22
2008-09-05 | 1:16 | 45 | 89.2 | 14.68 | 70 | 9.06 | 7.22
2008-09-05 | 1:31 | 45 | 89.3 | 14.62 | 70 | 9.08 | 7.2
2008-09-05 | 1:46 | 45 | 89.4 | 14.57 | 69 | 9.09 | 7.2
2008-09-05 | 2:01 | 45 | 89.4 | 14.53 | 69 | 9.1 | 7.19
2008-09-05 | 2:16 | 45 | 89.3 | 14.48 | 69 | 9.11 | 7.18
2008-09-05 | 2:31 | 45 | 89.3 | 14.45 | 69 | 9.11 | 7.18
2008-09-05 | 2:46 | 45 | 88.9 | 14.42 | 69 | 9.08 | 7.17
2008-09-05 | 3:01 | 45 | 88.8 | 14.4 | 69 | 9.07 | 7.18
2008-09-05 | 3:16 | 47 | 88.7 | 14.38 | 72 | 9.07 | 7.16
2008-09-05 | 3:31 | 46 | 88.6 | 14.36 | 71 | 9.06 | 7.17
2008-09-05 | 3:46 | 46 | 88.6 | 14.35 | 71 | 9.06 | 7.16
2008-09-05 | 4:01 | 46 | 88.8 | 14.34 | 71 | 9.09 | 7.15
2008-09-05 | 4:16 | 47 | 89 | 14.33 | 72 | 9.1 | 7.15
2008-09-05 | 4:31 | 47 | 88.7 | 14.33 | 72 | 9.07 | 7.15
2008-09-05 | 4:46 | 47 | 88.9 | 14.32 | 72 | 9.09 | 7.15
2008-09-05 | 5:01 | 47 | 88.6 | 14.32 | 72 | 9.07 | 7.15
2008-09-05 | 5:16 | 46 | 88.5 | 14.32 | 71 | 9.06 | 7.15
2008-09-05 | 5:31 | 46 | 88.4 | 14.31 | 71 | 9.05 | 7.15
2008-09-05 | 5:46 | 46 | 88.4 | 14.32 | 71 | 9.05 | 7.15
2008-09-05 | 6:01 | 46 | 88.4 | 14.32 | 71 | 9.04 | 7.14
2008-09-05 | 6:16 | 46 | 88.7 | 14.32 | 72 | 9.07 | 7.15
2008-09-05 | 6:31 | 46 | 88.5 | 14.31 | 71 | 9.05 | 7.14
2008-09-05 | 6:46 | 46 | 88.9 | 14.32 | 71 | 9.1 | 7.14
2008-09-05 | 7:01 | 46 | 89.2 | 14.3 | 71 | 9.13 | 7.14
2008-09-05 | 7:16 | 46 | 89.5 | 14.31 | 71 | 9.16 | 7.15
2008-09-05 | 7:31 | 46 | 90.5 | 14.33 | 71 | 9.26 | 7.16
2008-09-05 | 7:46 | 46 | 91.3 | 14.38 | 71 | 9.33 | 7.19
2008-09-05 | 8:01 | 46 | 92.3 | 14.43 | 71 | 9.42 | 7.21
2008-09-05 | 8:16 | 46 | 93.7 | 14.49 | 71 | 9.55 | 7.23
2008-09-05 | 8:31 | 46 | 94.6 | 14.55 | 71 | 9.63 | 7.25
2008-09-05 | 8:46 | 46 | 96 | 14.64 | 71 | 9.76 | 7.28
2008-09-05 | 9:01 | 46 | 97.4 | 14.73 | 71 | 9.88 | 7.32
2008-09-05 | 9:16 | 46 | 98.8 | 14.84 | 71 | 10 | 7.36
2008-09-05 | 9:31 | 46 | 100.2 | 14.95 | 70 | 10.12 | 7.4
2008-09-05 | 9:46 | 46 | 101.5 | 15.06 | 70 | 10.22 | 7.45
2008-09-05 | 10:01 | 45 | 103 | 15.22 | 70 | 10.33 | 7.5
2008-09-05 | 10:16 | 45 | 104.2 | 15.36 | 70 | 10.43 | 7.55
2008-09-05 | 10:31 | 45 | 105.1 | 15.53 | 70 | 10.48 | 7.62
2008-09-05 | 10:46 | 45 | 106.2 | 15.7 | 70 | 10.55 | 7.68
2008-09-05 | 11:01 | 45 | 107.4 | 15.88 | 70 | 10.63 | 7.76
2008-09-05 | 11:16 | 45 | 109 | 16.07 | 70 | 10.74 | 7.83
2008-09-05 | 11:31 | 45 | 110.2 | 16.26 | 70 | 10.82 | 7.9
2008-09-05 | 11:46 | 45 | 111.2 | 16.44 | 70 | 10.88 | 7.97
2008-09-05 | 12:01 | 45 | 112.1 | 16.62 | 70 | 10.92 | 8.03
2008-09-05 | 12:16 | 45 | 112.4 | 16.8 | 70 | 10.91 | 8.09
2008-09-05 | 12:31 | 45 | 113.3 | 16.98 | 70 | 10.95 | 8.14
2008-09-05 | 12:46 | 45 | 113.4 | 17.16 | 70 | 10.93 | 8.2
2008-09-05 | 13:00 | 45 | 113.6 | 17.33 | 70 | 10.9 | 8.25
2008-09-05 | 13:15 | 46 | 114.3 | 17.49 | 70 | 10.94 | 8.29
2008-09-05 | 13:30 | 46 | 114.3 | 17.66 | 71 | 10.9 | 8.33
2008-09-05 | 13:45 | 46 | 114.8 | 17.81 | 71 | 10.91 | 8.36
2008-09-05 | 14:00 | 46 | 115.1 | 17.95 | 71 | 10.91 | 8.39
2008-09-05 | 14:15 | 46 | 114.7 | 18.08 | 71 | 10.84 | 8.41
2008-09-05 | 14:30 | 46 | 115.1 | 18.19 | 71 | 10.86 | 8.44
2008-09-05 | 14:45 | 46 | 115.1 | 18.26 | 71 | 10.84 | 8.47
2008-09-05 | 15:00 | 45 | 114.6 | 18.33 | 70 | 10.77 | 8.49
2008-09-05 | 15:15 | 45 | 114.3 | 18.44 | 70 | 10.73 | 8.5
2008-09-05 | 15:30 | 45 | 114.1 | 18.54 | 70 | 10.69 | 8.52
2008-09-05 | 15:45 | 45 | 114.2 | 18.58 | 70 | 10.68 | 8.54
2008-09-05 | 16:00 | 45 | 114 | 18.62 | 70 | 10.66 | 8.55
2008-09-05 | 16:15 | 45 | 113.7 | 18.65 | 70 | 10.62 | 8.55
2008-09-05 | 16:30 | 45 | 113.3 | 18.68 | 70 | 10.58 | 8.56
2008-09-05 | 16:45 | 45 | 112.3 | 18.69 | 70 | 10.49 | 8.57
2008-09-05 | 17:00 | 46 | 111.5 | 18.67 | 71 | 10.41 | 8.57
2008-09-05 | 17:15 | 46 | 110.7 | 18.66 | 71 | 10.34 | 8.56
2008-09-05 | 17:30 | 46 | 110 | 18.62 | 71 | 10.28 | 8.56
2008-09-05 | 17:45 | 46 | 109.6 | 18.56 | 71 | 10.26 | 8.56
2008-09-05 | 18:00 | 46 | 108.9 | 18.49 | 70 | 10.21 | 8.54
2008-09-05 | 18:15 | 45 | 107.7 | 18.41 | 70 | 10.11 | 8.53
2008-09-05 | 18:30 | 45 | 106.4 | 18.31 | 70 | 10.01 | 8.51
2008-09-05 | 18:45 | 45 | 105.4 | 18.19 | 70 | 9.94 | 8.47
2008-09-05 | 19:00 | 45 | 104.1 | 18.06 | 70 | 9.84 | 8.43
2008-09-05 | 19:15 | 45 | 102.3 | 17.96 | 70 | 9.69 | 8.39
2008-09-05 | 19:30 | 46 | 101 | 17.84 | 71 | 9.59 | 8.34
2008-09-05 | 19:45 | 46 | 99.3 | 17.72 | 71 | 9.45 | 8.27
2008-09-05 | 20:00 | 46 | 97.9 | 17.6 | 71 | 9.34 | 8.2
2008-09-05 | 20:15 | 47 | 97.1 | 17.49 | 72 | 9.29 | 8.12
2008-09-05 | 20:30 | 47 | 95.9 | 17.36 | 72 | 9.2 | 8.04
2008-09-05 | 20:45 | 45 | 94.7 | 17.23 | 70 | 9.11 | 7.95
2008-09-05 | 21:00 | 45 | 94.3 | 17.1 | 70 | 9.09 | 7.86
2008-09-05 | 21:15 | 45 | 93.6 | 16.97 | 70 | 9.05 | 7.78
2008-09-05 | 21:30 | 45 | 92.6 | 16.85 | 70 | 8.98 | 7.69
2008-09-05 | 21:45 | 46 | 92 | 16.71 | 70 | 8.94 | 7.63
2008-09-05 | 22:00 | 46 | 91.5 | 16.57 | 70 | 8.92 | 7.57
2008-09-05 | 22:15 | 46 | 91 | 16.44 | 71 | 8.9 | 7.51
2008-09-05 | 22:30 | 46 | 90.9 | 16.31 | 71 | 8.92 | 7.47
2008-09-05 | 22:45 | 46 | 90.7 | 16.18 | 71 | 8.92 | 7.42
2008-09-05 | 23:00 | 46 | 90.4 | 16.06 | 71 | 8.91 | 7.39
2008-09-05 | 23:15 | 46 | 89.9 | 15.95 | 71 | 8.89 | 7.36
2008-09-05 | 23:30 | 46 | 89.9 | 15.84 | 71 | 8.9 | 7.33
2008-09-05 | 23:45 | 46 | 89.5 | 15.73 | 71 | 8.89 | 7.31
2008-09-06 | 0:00 | 46 | 89.7 | 15.63 | 71 | 8.92 | 7.29
2008-09-06 | 0:15 | 46 | 89.3 | 15.54 | 71 | 8.91 | 7.28
2008-09-06 | 0:30 | 46 | 89.5 | 15.45 | 71 | 8.94 | 7.26
2008-09-06 | 0:45 | 46 | 89.5 | 15.38 | 71 | 8.95 | 7.25
2008-09-06 | 1:01 | 46 | 89.5 | 15.3 | 71 | 8.97 | 7.24
2008-09-06 | 1:16 | 46 | 89.5 | 15.23 | 71 | 8.98 | 7.23
2008-09-06 | 1:31 | 46 | 89.5 | 15.17 | 71 | 8.99 | 7.22
2008-09-06 | 1:46 | 46 | 89.2 | 15.11 | 71 | 8.97 | 7.21
2008-09-06 | 2:01 | 46 | 89.3 | 15.06 | 71 | 8.99 | 7.21
2008-09-06 | 2:16 | 46 | 89.1 | 15.02 | 71 | 8.98 | 7.2
2008-09-06 | 2:31 | 46 | 89 | 14.99 | 71 | 8.98 | 7.2
2008-09-06 | 2:46 | 46 | 89.3 | 14.97 | 71 | 9.01 | 7.19
2008-09-06 | 3:01 | 46 | 89.1 | 14.96 | 71 | 8.99 | 7.19
2008-09-06 | 3:16 | 46 | 89.3 | 14.96 | 71 | 9.01 | 7.19
2008-09-06 | 3:31 | 46 | 89.4 | 14.96 | 71 | 9.02 | 7.18
2008-09-06 | 3:46 | 46 | 89 | 14.95 | 71 | 8.99 | 7.18
2008-09-06 | 4:01 | 46 | 89.2 | 14.97 | 71 | 9 | 7.19
2008-09-06 | 4:16 | 46 | 89.3 | 14.97 | 71 | 9.01 | 7.18
2008-09-06 | 4:31 | 46 | 88.9 | 14.99 | 71 | 8.97 | 7.18
2008-09-06 | 4:46 | 46 | 88.8 | 14.99 | 71 | 8.96 | 7.17
2008-09-06 | 5:01 | 46 | 88.8 | 15.01 | 70 | 8.95 | 7.18
2008-09-06 | 5:16 | 46 | 89 | 15.01 | 70 | 8.97 | 7.17
2008-09-06 | 5:31 | 46 | 88.7 | 15.01 | 70 | 8.94 | 7.17
2008-09-06 | 5:46 | 46 | 88.6 | 15.02 | 70 | 8.93 | 7.17
2008-09-06 | 6:01 | 45 | 88.8 | 15.02 | 70 | 8.95 | 7.16
2008-09-06 | 6:16 | 45 | 88.9 | 15.03 | 70 | 8.96 | 7.17
2008-09-06 | 6:31 | 45 | 89 | 15.02 | 70 | 8.97 | 7.17
2008-09-06 | 6:46 | 45 | 88.8 | 15.01 | 70 | 8.95 | 7.17
2008-09-06 | 7:01 | 45 | 89.3 | 15.01 | 70 | 9 | 7.17
2008-09-06 | 7:16 | 45 | 89.5 | 15.01 | 70 | 9.02 | 7.18
2008-09-06 | 7:31 | 45 | 90.1 | 15.03 | 70 | 9.08 | 7.18
2008-09-06 | 7:46 | 45 | 91.1 | 15.06 | 69 | 9.17 | 7.2
2008-09-06 | 8:01 | 45 | 92.1 | 15.1 | 69 | 9.26 | 7.22
2008-09-06 | 8:16 | 45 | 93.2 | 15.15 | 69 | 9.37 | 7.25
2008-09-06 | 8:31 | 45 | 94.7 | 15.2 | 69 | 9.51 | 7.27
2008-09-06 | 8:46 | 45 | 95.9 | 15.27 | 69 | 9.62 | 7.3
2008-09-06 | 9:01 | 45 | 97.3 | 15.35 | 69 | 9.74 | 7.33
2008-09-06 | 9:16 | 45 | 98.7 | 15.44 | 69 | 9.86 | 7.37
2008-09-06 | 9:31 | 45 | 99.7 | 15.56 | 69 | 9.94 | 7.41
2008-09-06 | 9:46 | 45 | 100.9 | 15.68 | 69 | 10.03 | 7.46
2008-09-06 | 10:01 | 45 | 102.6 | 15.81 | 69 | 10.17 | 7.51
2008-09-06 | 10:16 | 45 | 103.7 | 15.95 | 69 | 10.24 | 7.57
2008-09-06 | 10:31 | 45 | 105.2 | 16.1 | 69 | 10.36 | 7.62
2008-09-06 | 10:46 | 45 | 106.1 | 16.27 | 69 | 10.41 | 7.7
2008-09-06 | 11:01 | 45 | 107.6 | 16.43 | 69 | 10.52 | 7.76
2008-09-06 | 11:16 | 45 | 108.7 | 16.6 | 69 | 10.59 | 7.83
2008-09-06 | 11:31 | 45 | 109.3 | 16.78 | 69 | 10.61 | 7.89
2008-09-06 | 11:46 | 45 | 110.4 | 16.96 | 69 | 10.67 | 7.95
2008-09-06 | 12:01 | 45 | 111.2 | 17.13 | 69 | 10.71 | 8.01
2008-09-06 | 12:16 | 45 | 111.5 | 17.31 | 69 | 10.7 | 8.07
2008-09-06 | 12:31 | 45 | 112.4 | 17.49 | 69 | 10.75 | 8.12
2008-09-06 | 12:46 | 45 | 112.5 | 17.67 | 69 | 10.72 | 8.17
2008-09-06 | 13:00 | 45 | 112.9 | 17.84 | 69 | 10.72 | 8.22
2008-09-06 | 13:15 | 45 | 113.2 | 18.02 | 69 | 10.71 | 8.26
2008-09-06 | 13:30 | 45 | 113.8 | 18.17 | 69 | 10.73 | 8.29
2008-09-06 | 13:45 | 45 | 113.8 | 18.33 | 69 | 10.7 | 8.32
2008-09-06 | 14:00 | 44 | 113.8 | 18.48 | 68 | 10.67 | 8.35
2008-09-06 | 14:15 | 44 | 114.3 | 18.61 | 68 | 10.69 | 8.38
2008-09-06 | 14:30 | 44 | 114.6 | 18.73 | 68 | 10.69 | 8.42
2008-09-06 | 14:45 | 44 | 114.5 | 18.86 | 68 | 10.65 | 8.44
2008-09-06 | 15:00 | 44 | 114.5 | 18.97 | 68 | 10.62 | 8.46
2008-09-06 | 15:15 | 44 | 113.8 | 19.06 | 68 | 10.54 | 8.47
2008-09-06 | 15:30 | 44 | 113.3 | 19.13 | 68 | 10.49 | 8.49
2008-09-06 | 15:45 | 44 | 113.5 | 19.2 | 68 | 10.49 | 8.5
2008-09-06 | 16:00 | 44 | 112.8 | 19.25 | 68 | 10.41 | 8.52
2008-09-06 | 16:15 | 44 | 112.5 | 19.29 | 68 | 10.37 | 8.52
2008-09-06 | 16:30 | 44 | 112.2 | 19.33 | 68 | 10.34 | 8.53
2008-09-06 | 16:45 | 44 | 111.5 | 19.34 | 68 | 10.28 | 8.53
2008-09-06 | 17:00 | 44 | 110.5 | 19.33 | 68 | 10.18 | 8.53
2008-09-06 | 17:15 | 44 | 109.4 | 19.3 | 68 | 10.09 | 8.53
2008-09-06 | 17:30 | 44 | 108.6 | 19.27 | 68 | 10.01 | 8.52
2008-09-06 | 17:45 | 44 | 107.6 | 19.22 | 68 | 9.93 | 8.5
2008-09-06 | 18:00 | 44 | 107 | 19.15 | 68 | 9.89 | 8.49
2008-09-06 | 18:15 | 44 | 105.6 | 19.07 | 68 | 9.78 | 8.47
2008-09-06 | 18:30 | 44 | 104.6 | 18.97 | 68 | 9.7 | 8.43
2008-09-06 | 18:45 | 44 | 103.3 | 18.86 | 68 | 9.6 | 8.39
2008-09-06 | 19:00 | 44 | 101.5 | 18.75 | 68 | 9.46 | 8.34
2008-09-06 | 19:15 | 44 | 100.4 | 18.64 | 68 | 9.38 | 8.29
2008-09-06 | 19:30 | 44 | 99.2 | 18.53 | 68 | 9.29 | 8.22
2008-09-06 | 19:45 | 44 | 97.5 | 18.4 | 68 | 9.15 | 8.15
2008-09-06 | 20:00 | 44 | 96.1 | 18.28 | 68 | 9.05 | 8.07
2008-09-06 | 20:15 | 44 | 95.4 | 18.15 | 68 | 9 | 7.98
2008-09-06 | 20:30 | 44 | 94.3 | 18.03 | 68 | 8.92 | 7.88
2008-09-06 | 20:45 | 44 | 93.4 | 17.9 | 68 | 8.86 | 7.8
2008-09-06 | 21:00 | 45 | 92.9 | 17.77 | 69 | 8.84 | 7.71
2008-09-06 | 21:15 | 45 | 92.4 | 17.65 | 69 | 8.81 | 7.64
2008-09-06 | 21:30 | 45 | 91.8 | 17.53 | 69 | 8.77 | 7.57
2008-09-06 | 21:45 | 45 | 91 | 17.41 | 69 | 8.72 | 7.51
2008-09-06 | 22:00 | 45 | 90.5 | 17.29 | 69 | 8.69 | 7.47
2008-09-06 | 22:15 | 45 | 90 | 17.17 | 69 | 8.67 | 7.42
2008-09-06 | 22:30 | 45 | 90 | 17.05 | 69 | 8.69 | 7.39
2008-09-06 | 22:45 | 45 | 89.7 | 16.94 | 69 | 8.68 | 7.35
2008-09-06 | 23:00 | 45 | 89.2 | 16.82 | 69 | 8.65 | 7.32
2008-09-06 | 23:15 | 45 | 89 | 16.72 | 69 | 8.65 | 7.3
2008-09-06 | 23:30 | 45 | 89 | 16.62 | 69 | 8.67 | 7.28
2008-09-06 | 23:45 | 45 | 89 | 16.53 | 69 | 8.68 | 7.26
2008-09-07 | 0:00 | 45 | 89 | 16.44 | 69 | 8.7 | 7.24
2008-09-07 | 0:15 | 45 | 88.6 | 16.36 | 69 | 8.68 | 7.23
2008-09-07 | 0:30 | 45 | 88.5 | 16.28 | 69 | 8.68 | 7.22
2008-09-07 | 0:45 | 45 | 88.4 | 16.21 | 69 | 8.69 | 7.21
2008-09-07 | 1:01 | 45 | 88.4 | 16.15 | 69 | 8.7 | 7.2
2008-09-07 | 1:16 | 45 | 88.7 | 16.09 | 69 | 8.74 | 7.19
2008-09-07 | 1:31 | 45 | 88.9 | 16.04 | 69 | 8.76 | 7.18
2008-09-07 | 1:46 | 45 | 88.9 | 16 | 69 | 8.78 | 7.19
2008-09-07 | 2:01 | 45 | 89 | 15.97 | 69 | 8.79 | 7.18
2008-09-07 | 2:16 | 44 | 88.7 | 15.94 | 68 | 8.76 | 7.17
2008-09-07 | 2:31 | 44 | 88.9 | 15.92 | 68 | 8.78 | 7.17
2008-09-07 | 2:46 | 44 | 88.6 | 15.9 | 68 | 8.76 | 7.17
2008-09-07 | 3:01 | 44 | 88.7 | 15.9 | 68 | 8.78 | 7.17
2008-09-07 | 3:16 | 44 | 88.8 | 15.88 | 68 | 8.79 | 7.16
2008-09-07 | 3:31 | 44 | 88.8 | 15.88 | 68 | 8.79 | 7.16
2008-09-07 | 3:46 | 44 | 88.5 | 15.88 | 68 | 8.75 | 7.16
2008-09-07 | 4:01 | 44 | 88.6 | 15.88 | 68 | 8.77 | 7.15
2008-09-07 | 4:16 | 44 | 88.6 | 15.87 | 68 | 8.77 | 7.15
2008-09-07 | 4:31 | 44 | 88.6 | 15.88 | 67 | 8.77 | 7.15
2008-09-07 | 4:46 | 44 | 88.5 | 15.87 | 67 | 8.76 | 7.15
2008-09-07 | 5:01 | 44 | 88.2 | 15.89 | 67 | 8.73 | 7.15
2008-09-07 | 5:16 | 44 | 88 | 15.89 | 67 | 8.71 | 7.14
2008-09-07 | 5:31 | 44 | 87.9 | 15.89 | 67 | 8.7 | 7.14
2008-09-07 | 5:46 | 44 | 87.9 | 15.9 | 67 | 8.69 | 7.14
2008-09-07 | 6:01 | 44 | 88.2 | 15.9 | 67 | 8.72 | 7.14
2008-09-07 | 6:16 | 44 | 87.9 | 15.89 | 67 | 8.7 | 7.14
2008-09-07 | 6:31 | 44 | 88.2 | 15.88 | 67 | 8.72 | 7.14
2008-09-07 | 6:46 | 44 | 88.4 | 15.88 | 67 | 8.74 | 7.14
2008-09-07 | 7:01 | 44 | 88.5 | 15.87 | 67 | 8.76 | 7.14
2008-09-07 | 7:16 | 44 | 88.9 | 15.87 | 67 | 8.8 | 7.15
2008-09-07 | 7:31 | 44 | 89.5 | 15.87 | 67 | 8.86 | 7.16
2008-09-07 | 7:46 | 44 | 90.5 | 15.91 | 67 | 8.95 | 7.17
2008-09-07 | 8:01 | 44 | 91.6 | 15.94 | 67 | 9.06 | 7.2
2008-09-07 | 8:16 | 44 | 93.1 | 15.99 | 67 | 9.2 | 7.22
2008-09-07 | 8:31 | 44 | 94.1 | 16.05 | 67 | 9.28 | 7.25
2008-09-07 | 8:46 | 44 | 95.5 | 16.11 | 67 | 9.41 | 7.28
2008-09-07 | 9:01 | 44 | 96.6 | 16.18 | 68 | 9.5 | 7.31
2008-09-07 | 9:16 | 44 | 98.1 | 16.27 | 68 | 9.63 | 7.35
2008-09-07 | 9:31 | 44 | 99.7 | 16.37 | 68 | 9.76 | 7.4
2008-09-07 | 9:46 | 44 | 101.2 | 16.48 | 68 | 9.88 | 7.44
2008-09-07 | 10:01 | 44 | 102.2 | 16.6 | 68 | 9.96 | 7.49
2008-09-07 | 10:16 | 44 | 103.4 | 16.72 | 68 | 10.05 | 7.56
2008-09-07 | 10:31 | 44 | 104.6 | 16.85 | 68 | 10.14 | 7.61
2008-09-07 | 10:46 | 44 | 106.2 | 17 | 67 | 10.26 | 7.68
2008-09-07 | 11:01 | 44 | 107.1 | 17.15 | 67 | 10.32 | 7.75
2008-09-07 | 11:16 | 44 | 108.4 | 17.3 | 68 | 10.41 | 7.82
2008-09-07 | 11:31 | 44 | 109.1 | 17.45 | 67 | 10.44 | 7.88
2008-09-07 | 11:46 | 44 | 109.8 | 17.62 | 67 | 10.48 | 7.95
2008-09-07 | 12:01 | 44 | 111 | 17.78 | 68 | 10.55 | 8.01
2008-09-07 | 12:16 | 44 | 111.4 | 17.94 | 67 | 10.55 | 8.07
2008-09-07 | 12:31 | 44 | 112.4 | 18.1 | 67 | 10.61 | 8.13
2008-09-07 | 12:46 | 44 | 112.5 | 18.26 | 67 | 10.6 | 8.18
2008-09-07 | 13:00 | 44 | 113.2 | 18.41 | 67 | 10.63 | 8.22
2008-09-07 | 13:15 | 44 | 113.7 | 18.56 | 68 | 10.64 | 8.26
2008-09-07 | 13:30 | 44 | 113.6 | 18.7 | 67 | 10.6 | 8.31
2008-09-07 | 13:45 | 44 | 114.1 | 18.83 | 67 | 10.62 | 8.34
2008-09-07 | 14:00 | 44 | 114.4 | 18.95 | 67 | 10.62 | 8.37
2008-09-07 | 14:15 | 44 | 114.5 | 19.05 | 67 | 10.61 | 8.4
2008-09-07 | 14:30 | 44 | 114.1 | 19.14 | 67 | 10.55 | 8.42
2008-09-07 | 14:45 | 44 | 114.3 | 19.23 | 67 | 10.56 | 8.45
2008-09-07 | 15:00 | 44 | 114.3 | 19.31 | 67 | 10.54 | 8.47
2008-09-07 | 15:15 | 44 | 113.7 | 19.37 | 67 | 10.47 | 8.49
2008-09-07 | 15:30 | 44 | 113.7 | 19.42 | 67 | 10.45 | 8.5
2008-09-07 | 15:45 | 44 | 113.4 | 19.47 | 67 | 10.42 | 8.52
2008-09-07 | 16:00 | 44 | 112.7 | 19.49 | 67 | 10.35 | 8.53
2008-09-07 | 16:15 | 44 | 112.2 | 19.52 | 67 | 10.3 | 8.54
2008-09-07 | 16:30 | 44 | 111.9 | 19.52 | 67 | 10.27 | 8.55
2008-09-07 | 16:45 | 44 | 111 | 19.51 | 67 | 10.19 | 8.55
2008-09-07 | 17:00 | 44 | 110.3 | 19.46 | 67 | 10.13 | 8.55
2008-09-07 | 17:15 | 44 | 109.8 | 19.41 | 67 | 10.1 | 8.54
2008-09-07 | 17:30 | 44 | 109.2 | 19.35 | 67 | 10.05 | 8.54
2008-09-07 | 17:45 | 44 | 108 | 19.26 | 67 | 9.97 | 8.52
2008-09-07 | 18:00 | 44 | 107.2 | 19.17 | 67 | 9.91 | 8.51
2008-09-07 | 18:15 | 44 | 105.6 | 19.05 | 67 | 9.78 | 8.48
2008-09-07 | 18:30 | 44 | 104.5 | 18.94 | 67 | 9.7 | 8.45
2008-09-07 | 18:45 | 44 | 102.7 | 18.8 | 67 | 9.56 | 8.42
2008-09-07 | 19:00 | 44 | 101.6 | 18.67 | 68 | 9.48 | 8.36
2008-09-07 | 19:15 | 44 | 100.3 | 18.54 | 68 | 9.39 | 8.31
2008-09-07 | 19:30 | 44 | 99.1 | 18.42 | 68 | 9.3 | 8.25
2008-09-07 | 19:45 | 44 | 97.4 | 18.29 | 68 | 9.16 | 8.17
2008-09-07 | 20:00 | 44 | 96.1 | 18.16 | 68 | 9.07 | 8.09
2008-09-07 | 20:15 | 44 | 95.1 | 18.04 | 68 | 8.99 | 8.01
2008-09-07 | 20:30 | 44 | 94.5 | 17.91 | 68 | 8.96 | 7.92
2008-09-07 | 20:45 | 44 | 93.8 | 17.78 | 68 | 8.92 | 7.83
2008-09-07 | 21:00 | 44 | 92.8 | 17.65 | 68 | 8.85 | 7.75
2008-09-07 | 21:15 | 45 | 92.5 | 17.53 | 69 | 8.84 | 7.68
2008-09-07 | 21:30 | 45 | 91.7 | 17.4 | 69 | 8.79 | 7.61
2008-09-07 | 21:45 | 45 | 91.4 | 17.27 | 69 | 8.78 | 7.55
2008-09-07 | 22:00 | 45 | 91 | 17.14 | 69 | 8.77 | 7.49
2008-09-07 | 22:15 | 45 | 90.6 | 17.03 | 69 | 8.75 | 7.45
2008-09-07 | 22:30 | 45 | 90.3 | 16.91 | 69 | 8.75 | 7.41
2008-09-07 | 22:45 | 45 | 90.1 | 16.81 | 70 | 8.74 | 7.38
2008-09-07 | 23:00 | 45 | 89.8 | 16.71 | 69 | 8.73 | 7.35
2008-09-07 | 23:15 | 45 | 89.6 | 16.61 | 70 | 8.73 | 7.33
2008-09-07 | 23:30 | 45 | 89.5 | 16.52 | 70 | 8.73 | 7.31
2008-09-07 | 23:45 | 44 | 89.1 | 16.43 | 68 | 8.71 | 7.29
2008-09-08 | 0:00 | 45 | 88.9 | 16.35 | 69 | 8.71 | 7.27
2008-09-08 | 0:15 | 45 | 89.2 | 16.27 | 69 | 8.75 | 7.26
2008-09-08 | 0:30 | 45 | 89.2 | 16.19 | 69 | 8.77 | 7.24
2008-09-08 | 0:45 | 45 | 89 | 16.12 | 69 | 8.76 | 7.22
2008-09-08 | 1:01 | 45 | 89.2 | 16.05 | 69 | 8.79 | 7.22
2008-09-08 | 1:16 | 43 | 89.3 | 16 | 67 | 8.82 | 7.21
2008-09-08 | 1:31 | 45 | 89.3 | 15.95 | 69 | 8.82 | 7.21
2008-09-08 | 1:46 | 43 | 89.3 | 15.92 | 67 | 8.83 | 7.2
2008-09-08 | 2:01 | 45 | 89.3 | 15.9 | 69 | 8.83 | 7.2
2008-09-08 | 2:16 | 45 | 89 | 15.87 | 68 | 8.81 | 7.2
2008-09-08 | 2:31 | 44 | 89.2 | 15.85 | 68 | 8.83 | 7.19
2008-09-08 | 2:46 | 44 | 88.9 | 15.83 | 68 | 8.81 | 7.2
2008-09-08 | 3:01 | 44 | 89 | 15.82 | 68 | 8.82 | 7.19
2008-09-08 | 3:16 | 44 | 88.8 | 15.8 | 68 | 8.8 | 7.19
2008-09-08 | 3:31 | 44 | 88.9 | 15.79 | 67 | 8.82 | 7.18
2008-09-08 | 3:46 | 45 | 89 | 15.78 | 69 | 8.82 | 7.18
2008-09-08 | 4:01 | 43 | 88.6 | 15.76 | 67 | 8.79 | 7.17
2008-09-08 | 4:16 | 44 | 88.5 | 15.75 | 67 | 8.78 | 7.18
2008-09-08 | 4:31 | 43 | 88.7 | 15.73 | 67 | 8.8 | 7.17
2008-09-08 | 4:46 | 43 | 88.7 | 15.7 | 66 | 8.81 | 7.17
2008-09-08 | 5:01 | 44 | 88.7 | 15.69 | 67 | 8.81 | 7.17
2008-09-08 | 5:16 | 43 | 88.4 | 15.66 | 66 | 8.79 | 7.17
2008-09-08 | 5:31 | 43 | 88.6 | 15.64 | 67 | 8.81 | 7.17
2008-09-08 | 5:46 | 43 | 88.7 | 15.62 | 66 | 8.82 | 7.16
2008-09-08 | 6:01 | 44 | 88.5 | 15.58 | 67 | 8.82 | 7.16
2008-09-08 | 6:16 | 43 | 88.6 | 15.55 | 66 | 8.83 | 7.16
2008-09-08 | 6:31 | 44 | 88.4 | 15.51 | 67 | 8.81 | 7.16
2008-09-08 | 6:46 | 44 | 88.7 | 15.47 | 68 | 8.86 | 7.16
2008-09-08 | 7:01 | 43 | 88.8 | 15.42 | 67 | 8.87 | 7.16
2008-09-08 | 7:16 | 44 | 89.5 | 15.38 | 67 | 8.95 | 7.16
2008-09-08 | 7:31 | 44 | 90.3 | 15.36 | 68 | 9.03 | 7.17
2008-09-08 | 7:46 | 43 | 91.1 | 15.36 | 67 | 9.12 | 7.19
2008-09-08 | 8:01 | 44 | 92.3 | 15.36 | 68 | 9.24 | 7.21
2008-09-08 | 8:16 | 43 | 93.4 | 15.38 | 67 | 9.34 | 7.23
2008-09-08 | 8:31 | 43 | 94.8 | 15.4 | 66 | 9.47 | 7.26
2008-09-08 | 8:46 | 43 | 96 | 15.44 | 67 | 9.59 | 7.29
2008-09-08 | 9:01 | 42 | 97 | 15.5 | 65 | 9.68 | 7.32
2008-09-08 | 9:16 | 44 | 98.2 | 15.56 | 67 | 9.78 | 7.36
2008-09-08 | 9:31 | 44 | 99.4 | 15.64 | 67 | 9.88 | 7.4
2008-09-08 | 9:46 | 44 | 100.7 | 15.74 | 67 | 9.99 | 7.45
2008-09-08 | 10:01 | 44 | 102.5 | 15.86 | 67 | 10.14 | 7.5
2008-09-08 | 10:16 | 44 | 103.8 | 15.98 | 67 | 10.25 | 7.55
2008-09-08 | 10:31 | 44 | 105.2 | 16.12 | 67 | 10.35 | 7.61
2008-09-08 | 10:46 | 43 | 106 | 16.27 | 67 | 10.4 | 7.67
2008-09-08 | 11:01 | 44 | 107.3 | 16.43 | 67 | 10.5 | 7.74
2008-09-08 | 11:16 | 44 | 108.4 | 16.59 | 67 | 10.56 | 7.8
2008-09-08 | 11:31 | 43 | 108.9 | 16.75 | 67 | 10.58 | 7.87
2008-09-08 | 11:46 | 43 | 109.9 | 16.93 | 67 | 10.64 | 7.93
2008-09-08 | 12:01 | 43 | 110.8 | 17.1 | 67 | 10.69 | 8
2008-09-08 | 12:16 | 43 | 111.1 | 17.28 | 67 | 10.67 | 8.06
2008-09-08 | 12:31 | 43 | 111.5 | 17.45 | 67 | 10.67 | 8.11
2008-09-08 | 12:46 | 43 | 111.8 | 17.62 | 67 | 10.67 | 8.16
2008-09-08 | 13:00 | 43 | 112.6 | 17.79 | 67 | 10.71 | 8.2
2008-09-08 | 13:15 | 43 | 113.2 | 17.94 | 67 | 10.73 | 8.25
2008-09-08 | 13:30 | 43 | 113.1 | 18.1 | 67 | 10.69 | 8.28
2008-09-08 | 13:45 | 43 | 113.2 | 18.26 | 67 | 10.66 | 8.32
2008-09-08 | 14:00 | 43 | 113.7 | 18.4 | 67 | 10.68 | 8.36
2008-09-08 | 14:15 | 43 | 114 | 18.53 | 67 | 10.68 | 8.39
2008-09-08 | 14:30 | 43 | 114.1 | 18.64 | 67 | 10.66 | 8.42
2008-09-08 | 14:45 | 43 | 114.1 | 18.75 | 67 | 10.64 | 8.43
2008-09-08 | 15:00 | 43 | 113.5 | 18.85 | 66 | 10.56 | 8.45
2008-09-08 | 15:15 | 43 | 113.6 | 18.93 | 66 | 10.55 | 8.47
2008-09-08 | 15:30 | 43 | 113.5 | 19.02 | 66 | 10.53 | 8.49
2008-09-08 | 15:45 | 43 | 112.9 | 19.08 | 66 | 10.45 | 8.5
2008-09-08 | 16:00 | 43 | 112.7 | 19.12 | 66 | 10.43 | 8.5
2008-09-08 | 16:15 | 43 | 112.1 | 19.17 | 66 | 10.36 | 8.52
2008-09-08 | 16:30 | 43 | 111.7 | 19.18 | 66 | 10.32 | 8.52
2008-09-08 | 16:45 | 43 | 110.9 | 19.18 | 66 | 10.25 | 8.52
2008-09-08 | 17:00 | 43 | 110.6 | 19.17 | 66 | 10.22 | 8.54
2008-09-08 | 17:15 | 43 | 109.9 | 19.14 | 66 | 10.16 | 8.53
2008-09-08 | 17:30 | 43 | 108.7 | 19.1 | 66 | 10.07 | 8.53
2008-09-08 | 17:45 | 43 | 107.7 | 19.03 | 66 | 9.98 | 8.52
2008-09-08 | 18:00 | 43 | 107 | 18.95 | 66 | 9.93 | 8.49
2008-09-08 | 18:15 | 43 | 105.5 | 18.86 | 66 | 9.81 | 8.47
2008-09-08 | 18:30 | 43 | 104.2 | 18.75 | 66 | 9.71 | 8.44
2008-09-08 | 18:45 | 43 | 103.1 | 18.63 | 66 | 9.63 | 8.4
2008-09-08 | 19:00 | 43 | 101.5 | 18.51 | 67 | 9.51 | 8.36
2008-09-08 | 19:15 | 43 | 100.4 | 18.39 | 67 | 9.43 | 8.3
2008-09-08 | 19:30 | 43 | 99.3 | 18.26 | 67 | 9.35 | 8.24
2008-09-08 | 19:45 | 43 | 97.5 | 18.12 | 67 | 9.2 | 8.17
2008-09-08 | 20:00 | 44 | 96.6 | 17.98 | 67 | 9.15 | 8.09
2008-09-08 | 20:15 | 44 | 95.4 | 17.85 | 67 | 9.05 | 8
2008-09-08 | 20:30 | 44 | 94.8 | 17.72 | 67 | 9.02 | 7.91
2008-09-08 | 20:45 | 44 | 93.8 | 17.59 | 67 | 8.95 | 7.82
2008-09-08 | 21:00 | 44 | 93.2 | 17.45 | 68 | 8.92 | 7.74
2008-09-08 | 21:15 | 44 | 92.7 | 17.32 | 68 | 8.9 | 7.66
2008-09-08 | 21:30 | 44 | 92.1 | 17.18 | 68 | 8.87 | 7.6
2008-09-08 | 21:45 | 44 | 91.6 | 17.04 | 68 | 8.84 | 7.53
2008-09-08 | 22:00 | 44 | 91.1 | 16.91 | 68 | 8.82 | 7.48
2008-09-08 | 22:15 | 44 | 90.7 | 16.78 | 68 | 8.81 | 7.43
2008-09-08 | 22:30 | 44 | 90.2 | 16.65 | 68 | 8.78 | 7.4
2008-09-08 | 22:45 | 44 | 90 | 16.54 | 68 | 8.79 | 7.37
2008-09-08 | 23:00 | 44 | 89.6 | 16.42 | 68 | 8.76 | 7.34
2008-09-08 | 23:15 | 44 | 89.2 | 16.32 | 68 | 8.74 | 7.32
2008-09-08 | 23:30 | 44 | 89.3 | 16.23 | 68 | 8.77 | 7.3
2008-09-08 | 23:45 | 44 | 89.3 | 16.14 | 68 | 8.79 | 7.29
2008-09-09 | 0:00 | 44 | 89.3 | 16.06 | 68 | 8.8 | 7.26
2008-09-09 | 0:15 | 44 | 89 | 15.99 | 68 | 8.78 | 7.26
2008-09-09 | 0:30 | 44 | 89.1 | 15.92 | 68 | 8.81 | 7.25
2008-09-09 | 0:45 | 44 | 88.9 | 15.87 | 68 | 8.8 | 7.24
2008-09-09 | 1:01 | 44 | 89.1 | 15.82 | 68 | 8.83 | 7.23
2008-09-09 | 1:16 | 44 | 89.2 | 15.78 | 67 | 8.85 | 7.22
2008-09-09 | 1:31 | 44 | 89.2 | 15.75 | 67 | 8.85 | 7.2
2008-09-09 | 1:46 | 44 | 89 | 15.73 | 67 | 8.83 | 7.21
2008-09-09 | 2:01 | 44 | 89.1 | 15.72 | 67 | 8.85 | 7.21
2008-09-09 | 2:16 | 44 | 89.2 | 15.72 | 67 | 8.86 | 7.2
2008-09-09 | 2:31 | 44 | 89.2 | 15.72 | 67 | 8.85 | 7.2
2008-09-09 | 2:46 | 43 | 89.1 | 15.72 | 67 | 8.85 | 7.19
2008-09-09 | 3:01 | 43 | 88.8 | 15.74 | 67 | 8.82 | 7.19
2008-09-09 | 3:16 | 43 | 89 | 15.75 | 66 | 8.83 | 7.19
2008-09-09 | 3:31 | 43 | 88.7 | 15.75 | 66 | 8.8 | 7.19
2008-09-09 | 3:46 | 43 | 88.5 | 15.76 | 66 | 8.78 | 7.19
2008-09-09 | 4:01 | 43 | 88.4 | 15.76 | 66 | 8.77 | 7.18
2008-09-09 | 4:16 | 43 | 88.3 | 15.77 | 66 | 8.76 | 7.18
2008-09-09 | 4:31 | 43 | 88.3 | 15.75 | 65 | 8.76 | 7.18
2008-09-09 | 4:46 | 42 | 88.2 | 15.75 | 65 | 8.75 | 7.17
2008-09-09 | 5:01 | 42 | 88.2 | 15.74 | 65 | 8.76 | 7.18
2008-09-09 | 5:16 | 42 | 88.4 | 15.73 | 65 | 8.78 | 7.17
2008-09-09 | 5:31 | 42 | 88.3 | 15.72 | 65 | 8.77 | 7.16
2008-09-09 | 5:46 | 42 | 88.5 | 15.69 | 65 | 8.79 | 7.17
2008-09-09 | 6:01 | 42 | 88.2 | 15.66 | 65 | 8.77 | 7.16
2008-09-09 | 6:16 | 42 | 88.2 | 15.63 | 65 | 8.77 | 7.16
2008-09-09 | 6:31 | 42 | 88.1 | 15.6 | 65 | 8.77 | 7.16
2008-09-09 | 6:46 | 42 | 88.6 | 15.57 | 65 | 8.82 | 7.16
2008-09-09 | 7:01 | 43 | 89 | 15.54 | 65 | 8.87 | 7.16
2008-09-09 | 7:16 | 43 | 89.3 | 15.52 | 65 | 8.91 | 7.17
2008-09-09 | 7:31 | 43 | 90.3 | 15.49 | 66 | 9.01 | 7.18
2008-09-09 | 7:46 | 43 | 91.3 | 15.48 | 66 | 9.11 | 7.19
2008-09-09 | 8:01 | 43 | 92.5 | 15.46 | 66 | 9.24 | 7.21
2008-09-09 | 8:16 | 43 | 93.9 | 15.48 | 66 | 9.37 | 7.24
2008-09-09 | 8:31 | 43 | 94.4 | 15.47 | 66 | 9.42 | 7.27
2008-09-09 | 8:46 | 43 | 95 | 15.44 | 66 | 9.48 | 7.29
2008-09-09 | 9:01 | 43 | 96.3 | 15.45 | 66 | 9.61 | 7.3
2008-09-09 | 9:16 | 43 | 97.8 | 15.55 | 66 | 9.74 | 7.34
2008-09-09 | 9:31 | 43 | 99.5 | 15.64 | 66 | 9.9 | 7.39
2008-09-09 | 9:46 | 43 | 100.4 | 15.73 | 66 | 9.97 | 7.44
2008-09-09 | 10:01 | 43 | 101.5 | 15.82 | 66 | 10.06 | 7.48
2008-09-09 | 10:16 | 43 | 103.3 | 15.94 | 66 | 10.21 | 7.54
2008-09-09 | 10:31 | 43 | 104.8 | 16.07 | 66 | 10.32 | 7.6
2008-09-09 | 10:46 | 43 | 106.1 | 16.21 | 66 | 10.42 | 7.67
2008-09-09 | 11:01 | 43 | 106.9 | 16.37 | 66 | 10.47 | 7.74
2008-09-09 | 11:16 | 43 | 108.1 | 16.52 | 66 | 10.55 | 7.8
2008-09-09 | 11:31 | 43 | 109.1 | 16.67 | 66 | 10.62 | 7.87
2008-09-09 | 11:46 | 43 | 109.6 | 16.83 | 66 | 10.63 | 7.93
2008-09-09 | 12:01 | 43 | 110.2 | 17 | 66 | 10.65 | 7.99
2008-09-09 | 12:16 | 43 | 111.2 | 17.17 | 66 | 10.71 | 8.06
2008-09-09 | 12:31 | 43 | 111.9 | 17.35 | 66 | 10.74 | 8.11
2008-09-09 | 12:46 | 43 | 112.4 | 17.52 | 66 | 10.75 | 8.16
2008-09-09 | 13:00 | 43 | 112.4 | 17.69 | 66 | 10.71 | 8.2
2008-09-09 | 13:15 | 43 | 112.6 | 17.84 | 66 | 10.69 | 8.25
2008-09-09 | 13:30 | 43 | 112.9 | 18 | 66 | 10.69 | 8.3
2008-09-09 | 13:45 | 43 | 113.3 | 18.15 | 66 | 10.7 | 8.31
2008-09-09 | 14:00 | 43 | 114.3 | 18.29 | 66 | 10.75 | 8.33
2008-09-09 | 14:15 | 43 | 114.6 | 18.42 | 66 | 10.76 | 8.36
2008-09-09 | 14:30 | 43 | 114.8 | 18.55 | 66 | 10.74 | 8.4
2008-09-09 | 14:45 | 43 | 114.7 | 18.66 | 66 | 10.71 | 8.42
2008-09-09 | 15:00 | 43 | 114.6 | 18.76 | 66 | 10.69 | 8.45
2008-09-09 | 15:15 | 43 | 114.6 | 18.85 | 66 | 10.66 | 8.46
2008-09-09 | 15:30 | 43 | 114.4 | 18.92 | 66 | 10.63 | 8.48
2008-09-09 | 15:45 | 43 | 113.5 | 18.98 | 66 | 10.53 | 8.49
2008-09-09 | 16:00 | 43 | 113.4 | 19.03 | 66 | 10.52 | 8.51
2008-09-09 | 16:15 | 43 | 112.6 | 19.06 | 66 | 10.43 | 8.53
2008-09-09 | 16:30 | 43 | 111.9 | 19.07 | 66 | 10.37 | 8.53
2008-09-09 | 16:45 | 43 | 111.2 | 19.07 | 66 | 10.3 | 8.53
2008-09-09 | 17:00 | 43 | 110.9 | 19.04 | 66 | 10.28 | 8.53
2008-09-09 | 17:15 | 43 | 109.9 | 19 | 66 | 10.2 | 8.53
2008-09-09 | 17:30 | 43 | 109.4 | 18.95 | 66 | 10.16 | 8.51
2008-09-09 | 17:45 | 43 | 108.4 | 18.89 | 66 | 10.07 | 8.51
2008-09-09 | 18:00 | 43 | 107.4 | 18.8 | 66 | 10.01 | 8.49
2008-09-09 | 18:15 | 43 | 106.4 | 18.7 | 66 | 9.93 | 8.47
2008-09-09 | 18:30 | 43 | 105.1 | 18.57 | 66 | 9.83 | 8.43
2008-09-09 | 18:45 | 43 | 103.5 | 18.45 | 66 | 9.71 | 8.4
2008-09-09 | 19:00 | 43 | 101.8 | 18.35 | 66 | 9.57 | 8.35
2008-09-09 | 19:15 | 43 | 100.4 | 18.25 | 67 | 9.45 | 8.3
2008-09-09 | 19:30 | 43 | 99.3 | 18.14 | 67 | 9.38 | 8.24
2008-09-09 | 19:45 | 43 | 97.9 | 18.01 | 67 | 9.27 | 8.17
2008-09-09 | 20:00 | 43 | 97 | 17.88 | 67 | 9.21 | 8.08
2008-09-09 | 20:15 | 44 | 96.2 | 17.74 | 67 | 9.16 | 8.01
2008-09-09 | 20:30 | 44 | 95.2 | 17.61 | 67 | 9.08 | 7.92
2008-09-09 | 20:45 | 44 | 94.2 | 17.47 | 67 | 9.01 | 7.84
2008-09-09 | 21:00 | 44 | 93.7 | 17.33 | 68 | 9 | 7.75
2008-09-09 | 21:15 | 44 | 93.3 | 17.2 | 68 | 8.98 | 7.68
2008-09-09 | 21:30 | 44 | 92.7 | 17.05 | 68 | 8.95 | 7.61
2008-09-09 | 21:45 | 44 | 92.2 | 16.9 | 68 | 8.93 | 7.54
2008-09-09 | 22:00 | 44 | 91.5 | 16.77 | 68 | 8.88 | 7.49
2008-09-09 | 22:15 | 44 | 91.2 | 16.63 | 68 | 8.88 | 7.44
2008-09-09 | 22:30 | 44 | 90.6 | 16.5 | 68 | 8.85 | 7.41
2008-09-09 | 22:45 | 44 | 90.6 | 16.38 | 68 | 8.87 | 7.37
2008-09-09 | 23:00 | 45 | 90.1 | 16.26 | 69 | 8.85 | 7.35
2008-09-09 | 23:15 | 45 | 89.8 | 16.15 | 68 | 8.84 | 7.32
2008-09-09 | 23:30 | 44 | 90 | 16.04 | 68 | 8.87 | 7.3
2008-09-09 | 23:45 | 44 | 90 | 15.94 | 68 | 8.9 | 7.28
2008-09-10 | 0:00 | 44 | 89.7 | 15.84 | 68 | 8.88 | 7.26
2008-09-10 | 0:15 | 44 | 89.5 | 15.74 | 68 | 8.88 | 7.25
2008-09-10 | 0:30 | 44 | 89.4 | 15.66 | 68 | 8.89 | 7.24
2008-09-10 | 0:45 | 44 | 89.7 | 15.58 | 68 | 8.94 | 7.23
2008-09-10 | 1:01 | 44 | 89.8 | 15.51 | 68 | 8.96 | 7.22
2008-09-10 | 1:16 | 44 | 89.5 | 15.44 | 68 | 8.94 | 7.21
2008-09-10 | 1:31 | 44 | 89.8 | 15.39 | 68 | 8.98 | 7.2
2008-09-10 | 1:46 | 44 | 89.9 | 15.34 | 68 | 9 | 7.2
2008-09-10 | 2:01 | 44 | 89.7 | 15.3 | 68 | 8.98 | 7.19
2008-09-10 | 2:16 | 44 | 89.9 | 15.27 | 67 | 9.01 | 7.19
2008-09-10 | 2:31 | 44 | 89.7 | 15.24 | 67 | 9 | 7.18
2008-09-10 | 2:46 | 44 | 89.9 | 15.21 | 67 | 9.02 | 7.18
2008-09-10 | 3:01 | 44 | 90 | 15.2 | 67 | 9.04 | 7.18
2008-09-10 | 3:16 | 44 | 89.7 | 15.18 | 67 | 9.01 | 7.18
2008-09-10 | 3:31 | 44 | 89.6 | 15.15 | 67 | 9 | 7.17
2008-09-10 | 3:46 | 43 | 89.5 | 15.15 | 67 | 9 | 7.17
2008-09-10 | 4:01 | 43 | 89.8 | 15.13 | 67 | 9.03 | 7.17
2008-09-10 | 4:16 | 43 | 89.9 | 15.12 | 67 | 9.04 | 7.17
2008-09-10 | 4:31 | 43 | 90 | 15.11 | 67 | 9.05 | 7.17
2008-09-10 | 4:46 | 43 | 89.7 | 15.11 | 66 | 9.02 | 7.17
2008-09-10 | 5:01 | 43 | 89.8 | 15.09 | 66 | 9.04 | 7.16
2008-09-10 | 5:16 | 43 | 89.6 | 15.08 | 66 | 9.02 | 7.16
2008-09-10 | 5:31 | 43 | 89.7 | 15.06 | 66 | 9.04 | 7.16
2008-09-10 | 5:46 | 43 | 89.9 | 15.05 | 66 | 9.05 | 7.16
2008-09-10 | 6:01 | 43 | 89.9 | 15.04 | 66 | 9.06 | 7.16
2008-09-10 | 6:16 | 43 | 90 | 15.02 | 66 | 9.07 | 7.15
2008-09-10 | 6:31 | 43 | 89.7 | 15 | 66 | 9.04 | 7.15
2008-09-10 | 6:46 | 43 | 89.6 | 14.97 | 66 | 9.04 | 7.15
2008-09-10 | 7:01 | 43 | 90.2 | 14.94 | 66 | 9.11 | 7.16
2008-09-10 | 7:16 | 43 | 90.5 | 14.92 | 66 | 9.14 | 7.17
2008-09-10 | 7:31 | 43 | 91.1 | 14.92 | 66 | 9.21 | 7.18
2008-09-10 | 7:46 | 43 | 92.1 | 14.93 | 66 | 9.3 | 7.19
2008-09-10 | 8:01 | 43 | 93.6 | 14.95 | 67 | 9.45 | 7.21
2008-09-10 | 8:16 | 43 | 94.6 | 14.97 | 67 | 9.55 | 7.24
2008-09-10 | 8:31 | 43 | 95.7 | 15.01 | 67 | 9.65 | 7.27
2008-09-10 | 8:46 | 43 | 96.9 | 15.06 | 67 | 9.76 | 7.29
2008-09-10 | 9:01 | 43 | 98.6 | 15.12 | 67 | 9.91 | 7.33
2008-09-10 | 9:16 | 43 | 99.7 | 15.2 | 67 | 10.01 | 7.37
2008-09-10 | 9:31 | 43 | 101.3 | 15.29 | 67 | 10.15 | 7.41
2008-09-10 | 9:46 | 43 | 102.7 | 15.39 | 67 | 10.27 | 7.45
2008-09-10 | 10:01 | 43 | 103.7 | 15.5 | 67 | 10.35 | 7.5
2008-09-10 | 10:16 | 43 | 105 | 15.64 | 67 | 10.45 | 7.56
2008-09-10 | 10:31 | 43 | 106.2 | 15.78 | 67 | 10.53 | 7.61
2008-09-10 | 10:46 | 43 | 107.4 | 15.92 | 67 | 10.62 | 7.68
2008-09-10 | 11:01 | 43 | 108.5 | 16.07 | 67 | 10.69 | 7.74
2008-09-10 | 11:16 | 43 | 109.7 | 16.23 | 67 | 10.77 | 7.81
2008-09-10 | 11:31 | 43 | 111.1 | 16.39 | 67 | 10.88 | 7.87
2008-09-10 | 11:46 | 43 | 112.2 | 16.56 | 67 | 10.94 | 7.94
2008-09-10 | 12:01 | 43 | 113.3 | 16.74 | 67 | 11.01 | 7.99
2008-09-10 | 12:16 | 43 | 113.5 | 16.91 | 67 | 10.99 | 8.06
2008-09-10 | 12:31 | 43 | 114.3 | 17.08 | 67 | 11.03 | 8.11
2008-09-10 | 12:46 | 43 | 114.9 | 17.24 | 67 | 11.05 | 8.15
2008-09-10 | 13:00 | 43 | 114.9 | 17.39 | 67 | 11.01 | 8.21
2008-09-10 | 13:15 | 43 | 115.5 | 17.55 | 67 | 11.04 | 8.25
2008-09-10 | 13:30 | 43 | 115.9 | 17.69 | 67 | 11.04 | 8.28
2008-09-10 | 13:45 | 43 | 115.6 | 17.82 | 67 | 10.98 | 8.32
2008-09-10 | 14:00 | 43 | 115.6 | 17.94 | 67 | 10.96 | 8.36
2008-09-10 | 14:15 | 43 | 115.6 | 18.05 | 67 | 10.93 | 8.39
2008-09-10 | 14:30 | 43 | 115.6 | 18.15 | 67 | 10.91 | 8.42
2008-09-10 | 14:45 | 43 | 116 | 18.23 | 67 | 10.93 | 8.44
2008-09-10 | 15:00 | 43 | 115.5 | 18.31 | 67 | 10.87 | 8.46
2008-09-10 | 15:15 | 44 | 115.3 | 18.38 | 67 | 10.83 | 8.48
2008-09-10 | 15:30 | 44 | 115.5 | 18.42 | 67 | 10.84 | 8.49
2008-09-10 | 15:45 | 44 | 115.4 | 18.48 | 67 | 10.81 | 8.51
2008-09-10 | 16:00 | 44 | 114.5 | 18.51 | 67 | 10.73 | 8.52
2008-09-10 | 16:15 | 44 | 114.3 | 18.52 | 67 | 10.71 | 8.54
2008-09-10 | 16:30 | 43 | 113.8 | 18.53 | 67 | 10.66 | 8.54
2008-09-10 | 16:45 | 43 | 113.5 | 18.51 | 67 | 10.63 | 8.54
2008-09-10 | 17:00 | 43 | 112.2 | 18.5 | 67 | 10.52 | 8.54
2008-09-10 | 17:15 | 43 | 111.9 | 18.46 | 67 | 10.49 | 8.54
2008-09-10 | 17:30 | 43 | 110.6 | 18.4 | 67 | 10.39 | 8.53
2008-09-10 | 17:45 | 43 | 109.9 | 18.32 | 67 | 10.34 | 8.53
2008-09-10 | 18:00 | 43 | 109.2 | 18.22 | 67 | 10.29 | 8.51
2008-09-10 | 18:15 | 43 | 107.4 | 18.12 | 67 | 10.15 | 8.5
2008-09-10 | 18:30 | 43 | 106.1 | 17.99 | 67 | 10.05 | 8.46
2008-09-10 | 18:45 | 44 | 104.4 | 17.87 | 67 | 9.91 | 8.42
2008-09-10 | 19:00 | 44 | 103.7 | 17.75 | 67 | 9.86 | 8.38
2008-09-10 | 19:15 | 44 | 102.5 | 17.63 | 67 | 9.78 | 8.34
2008-09-10 | 19:30 | 44 | 100.8 | 17.49 | 67 | 9.65 | 8.28
2008-09-10 | 19:45 | 44 | 99.9 | 17.34 | 67 | 9.59 | 8.2
2008-09-10 | 20:00 | 44 | 98.8 | 17.2 | 67 | 9.51 | 8.12
2008-09-10 | 20:15 | 44 | 98 | 17.07 | 68 | 9.45 | 8.04
2008-09-10 | 20:30 | 44 | 97.2 | 16.94 | 68 | 9.4 | 7.95
2008-09-10 | 20:45 | 44 | 96.2 | 16.8 | 68 | 9.33 | 7.88
2008-09-10 | 21:00 | 44 | 95.7 | 16.67 | 68 | 9.32 | 7.79
2008-09-10 | 21:15 | 44 | 95.1 | 16.52 | 68 | 9.29 | 7.71
2008-09-10 | 21:30 | 44 | 94.7 | 16.39 | 68 | 9.27 | 7.64
2008-09-10 | 21:45 | 44 | 93.8 | 16.24 | 68 | 9.21 | 7.57
2008-09-10 | 22:00 | 45 | 93.3 | 16.1 | 69 | 9.19 | 7.52
2008-09-10 | 22:15 | 45 | 92.9 | 15.95 | 69 | 9.17 | 7.47
2008-09-10 | 22:30 | 45 | 92.8 | 15.81 | 69 | 9.2 | 7.43
2008-09-10 | 22:45 | 45 | 92.4 | 15.68 | 69 | 9.18 | 7.4
2008-09-10 | 23:00 | 45 | 92 | 15.55 | 69 | 9.17 | 7.37
2008-09-10 | 23:15 | 45 | 92.2 | 15.43 | 69 | 9.21 | 7.35
2008-09-10 | 23:30 | 45 | 92.1 | 15.3 | 69 | 9.23 | 7.32
2008-09-10 | 23:45 | 45 | 92.1 | 15.2 | 69 | 9.24 | 7.3
2008-09-11 | 0:00 | 45 | 91.7 | 15.1 | 69 | 9.23 | 7.29
2008-09-11 | 0:15 | 45 | 91.5 | 15.01 | 69 | 9.23 | 7.28
2008-09-11 | 0:30 | 45 | 91.5 | 14.92 | 69 | 9.24 | 7.27
2008-09-11 | 0:45 | 44 | 91.8 | 14.83 | 68 | 9.29 | 7.25
2008-09-11 | 1:01 | 44 | 92 | 14.76 | 68 | 9.32 | 7.25
2008-09-11 | 1:16 | 44 | 92.1 | 14.7 | 68 | 9.35 | 7.23
2008-09-11 | 1:31 | 44 | 92.2 | 14.64 | 68 | 9.37 | 7.23
2008-09-11 | 1:46 | 44 | 92.2 | 14.6 | 68 | 9.38 | 7.23
2008-09-11 | 2:01 | 44 | 91.9 | 14.56 | 68 | 9.35 | 7.22
2008-09-11 | 2:16 | 44 | 92.1 | 14.53 | 68 | 9.38 | 7.21
2008-09-11 | 2:31 | 44 | 92.2 | 14.52 | 68 | 9.39 | 7.21
2008-09-11 | 2:46 | 44 | 91.9 | 14.51 | 68 | 9.36 | 7.21
2008-09-11 | 3:01 | 44 | 92 | 14.49 | 67 | 9.38 | 7.2
2008-09-11 | 3:16 | 44 | 91.8 | 14.48 | 67 | 9.36 | 7.19
2008-09-11 | 3:31 | 44 | 92 | 14.47 | 67 | 9.38 | 7.2
2008-09-11 | 3:46 | 43 | 92.1 | 14.46 | 67 | 9.39 | 7.19
2008-09-11 | 4:01 | 43 | 91.7 | 14.45 | 67 | 9.36 | 7.2
2008-09-11 | 4:16 | 43 | 91.9 | 14.44 | 66 | 9.38 | 7.19
2008-09-11 | 4:31 | 43 | 91.6 | 14.43 | 66 | 9.36 | 7.19
2008-09-11 | 4:46 | 43 | 91.5 | 14.42 | 66 | 9.34 | 7.18
2008-09-11 | 5:01 | 43 | 91.8 | 14.4 | 66 | 9.38 | 7.18
2008-09-11 | 5:16 | 43 | 91.5 | 14.39 | 66 | 9.35 | 7.18
2008-09-11 | 5:31 | 43 | 91.8 | 14.36 | 66 | 9.38 | 7.18
2008-09-11 | 5:46 | 43 | 91.5 | 14.35 | 66 | 9.36 | 7.18
2008-09-11 | 6:01 | 43 | 91.7 | 14.32 | 66 | 9.39 | 7.17
2008-09-11 | 6:16 | 43 | 91.5 | 14.28 | 66 | 9.37 | 7.17
2008-09-11 | 6:31 | 43 | 91.4 | 14.25 | 66 | 9.37 | 7.17
2008-09-11 | 6:46 | 43 | 91.8 | 14.23 | 66 | 9.42 | 7.17
2008-09-11 | 7:01 | 43 | 91.9 | 14.2 | 66 | 9.43 | 7.17
2008-09-11 | 7:16 | 43 | 92.6 | 14.16 | 66 | 9.51 | 7.19
2008-09-11 | 7:31 | 43 | 93 | 14.13 | 66 | 9.55 | 7.19
2008-09-11 | 7:46 | 43 | 94.2 | 14.13 | 66 | 9.68 | 7.2
2008-09-11 | 8:01 | 43 | 95.4 | 14.13 | 66 | 9.8 | 7.22
2008-09-11 | 8:16 | 43 | 96.2 | 14.16 | 66 | 9.88 | 7.25
2008-09-11 | 8:31 | 43 | 97.3 | 14.19 | 67 | 9.98 | 7.28
2008-09-11 | 8:46 | 43 | 98.7 | 14.23 | 67 | 10.12 | 7.3
2008-09-11 | 9:01 | 43 | 100 | 14.28 | 67 | 10.24 | 7.33
2008-09-11 | 9:16 | 43 | 101 | 14.35 | 67 | 10.33 | 7.37
2008-09-11 | 9:31 | 43 | 102.2 | 14.45 | 67 | 10.42 | 7.41
2008-09-11 | 9:46 | 43 | 103.8 | 14.55 | 67 | 10.57 | 7.45
2008-09-11 | 10:01 | 43 | 105.1 | 14.67 | 67 | 10.67 | 7.5
2008-09-11 | 10:16 | 43 | 106.4 | 14.8 | 67 | 10.78 | 7.55
2008-09-11 | 10:31 | 43 | 107.7 | 14.94 | 67 | 10.87 | 7.6
2008-09-11 | 10:46 | 43 | 108.9 | 15.08 | 67 | 10.96 | 7.67
2008-09-11 | 11:01 | 43 | 110 | 15.23 | 67 | 11.04 | 7.72
2008-09-11 | 11:16 | 43 | 111 | 15.39 | 67 | 11.1 | 7.78
2008-09-11 | 11:31 | 43 | 112.1 | 15.56 | 67 | 11.16 | 7.84
2008-09-11 | 11:46 | 43 | 112.5 | 15.72 | 67 | 11.17 | 7.9
2008-09-11 | 12:01 | 43 | 113.6 | 15.89 | 67 | 11.24 | 7.97
2008-09-11 | 12:16 | 44 | 113.8 | 16.04 | 67 | 11.22 | 8.02
2008-09-11 | 12:31 | 43 | 114.7 | 16.21 | 66 | 11.27 | 8.07
2008-09-11 | 12:46 | 43 | 115.2 | 16.38 | 67 | 11.28 | 8.12
2008-09-11 | 13:00 | 43 | 115.2 | 16.54 | 67 | 11.24 | 8.17
2008-09-11 | 13:15 | 43 | 115.8 | 16.71 | 67 | 11.26 | 8.21
2008-09-11 | 13:30 | 43 | 116.2 | 16.85 | 67 | 11.26 | 8.26
2008-09-11 | 13:45 | 43 | 116 | 16.99 | 66 | 11.21 | 8.29
2008-09-11 | 14:00 | 43 | 116 | 17.13 | 67 | 11.18 | 8.33
2008-09-11 | 14:15 | 43 | 116.1 | 17.24 | 66 | 11.16 | 8.35
2008-09-11 | 14:30 | 43 | 116.1 | 17.35 | 66 | 11.14 | 8.39
2008-09-11 | 14:45 | 43 | 116 | 17.45 | 66 | 11.11 | 8.41
2008-09-11 | 15:00 | 43 | 116 | 17.52 | 66 | 11.09 | 8.43
2008-09-11 | 15:15 | 43 | 116.4 | 17.59 | 66 | 11.11 | 8.45
2008-09-11 | 15:30 | 43 | 116.3 | 17.65 | 66 | 11.09 | 8.47
2008-09-11 | 15:45 | 43 | 115.6 | 17.71 | 66 | 11.01 | 8.49
2008-09-11 | 16:00 | 43 | 115.1 | 17.75 | 66 | 10.95 | 8.51
2008-09-11 | 16:15 | 43 | 114.6 | 17.77 | 66 | 10.9 | 8.52
2008-09-11 | 16:30 | 43 | 114.6 | 17.77 | 66 | 10.9 | 8.53
2008-09-11 | 16:45 | 43 | 113.7 | 17.78 | 66 | 10.81 | 8.53
2008-09-11 | 17:00 | 43 | 113 | 17.75 | 66 | 10.75 | 8.53
2008-09-11 | 17:15 | 43 | 112.1 | 17.72 | 66 | 10.67 | 8.53
2008-09-11 | 17:30 | 43 | 111.8 | 17.68 | 66 | 10.65 | 8.52
2008-09-11 | 17:45 | 43 | 111 | 17.61 | 66 | 10.59 | 8.51
2008-09-11 | 18:00 | 43 | 110 | 17.53 | 66 | 10.51 | 8.49
2008-09-11 | 18:15 | 43 | 108.5 | 17.43 | 66 | 10.39 | 8.47
2008-09-11 | 18:30 | 43 | 107 | 17.33 | 66 | 10.27 | 8.44
2008-09-11 | 18:45 | 43 | 106 | 17.21 | 66 | 10.2 | 8.4
2008-09-11 | 19:00 | 43 | 104.8 | 17.1 | 67 | 10.11 | 8.36
2008-09-11 | 19:15 | 43 | 103.6 | 16.98 | 67 | 10.01 | 8.3
2008-09-11 | 19:30 | 43 | 101.9 | 16.85 | 67 | 9.88 | 8.23
2008-09-11 | 19:45 | 43 | 100.7 | 16.72 | 67 | 9.79 | 8.17
2008-09-11 | 20:00 | 43 | 99.9 | 16.58 | 67 | 9.74 | 8.08
2008-09-11 | 20:15 | 44 | 99.2 | 16.46 | 67 | 9.7 | 8
2008-09-11 | 20:30 | 44 | 98 | 16.32 | 67 | 9.6 | 7.92
2008-09-11 | 20:45 | 44 | 97.2 | 16.17 | 67 | 9.56 | 7.83
2008-09-11 | 21:00 | 44 | 96.4 | 16.03 | 68 | 9.51 | 7.75
2008-09-11 | 21:15 | 44 | 96.2 | 15.88 | 68 | 9.51 | 7.67
2008-09-11 | 21:30 | 44 | 95.7 | 15.74 | 68 | 9.5 | 7.6
2008-09-11 | 21:45 | 44 | 95 | 15.6 | 68 | 9.46 | 7.55
2008-09-11 | 22:00 | 44 | 94.8 | 15.46 | 68 | 9.46 | 7.5
2008-09-11 | 22:15 | 44 | 94.4 | 15.32 | 68 | 9.45 | 7.46
2008-09-11 | 22:30 | 44 | 94.1 | 15.19 | 68 | 9.45 | 7.42
2008-09-11 | 22:45 | 44 | 93.6 | 15.07 | 68 | 9.42 | 7.39
2008-09-11 | 23:00 | 44 | 93.6 | 14.95 | 68 | 9.45 | 7.36
2008-09-11 | 23:15 | 45 | 93.1 | 14.84 | 68 | 9.42 | 7.34
2008-09-11 | 23:30 | 44 | 93.2 | 14.75 | 68 | 9.45 | 7.32
2008-09-11 | 23:45 | 44 | 93 | 14.66 | 68 | 9.44 | 7.3
2008-09-12 | 0:00 | 44 | 93.1 | 14.57 | 68 | 9.47 | 7.29
2008-09-12 | 0:15 | 44 | 93.1 | 14.49 | 68 | 9.49 | 7.28
2008-09-12 | 0:30 | 44 | 93.1 | 14.41 | 68 | 9.5 | 7.27
2008-09-12 | 0:45 | 44 | 92.7 | 14.34 | 68 | 9.48 | 7.26
2008-09-12 | 1:01 | 44 | 92.9 | 14.27 | 68 | 9.52 | 7.25
2008-09-12 | 1:16 | 44 | 92.9 | 14.2 | 68 | 9.53 | 7.24
2008-09-12 | 1:31 | 44 | 92.9 | 14.15 | 68 | 9.55 | 7.23
2008-09-12 | 1:46 | 44 | 92.9 | 14.09 | 68 | 9.55 | 7.23
2008-09-12 | 2:01 | 44 | 92.9 | 14.04 | 68 | 9.56 | 7.23
2008-09-12 | 2:16 | 44 | 92.8 | 14 | 68 | 9.57 | 7.21
2008-09-12 | 2:31 | 44 | 92.8 | 13.96 | 68 | 9.57 | 7.22
2008-09-12 | 2:46 | 44 | 92.5 | 13.92 | 68 | 9.55 | 7.21
2008-09-12 | 3:01 | 44 | 92.7 | 13.9 | 68 | 9.58 | 7.21
2008-09-12 | 3:16 | 44 | 92.5 | 13.88 | 68 | 9.55 | 7.21
2008-09-12 | 3:31 | 44 | 92.4 | 13.86 | 67 | 9.55 | 7.2
2008-09-12 | 3:46 | 44 | 92.6 | 13.83 | 67 | 9.57 | 7.2
2008-09-12 | 4:01 | 44 | 92.6 | 13.82 | 67 | 9.58 | 7.2
2008-09-12 | 4:16 | 44 | 92.4 | 13.82 | 67 | 9.56 | 7.19
2008-09-12 | 4:31 | 44 | 92.2 | 13.81 | 67 | 9.54 | 7.2
2008-09-12 | 4:46 | 44 | 92.5 | 13.81 | 67 | 9.57 | 7.19
2008-09-12 | 5:01 | 44 | 92.6 | 13.81 | 67 | 9.58 | 7.19
2008-09-12 | 5:16 | 44 | 92.2 | 13.79 | 67 | 9.55 | 7.19
2008-09-12 | 5:31 | 44 | 92.1 | 13.8 | 67 | 9.54 | 7.19
2008-09-12 | 5:46 | 44 | 92.4 | 13.8 | 67 | 9.57 | 7.18
2008-09-12 | 6:01 | 44 | 92.5 | 13.78 | 67 | 9.58 | 7.18
2008-09-12 | 6:16 | 44 | 92.2 | 13.78 | 67 | 9.55 | 7.19
2008-09-12 | 6:31 | 44 | 92.5 | 13.77 | 67 | 9.58 | 7.18
2008-09-12 | 6:46 | 44 | 92.6 | 13.76 | 67 | 9.6 | 7.18
2008-09-12 | 7:01 | 44 | 93 | 13.75 | 67 | 9.63 | 7.18
2008-09-12 | 7:16 | 44 | 93.4 | 13.74 | 67 | 9.68 | 7.19
2008-09-12 | 7:31 | 44 | 94 | 13.75 | 67 | 9.74 | 7.2
2008-09-12 | 7:46 | 44 | 94.6 | 13.76 | 67 | 9.8 | 7.22
2008-09-12 | 8:01 | 44 | 95.6 | 13.8 | 67 | 9.9 | 7.23
2008-09-12 | 8:16 | 44 | 97.1 | 13.84 | 67 | 10.04 | 7.26
2008-09-12 | 8:31 | 44 | 98.3 | 13.89 | 67 | 10.15 | 7.28
2008-09-12 | 8:46 | 44 | 99.4 | 13.95 | 67 | 10.26 | 7.32
2008-09-12 | 9:01 | 44 | 100.7 | 14.04 | 67 | 10.36 | 7.34
2008-09-12 | 9:16 | 44 | 101.9 | 14.13 | 67 | 10.47 | 7.38
2008-09-12 | 9:31 | 44 | 102.9 | 14.24 | 67 | 10.55 | 7.42
2008-09-12 | 9:46 | 44 | 104.5 | 14.36 | 67 | 10.68 | 7.46
2008-09-12 | 10:01 | 44 | 105.7 | 14.49 | 67 | 10.78 | 7.5
2008-09-12 | 10:16 | 44 | 106.6 | 14.63 | 67 | 10.84 | 7.55
2008-09-12 | 10:31 | 44 | 108.1 | 14.79 | 67 | 10.95 | 7.6
2008-09-12 | 10:46 | 44 | 109.5 | 14.96 | 67 | 11.05 | 7.67
2008-09-12 | 11:01 | 44 | 110.3 | 15.13 | 67 | 11.09 | 7.73
2008-09-12 | 11:16 | 44 | 111.3 | 15.3 | 68 | 11.15 | 7.79
2008-09-12 | 11:31 | 44 | 112.3 | 15.48 | 67 | 11.2 | 7.86
2008-09-12 | 11:46 | 44 | 113.2 | 15.67 | 67 | 11.25 | 7.92
2008-09-12 | 12:01 | 44 | 114.4 | 15.85 | 67 | 11.33 | 7.98
2008-09-12 | 12:16 | 44 | 114.9 | 16.03 | 67 | 11.33 | 8.04
2008-09-12 | 12:31 | 44 | 115.8 | 16.21 | 67 | 11.38 | 8.1
2008-09-12 | 12:46 | 44 | 115.9 | 16.38 | 67 | 11.34 | 8.15
2008-09-12 | 13:00 | 44 | 116.6 | 16.55 | 67 | 11.38 | 8.2
2008-09-12 | 13:15 | 44 | 116.6 | 16.71 | 67 | 11.34 | 8.25
2008-09-12 | 13:30 | 44 | 117.2 | 16.88 | 67 | 11.36 | 8.29
2008-09-12 | 13:45 | 44 | 117.6 | 17.02 | 67 | 11.36 | 8.33
2008-09-12 | 14:00 | 44 | 117.7 | 17.17 | 67 | 11.34 | 8.36
2008-09-12 | 14:15 | 44 | 117.9 | 17.28 | 67 | 11.33 | 8.39
2008-09-12 | 14:30 | 44 | 117.5 | 17.39 | 67 | 11.26 | 8.42
2008-09-12 | 14:45 | 44 | 117.3 | 17.5 | 67 | 11.22 | 8.44