## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude
Number | 541-382-6103 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AB against 556, NIST, & Titration
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-14 |  | 2008-07-14 |  | A
Time | Computer set to atomic time |  | 10:42 |  | 10:42 |  | A
Temp | VWR NIST traceable surface water |  | 7.6 |  | 7.7 |  | A
DO | Winkler titration of surface water |  | 11.6 |  | 11.29 |  | B
SpC | In Situ µS/mL |  | 24 |  | 24 |  | A
Ph | In Situ (Aginst AA) pH 7 |  | 7.23 |  | 8.24 |  | Fail
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2008-07-14 15:00 | 16.69 | 6.68 | 9.78 | 100.6 | 35 | 25
Sonde | 2008-07-14 15:00 | 16.75 | 7.35 | 9.66 | 99.5 | 35 | 23
Grade |  | A | Fail | A |  | A
Audit 2 | 2008-07-18 14:30 | 17.24 | 6.51 | 9.76 | 101.5 | 40 | 28
Sonde | 7/18/14:30 | 17.33 | 7.3 | 9.56 | 99.6 | 39 | 25
Grade |  | A | Fail | A |  | A
Audit 3 | 2008-07-22 14:15 | 18.9 | 6.93 | 9.25 | 99.6 | 41 | 29
Sonde | 2008-07-22 14:15 | 18.95 | 7.44 | 9.25 | 100.2 | 41 | 26
Grade |  | A | Fail | A |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AB against 556, NIST, & Titration
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-22 |  | 2008-07-22 |  | A
Time | Computer set to atomic time |  | 14:15 |  | 14:15 |  | A
Temp | VWR NIST traceable surface water |  | 18.7 |  | 18.95 |  | A
DO | Winkler titration of surface water |  | 9.1 |  | 9.3 |  | A
SpC | In Situ µS/mL |  | 41 |  | 41 |  | A
Ph | pH In Situ |  | 6.93 |  | 7.44 |  | Fail
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Sonde AB File Name: 071408AB pH was checked against 556 and Sonde value should be trusted |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS & Sonde AB
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-14 |  | 2008-07-14 |  | A
Time | Computer set to atomic time |  | 10:52 |  | 10:52 |  | A
Temp | VWR NIST traceable surface water |  | 7.7 |  | 7.82 |  | A
DO | Winkler titration of surface water |  | 11.6 |  | 11.87 |  | A
SpC | 148 µS/mL |  | 24 |  | 24 |  | A
Ph | pH 7 standard solution |  | 7.81 |  | 4.95 |  | Fail
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-07-14 15:00 | 16.69 | 6.68 | 9.78 | 100.6 | 35 | 25
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-14 |  | 2008-07-14 |  | A
Time | Computer set to atomic time |  | 16:30 |  | 16:30 |  | A
Temp | VWR NIST traceable surface water |  | 22.1 |  | 22.05 |  | A
DO | Winkler titration of surface water |  | 9.39 |  | 9.22 |  | A
SpC | 148 µS/mL |  | 55 |  | 56 |  | A
Ph | pH 7 standard solution |  | 8.16 |  | 8.3 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 File: 071408AB |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS & Sonde AB
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-17 |  | 2008-07-17 |  | A
Time | Computer set to atomic time |  | 14:15 |  | 14:15 |  | A
Temp | VWR NIST traceable surface water |  | 17.6 |  | 17.81 |  | A
DO | Winkler titration of surface water |  | 11.5 |  | 11.2 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 142 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.19 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-07-18 14:30 | 17.24 | 6.51 | 9.76 | 101.5 | 40 | 28
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-21 |  | 2008-07-21 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | VWR NIST traceable surface water |  | 18.1 |  | 18.34 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 9.45 |  | B
SpC | 148 µS/mL |  | 149.3 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.21 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Lesley Jones |  | Latitude |  | 0
Number | 541-382-6103 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS & Sonde AB
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-21 |  | 2008-07-21 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | VWR NIST traceable surface water |  | 18.1 |  | 18.34 |  | A
DO | Winkler titration of surface water |  | 8.9 |  | 9.45 |  | B
SpC | 148 µS/mL |  | 149.3 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.21 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-07-22 14:15 | 18.9 | 6.93 | 9.25 | 99.6 | 41 | 29
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-07-22 |  | 2008-07-22 |  | A
Time | Computer set to atomic time |  | 14:15 |  | 14:15 |  | A
Temp | VWR NIST traceable surface water |  | 18.7 |  | 18.9 |  | A
DO | Winkler titration of surface water |  | 9.1 |  | 9.25 |  | A
SpC | 148 µS/mL |  | 149.3 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.16 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
DateTime | Time | TDS | DO% | Temp | SpCond | DO Conc | pH
M/D/Y |  | mg/L | % | ˚C | uS/cm | mg/L
2008-07-14 | 15:00 | 23 | 99.5 | 16.75 | 35 | 9.66 | 7.35
2008-07-14 | 15:15 | 23 | 98.4 | 16.85 | 35 | 9.54 | 7.24
2008-07-14 | 15:30 | 23 | 97.4 | 16.9 | 35 | 9.44 | 7.21
2008-07-14 | 15:45 | 23 | 97 | 16.9 | 35 | 9.39 | 7.19
2008-07-14 | 16:00 | 23 | 97.1 | 16.98 | 35 | 9.39 | 7.18
2008-07-14 | 16:15 | 23 | 97.1 | 17.15 | 35 | 9.36 | 7.17
2008-07-14 | 16:30 | 23 | 97.1 | 17.32 | 35 | 9.32 | 7.17
2008-07-14 | 16:45 | 23 | 96.9 | 17.38 | 35 | 9.29 | 7.15
2008-07-14 | 17:00 | 23 | 96.3 | 17.43 | 35 | 9.23 | 7.16
2008-07-14 | 17:15 | 23 | 96.4 | 17.39 | 35 | 9.24 | 7.15
2008-07-14 | 17:30 | 23 | 96 | 17.37 | 35 | 9.21 | 7.14
2008-07-14 | 17:45 | 23 | 95.8 | 17.33 | 35 | 9.2 | 7.14
2008-07-14 | 18:00 | 23 | 96.2 | 17.31 | 35 | 9.24 | 7.13
2008-07-14 | 18:15 | 23 | 95.9 | 17.35 | 35 | 9.2 | 7.14
2008-07-14 | 18:30 | 23 | 95.8 | 17.4 | 35 | 9.18 | 7.14
2008-07-14 | 18:45 | 23 | 95.7 | 17.42 | 35 | 9.17 | 7.15
2008-07-14 | 19:00 | 23 | 95.9 | 17.43 | 35 | 9.18 | 7.14
2008-07-14 | 19:15 | 23 | 95.5 | 17.39 | 36 | 9.16 | 7.15
2008-07-14 | 19:30 | 23 | 95.9 | 17.3 | 36 | 9.21 | 7.15
2008-07-14 | 19:45 | 23 | 96 | 17.17 | 36 | 9.24 | 7.14
2008-07-14 | 20:00 | 23 | 95.8 | 17.1 | 36 | 9.24 | 7.14
2008-07-14 | 20:15 | 23 | 95.7 | 17.03 | 36 | 9.24 | 7.14
2008-07-14 | 20:30 | 23 | 96 | 16.92 | 36 | 9.29 | 7.14
2008-07-14 | 20:45 | 23 | 95.7 | 16.82 | 36 | 9.29 | 7.14
2008-07-14 | 21:00 | 23 | 95.6 | 16.69 | 36 | 9.31 | 7.13
2008-07-14 | 21:15 | 23 | 95.6 | 16.58 | 36 | 9.32 | 7.13
2008-07-14 | 21:30 | 24 | 96 | 16.48 | 36 | 9.38 | 7.13
2008-07-14 | 21:45 | 24 | 96.2 | 16.38 | 36 | 9.42 | 7.12
2008-07-14 | 22:00 | 24 | 96.3 | 16.27 | 36 | 9.45 | 7.12
2008-07-14 | 22:15 | 24 | 96.3 | 16.18 | 36 | 9.47 | 7.12
2008-07-14 | 22:30 | 24 | 96 | 16.07 | 37 | 9.46 | 7.12
2008-07-14 | 22:45 | 24 | 96.3 | 15.97 | 37 | 9.51 | 7.12
2008-07-14 | 23:00 | 24 | 96.1 | 15.87 | 37 | 9.51 | 7.12
2008-07-14 | 23:15 | 24 | 96.3 | 15.78 | 37 | 9.55 | 7.12
2008-07-14 | 23:30 | 24 | 96.5 | 15.68 | 37 | 9.59 | 7.11
2008-07-14 | 23:45 | 24 | 96.6 | 15.57 | 37 | 9.62 | 7.11
2008-07-15 | 0:00 | 24 | 96.3 | 15.44 | 37 | 9.62 | 7.11
2008-07-15 | 0:15 | 24 | 96.3 | 15.3 | 37 | 9.64 | 7.11
2008-07-15 | 0:30 | 24 | 96.6 | 15.18 | 37 | 9.71 | 7.12
2008-07-15 | 0:45 | 24 | 96.8 | 15.07 | 37 | 9.74 | 7.12
2008-07-15 | 1:01 | 24 | 96.9 | 14.97 | 37 | 9.77 | 7.12
2008-07-15 | 1:16 | 24 | 96.9 | 14.86 | 37 | 9.8 | 7.11
2008-07-15 | 1:31 | 24 | 97 | 14.75 | 37 | 9.83 | 7.11
2008-07-15 | 1:46 | 24 | 97 | 14.65 | 37 | 9.86 | 7.1
2008-07-15 | 2:01 | 24 | 97.1 | 14.54 | 37 | 9.89 | 7.1
2008-07-15 | 2:16 | 24 | 96.8 | 14.43 | 37 | 9.88 | 7.1
2008-07-15 | 2:31 | 24 | 96.6 | 14.31 | 36 | 9.89 | 7.09
2008-07-15 | 2:46 | 24 | 97 | 14.18 | 36 | 9.95 | 7.09
2008-07-15 | 3:01 | 23 | 97.1 | 14.05 | 36 | 10 | 7.1
2008-07-15 | 3:16 | 23 | 96.9 | 13.92 | 36 | 10 | 7.08
2008-07-15 | 3:31 | 23 | 96.8 | 13.78 | 36 | 10.02 | 7.08
2008-07-15 | 3:46 | 23 | 97.1 | 13.65 | 36 | 10.09 | 7.07
2008-07-15 | 4:01 | 23 | 96.9 | 13.53 | 35 | 10.1 | 7.06
2008-07-15 | 4:16 | 23 | 96.9 | 13.42 | 35 | 10.11 | 7.06
2008-07-15 | 4:31 | 23 | 96.8 | 13.32 | 35 | 10.13 | 7.06
2008-07-15 | 4:46 | 23 | 96.8 | 13.23 | 35 | 10.15 | 7.06
2008-07-15 | 5:01 | 23 | 96.9 | 13.14 | 35 | 10.18 | 7.05
2008-07-15 | 5:16 | 22 | 97.3 | 13.04 | 34 | 10.25 | 7.05
2008-07-15 | 5:31 | 22 | 97.5 | 12.93 | 34 | 10.29 | 7.05
2008-07-15 | 5:46 | 22 | 97.6 | 12.84 | 34 | 10.33 | 7.04
2008-07-15 | 6:01 | 22 | 97.4 | 12.74 | 34 | 10.32 | 7.04
2008-07-15 | 6:16 | 22 | 97.3 | 12.64 | 34 | 10.33 | 7.04
2008-07-15 | 6:31 | 22 | 97.3 | 12.55 | 34 | 10.36 | 7.04
2008-07-15 | 6:46 | 22 | 97.4 | 12.47 | 34 | 10.39 | 7.04
2008-07-15 | 7:01 | 22 | 97.5 | 12.39 | 34 | 10.41 | 7.05
2008-07-15 | 7:16 | 22 | 98 | 12.33 | 34 | 10.49 | 7.05
2008-07-15 | 7:31 | 22 | 98.3 | 12.27 | 34 | 10.53 | 7.05
2008-07-15 | 7:46 | 22 | 98.5 | 12.22 | 34 | 10.56 | 7.06
2008-07-15 | 8:01 | 22 | 98.7 | 12.18 | 34 | 10.59 | 7.06
2008-07-15 | 8:16 | 22 | 98.4 | 12.14 | 34 | 10.58 | 7.06
2008-07-15 | 8:31 | 22 | 98.8 | 12.1 | 34 | 10.62 | 7.08
2008-07-15 | 8:46 | 22 | 98.6 | 12.1 | 34 | 10.61 | 7.08
2008-07-15 | 9:01 | 22 | 99.1 | 12.08 | 34 | 10.66 | 7.08
2008-07-15 | 9:16 | 22 | 99 | 12.11 | 34 | 10.64 | 7.09
2008-07-15 | 9:31 | 22 | 99.4 | 12.13 | 34 | 10.68 | 7.09
2008-07-15 | 9:46 | 22 | 99.2 | 12.19 | 34 | 10.64 | 7.1
2008-07-15 | 10:01 | 22 | 99.5 | 12.25 | 34 | 10.67 | 7.11
2008-07-15 | 10:16 | 22 | 99.8 | 12.38 | 34 | 10.66 | 7.12
2008-07-15 | 10:31 | 22 | 99.9 | 12.5 | 34 | 10.65 | 7.13
2008-07-15 | 10:46 | 22 | 100 | 12.64 | 34 | 10.62 | 7.13
2008-07-15 | 11:01 | 22 | 100.1 | 12.82 | 34 | 10.59 | 7.14
2008-07-15 | 11:16 | 22 | 100.3 | 12.99 | 34 | 10.57 | 7.14
2008-07-15 | 11:31 | 22 | 100.4 | 13.14 | 34 | 10.54 | 7.15
2008-07-15 | 11:46 | 22 | 100 | 13.29 | 34 | 10.47 | 7.15
2008-07-15 | 12:01 | 22 | 99.9 | 13.46 | 34 | 10.42 | 7.16
2008-07-15 | 12:16 | 22 | 99.8 | 13.66 | 34 | 10.37 | 7.16
2008-07-15 | 12:31 | 22 | 100.2 | 13.98 | 34 | 10.33 | 7.17
2008-07-15 | 12:46 | 22 | 100.4 | 14.3 | 34 | 10.28 | 7.17
2008-07-15 | 13:00 | 22 | 100 | 14.62 | 34 | 10.17 | 7.17
2008-07-15 | 13:15 | 22 | 99.8 | 14.94 | 34 | 10.08 | 7.19
2008-07-15 | 13:30 | 22 | 100 | 15.22 | 34 | 10.04 | 7.18
2008-07-15 | 13:45 | 22 | 100 | 15.48 | 34 | 9.98 | 7.19
2008-07-15 | 14:00 | 22 | 99.5 | 15.74 | 34 | 9.88 | 7.2
2008-07-15 | 14:15 | 22 | 99.7 | 16 | 34 | 9.84 | 7.2
2008-07-15 | 14:30 | 22 | 99.2 | 16.28 | 35 | 9.73 | 7.2
2008-07-15 | 14:45 | 22 | 99.3 | 16.55 | 35 | 9.68 | 7.21
2008-07-15 | 15:00 | 23 | 99 | 16.68 | 35 | 9.63 | 7.21
2008-07-15 | 15:15 | 23 | 98.8 | 16.77 | 35 | 9.6 | 7.21
2008-07-15 | 15:30 | 23 | 98.8 | 16.91 | 35 | 9.57 | 7.21
2008-07-15 | 15:45 | 23 | 98.5 | 16.93 | 35 | 9.54 | 7.21
2008-07-15 | 16:00 | 23 | 97.9 | 16.8 | 35 | 9.5 | 7.2
2008-07-15 | 16:15 | 23 | 97.7 | 16.8 | 35 | 9.48 | 7.2
2008-07-15 | 16:30 | 23 | 97.7 | 16.9 | 35 | 9.46 | 7.21
2008-07-15 | 16:45 | 23 | 97.6 | 16.93 | 35 | 9.45 | 7.2
2008-07-15 | 17:00 | 23 | 97.8 | 16.91 | 35 | 9.47 | 7.21
2008-07-15 | 17:15 | 23 | 97.6 | 16.8 | 35 | 9.47 | 7.2
2008-07-15 | 17:30 | 23 | 97.4 | 16.73 | 35 | 9.47 | 7.2
2008-07-15 | 17:45 | 23 | 97.3 | 16.73 | 35 | 9.45 | 7.19
2008-07-15 | 18:00 | 23 | 97.1 | 16.7 | 35 | 9.45 | 7.19
2008-07-15 | 18:15 | 23 | 97.4 | 16.69 | 35 | 9.47 | 7.19
2008-07-15 | 18:30 | 23 | 97.5 | 16.74 | 35 | 9.47 | 7.18
2008-07-15 | 18:45 | 23 | 97.5 | 16.7 | 36 | 9.49 | 7.19
2008-07-15 | 19:00 | 23 | 97.5 | 16.64 | 36 | 9.5 | 7.18
2008-07-15 | 19:15 | 23 | 97.5 | 16.58 | 36 | 9.51 | 7.18
2008-07-15 | 19:30 | 23 | 97.5 | 16.51 | 36 | 9.52 | 7.18
2008-07-15 | 19:45 | 23 | 97 | 16.47 | 36 | 9.48 | 7.18
2008-07-15 | 20:00 | 23 | 96.8 | 16.41 | 36 | 9.47 | 7.17
2008-07-15 | 20:15 | 23 | 97.1 | 16.35 | 36 | 9.51 | 7.16
2008-07-15 | 20:30 | 23 | 97.1 | 16.35 | 36 | 9.52 | 7.16
2008-07-15 | 20:45 | 23 | 97.1 | 16.35 | 36 | 9.51 | 7.15
2008-07-15 | 21:00 | 24 | 96.7 | 16.33 | 36 | 9.48 | 7.15
2008-07-15 | 21:15 | 24 | 96.9 | 16.29 | 36 | 9.5 | 7.15
2008-07-15 | 21:30 | 24 | 97 | 16.21 | 36 | 9.53 | 7.14
2008-07-15 | 21:45 | 24 | 97 | 16.09 | 37 | 9.56 | 7.14
2008-07-15 | 22:00 | 24 | 97.1 | 15.97 | 37 | 9.59 | 7.14
2008-07-15 | 22:15 | 24 | 96.8 | 15.85 | 37 | 9.59 | 7.13
2008-07-15 | 22:30 | 24 | 96.7 | 15.73 | 37 | 9.6 | 7.13
2008-07-15 | 22:45 | 24 | 97 | 15.61 | 37 | 9.66 | 7.12
2008-07-15 | 23:00 | 24 | 97.2 | 15.5 | 37 | 9.7 | 7.12
2008-07-15 | 23:15 | 24 | 96.9 | 15.41 | 37 | 9.69 | 7.13
2008-07-15 | 23:30 | 24 | 96.8 | 15.32 | 37 | 9.7 | 7.13
2008-07-15 | 23:45 | 24 | 96.8 | 15.23 | 37 | 9.71 | 7.12
2008-07-16 | 0:00 | 24 | 96.8 | 15.14 | 37 | 9.73 | 7.12
2008-07-16 | 0:15 | 24 | 97.2 | 15.04 | 37 | 9.79 | 7.12
2008-07-16 | 0:30 | 24 | 97 | 14.94 | 37 | 9.8 | 7.12
2008-07-16 | 0:45 | 24 | 97 | 14.84 | 37 | 9.81 | 7.12
2008-07-16 | 1:01 | 24 | 97 | 14.72 | 37 | 9.84 | 7.12
2008-07-16 | 1:16 | 24 | 97 | 14.6 | 37 | 9.87 | 7.12
2008-07-16 | 1:31 | 24 | 97.4 | 14.46 | 37 | 9.94 | 7.12
2008-07-16 | 1:46 | 24 | 97.6 | 14.33 | 37 | 9.99 | 7.11
2008-07-16 | 2:01 | 24 | 97.8 | 14.19 | 37 | 10.03 | 7.11
2008-07-16 | 2:16 | 24 | 97.5 | 14.06 | 37 | 10.03 | 7.1
2008-07-16 | 2:31 | 24 | 97.4 | 13.93 | 37 | 10.05 | 7.11
2008-07-16 | 2:46 | 24 | 97.4 | 13.8 | 37 | 10.08 | 7.1
2008-07-16 | 3:01 | 24 | 97.4 | 13.68 | 37 | 10.1 | 7.1
2008-07-16 | 3:16 | 24 | 97.4 | 13.57 | 37 | 10.13 | 7.1
2008-07-16 | 3:31 | 24 | 97.4 | 13.47 | 37 | 10.16 | 7.1
2008-07-16 | 3:46 | 24 | 97.8 | 13.38 | 36 | 10.22 | 7.09
2008-07-16 | 4:01 | 23 | 97.6 | 13.29 | 36 | 10.22 | 7.09
2008-07-16 | 4:16 | 23 | 98 | 13.21 | 36 | 10.27 | 7.09
2008-07-16 | 4:31 | 23 | 98.1 | 13.12 | 36 | 10.31 | 7.09
2008-07-16 | 4:46 | 23 | 98.2 | 13.03 | 35 | 10.34 | 7.08
2008-07-16 | 5:01 | 23 | 97.9 | 12.93 | 35 | 10.33 | 7.07
2008-07-16 | 5:16 | 23 | 97.8 | 12.83 | 35 | 10.34 | 7.08
2008-07-16 | 5:31 | 23 | 98.2 | 12.7 | 35 | 10.41 | 7.07
2008-07-16 | 5:46 | 23 | 97.9 | 12.56 | 35 | 10.42 | 7.07
2008-07-16 | 6:01 | 23 | 98 | 12.42 | 35 | 10.46 | 7.07
2008-07-16 | 6:16 | 22 | 98.4 | 12.28 | 35 | 10.54 | 7.06
2008-07-16 | 6:31 | 22 | 98.7 | 12.15 | 34 | 10.6 | 7.06
2008-07-16 | 6:46 | 22 | 98.8 | 12.04 | 34 | 10.64 | 7.06
2008-07-16 | 7:01 | 22 | 99 | 11.95 | 34 | 10.68 | 7.06
2008-07-16 | 7:16 | 22 | 99.1 | 11.88 | 34 | 10.71 | 7.05
2008-07-16 | 7:31 | 22 | 98.9 | 11.83 | 34 | 10.7 | 7.06
2008-07-16 | 7:46 | 22 | 99.3 | 11.8 | 34 | 10.75 | 7.06
2008-07-16 | 8:01 | 22 | 99 | 11.76 | 34 | 10.73 | 7.06
2008-07-16 | 8:16 | 22 | 99 | 11.73 | 34 | 10.73 | 7.07
2008-07-16 | 8:31 | 22 | 99.5 | 11.72 | 34 | 10.79 | 7.07
2008-07-16 | 8:46 | 22 | 99.8 | 11.73 | 34 | 10.82 | 7.07
2008-07-16 | 9:01 | 22 | 99.6 | 11.74 | 34 | 10.8 | 7.08
2008-07-16 | 9:16 | 22 | 99.7 | 11.78 | 34 | 10.8 | 7.08
2008-07-16 | 9:31 | 22 | 100.1 | 11.83 | 34 | 10.84 | 7.09
2008-07-16 | 9:46 | 22 | 100.3 | 11.9 | 34 | 10.84 | 7.1
2008-07-16 | 10:01 | 22 | 100.4 | 12.02 | 34 | 10.81 | 7.11
2008-07-16 | 10:16 | 22 | 100.4 | 12.18 | 34 | 10.78 | 7.11
2008-07-16 | 10:31 | 22 | 100 | 12.39 | 34 | 10.68 | 7.11
2008-07-16 | 10:46 | 22 | 100.2 | 12.62 | 34 | 10.65 | 7.13
2008-07-16 | 11:01 | 22 | 100.5 | 12.86 | 34 | 10.62 | 7.13
2008-07-16 | 11:16 | 22 | 100.8 | 13.11 | 34 | 10.6 | 7.13
2008-07-16 | 11:31 | 22 | 100.6 | 13.38 | 34 | 10.51 | 7.14
2008-07-16 | 11:46 | 22 | 100.6 | 13.66 | 34 | 10.44 | 7.15
2008-07-16 | 12:01 | 22 | 100.6 | 13.95 | 34 | 10.37 | 7.16
2008-07-16 | 12:16 | 22 | 101 | 14.27 | 34 | 10.35 | 7.16
2008-07-16 | 12:31 | 22 | 101 | 14.51 | 34 | 10.29 | 7.16
2008-07-16 | 12:46 | 22 | 100.5 | 14.58 | 34 | 10.23 | 7.16
2008-07-16 | 13:00 | 22 | 100.3 | 14.88 | 34 | 10.14 | 7.17
2008-07-16 | 13:15 | 22 | 100.7 | 15.3 | 34 | 10.09 | 7.18
2008-07-16 | 13:30 | 22 | 100 | 15.49 | 34 | 9.98 | 7.18
2008-07-16 | 13:45 | 22 | 100.3 | 15.73 | 34 | 9.96 | 7.19
2008-07-16 | 14:00 | 22 | 100.5 | 16 | 34 | 9.92 | 7.19
2008-07-16 | 14:15 | 22 | 100.3 | 16.13 | 35 | 9.87 | 7.2
2008-07-16 | 14:30 | 23 | 100 | 16.33 | 35 | 9.8 | 7.2
2008-07-16 | 14:45 | 23 | 99.6 | 16.6 | 35 | 9.71 | 7.21
2008-07-16 | 15:00 | 23 | 99.7 | 16.96 | 35 | 9.64 | 7.2
2008-07-16 | 15:15 | 23 | 99.2 | 17.17 | 35 | 9.55 | 7.22
2008-07-16 | 15:30 | 23 | 98.9 | 17.35 | 35 | 9.49 | 7.22
2008-07-16 | 15:45 | 23 | 98.6 | 17.51 | 35 | 9.43 | 7.22
2008-07-16 | 16:00 | 23 | 98.7 | 17.63 | 35 | 9.42 | 7.22
2008-07-16 | 16:15 | 23 | 98.2 | 17.8 | 35 | 9.34 | 7.22
2008-07-16 | 16:30 | 23 | 98 | 17.92 | 35 | 9.29 | 7.22
2008-07-16 | 16:45 | 23 | 98.2 | 17.94 | 35 | 9.3 | 7.22
2008-07-16 | 17:00 | 23 | 98.2 | 17.78 | 35 | 9.34 | 7.22
2008-07-16 | 17:15 | 23 | 97.8 | 17.66 | 35 | 9.32 | 7.22
2008-07-16 | 17:30 | 23 | 97.9 | 17.57 | 35 | 9.36 | 7.21
2008-07-16 | 17:45 | 23 | 97.9 | 17.52 | 35 | 9.36 | 7.21
2008-07-16 | 18:00 | 23 | 97.8 | 17.55 | 35 | 9.34 | 7.2
2008-07-16 | 18:15 | 23 | 97.3 | 17.61 | 36 | 9.29 | 7.2
2008-07-16 | 18:30 | 23 | 97.5 | 17.62 | 36 | 9.3 | 7.19
2008-07-16 | 18:45 | 23 | 97.1 | 17.6 | 35 | 9.27 | 7.19
2008-07-16 | 19:00 | 23 | 97.3 | 17.56 | 35 | 9.3 | 7.19
2008-07-16 | 19:15 | 23 | 97.4 | 17.52 | 36 | 9.32 | 7.19
2008-07-16 | 19:30 | 23 | 97.4 | 17.47 | 36 | 9.32 | 7.18
2008-07-16 | 19:45 | 23 | 97.1 | 17.41 | 36 | 9.3 | 7.19
2008-07-16 | 20:00 | 23 | 97.3 | 17.36 | 36 | 9.33 | 7.18
2008-07-16 | 20:15 | 23 | 97.3 | 17.26 | 36 | 9.36 | 7.18
2008-07-16 | 20:30 | 24 | 97.4 | 17.16 | 36 | 9.38 | 7.18
2008-07-16 | 20:45 | 24 | 97.4 | 17.03 | 36 | 9.4 | 7.17
2008-07-16 | 21:00 | 24 | 97 | 16.91 | 36 | 9.39 | 7.16
2008-07-16 | 21:15 | 24 | 96.9 | 16.79 | 36 | 9.4 | 7.17
2008-07-16 | 21:30 | 24 | 97.2 | 16.68 | 36 | 9.45 | 7.16
2008-07-16 | 21:45 | 24 | 96.9 | 16.56 | 37 | 9.46 | 7.16
2008-07-16 | 22:00 | 24 | 96.9 | 16.44 | 37 | 9.47 | 7.16
2008-07-16 | 22:15 | 24 | 96.9 | 16.31 | 37 | 9.5 | 7.16
2008-07-16 | 22:30 | 24 | 97.3 | 16.19 | 37 | 9.56 | 7.15
2008-07-16 | 22:45 | 24 | 97.1 | 16.06 | 37 | 9.57 | 7.15
2008-07-16 | 23:00 | 24 | 97.4 | 15.93 | 37 | 9.63 | 7.15
2008-07-16 | 23:15 | 24 | 97.5 | 15.8 | 37 | 9.67 | 7.15
2008-07-16 | 23:30 | 24 | 97.7 | 15.67 | 37 | 9.71 | 7.14
2008-07-16 | 23:45 | 24 | 97.4 | 15.55 | 37 | 9.7 | 7.15
2008-07-17 | 0:00 | 24 | 97.2 | 15.43 | 37 | 9.71 | 7.15
2008-07-17 | 0:15 | 24 | 97.6 | 15.29 | 37 | 9.78 | 7.15
2008-07-17 | 0:30 | 24 | 97.4 | 15.14 | 37 | 9.79 | 7.15
2008-07-17 | 0:45 | 24 | 97.4 | 14.98 | 37 | 9.82 | 7.14
2008-07-17 | 1:01 | 24 | 97.4 | 14.81 | 37 | 9.86 | 7.15
2008-07-17 | 1:16 | 24 | 97.4 | 14.63 | 37 | 9.9 | 7.14
2008-07-17 | 1:31 | 24 | 97.8 | 14.45 | 37 | 9.98 | 7.14
2008-07-17 | 1:46 | 24 | 97.7 | 14.28 | 37 | 10 | 7.14
2008-07-17 | 2:01 | 24 | 97.6 | 14.12 | 37 | 10.04 | 7.13
2008-07-17 | 2:16 | 24 | 98 | 13.97 | 37 | 10.11 | 7.14
2008-07-17 | 2:31 | 24 | 98.2 | 13.8 | 37 | 10.17 | 7.13
2008-07-17 | 2:46 | 24 | 98.4 | 13.64 | 37 | 10.22 | 7.12
2008-07-17 | 3:01 | 24 | 98.5 | 13.47 | 37 | 10.27 | 7.12
2008-07-17 | 3:16 | 24 | 98.5 | 13.31 | 36 | 10.31 | 7.11
2008-07-17 | 3:31 | 24 | 98.2 | 13.14 | 36 | 10.32 | 7.11
2008-07-17 | 3:46 | 24 | 98.5 | 12.98 | 36 | 10.38 | 7.11
2008-07-17 | 4:01 | 23 | 98.2 | 12.83 | 36 | 10.39 | 7.1
2008-07-17 | 4:16 | 23 | 98.1 | 12.69 | 36 | 10.41 | 7.1
2008-07-17 | 4:31 | 23 | 98.5 | 12.54 | 36 | 10.49 | 7.09
2008-07-17 | 4:46 | 23 | 98.3 | 12.4 | 35 | 10.5 | 7.08
2008-07-17 | 5:01 | 23 | 98.3 | 12.26 | 35 | 10.53 | 7.08
2008-07-17 | 5:16 | 23 | 98.7 | 12.13 | 35 | 10.6 | 7.08
2008-07-17 | 5:31 | 23 | 98.5 | 12 | 35 | 10.61 | 7.07
2008-07-17 | 5:46 | 22 | 98.4 | 11.89 | 35 | 10.63 | 7.07
2008-07-17 | 6:01 | 22 | 98.8 | 11.8 | 34 | 10.7 | 7.06
2008-07-17 | 6:16 | 22 | 98.6 | 11.71 | 34 | 10.7 | 7.07
2008-07-17 | 6:31 | 22 | 98.6 | 11.62 | 34 | 10.72 | 7.07
2008-07-17 | 6:46 | 22 | 98.6 | 11.53 | 34 | 10.74 | 7.07
2008-07-17 | 7:01 | 22 | 98.7 | 11.45 | 34 | 10.77 | 7.07
2008-07-17 | 7:16 | 22 | 98.8 | 11.39 | 34 | 10.8 | 7.07
2008-07-17 | 7:31 | 22 | 99.3 | 11.32 | 34 | 10.87 | 7.06
2008-07-17 | 7:46 | 22 | 99.2 | 11.27 | 34 | 10.87 | 7.07
2008-07-17 | 8:01 | 22 | 99.6 | 11.22 | 34 | 10.93 | 7.06
2008-07-17 | 8:16 | 22 | 99.9 | 11.16 | 34 | 10.97 | 7.07
2008-07-17 | 8:31 | 22 | 100.1 | 11.12 | 34 | 11.01 | 7.06
2008-07-17 | 8:46 | 22 | 100 | 11.07 | 34 | 11 | 7.08
2008-07-17 | 9:01 | 22 | 100.4 | 11.04 | 34 | 11.06 | 7.08
2008-07-17 | 9:16 | 22 | 100.4 | 11.04 | 34 | 11.06 | 7.09
2008-07-17 | 9:31 | 22 | 100.7 | 11.07 | 34 | 11.09 | 7.09
2008-07-17 | 9:46 | 22 | 100.8 | 11.13 | 34 | 11.09 | 7.09
2008-07-17 | 10:01 | 22 | 100.5 | 11.25 | 34 | 11.02 | 7.1
2008-07-17 | 10:16 | 22 | 100.7 | 11.4 | 34 | 11 | 7.11
2008-07-17 | 10:31 | 22 | 100.3 | 11.6 | 34 | 10.91 | 7.11
2008-07-17 | 10:46 | 22 | 100.6 | 11.83 | 34 | 10.89 | 7.12
2008-07-17 | 11:01 | 22 | 100.5 | 12.07 | 34 | 10.81 | 7.13
2008-07-17 | 11:16 | 22 | 101 | 12.33 | 34 | 10.8 | 7.13
2008-07-17 | 11:31 | 22 | 100.9 | 12.6 | 34 | 10.72 | 7.14
2008-07-17 | 11:46 | 22 | 100.8 | 12.9 | 34 | 10.65 | 7.15
2008-07-17 | 12:01 | 22 | 100.8 | 13.21 | 34 | 10.58 | 7.16
2008-07-17 | 12:16 | 22 | 101.2 | 13.53 | 34 | 10.54 | 7.16
2008-07-17 | 12:31 | 22 | 101 | 13.85 | 34 | 10.44 | 7.17
2008-07-17 | 12:46 | 22 | 100.8 | 14.19 | 34 | 10.35 | 7.18
2008-07-17 | 13:00 | 22 | 101.1 | 14.55 | 35 | 10.3 | 7.18
2008-07-17 | 13:15 | 22 | 101.1 | 14.9 | 35 | 10.22 | 7.19
2008-07-17 | 13:30 | 23 | 101.1 | 15.25 | 35 | 10.14 | 7.19
2008-07-17 | 13:45 | 23 | 100.6 | 15.58 | 35 | 10.01 | 7.2
2008-07-17 | 14:00 | 23 | 100.6 | 15.91 | 35 | 9.95 | 7.21
2008-07-17 | 14:15 | 23 | 100.1 | 16.2 | 35 | 9.83 | 7.22
2008-07-17 | 14:30 | 23 | 99.7 | 16.48 | 35 | 9.74 | 7.23
2008-07-17 | 14:45 | 23 | 99.8 | 16.72 | 35 | 9.7 | 7.23
2008-07-17 | 15:00 | 23 | 99.6 | 16.95 | 35 | 9.64 | 7.23
2008-07-17 | 15:15 | 23 | 99.5 | 17.15 | 35 | 9.58 | 7.23
2008-07-17 | 15:30 | 23 | 99.3 | 17.31 | 35 | 9.53 | 7.23
2008-07-17 | 15:45 | 23 | 99.1 | 17.42 | 35 | 9.49 | 7.24
2008-07-17 | 16:00 | 23 | 98.9 | 17.53 | 35 | 9.45 | 7.24
2008-07-17 | 16:15 | 23 | 98.3 | 17.6 | 35 | 9.39 | 7.24
2008-07-17 | 16:30 | 23 | 98.4 | 17.67 | 35 | 9.38 | 7.24
2008-07-17 | 16:45 | 23 | 98.3 | 17.7 | 35 | 9.36 | 7.24
2008-07-17 | 17:00 | 23 | 97.8 | 17.72 | 36 | 9.31 | 7.24
2008-07-17 | 17:15 | 23 | 97.9 | 17.7 | 36 | 9.32 | 7.23
2008-07-17 | 17:30 | 23 | 97.5 | 17.69 | 36 | 9.29 | 7.23
2008-07-17 | 17:45 | 23 | 97.2 | 17.65 | 36 | 9.27 | 7.22
2008-07-17 | 18:00 | 23 | 97.5 | 17.62 | 36 | 9.3 | 7.23
2008-07-17 | 18:15 | 23 | 97.5 | 17.59 | 36 | 9.31 | 7.22
2008-07-17 | 18:30 | 23 | 97.1 | 17.55 | 36 | 9.28 | 7.22
2008-07-17 | 18:45 | 23 | 96.9 | 17.51 | 36 | 9.27 | 7.22
2008-07-17 | 19:00 | 23 | 97.2 | 17.47 | 36 | 9.3 | 7.21
2008-07-17 | 19:15 | 23 | 96.9 | 17.41 | 36 | 9.29 | 7.21
2008-07-17 | 19:30 | 24 | 96.8 | 17.35 | 36 | 9.29 | 7.2
2008-07-17 | 19:45 | 24 | 97.1 | 17.28 | 36 | 9.33 | 7.2
2008-07-17 | 20:00 | 24 | 97.3 | 17.18 | 36 | 9.37 | 7.2
2008-07-17 | 20:15 | 24 | 97 | 17.05 | 36 | 9.36 | 7.19
2008-07-17 | 20:30 | 24 | 96.8 | 16.94 | 37 | 9.37 | 7.19
2008-07-17 | 20:45 | 24 | 96.8 | 16.79 | 37 | 9.4 | 7.19
2008-07-17 | 21:00 | 24 | 97.1 | 16.63 | 37 | 9.46 | 7.18
2008-07-17 | 21:15 | 24 | 96.9 | 16.48 | 37 | 9.47 | 7.18
2008-07-17 | 21:30 | 24 | 97.2 | 16.33 | 37 | 9.53 | 7.18
2008-07-17 | 21:45 | 24 | 97 | 16.2 | 37 | 9.53 | 7.18
2008-07-17 | 22:00 | 24 | 97.3 | 16.07 | 37 | 9.59 | 7.17
2008-07-17 | 22:15 | 24 | 97.1 | 15.95 | 37 | 9.59 | 7.17
2008-07-17 | 22:30 | 24 | 97 | 15.81 | 37 | 9.61 | 7.18
2008-07-17 | 22:45 | 24 | 97.3 | 15.7 | 37 | 9.67 | 7.18
2008-07-17 | 23:00 | 24 | 97.5 | 15.57 | 37 | 9.71 | 7.17
2008-07-17 | 23:15 | 24 | 97.2 | 15.46 | 37 | 9.71 | 7.18
2008-07-17 | 23:30 | 24 | 97.5 | 15.33 | 38 | 9.76 | 7.17
2008-07-17 | 23:45 | 24 | 97.6 | 15.2 | 38 | 9.8 | 7.18
2008-07-18 | 0:00 | 25 | 97.7 | 15.06 | 38 | 9.84 | 7.18
2008-07-18 | 0:15 | 25 | 97.8 | 14.91 | 38 | 9.88 | 7.18
2008-07-18 | 0:30 | 25 | 97.9 | 14.76 | 38 | 9.92 | 7.18
2008-07-18 | 0:45 | 25 | 97.9 | 14.61 | 38 | 9.96 | 7.18
2008-07-18 | 1:01 | 25 | 97.6 | 14.45 | 38 | 9.96 | 7.18
2008-07-18 | 1:16 | 25 | 97.9 | 14.29 | 38 | 10.02 | 7.17
2008-07-18 | 1:31 | 25 | 98 | 14.12 | 38 | 10.08 | 7.18
2008-07-18 | 1:46 | 25 | 98.2 | 13.95 | 38 | 10.13 | 7.17
2008-07-18 | 2:01 | 25 | 98.2 | 13.8 | 38 | 10.17 | 7.17
2008-07-18 | 2:16 | 25 | 97.9 | 13.64 | 38 | 10.17 | 7.18
2008-07-18 | 2:31 | 25 | 98.2 | 13.5 | 38 | 10.23 | 7.17
2008-07-18 | 2:46 | 25 | 98.3 | 13.36 | 38 | 10.28 | 7.17
2008-07-18 | 3:01 | 25 | 98.5 | 13.22 | 39 | 10.32 | 7.17
2008-07-18 | 3:16 | 25 | 98.1 | 13.08 | 39 | 10.32 | 7.18
2008-07-18 | 3:31 | 25 | 98 | 12.95 | 39 | 10.34 | 7.18
2008-07-18 | 3:46 | 25 | 98.4 | 12.81 | 39 | 10.41 | 7.18
2008-07-18 | 4:01 | 26 | 98.6 | 12.68 | 39 | 10.46 | 7.18
2008-07-18 | 4:16 | 26 | 98.7 | 12.55 | 39 | 10.5 | 7.18
2008-07-18 | 4:31 | 26 | 98.8 | 12.42 | 39 | 10.54 | 7.18
2008-07-18 | 4:46 | 25 | 98.4 | 12.29 | 39 | 10.53 | 7.18
2008-07-18 | 5:01 | 25 | 98.7 | 12.17 | 39 | 10.6 | 7.17
2008-07-18 | 5:16 | 25 | 98.4 | 12.05 | 39 | 10.59 | 7.17
2008-07-18 | 5:31 | 25 | 98.7 | 11.93 | 39 | 10.66 | 7.17
2008-07-18 | 5:46 | 25 | 98.5 | 11.81 | 39 | 10.66 | 7.17
2008-07-18 | 6:01 | 25 | 98.9 | 11.7 | 39 | 10.73 | 7.16
2008-07-18 | 6:16 | 25 | 98.7 | 11.58 | 39 | 10.73 | 7.17
2008-07-18 | 6:31 | 25 | 99.1 | 11.48 | 39 | 10.81 | 7.17
2008-07-18 | 6:46 | 25 | 99.3 | 11.38 | 39 | 10.85 | 7.17
2008-07-18 | 7:01 | 25 | 99.4 | 11.3 | 39 | 10.89 | 7.16
2008-07-18 | 7:16 | 25 | 99.1 | 11.23 | 39 | 10.87 | 7.17
2008-07-18 | 7:31 | 25 | 99.5 | 11.18 | 38 | 10.92 | 7.17
2008-07-18 | 7:46 | 25 | 99.3 | 11.15 | 38 | 10.91 | 7.17
2008-07-18 | 8:01 | 25 | 99.7 | 11.13 | 38 | 10.96 | 7.17
2008-07-18 | 8:16 | 25 | 99.9 | 11.12 | 38 | 10.98 | 7.17
2008-07-18 | 8:31 | 25 | 99.7 | 11.12 | 38 | 10.97 | 7.17
2008-07-18 | 8:46 | 25 | 100.1 | 11.12 | 38 | 11.01 | 7.17
2008-07-18 | 9:01 | 25 | 100.1 | 11.12 | 38 | 11 | 7.18
2008-07-18 | 9:16 | 25 | 100.6 | 11.16 | 38 | 11.05 | 7.18
2008-07-18 | 9:31 | 25 | 100.8 | 11.23 | 38 | 11.05 | 7.19
2008-07-18 | 9:46 | 25 | 100.9 | 11.34 | 38 | 11.04 | 7.18
2008-07-18 | 10:01 | 25 | 100.4 | 11.48 | 38 | 10.95 | 7.19
2008-07-18 | 10:16 | 25 | 100.7 | 11.66 | 38 | 10.93 | 7.2
2008-07-18 | 10:31 | 25 | 100.3 | 11.89 | 38 | 10.83 | 7.21
2008-07-18 | 10:46 | 25 | 100.6 | 12.17 | 38 | 10.8 | 7.21
2008-07-18 | 11:01 | 25 | 100.9 | 12.47 | 38 | 10.76 | 7.23
2008-07-18 | 11:16 | 25 | 100.7 | 12.79 | 38 | 10.66 | 7.23
2008-07-18 | 11:31 | 25 | 101.1 | 13.12 | 38 | 10.62 | 7.23
2008-07-18 | 11:46 | 25 | 100.8 | 13.46 | 38 | 10.52 | 7.24
2008-07-18 | 12:01 | 25 | 101.2 | 13.81 | 38 | 10.47 | 7.25
2008-07-18 | 12:16 | 25 | 101.3 | 14.18 | 38 | 10.4 | 7.26
2008-07-18 | 12:31 | 25 | 101.3 | 14.54 | 39 | 10.31 | 7.26
2008-07-18 | 12:46 | 25 | 101.2 | 14.93 | 39 | 10.22 | 7.26
2008-07-18 | 13:00 | 25 | 100.8 | 15.29 | 39 | 10.1 | 7.28
2008-07-18 | 13:15 | 25 | 100.8 | 15.67 | 39 | 10.02 | 7.28
2008-07-18 | 13:30 | 25 | 100.8 | 16.03 | 39 | 9.94 | 7.29
2008-07-18 | 13:45 | 25 | 100.7 | 16.4 | 39 | 9.86 | 7.29
2008-07-18 | 14:00 | 25 | 100.1 | 16.74 | 39 | 9.73 | 7.3
2008-07-18 | 14:15 | 25 | 100.1 | 17.04 | 39 | 9.66 | 7.3
2008-07-18 | 14:30 | 25 | 99.6 | 17.33 | 39 | 9.56 | 7.3
2008-07-18 | 14:45 | 25 | 99.6 | 17.58 | 39 | 9.51 | 7.32
2008-07-18 | 15:00 | 25 | 99.2 | 17.81 | 39 | 9.43 | 7.32
2008-07-18 | 15:15 | 26 | 98.9 | 17.98 | 39 | 9.36 | 7.32
2008-07-18 | 15:30 | 26 | 98.6 | 18.12 | 39 | 9.31 | 7.32
2008-07-18 | 15:45 | 26 | 98.4 | 18.23 | 39 | 9.27 | 7.33
2008-07-18 | 16:00 | 26 | 98.1 | 18.29 | 39 | 9.24 | 7.33
2008-07-18 | 16:15 | 26 | 98.3 | 18.33 | 39 | 9.25 | 7.33
2008-07-18 | 16:30 | 26 | 97.9 | 18.35 | 40 | 9.21 | 7.33
2008-07-18 | 16:45 | 26 | 98 | 18.34 | 40 | 9.22 | 7.32
2008-07-18 | 17:00 | 26 | 98 | 18.3 | 40 | 9.22 | 7.32
2008-07-18 | 17:15 | 26 | 97.9 | 18.25 | 40 | 9.22 | 7.31
2008-07-18 | 17:30 | 26 | 97.8 | 18.2 | 40 | 9.22 | 7.3
2008-07-18 | 17:45 | 26 | 97.3 | 18.14 | 40 | 9.19 | 7.3
2008-07-18 | 18:00 | 26 | 97.4 | 18.06 | 40 | 9.21 | 7.29
2008-07-18 | 18:15 | 26 | 97.1 | 18 | 40 | 9.19 | 7.29
2008-07-18 | 18:30 | 26 | 96.9 | 17.92 | 40 | 9.19 | 7.28
2008-07-18 | 18:45 | 26 | 96.8 | 17.85 | 40 | 9.19 | 7.28
2008-07-18 | 19:00 | 26 | 97.1 | 17.8 | 40 | 9.23 | 7.28
2008-07-18 | 19:15 | 26 | 97.2 | 17.75 | 40 | 9.25 | 7.27
2008-07-18 | 19:30 | 26 | 97.2 | 17.68 | 40 | 9.27 | 7.27
2008-07-18 | 19:45 | 26 | 97.3 | 17.59 | 40 | 9.29 | 7.27
2008-07-18 | 20:00 | 26 | 96.9 | 17.47 | 40 | 9.27 | 7.27
2008-07-18 | 20:15 | 26 | 97.1 | 17.34 | 40 | 9.32 | 7.26
2008-07-18 | 20:30 | 26 | 97.2 | 17.21 | 40 | 9.35 | 7.25
2008-07-18 | 20:45 | 26 | 96.9 | 17.04 | 40 | 9.36 | 7.26
2008-07-18 | 21:00 | 26 | 97.1 | 16.85 | 41 | 9.41 | 7.25
2008-07-18 | 21:15 | 26 | 96.9 | 16.67 | 41 | 9.43 | 7.25
2008-07-18 | 21:30 | 26 | 97.2 | 16.49 | 41 | 9.49 | 7.24
2008-07-18 | 21:45 | 26 | 97.3 | 16.32 | 41 | 9.54 | 7.24
2008-07-18 | 22:00 | 26 | 97.4 | 16.15 | 41 | 9.58 | 7.24
2008-07-18 | 22:15 | 27 | 97.1 | 16 | 41 | 9.58 | 7.24
2008-07-18 | 22:30 | 27 | 97 | 15.84 | 41 | 9.6 | 7.24
2008-07-18 | 22:45 | 27 | 97.3 | 15.69 | 41 | 9.67 | 7.24
2008-07-18 | 23:00 | 27 | 97.5 | 15.55 | 41 | 9.72 | 7.23
2008-07-18 | 23:15 | 27 | 97.2 | 15.41 | 41 | 9.71 | 7.23
2008-07-18 | 23:30 | 27 | 97.1 | 15.26 | 41 | 9.74 | 7.24
2008-07-18 | 23:45 | 27 | 97.5 | 15.11 | 41 | 9.81 | 7.24
2008-07-19 | 0:00 | 27 | 97.3 | 14.95 | 41 | 9.82 | 7.23
2008-07-19 | 0:15 | 27 | 97.6 | 14.78 | 41 | 9.89 | 7.23
2008-07-19 | 0:30 | 27 | 97.4 | 14.62 | 41 | 9.9 | 7.23
2008-07-19 | 0:45 | 27 | 97.4 | 14.46 | 41 | 9.93 | 7.23
2008-07-19 | 1:01 | 27 | 97.7 | 14.29 | 41 | 10.01 | 7.23
2008-07-19 | 1:16 | 27 | 97.9 | 14.14 | 42 | 10.06 | 7.23
2008-07-19 | 1:31 | 27 | 98 | 13.98 | 42 | 10.11 | 7.22
2008-07-19 | 1:46 | 27 | 97.8 | 13.82 | 42 | 10.12 | 7.23
2008-07-19 | 2:01 | 27 | 97.7 | 13.65 | 42 | 10.14 | 7.22
2008-07-19 | 2:16 | 27 | 97.6 | 13.48 | 42 | 10.18 | 7.22
2008-07-19 | 2:31 | 27 | 98.1 | 13.32 | 42 | 10.26 | 7.22
2008-07-19 | 2:46 | 27 | 98.3 | 13.16 | 42 | 10.32 | 7.21
2008-07-19 | 3:01 | 27 | 98 | 13 | 42 | 10.33 | 7.21
2008-07-19 | 3:16 | 27 | 98 | 12.84 | 42 | 10.36 | 7.21
2008-07-19 | 3:31 | 27 | 98.4 | 12.68 | 42 | 10.44 | 7.21
2008-07-19 | 3:46 | 27 | 98.6 | 12.53 | 42 | 10.49 | 7.21
2008-07-19 | 4:01 | 27 | 98.3 | 12.38 | 42 | 10.5 | 7.21
2008-07-19 | 4:16 | 27 | 98.6 | 12.23 | 42 | 10.57 | 7.21
2008-07-19 | 4:31 | 27 | 98.3 | 12.08 | 42 | 10.58 | 7.21
2008-07-19 | 4:46 | 27 | 98.3 | 11.93 | 42 | 10.61 | 7.2
2008-07-19 | 5:01 | 27 | 98.7 | 11.78 | 42 | 10.69 | 7.2
2008-07-19 | 5:16 | 27 | 98.5 | 11.63 | 42 | 10.7 | 7.21
2008-07-19 | 5:31 | 27 | 98.9 | 11.49 | 42 | 10.78 | 7.2
2008-07-19 | 5:46 | 27 | 99 | 11.35 | 42 | 10.83 | 7.2
2008-07-19 | 6:01 | 27 | 99.2 | 11.21 | 42 | 10.89 | 7.2
2008-07-19 | 6:16 | 27 | 98.9 | 11.07 | 42 | 10.89 | 7.19
2008-07-19 | 6:31 | 27 | 99.3 | 10.94 | 41 | 10.96 | 7.2
2008-07-19 | 6:46 | 27 | 99.1 | 10.82 | 41 | 10.97 | 7.19
2008-07-19 | 7:01 | 27 | 99.1 | 10.72 | 41 | 11 | 7.19
2008-07-19 | 7:16 | 27 | 99.6 | 10.65 | 41 | 11.07 | 7.19
2008-07-19 | 7:31 | 27 | 99.4 | 10.58 | 41 | 11.07 | 7.19
2008-07-19 | 7:46 | 27 | 99.4 | 10.54 | 41 | 11.08 | 7.2
2008-07-19 | 8:01 | 27 | 99.9 | 10.52 | 41 | 11.14 | 7.19
2008-07-19 | 8:16 | 27 | 100.2 | 10.49 | 41 | 11.18 | 7.2
2008-07-19 | 8:31 | 27 | 100 | 10.49 | 41 | 11.16 | 7.2
2008-07-19 | 8:46 | 26 | 100.5 | 10.49 | 41 | 11.22 | 7.2
2008-07-19 | 9:01 | 26 | 100.8 | 10.5 | 41 | 11.25 | 7.2
2008-07-19 | 9:16 | 26 | 101.1 | 10.52 | 41 | 11.28 | 7.2
2008-07-19 | 9:31 | 26 | 101.3 | 10.57 | 41 | 11.28 | 7.21
2008-07-19 | 9:46 | 26 | 100.9 | 10.65 | 41 | 11.21 | 7.21
2008-07-19 | 10:01 | 26 | 100.7 | 10.77 | 41 | 11.16 | 7.22
2008-07-19 | 10:16 | 26 | 100.6 | 10.93 | 40 | 11.11 | 7.22
2008-07-19 | 10:31 | 26 | 100.9 | 11.15 | 41 | 11.09 | 7.23
2008-07-19 | 10:46 | 26 | 100.6 | 11.45 | 41 | 10.98 | 7.24
2008-07-19 | 11:01 | 26 | 101.2 | 11.78 | 40 | 10.96 | 7.25
2008-07-19 | 11:16 | 26 | 101.5 | 12.14 | 40 | 10.91 | 7.25
2008-07-19 | 11:31 | 26 | 101.3 | 12.5 | 40 | 10.8 | 7.26
2008-07-19 | 11:46 | 26 | 101.7 | 12.87 | 40 | 10.74 | 7.27
2008-07-19 | 12:01 | 26 | 101.8 | 13.26 | 40 | 10.67 | 7.27
2008-07-19 | 12:16 | 26 | 101.4 | 13.67 | 40 | 10.53 | 7.28
2008-07-19 | 12:31 | 26 | 101.2 | 14.07 | 41 | 10.41 | 7.29
2008-07-19 | 12:46 | 26 | 101.1 | 14.5 | 41 | 10.3 | 7.29
2008-07-19 | 13:00 | 26 | 101.3 | 14.93 | 41 | 10.23 | 7.3
2008-07-19 | 13:15 | 26 | 101.3 | 15.36 | 41 | 10.14 | 7.3
2008-07-19 | 13:30 | 26 | 100.8 | 15.8 | 41 | 10 | 7.31
2008-07-19 | 13:45 | 27 | 100.6 | 16.2 | 41 | 9.88 | 7.31
2008-07-19 | 14:00 | 27 | 100.2 | 16.61 | 41 | 9.77 | 7.33
2008-07-19 | 14:15 | 27 | 100.4 | 16.97 | 41 | 9.71 | 7.34
2008-07-19 | 14:30 | 27 | 100.2 | 17.3 | 41 | 9.62 | 7.33
2008-07-19 | 14:45 | 27 | 100 | 17.6 | 41 | 9.54 | 7.34
2008-07-19 | 15:00 | 27 | 99.4 | 17.87 | 41 | 9.44 | 7.35
2008-07-19 | 15:15 | 27 | 99.3 | 18.1 | 41 | 9.38 | 7.36
2008-07-19 | 15:30 | 27 | 98.7 | 18.27 | 41 | 9.29 | 7.36
2008-07-19 | 15:45 | 27 | 98.7 | 18.39 | 41 | 9.27 | 7.36
2008-07-19 | 16:00 | 27 | 98 | 18.48 | 41 | 9.19 | 7.36
2008-07-19 | 16:15 | 27 | 98.3 | 18.53 | 41 | 9.2 | 7.35
2008-07-19 | 16:30 | 27 | 98.2 | 18.56 | 41 | 9.19 | 7.35
2008-07-19 | 16:45 | 27 | 97.7 | 18.56 | 41 | 9.14 | 7.35
2008-07-19 | 17:00 | 27 | 97.7 | 18.53 | 41 | 9.15 | 7.34
2008-07-19 | 17:15 | 27 | 97.6 | 18.47 | 41 | 9.15 | 7.34
2008-07-19 | 17:30 | 27 | 97.1 | 18.43 | 41 | 9.12 | 7.33
2008-07-19 | 17:45 | 27 | 96.9 | 18.38 | 41 | 9.1 | 7.32
2008-07-19 | 18:00 | 27 | 97.1 | 18.32 | 41 | 9.13 | 7.32
2008-07-19 | 18:15 | 27 | 97.1 | 18.26 | 41 | 9.14 | 7.31
2008-07-19 | 18:30 | 27 | 97.1 | 18.19 | 41 | 9.15 | 7.31
2008-07-19 | 18:45 | 27 | 97 | 18.15 | 41 | 9.16 | 7.31
2008-07-19 | 19:00 | 27 | 96.6 | 18.09 | 41 | 9.13 | 7.3
2008-07-19 | 19:15 | 27 | 96.8 | 18.05 | 42 | 9.16 | 7.29
2008-07-19 | 19:30 | 27 | 96.5 | 18 | 42 | 9.14 | 7.29
2008-07-19 | 19:45 | 27 | 96.4 | 17.92 | 42 | 9.14 | 7.29
2008-07-19 | 20:00 | 27 | 96.7 | 17.83 | 42 | 9.18 | 7.29
2008-07-19 | 20:15 | 27 | 96.4 | 17.72 | 42 | 9.18 | 7.28
2008-07-19 | 20:30 | 27 | 96.7 | 17.58 | 42 | 9.23 | 7.28
2008-07-19 | 20:45 | 27 | 96.8 | 17.45 | 42 | 9.27 | 7.28
2008-07-19 | 21:00 | 27 | 96.5 | 17.29 | 42 | 9.27 | 7.27
2008-07-19 | 21:15 | 27 | 96.4 | 17.11 | 42 | 9.3 | 7.26
2008-07-19 | 21:30 | 27 | 96.8 | 16.93 | 42 | 9.37 | 7.26
2008-07-19 | 21:45 | 27 | 97 | 16.76 | 42 | 9.42 | 7.25
2008-07-19 | 22:00 | 27 | 96.7 | 16.61 | 42 | 9.42 | 7.25
2008-07-19 | 22:15 | 27 | 96.6 | 16.47 | 42 | 9.44 | 7.25
2008-07-19 | 22:30 | 28 | 96.9 | 16.35 | 42 | 9.5 | 7.24
2008-07-19 | 22:45 | 28 | 97.1 | 16.24 | 42 | 9.54 | 7.24
2008-07-19 | 23:00 | 28 | 97.2 | 16.13 | 42 | 9.57 | 7.24
2008-07-19 | 23:15 | 28 | 96.9 | 16.01 | 43 | 9.56 | 7.25
2008-07-19 | 23:30 | 28 | 97.2 | 15.89 | 43 | 9.62 | 7.25
2008-07-19 | 23:45 | 28 | 97.4 | 15.75 | 43 | 9.66 | 7.25
2008-07-20 | 0:00 | 28 | 97.4 | 15.62 | 43 | 9.7 | 7.25
2008-07-20 | 0:15 | 28 | 97.1 | 15.47 | 43 | 9.7 | 7.25
2008-07-20 | 0:30 | 28 | 97 | 15.34 | 43 | 9.71 | 7.25
2008-07-20 | 0:45 | 28 | 97.4 | 15.21 | 43 | 9.78 | 7.25
2008-07-20 | 1:01 | 28 | 97.2 | 15.08 | 43 | 9.78 | 7.25
2008-07-20 | 1:16 | 28 | 97.5 | 14.94 | 43 | 9.85 | 7.25
2008-07-20 | 1:31 | 28 | 97.3 | 14.79 | 43 | 9.85 | 7.25
2008-07-20 | 1:46 | 28 | 97.6 | 14.65 | 43 | 9.92 | 7.25
2008-07-20 | 2:01 | 28 | 97.8 | 14.5 | 43 | 9.97 | 7.24
2008-07-20 | 2:16 | 28 | 97.9 | 14.36 | 43 | 10.01 | 7.24
2008-07-20 | 2:31 | 28 | 98 | 14.21 | 43 | 10.05 | 7.25
2008-07-20 | 2:46 | 28 | 97.7 | 14.07 | 43 | 10.05 | 7.25
2008-07-20 | 3:01 | 28 | 98 | 13.93 | 43 | 10.11 | 7.24
2008-07-20 | 3:16 | 28 | 98.1 | 13.8 | 43 | 10.16 | 7.24
2008-07-20 | 3:31 | 28 | 98.2 | 13.66 | 43 | 10.2 | 7.24
2008-07-20 | 3:46 | 28 | 98.3 | 13.53 | 43 | 10.24 | 7.24
2008-07-20 | 4:01 | 28 | 98 | 13.39 | 44 | 10.23 | 7.24
2008-07-20 | 4:16 | 28 | 98.2 | 13.26 | 44 | 10.29 | 7.24
2008-07-20 | 4:31 | 28 | 98 | 13.13 | 44 | 10.29 | 7.24
2008-07-20 | 4:46 | 28 | 98.3 | 13 | 44 | 10.36 | 7.23
2008-07-20 | 5:01 | 28 | 98.1 | 12.87 | 43 | 10.36 | 7.23
2008-07-20 | 5:16 | 28 | 98.4 | 12.75 | 43 | 10.43 | 7.23
2008-07-20 | 5:31 | 28 | 98.2 | 12.62 | 43 | 10.43 | 7.23
2008-07-20 | 5:46 | 28 | 98.6 | 12.49 | 43 | 10.5 | 7.23
2008-07-20 | 6:01 | 28 | 98.4 | 12.36 | 43 | 10.51 | 7.22
2008-07-20 | 6:16 | 28 | 98.7 | 12.22 | 43 | 10.59 | 7.23
2008-07-20 | 6:31 | 28 | 99 | 12.11 | 43 | 10.64 | 7.23
2008-07-20 | 6:46 | 28 | 99.2 | 12 | 43 | 10.69 | 7.22
2008-07-20 | 7:01 | 28 | 99.4 | 11.91 | 43 | 10.74 | 7.23
2008-07-20 | 7:16 | 28 | 99.6 | 11.83 | 43 | 10.77 | 7.23
2008-07-20 | 7:31 | 28 | 99.3 | 11.77 | 43 | 10.76 | 7.23
2008-07-20 | 7:46 | 28 | 99.3 | 11.73 | 43 | 10.77 | 7.24
2008-07-20 | 8:01 | 28 | 99.8 | 11.71 | 42 | 10.83 | 7.24
2008-07-20 | 8:16 | 28 | 100.1 | 11.7 | 42 | 10.86 | 7.25
2008-07-20 | 8:31 | 27 | 100.4 | 11.71 | 42 | 10.89 | 7.24
2008-07-20 | 8:46 | 27 | 100.1 | 11.73 | 42 | 10.86 | 7.25
2008-07-20 | 9:01 | 27 | 100.6 | 11.76 | 42 | 10.9 | 7.26
2008-07-20 | 9:16 | 27 | 100.8 | 11.8 | 42 | 10.92 | 7.26
2008-07-20 | 9:31 | 27 | 101 | 11.86 | 42 | 10.92 | 7.27
2008-07-20 | 9:46 | 27 | 101.1 | 11.95 | 42 | 10.91 | 7.27
2008-07-20 | 10:01 | 27 | 101.2 | 12.06 | 42 | 10.89 | 7.27
2008-07-20 | 10:16 | 27 | 101.3 | 12.2 | 42 | 10.87 | 7.28
2008-07-20 | 10:31 | 27 | 100.9 | 12.38 | 42 | 10.78 | 7.29
2008-07-20 | 10:46 | 27 | 101.2 | 12.62 | 42 | 10.75 | 7.29
2008-07-20 | 11:01 | 27 | 101.4 | 12.89 | 42 | 10.71 | 7.29
2008-07-20 | 11:16 | 27 | 101.5 | 13.18 | 42 | 10.65 | 7.31
2008-07-20 | 11:31 | 27 | 101.6 | 13.48 | 42 | 10.59 | 7.32
2008-07-20 | 11:46 | 27 | 101.6 | 13.8 | 42 | 10.52 | 7.32
2008-07-20 | 12:01 | 27 | 101.6 | 14.13 | 42 | 10.44 | 7.33
2008-07-20 | 12:16 | 27 | 101.5 | 14.48 | 42 | 10.36 | 7.33
2008-07-20 | 12:31 | 27 | 101.5 | 14.83 | 42 | 10.27 | 7.35
2008-07-20 | 12:46 | 27 | 101.4 | 15.19 | 42 | 10.18 | 7.34
2008-07-20 | 13:00 | 27 | 101.3 | 15.56 | 42 | 10.09 | 7.35
2008-07-20 | 13:15 | 27 | 101.1 | 15.92 | 42 | 10 | 7.36
2008-07-20 | 13:30 | 27 | 101 | 16.28 | 42 | 9.91 | 7.36
2008-07-20 | 13:45 | 27 | 100.9 | 16.63 | 42 | 9.83 | 7.37
2008-07-20 | 14:00 | 27 | 100.7 | 16.96 | 42 | 9.74 | 7.37
2008-07-20 | 14:15 | 27 | 100.5 | 17.27 | 42 | 9.66 | 7.38
2008-07-20 | 14:30 | 27 | 99.9 | 17.55 | 42 | 9.55 | 7.38
2008-07-20 | 14:45 | 27 | 99.5 | 17.83 | 42 | 9.46 | 7.38
2008-07-20 | 15:00 | 27 | 99.6 | 18.07 | 42 | 9.42 | 7.38
2008-07-20 | 15:15 | 27 | 99.1 | 18.29 | 42 | 9.33 | 7.39
2008-07-20 | 15:30 | 27 | 99.2 | 18.48 | 42 | 9.3 | 7.38
2008-07-20 | 15:45 | 27 | 98.7 | 18.62 | 42 | 9.23 | 7.39
2008-07-20 | 16:00 | 27 | 98.8 | 18.74 | 42 | 9.22 | 7.39
2008-07-20 | 16:15 | 27 | 98.4 | 18.82 | 42 | 9.16 | 7.38
2008-07-20 | 16:30 | 27 | 98.2 | 18.85 | 42 | 9.14 | 7.39
2008-07-20 | 16:45 | 27 | 98 | 18.87 | 42 | 9.11 | 7.39
2008-07-20 | 17:00 | 27 | 98.2 | 18.88 | 42 | 9.13 | 7.38
2008-07-20 | 17:15 | 27 | 98.1 | 18.88 | 42 | 9.13 | 7.37
2008-07-20 | 17:30 | 27 | 97.7 | 18.85 | 42 | 9.09 | 7.36
2008-07-20 | 17:45 | 27 | 97.4 | 18.84 | 42 | 9.06 | 7.36
2008-07-20 | 18:00 | 27 | 97.6 | 18.8 | 42 | 9.09 | 7.36
2008-07-20 | 18:15 | 27 | 97.6 | 18.78 | 42 | 9.09 | 7.35
2008-07-20 | 18:30 | 27 | 97.2 | 18.74 | 42 | 9.06 | 7.34
2008-07-20 | 18:45 | 27 | 96.9 | 18.71 | 42 | 9.04 | 7.34
2008-07-20 | 19:00 | 27 | 97.2 | 18.67 | 42 | 9.07 | 7.34
2008-07-20 | 19:15 | 27 | 97.2 | 18.64 | 42 | 9.09 | 7.33
2008-07-20 | 19:30 | 27 | 97.2 | 18.6 | 42 | 9.09 | 7.32
2008-07-20 | 19:45 | 27 | 97.2 | 18.54 | 42 | 9.1 | 7.31
2008-07-20 | 20:00 | 27 | 97.1 | 18.48 | 42 | 9.1 | 7.31
2008-07-20 | 20:15 | 27 | 96.7 | 18.38 | 42 | 9.08 | 7.3
2008-07-20 | 20:30 | 28 | 96.5 | 18.27 | 42 | 9.08 | 7.29
2008-07-20 | 20:45 | 28 | 96.8 | 18.15 | 42 | 9.13 | 7.28
2008-07-20 | 21:00 | 28 | 96.5 | 18.03 | 43 | 9.13 | 7.28
2008-07-20 | 21:15 | 28 | 96.8 | 17.88 | 43 | 9.18 | 7.27
2008-07-20 | 21:30 | 28 | 96.6 | 17.74 | 43 | 9.19 | 7.27
2008-07-20 | 21:45 | 28 | 96.8 | 17.61 | 43 | 9.24 | 7.26
2008-07-20 | 22:00 | 28 | 96.6 | 17.48 | 43 | 9.24 | 7.26
2008-07-20 | 22:15 | 28 | 96.6 | 17.36 | 43 | 9.26 | 7.26
2008-07-20 | 22:30 | 28 | 96.5 | 17.24 | 43 | 9.28 | 7.26
2008-07-20 | 22:45 | 28 | 96.6 | 17.12 | 43 | 9.31 | 7.26
2008-07-20 | 23:00 | 28 | 96.6 | 17 | 43 | 9.34 | 7.26
2008-07-20 | 23:15 | 28 | 97 | 16.89 | 43 | 9.4 | 7.26
2008-07-20 | 23:30 | 28 | 97.2 | 16.78 | 43 | 9.44 | 7.26
2008-07-20 | 23:45 | 28 | 97 | 16.67 | 43 | 9.44 | 7.26
2008-07-21 | 0:00 | 28 | 97.3 | 16.55 | 43 | 9.49 | 7.26
2008-07-21 | 0:15 | 28 | 97.4 | 16.44 | 43 | 9.52 | 7.27
2008-07-21 | 0:30 | 28 | 97.5 | 16.32 | 44 | 9.56 | 7.26
2008-07-21 | 0:45 | 28 | 97.2 | 16.19 | 44 | 9.56 | 7.26
2008-07-21 | 1:01 | 28 | 97.1 | 16.06 | 44 | 9.57 | 7.26
2008-07-21 | 1:16 | 28 | 97.1 | 15.93 | 44 | 9.6 | 7.26
2008-07-21 | 1:31 | 28 | 97.2 | 15.8 | 44 | 9.63 | 7.26
2008-07-21 | 1:46 | 29 | 97.6 | 15.66 | 44 | 9.7 | 7.26
2008-07-21 | 2:01 | 29 | 97.8 | 15.52 | 44 | 9.75 | 7.26
2008-07-21 | 2:16 | 29 | 97.9 | 15.39 | 44 | 9.79 | 7.25
2008-07-21 | 2:31 | 29 | 98 | 15.28 | 44 | 9.82 | 7.26
2008-07-21 | 2:46 | 29 | 98.1 | 15.17 | 44 | 9.85 | 7.26
2008-07-21 | 3:01 | 29 | 97.7 | 15.07 | 44 | 9.84 | 7.26
2008-07-21 | 3:16 | 29 | 97.6 | 14.97 | 44 | 9.85 | 7.26
2008-07-21 | 3:31 | 29 | 97.6 | 14.86 | 44 | 9.87 | 7.25
2008-07-21 | 3:46 | 29 | 98 | 14.76 | 44 | 9.93 | 7.25
2008-07-21 | 4:01 | 29 | 97.8 | 14.67 | 44 | 9.93 | 7.25
2008-07-21 | 4:16 | 29 | 97.7 | 14.58 | 44 | 9.95 | 7.25
2008-07-21 | 4:31 | 29 | 98.1 | 14.48 | 44 | 10.01 | 7.25
2008-07-21 | 4:46 | 29 | 98.3 | 14.39 | 44 | 10.04 | 7.25
2008-07-21 | 5:01 | 29 | 98.4 | 14.29 | 44 | 10.08 | 7.25
2008-07-21 | 5:16 | 29 | 98.1 | 14.19 | 44 | 10.07 | 7.25
2008-07-21 | 5:31 | 29 | 98.4 | 14.08 | 44 | 10.12 | 7.25
2008-07-21 | 5:46 | 29 | 98.1 | 13.98 | 44 | 10.12 | 7.25
2008-07-21 | 6:01 | 29 | 98 | 13.87 | 44 | 10.13 | 7.25
2008-07-21 | 6:16 | 29 | 98.5 | 13.76 | 44 | 10.2 | 7.24
2008-07-21 | 6:31 | 28 | 98.4 | 13.65 | 44 | 10.22 | 7.25
2008-07-21 | 6:46 | 28 | 98.8 | 13.55 | 44 | 10.28 | 7.25
2008-07-21 | 7:01 | 28 | 99.1 | 13.46 | 43 | 10.33 | 7.25
2008-07-21 | 7:16 | 28 | 98.9 | 13.38 | 43 | 10.33 | 7.24
2008-07-21 | 7:31 | 28 | 99.3 | 13.31 | 43 | 10.39 | 7.25
2008-07-21 | 7:46 | 28 | 99.1 | 13.26 | 43 | 10.39 | 7.26
2008-07-21 | 8:01 | 28 | 99.6 | 13.22 | 43 | 10.44 | 7.26
2008-07-21 | 8:16 | 28 | 99.8 | 13.19 | 43 | 10.48 | 7.26
2008-07-21 | 8:31 | 28 | 100.1 | 13.18 | 43 | 10.5 | 7.27
2008-07-21 | 8:46 | 28 | 100.3 | 13.17 | 43 | 10.52 | 7.27
2008-07-21 | 9:01 | 28 | 100.5 | 13.18 | 43 | 10.54 | 7.27
2008-07-21 | 9:16 | 28 | 100.6 | 13.21 | 42 | 10.56 | 7.28
2008-07-21 | 9:31 | 27 | 100.8 | 13.24 | 42 | 10.56 | 7.28
2008-07-21 | 9:46 | 27 | 100.5 | 13.31 | 42 | 10.51 | 7.3
2008-07-21 | 10:01 | 27 | 100.8 | 13.4 | 42 | 10.52 | 7.29
2008-07-21 | 10:16 | 27 | 100.5 | 13.53 | 42 | 10.47 | 7.3
2008-07-21 | 10:31 | 27 | 100.4 | 13.7 | 42 | 10.41 | 7.3
2008-07-21 | 10:46 | 27 | 100.4 | 13.92 | 42 | 10.36 | 7.31
2008-07-21 | 11:01 | 27 | 100.9 | 14.19 | 42 | 10.35 | 7.32
2008-07-21 | 11:16 | 27 | 100.7 | 14.49 | 42 | 10.27 | 7.33
2008-07-21 | 11:31 | 27 | 100.7 | 14.8 | 42 | 10.2 | 7.33
2008-07-21 | 11:46 | 27 | 100.7 | 15.13 | 42 | 10.13 | 7.34
2008-07-21 | 12:01 | 27 | 100.7 | 15.48 | 42 | 10.05 | 7.35
2008-07-21 | 12:16 | 27 | 101 | 15.83 | 42 | 10.01 | 7.35
2008-07-21 | 12:31 | 27 | 100.7 | 16.2 | 42 | 9.9 | 7.36
2008-07-21 | 12:46 | 27 | 100.5 | 16.58 | 42 | 9.8 | 7.37
2008-07-21 | 13:00 | 27 | 100.4 | 16.97 | 42 | 9.71 | 7.38
2008-07-21 | 13:15 | 27 | 100.6 | 17.36 | 42 | 9.65 | 7.38
2008-07-21 | 13:30 | 27 | 100.5 | 17.74 | 42 | 9.57 | 7.39
2008-07-21 | 13:45 | 27 | 100.5 | 18.11 | 42 | 9.49 | 7.39
2008-07-21 | 14:00 | 27 | 100.3 | 18.46 | 42 | 9.41 | 7.4
2008-07-21 | 14:15 | 27 | 99.8 | 18.78 | 42 | 9.3 | 7.41
2008-07-21 | 14:30 | 27 | 99.8 | 19.07 | 42 | 9.25 | 7.41
2008-07-21 | 14:45 | 27 | 99.3 | 19.32 | 42 | 9.15 | 7.42
2008-07-21 | 15:00 | 27 | 99.3 | 19.56 | 42 | 9.11 | 7.42
2008-07-21 | 15:15 | 27 | 98.8 | 19.72 | 42 | 9.04 | 7.42
2008-07-21 | 15:30 | 27 | 98.8 | 19.76 | 42 | 9.02 | 7.42
2008-07-21 | 15:45 | 27 | 98.6 | 19.77 | 42 | 9.01 | 7.41
2008-07-21 | 16:00 | 27 | 98.6 | 19.85 | 42 | 8.99 | 7.41
2008-07-21 | 16:15 | 27 | 98.2 | 19.91 | 42 | 8.95 | 7.41
2008-07-21 | 16:30 | 27 | 98.4 | 20.07 | 42 | 8.94 | 7.41
2008-07-21 | 16:45 | 27 | 98.1 | 20.01 | 42 | 8.92 | 7.42
2008-07-21 | 17:00 | 27 | 98.2 | 19.89 | 42 | 8.95 | 7.42
2008-07-21 | 17:15 | 27 | 98.2 | 19.7 | 42 | 8.98 | 7.41
2008-07-21 | 17:30 | 27 | 98.2 | 19.61 | 42 | 8.99 | 7.4
2008-07-21 | 17:45 | 27 | 98 | 19.53 | 42 | 8.99 | 7.39
2008-07-21 | 18:00 | 27 | 97.4 | 19.46 | 42 | 8.95 | 7.38
2008-07-21 | 18:15 | 27 | 97 | 19.43 | 42 | 8.92 | 7.38
2008-07-21 | 18:30 | 27 | 96.7 | 19.39 | 42 | 8.9 | 7.37
2008-07-21 | 18:45 | 28 | 96.9 | 19.38 | 42 | 8.92 | 7.36
2008-07-21 | 19:00 | 27 | 97.1 | 19.37 | 42 | 8.94 | 7.35
2008-07-21 | 19:15 | 28 | 96.8 | 19.34 | 42 | 8.92 | 7.35
2008-07-21 | 19:30 | 28 | 97.1 | 19.29 | 42 | 8.96 | 7.35
2008-07-21 | 19:45 | 28 | 96.8 | 19.2 | 42 | 8.95 | 7.35
2008-07-21 | 20:00 | 28 | 97 | 19.12 | 43 | 8.98 | 7.34
2008-07-21 | 20:15 | 28 | 97 | 19.03 | 43 | 8.99 | 7.34
2008-07-21 | 20:30 | 28 | 96.9 | 18.9 | 43 | 9.01 | 7.33
2008-07-21 | 20:45 | 28 | 96.6 | 18.76 | 43 | 9 | 7.33
2008-07-21 | 21:00 | 28 | 96.7 | 18.64 | 43 | 9.04 | 7.33
2008-07-21 | 21:15 | 28 | 96.4 | 18.54 | 43 | 9.03 | 7.31
2008-07-21 | 21:30 | 28 | 96.7 | 18.44 | 43 | 9.07 | 7.31
2008-07-21 | 21:45 | 28 | 96.8 | 18.34 | 43 | 9.1 | 7.3
2008-07-21 | 22:00 | 28 | 96.5 | 18.24 | 43 | 9.1 | 7.3
2008-07-21 | 22:15 | 28 | 96.5 | 18.13 | 43 | 9.11 | 7.3
2008-07-21 | 22:30 | 28 | 96.8 | 18.01 | 43 | 9.16 | 7.3
2008-07-21 | 22:45 | 28 | 97 | 17.88 | 43 | 9.2 | 7.29
2008-07-21 | 23:00 | 28 | 96.8 | 17.73 | 43 | 9.21 | 7.3
2008-07-21 | 23:15 | 28 | 96.7 | 17.57 | 43 | 9.24 | 7.3
2008-07-21 | 23:30 | 28 | 97.1 | 17.41 | 44 | 9.3 | 7.3
2008-07-21 | 23:45 | 28 | 97.3 | 17.26 | 44 | 9.35 | 7.3
2008-07-22 | 0:00 | 28 | 97.3 | 17.12 | 44 | 9.38 | 7.29
2008-07-22 | 0:15 | 28 | 97.4 | 17 | 44 | 9.41 | 7.3
2008-07-22 | 0:30 | 28 | 97.5 | 16.89 | 44 | 9.44 | 7.29
2008-07-22 | 0:45 | 28 | 97.5 | 16.78 | 44 | 9.47 | 7.29
2008-07-22 | 1:01 | 28 | 97.5 | 16.68 | 44 | 9.49 | 7.29
2008-07-22 | 1:16 | 29 | 97.2 | 16.56 | 44 | 9.48 | 7.29
2008-07-22 | 1:31 | 29 | 97.1 | 16.45 | 44 | 9.49 | 7.28
2008-07-22 | 1:46 | 29 | 97.5 | 16.34 | 44 | 9.55 | 7.28
2008-07-22 | 2:01 | 29 | 97.7 | 16.23 | 44 | 9.59 | 7.28
2008-07-22 | 2:16 | 29 | 97.4 | 16.13 | 44 | 9.59 | 7.28
2008-07-22 | 2:31 | 29 | 97.3 | 16.04 | 44 | 9.6 | 7.28
2008-07-22 | 2:46 | 29 | 97.3 | 15.94 | 44 | 9.62 | 7.28
2008-07-22 | 3:01 | 29 | 97.8 | 15.85 | 44 | 9.68 | 7.28
2008-07-22 | 3:16 | 29 | 98 | 15.76 | 44 | 9.72 | 7.28
2008-07-22 | 3:31 | 29 | 98.1 | 15.67 | 44 | 9.75 | 7.28
2008-07-22 | 3:46 | 29 | 98.2 | 15.58 | 44 | 9.78 | 7.27
2008-07-22 | 4:01 | 29 | 97.8 | 15.49 | 44 | 9.76 | 7.27
2008-07-22 | 4:16 | 29 | 98.1 | 15.39 | 44 | 9.81 | 7.27
2008-07-22 | 4:31 | 29 | 98.2 | 15.28 | 44 | 9.85 | 7.27
2008-07-22 | 4:46 | 29 | 97.9 | 15.15 | 44 | 9.84 | 7.27
2008-07-22 | 5:01 | 29 | 97.8 | 15.01 | 44 | 9.86 | 7.27
2008-07-22 | 5:16 | 29 | 98.2 | 14.87 | 44 | 9.93 | 7.27
2008-07-22 | 5:31 | 28 | 98.4 | 14.73 | 44 | 9.98 | 7.27
2008-07-22 | 5:46 | 28 | 98.1 | 14.59 | 43 | 9.98 | 7.26
2008-07-22 | 6:01 | 28 | 98.1 | 14.44 | 43 | 10.01 | 7.26
2008-07-22 | 6:16 | 28 | 98.1 | 14.3 | 43 | 10.05 | 7.25
2008-07-22 | 6:31 | 28 | 98.2 | 14.17 | 43 | 10.09 | 7.25
2008-07-22 | 6:46 | 28 | 98.7 | 14.04 | 42 | 10.17 | 7.25
2008-07-22 | 7:01 | 27 | 98.6 | 13.92 | 42 | 10.19 | 7.25
2008-07-22 | 7:16 | 27 | 98.7 | 13.81 | 42 | 10.22 | 7.25
2008-07-22 | 7:31 | 27 | 98.8 | 13.72 | 42 | 10.25 | 7.25
2008-07-22 | 7:46 | 27 | 99.3 | 13.65 | 42 | 10.32 | 7.25
2008-07-22 | 8:01 | 27 | 99.7 | 13.6 | 41 | 10.36 | 7.25
2008-07-22 | 8:16 | 27 | 99.5 | 13.56 | 41 | 10.35 | 7.26
2008-07-22 | 8:31 | 27 | 100 | 13.55 | 41 | 10.4 | 7.26
2008-07-22 | 8:46 | 27 | 100.3 | 13.56 | 41 | 10.43 | 7.26
2008-07-22 | 9:01 | 27 | 100.4 | 13.58 | 41 | 10.44 | 7.27
2008-07-22 | 9:16 | 26 | 100.6 | 13.61 | 41 | 10.46 | 7.28
2008-07-22 | 9:31 | 26 | 100.3 | 13.65 | 41 | 10.42 | 7.28
2008-07-22 | 9:46 | 26 | 100.3 | 13.69 | 40 | 10.4 | 7.29
2008-07-22 | 10:01 | 26 | 100.7 | 13.76 | 40 | 10.43 | 7.29
2008-07-22 | 10:16 | 26 | 100.8 | 13.87 | 40 | 10.42 | 7.29
2008-07-22 | 10:31 | 26 | 100.5 | 14.03 | 40 | 10.35 | 7.3
2008-07-22 | 10:46 | 26 | 100.8 | 14.25 | 40 | 10.33 | 7.3
2008-07-22 | 11:01 | 26 | 100.6 | 14.51 | 40 | 10.26 | 7.31
2008-07-22 | 11:16 | 26 | 100.7 | 14.81 | 40 | 10.19 | 7.31
2008-07-22 | 11:31 | 26 | 101.1 | 15.11 | 40 | 10.17 | 7.32
2008-07-22 | 11:46 | 26 | 100.9 | 15.43 | 40 | 10.08 | 7.33
2008-07-22 | 12:01 | 26 | 101.3 | 15.78 | 40 | 10.04 | 7.34
2008-07-22 | 12:16 | 26 | 101.3 | 16.13 | 40 | 9.97 | 7.35
2008-07-22 | 12:31 | 26 | 101.3 | 16.49 | 40 | 9.9 | 7.35
2008-07-22 | 12:46 | 26 | 100.9 | 16.86 | 40 | 9.78 | 7.36
2008-07-22 | 13:00 | 26 | 100.6 | 17.23 | 40 | 9.68 | 7.37
2008-07-22 | 13:15 | 26 | 100.8 | 17.62 | 40 | 9.62 | 7.38
2008-07-22 | 13:30 | 26 | 100.8 | 17.98 | 40 | 9.55 | 7.39
2008-07-22 | 13:45 | 26 | 100.7 | 18.32 | 40 | 9.47 | 7.4
2008-07-22 | 14:00 | 26 | 100.6 | 18.65 | 40 | 9.39 | 7.41
2008-07-22 | 14:15 | 26 | 100.2 | 18.95 | 41 | 9.3 | 7.44