## month year sampled
|  | Source |  | state | state | state | state | state | state | state*
|  | Benficial use |  | aquatic | aquatic | aquatic | aquatic | aquatic | aquatic | aquatic
|  | Calculation |  | 7-day | acute | acute | acute | acute | acute | acute
|  | Criteria / guideline |  | 18C | 6.5 - 8.5 | 8.0 mg/L or 90% if temp/elev precludes | 110% | µS/cm | 500.0 mg/L | 6 NTU
Station ID | System | Description | date/time | temp | ph | DO | %Sat | Sp. Cond | calc TDS | Turb
WC 038.00 | Whychus Creek | Rd 1514 | 2008-06-06 16:22 | 3.04 | 5.71 | 14.3 | 106 | 24 | 20
WC 006.00 | Whychus Creek | Rd 6360 | 2008-06-06 14:53 | 8.93 | 6.78 | 12.56 | 109 | 37 | 20
* fish: based on state summer background median NTU = 3, and activities = 20 / 30 days add 3 NTU. recreation: unsuitable is 3+ NTU Source: draft turbidity criteria DEQ

## master equipment audit
Organization | UDWC
Program Name | WQMP
Contact Name | Lesley Jones
Number
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  |  |  |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  |  |  |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  |  |  |  | A
Post-sampling Field Equipment Check
Parameter | Audit reference | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  |  |  |  | A
DO (mg/L) | Winkler titration of surface water | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  |  |  |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  |  |  |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  |  |  |  | A
Comments: |  | Grade | Temp | DO | SpC | pH | Turb
|  |  | C | mg/L | µS/mL | SU | NTU
|  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Station ID 1
Organization | UDWC |  | Station ID |  | WC 038.00
Program Name | WQMP |  | System |  | Whychus Creek
Contact Name | Lesley Jones |  | Description |  | Rd 1514
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time |  | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL |  | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution |  | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-06-06 16:22 | 3.04 | 5.71 | 14.3 | 106 | 24 | 20
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 0 |  | 0 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Station ID 2
Organization | UDWC |  | Station ID |  | WC 006.00
Program Name | WQMP |  | System |  | Whychus Creek
Contact Name | Lesley Jones |  | Description |  | Rd 6360
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time |  | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL |  | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution |  | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-06-06 14:53 | 8.93 | 6.78 | 12.56 | 109 | 37 | 20
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 0 |  | 0 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Station ID 3
Organization | UDWC |  | Station ID |  | 0
Program Name | WQMP |  | System |  | 0
Contact Name | Lesley Jones |  | Description |  | 0
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time |  | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL |  | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution |  | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 0 | 0 | 0 | 0 | 0 | 0 | 0
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 0 |  | 0 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Station ID 4
Organization | UDWC |  | Station ID |  | 0
Program Name | WQMP |  | System |  | 0
Contact Name | Lesley Jones |  | Description |  | 0
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time |  | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL |  | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution |  | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 0 | 0 | 0 | 0 | 0 | 0 | 0
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 0 |  | 0 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Station ID 5
Organization | UDWC |  | Station ID |  | 0
Program Name | WQMP |  | System |  | 0
Contact Name | Lesley Jones |  | Description |  | 0
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time |  | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL |  | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution |  | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 0 | 0 | 0 | 0 | 0 | 0 | 0
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 0 |  | 0 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Station ID 6
Organization | UDWC |  | Station ID |  | 0
Program Name | WQMP |  | System |  | 0
Contact Name | Lesley Jones |  | Description |  | 0
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time |  | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL |  | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution |  | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 0 | 0 | 0 | 0 | 0 | 0 | 0
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 0 |  | 0 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Station ID 7
Organization | UDWC |  | Station ID |  | 0
Program Name | WQMP |  | System |  | 0
Contact Name | Lesley Jones |  | Description |  | 0
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time |  | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL |  | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution |  | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 0 | 0 | 0 | 0 | 0 | 0 | 0
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 0 |  | 0 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Station ID 8
Organization | UDWC |  | Station ID |  | 0
Program Name | WQMP |  | System |  | 0
Contact Name | Lesley Jones |  | Description |  | 0
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time |  | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL |  | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution |  | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 0 | 0 | 0 | 0 | 0 | 0 | 0
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 0 |  | 0 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Station ID 9
Organization | UDWC |  | Station ID |  | 0
Program Name | WQMP |  | System |  | 0
Contact Name | Lesley Jones |  | Description |  | 0
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-05 |  | 2008-06-05 |  | A
Time | Computer set to atomic time |  | 17:53 |  | 17:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 13.8 |  | 13.99 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.96 |  | B
SpC (uS/ml) | 148 µS/mL |  | 63 |  | 63 |  | A
ph (SU) | pH 7 standard solution |  | 7.94 |  | 8.44 |  | B
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 0 | 0 | 0 | 0 | 0 | 0 | 0
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-06 |  | 2008-06-06 |  | A
Time | Computer set to atomic time |  | 14:53 |  | 14:53 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 0 |  | 0 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.6 |  | 12.56 |  | Fail
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 150 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1