## master equipment audit
Organization | Upper Deschutes Watershed Council
Program Name | Water Quality Monitoring Program
Contact Name | Lesley Jones
Number | 541-382-6103
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time | 2008-08-11 |  | 2008-08-11 |  | A
Time | Computer set to atomic time | 16:22 |  | 16:22 |  | A
Temp (°C) | VWR NIST traceable surface water | 18.4 |  | 18.15 |  | A
DO (mg/L) | Winkler titration of surface water | 10.6 |  | 10.52 |  | A
SpC (uS/ml) | 148 µS/mL | 149.3 |  | 149 |  | A
ph (SU) | pH 7 standard solution | 7.02 |  | 7.04 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  |  |  |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  |  |  |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  |  |  |  | A
Comments: |  | Grade | Temp | DO | SpC | pH | Turb
|  |  | C | mg/L | µS/mL | SU | NTU
|  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1