## WC 006.00
Organization | UDWC |  | Station ID |  | WC 006.00
Program Name | WQMP |  | System |  | Whychus Creek
Contact Name | Lesley Jones/Mike Logan |  | Description |  | Rd 6360 Crossing
Number | 0 |  | Latitude / Longitude
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-25 |  | 2008-06-25 |  | A
Time | Computer set to atomic time |  | 12:04 |  | 12:04 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 14.8 |  | 14.96 |  | A
DO (mg/L) | Winkler titration of surface water |  | 10.5 |  | 10.58 |  | A
SpC (uS/ml) | 148 µS/mL |  | 59 |  | 60 |  | A
ph (SU) | pH 7 standard solution |  | 7.78 |  | 6.75 |  | Fail
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2008-06-26 14:15 | 12.57 | 6.47 | 10.73 | 101 | 30 | 20
0
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2008-06-30 |  | 2008-06-30 |  | A
Time | Computer set to atomic time |  | 14:40 |  | 14:40 |  | A
Temp (°C) | VWR NIST traceable surface water |  | 18.2 |  | 18.45 |  | A
DO (mg/L) | Winkler titration of surface water |  | 9.2 |  | 9.34 |  | A
SpC (uS/ml) | 148 µS/mL |  | 147 |  | 144 |  | A
ph (SU) | pH 7 standard solution |  | 7.02 |  | 7.23 |  | A
Turb (NTU) | Gelex secondary 0-5 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-500 NTU |  | 0 |  | 0 |  | A
Turb (NTU) | Gelex secondary 0-1000 NTU |  | 0 |  | 0 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1
Formulas:
Audit value = '[2]master equipment audit'!D9:E9
Field equip. value = '[2]master equipment audit'!F9:G9
Audit reference = '[2]month year sampled'!D8