## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AA (Audited against Sonde AB, NIST & Pro-Sereis (loner))
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-17 |  | 2009-06-17 |  | A
Time | Computer set to atomic time |  | 11:17 |  | 11:17 |  | A
Temp | In Situ |  | 16.2 |  | 16.41 |  | A
DO | In Situ |  | 9.89 |  | 10.24 |  | B
SpC | In Situ |  | 55.9 |  | 57 |  | A
Ph | In Situ |  | 8 |  | 8.15 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-06-18 12:00 | 10.2 | 7.54 | 11 | 97.9 | 30.5 | 19.5
Sonde | 2009-06-18 12:00 | 10.21 | 7.64 | 11.53 | 102.7 | 30 | 19
Grade |  | A | A | B |  | A
Audit 2 | 2009-06-23 10:59 | 10.4 | 7.68 | 10.98 | 98.2 | 41.2 | 26.7
Sonde | 2009-06-23 11:00 | 10.44 | 7.84 | 11.89 | 106.4 | 40 | 26
Grade |  | A | A | B |  | A
Audit 3 | 2009-06-29 12:04 | 12.9 | 7.66 | 10.13 | 95.9 | 35.2 | 22.75
Sonde | 2009-06-29 12:00 | 12.83 | 7.77 | 11.61 | 109.8 | 36 | 23
Grade |  | A | A | Fail |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AA (Audited against Sonde AB, NIST & Pro-Sereis (loner))
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-29 |  | 2009-06-29 |  | A
Time | Computer set to atomic time |  | 12:10 |  | 12:10 |  | A
Temp | VWR NIST traceable surface water |  | 12.9 |  | 13.04 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 11.02 |  | B
SpC | In Situ |  | 35.5 |  | 37 |  | A
Ph | In Situ |  | 7.68 |  | 7.78 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
DO data can be kept considering it passed on the Pre and Post Deployment audits and data didn't indicate any standards were violated |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-15 |  | 2009-06-15 |  | A
Time | Computer set to atomic time |  | 14:48 |  | 14:48 |  | A
Temp | VWR NIST traceable surface water |  | 11.3 |  | 11.5 |  | A
DO | Winkler titration of surface water |  | 10.2 |  | 10.14 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 154.3 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-06-18 12:00 | 10.2 | 7.54 | 11 | 97.9 | 30.5 | 19.5
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-18 |  | 2009-06-18 |  | A
Time | Computer set to atomic time |  | 15:12 |  | 15:12 |  | A
Temp | VWR NIST traceable surface water |  | 10.9 |  | 11 |  | A
DO | Winkler titration of surface water |  | 9.8 |  | 10.31 |  | B
SpC | 148 µS/mL |  | 156.7 |  | 156.2 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.05 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-18 |  | 2009-06-18 |  | A
Time | Computer set to atomic time |  | 15:12 |  | 15:12 |  | A
Temp | VWR NIST traceable surface water |  | 10.9 |  | 11 |  | A
DO | Winkler titration of surface water |  | 9.8 |  | 10.31 |  | B
SpC | 148 µS/mL |  | 156.7 |  | 156.2 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.05 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-06-23 10:59 | 10.4 | 7.68 | 10.98 | 98.2 | 41.2 | 26.7
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-23 |  | 2009-06-23 |  | A
Time | Computer set to atomic time |  | 11:51 |  | 11:51 |  | A
Temp | VWR NIST traceable surface water |  | 17 |  | 17.3 |  | A
DO | Winkler titration of surface water |  | 10 |  | 9.79 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 159.7 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.03 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-23 |  | 2009-06-23 |  | A
Time | Computer set to atomic time |  | 11:51 |  | 11:51 |  | A
Temp | VWR NIST traceable surface water |  | 17 |  | 17.3 |  | A
DO | Winkler titration of surface water |  | 10 |  | 9.79 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 159.7 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.03 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-06-29 12:04 | 12.9 | 7.66 | 10.13 | 95.9 | 35.2 | 22.75
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-29 |  | 2009-06-29 |  | A
Time | Computer set to atomic time |  | 12:15 |  | 12:15 |  | A
Temp | VWR NIST traceable surface water |  | 13 |  | 13.1 |  | A
DO | Winkler titration of surface water |  | 10.1 |  | 10.07 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 182.6 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.01 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
SpC readings in field matched Sonde and seemed very accurate |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | TDS | SpCond | DO Conc | DO% | pH
M/D/Y |  | °C | mg/L | uS/cm | mg/L | %
2009-06-18 | 12:00 | 10.21 | 19 | 30 | 11.53 | 102.7 | 7.64
2009-06-18 | 12:16 | 10.49 | 19 | 30 | 11.9 | 106.6 | 7.66
2009-06-18 | 12:31 | 10.77 | 19 | 30 | 11.84 | 106.8 | 7.68
2009-06-18 | 12:46 | 11.07 | 19 | 30 | 11.75 | 106.7 | 7.68
2009-06-18 | 13:00 | 11.36 | 19 | 30 | 11.62 | 106.2 | 7.7
2009-06-18 | 13:15 | 11.67 | 20 | 30 | 11.5 | 105.9 | 7.71
2009-06-18 | 13:30 | 11.9 | 20 | 30 | 11.47 | 106.2 | 7.71
2009-06-18 | 13:45 | 11.98 | 20 | 30 | 11.35 | 105.3 | 7.72
2009-06-18 | 14:00 | 12.4 | 20 | 30 | 11.29 | 105.7 | 7.72
2009-06-18 | 14:15 | 12.82 | 20 | 30 | 11.12 | 105.1 | 7.73
2009-06-18 | 14:30 | 12.91 | 20 | 30 | 11.03 | 104.5 | 7.73
2009-06-18 | 14:45 | 13.01 | 20 | 30 | 11.04 | 104.8 | 7.73
2009-06-18 | 15:00 | 13.36 | 20 | 30 | 10.97 | 105 | 7.74
2009-06-18 | 15:15 | 13.37 | 20 | 31 | 10.89 | 104.2 | 7.75
2009-06-18 | 15:30 | 13.39 | 20 | 31 | 10.9 | 104.3 | 7.75
2009-06-18 | 15:45 | 13.45 | 20 | 31 | 10.87 | 104.2 | 7.74
2009-06-18 | 16:00 | 13.53 | 20 | 31 | 10.84 | 104.1 | 7.74
2009-06-18 | 16:15 | 13.59 | 20 | 31 | 10.75 | 103.4 | 7.75
2009-06-18 | 16:30 | 13.6 | 20 | 31 | 10.75 | 103.4 | 7.74
2009-06-18 | 16:45 | 13.56 | 20 | 31 | 10.71 | 102.9 | 7.74
2009-06-18 | 17:00 | 13.44 | 20 | 31 | 10.76 | 103.1 | 7.75
2009-06-18 | 17:15 | 13.35 | 20 | 31 | 10.77 | 103 | 7.74
2009-06-18 | 17:30 | 13.32 | 20 | 31 | 10.76 | 102.9 | 7.74
2009-06-18 | 17:45 | 13.32 | 20 | 31 | 10.69 | 102.2 | 7.74
2009-06-18 | 18:00 | 13.33 | 20 | 31 | 10.65 | 101.8 | 7.74
2009-06-18 | 18:15 | 13.34 | 20 | 31 | 10.62 | 101.5 | 7.74
2009-06-18 | 18:30 | 13.39 | 20 | 31 | 10.59 | 101.4 | 7.74
2009-06-18 | 18:45 | 13.4 | 20 | 31 | 10.63 | 101.8 | 7.73
2009-06-18 | 19:00 | 13.37 | 21 | 32 | 10.6 | 101.5 | 7.73
2009-06-18 | 19:15 | 13.32 | 21 | 32 | 10.59 | 101.2 | 7.73
2009-06-18 | 19:30 | 13.27 | 21 | 32 | 10.58 | 101 | 7.73
2009-06-18 | 19:45 | 13.26 | 21 | 32 | 10.56 | 100.9 | 7.73
2009-06-18 | 20:00 | 13.23 | 21 | 32 | 10.6 | 101.1 | 7.72
2009-06-18 | 20:15 | 13.15 | 21 | 32 | 10.58 | 100.7 | 7.72
2009-06-18 | 20:30 | 13.06 | 21 | 32 | 10.62 | 100.9 | 7.72
2009-06-18 | 20:45 | 12.97 | 21 | 32 | 10.59 | 100.5 | 7.71
2009-06-18 | 21:00 | 12.9 | 21 | 32 | 10.58 | 100.2 | 7.71
2009-06-18 | 21:15 | 12.82 | 21 | 32 | 10.64 | 100.5 | 7.7
2009-06-18 | 21:30 | 12.74 | 21 | 32 | 10.61 | 100.1 | 7.7
2009-06-18 | 21:45 | 12.66 | 21 | 32 | 10.61 | 100 | 7.7
2009-06-18 | 22:00 | 12.59 | 21 | 32 | 10.65 | 100.1 | 7.7
2009-06-18 | 22:15 | 12.51 | 21 | 33 | 10.65 | 100 | 7.69
2009-06-18 | 22:30 | 12.43 | 21 | 33 | 10.64 | 99.7 | 7.7
2009-06-18 | 22:45 | 12.34 | 21 | 33 | 10.67 | 99.8 | 7.69
2009-06-18 | 23:00 | 12.23 | 21 | 33 | 10.69 | 99.8 | 7.69
2009-06-18 | 23:15 | 12.14 | 21 | 33 | 10.71 | 99.7 | 7.7
2009-06-18 | 23:30 | 12.08 | 21 | 33 | 10.77 | 100.1 | 7.7
2009-06-18 | 23:45 | 12.04 | 21 | 33 | 10.8 | 100.3 | 7.69
2009-06-19 | 0:00 | 11.98 | 21 | 33 | 10.82 | 100.4 | 7.69
2009-06-19 | 0:15 | 11.92 | 21 | 33 | 10.8 | 100 | 7.69
2009-06-19 | 0:30 | 11.86 | 21 | 33 | 10.79 | 99.8 | 7.69
2009-06-19 | 0:45 | 11.8 | 21 | 33 | 10.79 | 99.7 | 7.69
2009-06-19 | 1:01 | 11.76 | 21 | 33 | 10.85 | 100.2 | 7.69
2009-06-19 | 1:16 | 11.71 | 21 | 32 | 10.83 | 99.9 | 7.69
2009-06-19 | 1:31 | 11.65 | 21 | 32 | 10.85 | 99.9 | 7.69
2009-06-19 | 1:46 | 11.6 | 21 | 32 | 10.85 | 99.7 | 7.69
2009-06-19 | 2:01 | 11.55 | 21 | 32 | 10.91 | 100.2 | 7.68
2009-06-19 | 2:16 | 11.51 | 21 | 32 | 10.93 | 100.3 | 7.69
2009-06-19 | 2:31 | 11.48 | 21 | 32 | 10.95 | 100.4 | 7.68
2009-06-19 | 2:46 | 11.45 | 21 | 32 | 10.96 | 100.5 | 7.68
2009-06-19 | 3:01 | 11.42 | 21 | 32 | 10.93 | 100 | 7.68
2009-06-19 | 3:16 | 11.4 | 20 | 31 | 10.91 | 99.9 | 7.68
2009-06-19 | 3:31 | 11.38 | 20 | 31 | 10.96 | 100.2 | 7.67
2009-06-19 | 3:46 | 11.36 | 20 | 31 | 10.98 | 100.4 | 7.66
2009-06-19 | 4:01 | 11.32 | 20 | 31 | 11 | 100.5 | 7.66
2009-06-19 | 4:16 | 11.26 | 20 | 31 | 10.97 | 100.1 | 7.66
2009-06-19 | 4:31 | 11.19 | 20 | 31 | 10.97 | 99.9 | 7.66
2009-06-19 | 4:46 | 11.1 | 20 | 31 | 11.04 | 100.4 | 7.65
2009-06-19 | 5:01 | 11 | 20 | 31 | 11.04 | 100.1 | 7.65
2009-06-19 | 5:16 | 10.9 | 20 | 31 | 11.06 | 100.1 | 7.66
2009-06-19 | 5:31 | 10.78 | 20 | 30 | 11.14 | 100.5 | 7.65
2009-06-19 | 5:46 | 10.67 | 20 | 30 | 11.15 | 100.3 | 7.65
2009-06-19 | 6:01 | 10.55 | 20 | 30 | 11.17 | 100.2 | 7.65
2009-06-19 | 6:16 | 10.43 | 20 | 30 | 11.21 | 100.3 | 7.65
2009-06-19 | 6:31 | 10.31 | 20 | 30 | 11.31 | 101 | 7.64
2009-06-19 | 6:46 | 10.21 | 20 | 30 | 11.32 | 100.8 | 7.65
2009-06-19 | 7:01 | 10.13 | 20 | 30 | 11.35 | 100.9 | 7.65
2009-06-19 | 7:16 | 10.05 | 19 | 30 | 11.39 | 101 | 7.65
2009-06-19 | 7:31 | 10 | 19 | 30 | 11.42 | 101.1 | 7.66
2009-06-19 | 7:46 | 9.96 | 19 | 30 | 11.45 | 101.4 | 7.66
2009-06-19 | 8:01 | 9.94 | 19 | 30 | 11.54 | 102.1 | 7.66
2009-06-19 | 8:16 | 9.92 | 19 | 30 | 11.58 | 102.4 | 7.66
2009-06-19 | 8:31 | 9.91 | 19 | 30 | 11.59 | 102.5 | 7.67
2009-06-19 | 8:46 | 9.9 | 19 | 30 | 11.68 | 103.2 | 7.67
2009-06-19 | 9:01 | 9.87 | 19 | 30 | 11.67 | 103.1 | 7.67
2009-06-19 | 9:16 | 9.86 | 19 | 30 | 11.72 | 103.5 | 7.67
2009-06-19 | 9:31 | 9.88 | 19 | 30 | 11.68 | 103.2 | 7.67
2009-06-19 | 9:46 | 9.99 | 19 | 30 | 11.71 | 103.7 | 7.68
2009-06-19 | 10:01 | 10.15 | 19 | 30 | 11.7 | 104 | 7.68
2009-06-19 | 10:16 | 10.3 | 19 | 30 | 11.66 | 104 | 7.69
2009-06-19 | 10:31 | 10.44 | 19 | 30 | 11.58 | 103.6 | 7.69
2009-06-19 | 10:46 | 10.55 | 19 | 30 | 11.59 | 104 | 7.69
2009-06-19 | 11:01 | 10.65 | 19 | 30 | 11.58 | 104.2 | 7.7
2009-06-19 | 11:16 | 10.71 | 20 | 30 | 11.6 | 104.5 | 7.7
2009-06-19 | 11:31 | 10.82 | 20 | 30 | 11.56 | 104.3 | 7.7
2009-06-19 | 11:46 | 11.04 | 20 | 30 | 11.54 | 104.7 | 7.72
2009-06-19 | 12:01 | 11.28 | 20 | 30 | 11.54 | 105.3 | 7.71
2009-06-19 | 12:16 | 11.7 | 20 | 30 | 11.48 | 105.8 | 7.71
2009-06-19 | 12:31 | 11.98 | 20 | 30 | 11.38 | 105.6 | 7.72
2009-06-19 | 12:46 | 12.22 | 20 | 30 | 11.32 | 105.6 | 7.72
2009-06-19 | 13:00 | 12.57 | 20 | 30 | 11.26 | 105.8 | 7.73
2009-06-19 | 13:15 | 12.89 | 20 | 30 | 11.12 | 105.3 | 7.73
2009-06-19 | 13:30 | 13.2 | 20 | 30 | 11.07 | 105.5 | 7.74
2009-06-19 | 13:45 | 13.52 | 20 | 30 | 10.99 | 105.5 | 7.73
2009-06-19 | 14:00 | 13.75 | 20 | 30 | 10.92 | 105.4 | 7.75
2009-06-19 | 14:15 | 13.88 | 20 | 30 | 10.86 | 105.1 | 7.74
2009-06-19 | 14:30 | 13.98 | 20 | 30 | 10.81 | 104.9 | 7.75
2009-06-19 | 14:45 | 14.12 | 20 | 30 | 10.75 | 104.6 | 7.75
2009-06-19 | 15:00 | 14.21 | 20 | 31 | 10.7 | 104.3 | 7.74
2009-06-19 | 15:15 | 14.4 | 20 | 31 | 10.59 | 103.6 | 7.74
2009-06-19 | 15:30 | 14.63 | 20 | 31 | 10.47 | 103.1 | 7.75
2009-06-19 | 15:45 | 14.88 | 20 | 31 | 10.42 | 103.1 | 7.76
2009-06-19 | 16:00 | 15.02 | 20 | 31 | 10.38 | 102.9 | 7.75
2009-06-19 | 16:15 | 15.01 | 20 | 31 | 10.36 | 102.8 | 7.75
2009-06-19 | 16:30 | 14.98 | 20 | 31 | 10.36 | 102.7 | 7.75
2009-06-19 | 16:45 | 14.98 | 20 | 31 | 10.34 | 102.5 | 7.75
2009-06-19 | 17:00 | 14.98 | 20 | 31 | 10.32 | 102.3 | 7.75
2009-06-19 | 17:15 | 14.99 | 20 | 31 | 10.29 | 102.1 | 7.74
2009-06-19 | 17:30 | 14.97 | 20 | 31 | 10.28 | 101.9 | 7.75
2009-06-19 | 17:45 | 14.94 | 20 | 31 | 10.27 | 101.8 | 7.74
2009-06-19 | 18:00 | 14.92 | 20 | 31 | 10.26 | 101.6 | 7.74
2009-06-19 | 18:15 | 14.92 | 20 | 31 | 10.25 | 101.5 | 7.74
2009-06-19 | 18:30 | 14.89 | 20 | 31 | 10.25 | 101.4 | 7.74
2009-06-19 | 18:45 | 14.83 | 20 | 31 | 10.21 | 100.8 | 7.74
2009-06-19 | 19:00 | 14.79 | 20 | 31 | 10.22 | 100.9 | 7.73
2009-06-19 | 19:15 | 14.75 | 20 | 31 | 10.24 | 101 | 7.73
2009-06-19 | 19:30 | 14.67 | 20 | 31 | 10.25 | 101 | 7.73
2009-06-19 | 19:45 | 14.59 | 20 | 31 | 10.27 | 100.9 | 7.72
2009-06-19 | 20:00 | 14.47 | 20 | 31 | 10.29 | 100.9 | 7.72
2009-06-19 | 20:15 | 14.33 | 20 | 31 | 10.32 | 100.8 | 7.71
2009-06-19 | 20:30 | 14.23 | 20 | 31 | 10.28 | 100.2 | 7.72
2009-06-19 | 20:45 | 14.15 | 20 | 31 | 10.32 | 100.5 | 7.71
2009-06-19 | 21:00 | 14.05 | 20 | 31 | 10.3 | 100 | 7.7
2009-06-19 | 21:15 | 13.95 | 20 | 31 | 10.35 | 100.3 | 7.7
2009-06-19 | 21:30 | 13.84 | 20 | 31 | 10.33 | 99.9 | 7.69
2009-06-19 | 21:45 | 13.71 | 20 | 31 | 10.4 | 100.2 | 7.69
2009-06-19 | 22:00 | 13.59 | 20 | 31 | 10.39 | 99.9 | 7.69
2009-06-19 | 22:15 | 13.46 | 20 | 31 | 10.46 | 100.3 | 7.68
2009-06-19 | 22:30 | 13.34 | 20 | 31 | 10.45 | 99.9 | 7.68
2009-06-19 | 22:45 | 13.19 | 20 | 31 | 10.53 | 100.4 | 7.67
2009-06-19 | 23:00 | 13.03 | 20 | 31 | 10.54 | 100.1 | 7.67
2009-06-19 | 23:15 | 12.86 | 20 | 31 | 10.57 | 100 | 7.66
2009-06-19 | 23:30 | 12.68 | 20 | 31 | 10.61 | 100 | 7.66
2009-06-19 | 23:45 | 12.51 | 20 | 30 | 10.7 | 100.4 | 7.65
2009-06-20 | 0:00 | 12.33 | 20 | 30 | 10.72 | 100.2 | 7.65
2009-06-20 | 0:15 | 12.17 | 20 | 30 | 10.76 | 100.2 | 7.65
2009-06-20 | 0:30 | 12.02 | 19 | 30 | 10.79 | 100.2 | 7.64
2009-06-20 | 0:45 | 11.9 | 19 | 30 | 10.82 | 100.2 | 7.64
2009-06-20 | 1:01 | 11.77 | 19 | 30 | 10.85 | 100.2 | 7.64
2009-06-20 | 1:16 | 11.65 | 19 | 29 | 10.88 | 100.2 | 7.63
2009-06-20 | 1:31 | 11.52 | 19 | 29 | 10.98 | 100.7 | 7.63
2009-06-20 | 1:46 | 11.38 | 19 | 29 | 11.03 | 100.9 | 7.62
2009-06-20 | 2:01 | 11.24 | 19 | 29 | 11.09 | 101.1 | 7.61
2009-06-20 | 2:16 | 11.09 | 19 | 29 | 11.08 | 100.7 | 7.62
2009-06-20 | 2:31 | 10.94 | 19 | 29 | 11.16 | 101.1 | 7.61
2009-06-20 | 2:46 | 10.8 | 19 | 29 | 11.16 | 100.7 | 7.61
2009-06-20 | 3:01 | 10.65 | 19 | 29 | 11.19 | 100.7 | 7.61
2009-06-20 | 3:16 | 10.5 | 19 | 29 | 11.23 | 100.6 | 7.6
2009-06-20 | 3:31 | 10.34 | 18 | 28 | 11.28 | 100.7 | 7.6
2009-06-20 | 3:46 | 10.2 | 18 | 28 | 11.32 | 100.7 | 7.6
2009-06-20 | 4:01 | 10.06 | 18 | 28 | 11.41 | 101.2 | 7.59
2009-06-20 | 4:16 | 9.92 | 18 | 28 | 11.43 | 101.1 | 7.58
2009-06-20 | 4:31 | 9.79 | 18 | 28 | 11.46 | 101 | 7.58
2009-06-20 | 4:46 | 9.66 | 18 | 28 | 11.51 | 101.1 | 7.58
2009-06-20 | 5:01 | 9.54 | 18 | 28 | 11.59 | 101.6 | 7.59
2009-06-20 | 5:16 | 9.42 | 18 | 28 | 11.59 | 101.3 | 7.58
2009-06-20 | 5:31 | 9.31 | 18 | 28 | 11.62 | 101.3 | 7.58
2009-06-20 | 5:46 | 9.19 | 18 | 28 | 11.63 | 101.1 | 7.57
2009-06-20 | 6:01 | 9.08 | 18 | 28 | 11.68 | 101.3 | 7.58
2009-06-20 | 6:16 | 8.98 | 18 | 28 | 11.77 | 101.8 | 7.58
2009-06-20 | 6:31 | 8.88 | 18 | 28 | 11.84 | 102.1 | 7.58
2009-06-20 | 6:46 | 8.78 | 18 | 28 | 11.83 | 101.8 | 7.57
2009-06-20 | 7:01 | 8.69 | 18 | 28 | 11.86 | 101.8 | 7.57
2009-06-20 | 7:16 | 8.62 | 18 | 28 | 11.89 | 101.9 | 7.58
2009-06-20 | 7:31 | 8.56 | 18 | 28 | 11.97 | 102.5 | 7.59
2009-06-20 | 7:46 | 8.51 | 18 | 28 | 12.01 | 102.7 | 7.6
2009-06-20 | 8:01 | 8.48 | 18 | 28 | 12.07 | 103.1 | 7.6
2009-06-20 | 8:16 | 8.46 | 18 | 28 | 12.04 | 102.8 | 7.61
2009-06-20 | 8:31 | 8.45 | 18 | 28 | 12.05 | 102.8 | 7.61
2009-06-20 | 8:46 | 8.45 | 18 | 28 | 12.03 | 102.7 | 7.62
2009-06-20 | 9:01 | 8.45 | 18 | 28 | 12.05 | 102.9 | 7.62
2009-06-20 | 9:16 | 8.44 | 18 | 28 | 12.14 | 103.6 | 7.63
2009-06-20 | 9:31 | 8.44 | 18 | 28 | 12.19 | 104 | 7.63
2009-06-20 | 9:46 | 8.44 | 18 | 28 | 12.22 | 104.3 | 7.63
2009-06-20 | 10:01 | 8.45 | 18 | 28 | 12.25 | 104.6 | 7.64
2009-06-20 | 10:16 | 8.46 | 18 | 28 | 12.26 | 104.7 | 7.64
2009-06-20 | 10:31 | 8.48 | 18 | 28 | 12.21 | 104.3 | 7.65
2009-06-20 | 10:46 | 8.51 | 18 | 28 | 12.25 | 104.7 | 7.64
2009-06-20 | 11:01 | 8.57 | 18 | 28 | 12.21 | 104.5 | 7.65
2009-06-20 | 11:16 | 8.62 | 19 | 29 | 12.26 | 105.1 | 7.66
2009-06-20 | 11:31 | 8.67 | 19 | 29 | 12.28 | 105.3 | 7.65
2009-06-20 | 11:46 | 8.76 | 19 | 29 | 12.26 | 105.5 | 7.66
2009-06-20 | 12:01 | 8.88 | 19 | 29 | 12.22 | 105.4 | 7.67
2009-06-20 | 12:16 | 8.92 | 19 | 29 | 12.22 | 105.6 | 7.66
2009-06-20 | 12:31 | 9.12 | 19 | 29 | 12.21 | 105.9 | 7.66
2009-06-20 | 12:46 | 9.18 | 19 | 29 | 12.19 | 105.9 | 7.67
2009-06-20 | 13:00 | 9.38 | 19 | 29 | 12.08 | 105.5 | 7.68
2009-06-20 | 13:15 | 9.62 | 19 | 29 | 11.99 | 105.3 | 7.68
2009-06-20 | 13:30 | 9.63 | 19 | 29 | 11.96 | 105 | 7.69
2009-06-20 | 13:45 | 9.64 | 19 | 29 | 11.98 | 105.2 | 7.68
2009-06-20 | 14:00 | 9.87 | 19 | 29 | 11.94 | 105.4 | 7.69
2009-06-20 | 14:15 | 9.84 | 19 | 29 | 11.94 | 105.4 | 7.69
2009-06-20 | 14:30 | 10.21 | 19 | 29 | 11.82 | 105.3 | 7.7
2009-06-20 | 14:45 | 10.59 | 19 | 29 | 11.73 | 105.3 | 7.71
2009-06-20 | 15:00 | 10.91 | 19 | 30 | 11.64 | 105.3 | 7.7
2009-06-20 | 15:15 | 11.22 | 19 | 30 | 11.49 | 104.7 | 7.72
2009-06-20 | 15:30 | 11.4 | 19 | 30 | 11.44 | 104.7 | 7.71
2009-06-20 | 15:45 | 11.53 | 19 | 30 | 11.38 | 104.4 | 7.71
2009-06-20 | 16:00 | 11.62 | 19 | 30 | 11.3 | 104 | 7.71
2009-06-20 | 16:15 | 11.9 | 19 | 30 | 11.23 | 104 | 7.71
2009-06-20 | 16:30 | 11.99 | 20 | 30 | 11.16 | 103.5 | 7.72
2009-06-20 | 16:45 | 11.98 | 20 | 30 | 11.12 | 103.1 | 7.72
2009-06-20 | 17:00 | 11.94 | 20 | 30 | 11.16 | 103.5 | 7.72
2009-06-20 | 17:15 | 11.89 | 20 | 30 | 11.14 | 103.1 | 7.72
2009-06-20 | 17:30 | 11.85 | 20 | 30 | 11.13 | 102.9 | 7.72
2009-06-20 | 17:45 | 11.94 | 20 | 30 | 11.09 | 102.8 | 7.72
2009-06-20 | 18:00 | 12.02 | 20 | 30 | 11.09 | 103 | 7.72
2009-06-20 | 18:15 | 12.04 | 20 | 30 | 11.08 | 102.9 | 7.72
2009-06-20 | 18:30 | 12.01 | 20 | 31 | 11.03 | 102.4 | 7.71
2009-06-20 | 18:45 | 12.01 | 20 | 31 | 11.04 | 102.4 | 7.71
2009-06-20 | 19:00 | 12.02 | 20 | 31 | 10.98 | 102 | 7.71
2009-06-20 | 19:15 | 12.05 | 20 | 31 | 10.95 | 101.7 | 7.71
2009-06-20 | 19:30 | 12.06 | 20 | 31 | 10.98 | 102 | 7.71
2009-06-20 | 19:45 | 12 | 20 | 31 | 11 | 102.1 | 7.71
2009-06-20 | 20:00 | 11.88 | 20 | 31 | 11.03 | 102.1 | 7.71
2009-06-20 | 20:15 | 11.76 | 20 | 31 | 11.06 | 102 | 7.7
2009-06-20 | 20:30 | 11.65 | 20 | 31 | 11.03 | 101.5 | 7.7
2009-06-20 | 20:45 | 11.57 | 20 | 31 | 11.06 | 101.6 | 7.7
2009-06-20 | 21:00 | 11.49 | 20 | 31 | 11.09 | 101.7 | 7.69
2009-06-20 | 21:15 | 11.39 | 20 | 31 | 11.06 | 101.2 | 7.69
2009-06-20 | 21:30 | 11.29 | 20 | 31 | 11.07 | 101 | 7.69
2009-06-20 | 21:45 | 11.18 | 20 | 31 | 11.08 | 100.9 | 7.69
2009-06-20 | 22:00 | 11.11 | 21 | 32 | 11.09 | 100.8 | 7.68
2009-06-20 | 22:15 | 11.06 | 21 | 32 | 11.1 | 100.8 | 7.68
2009-06-20 | 22:30 | 11.01 | 21 | 32 | 11.16 | 101.2 | 7.67
2009-06-20 | 22:45 | 10.99 | 21 | 32 | 11.18 | 101.4 | 7.68
2009-06-20 | 23:00 | 10.98 | 21 | 32 | 11.19 | 101.4 | 7.68
2009-06-20 | 23:15 | 10.94 | 21 | 32 | 11.15 | 101 | 7.68
2009-06-20 | 23:30 | 10.88 | 21 | 32 | 11.16 | 100.9 | 7.68
2009-06-20 | 23:45 | 10.78 | 21 | 32 | 11.17 | 100.8 | 7.68
2009-06-21 | 0:00 | 10.66 | 21 | 32 | 11.25 | 101.2 | 7.68
2009-06-21 | 0:15 | 10.54 | 21 | 32 | 11.32 | 101.5 | 7.68
2009-06-21 | 0:30 | 10.41 | 21 | 32 | 11.31 | 101.2 | 7.68
2009-06-21 | 0:45 | 10.3 | 21 | 32 | 11.38 | 101.5 | 7.67
2009-06-21 | 1:01 | 10.19 | 21 | 32 | 11.43 | 101.7 | 7.67
2009-06-21 | 1:16 | 10.1 | 21 | 32 | 11.46 | 101.8 | 7.67
2009-06-21 | 1:31 | 10.03 | 21 | 32 | 11.49 | 101.9 | 7.67
2009-06-21 | 1:46 | 9.96 | 21 | 32 | 11.46 | 101.5 | 7.67
2009-06-21 | 2:01 | 9.9 | 21 | 32 | 11.46 | 101.3 | 7.66
2009-06-21 | 2:16 | 9.83 | 21 | 32 | 11.48 | 101.3 | 7.67
2009-06-21 | 2:31 | 9.75 | 21 | 32 | 11.49 | 101.2 | 7.66
2009-06-21 | 2:46 | 9.65 | 21 | 32 | 11.57 | 101.7 | 7.66
2009-06-21 | 3:01 | 9.56 | 21 | 32 | 11.64 | 102 | 7.66
2009-06-21 | 3:16 | 9.46 | 21 | 32 | 11.62 | 101.7 | 7.65
2009-06-21 | 3:31 | 9.36 | 20 | 32 | 11.64 | 101.5 | 7.65
2009-06-21 | 3:46 | 9.26 | 20 | 31 | 11.66 | 101.5 | 7.65
2009-06-21 | 4:01 | 9.16 | 20 | 31 | 11.69 | 101.5 | 7.65
2009-06-21 | 4:16 | 9.05 | 20 | 31 | 11.77 | 102 | 7.65
2009-06-21 | 4:31 | 8.96 | 20 | 31 | 11.77 | 101.8 | 7.65
2009-06-21 | 4:46 | 8.87 | 20 | 31 | 11.79 | 101.7 | 7.65
2009-06-21 | 5:01 | 8.79 | 20 | 31 | 11.87 | 102.2 | 7.64
2009-06-21 | 5:16 | 8.71 | 20 | 31 | 11.91 | 102.3 | 7.65
2009-06-21 | 5:31 | 8.63 | 20 | 31 | 11.9 | 102 | 7.64
2009-06-21 | 5:46 | 8.56 | 20 | 31 | 11.91 | 101.9 | 7.65
2009-06-21 | 6:01 | 8.5 | 20 | 31 | 11.98 | 102.4 | 7.65
2009-06-21 | 6:16 | 8.44 | 20 | 31 | 12.01 | 102.5 | 7.64
2009-06-21 | 6:31 | 8.39 | 20 | 31 | 12 | 102.3 | 7.65
2009-06-21 | 6:46 | 8.33 | 20 | 31 | 12.07 | 102.7 | 7.65
2009-06-21 | 7:01 | 8.28 | 20 | 31 | 12.12 | 103 | 7.65
2009-06-21 | 7:16 | 8.22 | 20 | 31 | 12.16 | 103.2 | 7.66
2009-06-21 | 7:31 | 8.17 | 20 | 31 | 12.21 | 103.5 | 7.66
2009-06-21 | 7:46 | 8.13 | 20 | 31 | 12.19 | 103.3 | 7.66
2009-06-21 | 8:01 | 8.1 | 20 | 31 | 12.21 | 103.4 | 7.67
2009-06-21 | 8:16 | 8.08 | 20 | 31 | 12.25 | 103.6 | 7.68
2009-06-21 | 8:31 | 8.04 | 20 | 31 | 12.35 | 104.4 | 7.67
2009-06-21 | 8:46 | 8.02 | 20 | 31 | 12.38 | 104.6 | 7.68
2009-06-21 | 9:01 | 8 | 20 | 31 | 12.43 | 104.9 | 7.69
2009-06-21 | 9:16 | 8.02 | 20 | 31 | 12.53 | 105.9 | 7.68
2009-06-21 | 9:31 | 8.08 | 20 | 31 | 12.49 | 105.7 | 7.69
2009-06-21 | 9:46 | 8.12 | 20 | 31 | 12.47 | 105.6 | 7.69
2009-06-21 | 10:01 | 8.14 | 20 | 31 | 12.47 | 105.6 | 7.7
2009-06-21 | 10:16 | 8.2 | 20 | 31 | 12.46 | 105.7 | 7.7
2009-06-21 | 10:31 | 8.24 | 20 | 31 | 12.46 | 105.8 | 7.7
2009-06-21 | 10:46 | 8.33 | 20 | 31 | 12.46 | 106.1 | 7.7
2009-06-21 | 11:01 | 8.49 | 20 | 31 | 12.45 | 106.4 | 7.7
2009-06-21 | 11:16 | 8.64 | 20 | 31 | 12.47 | 107 | 7.72
2009-06-21 | 11:31 | 8.73 | 20 | 31 | 12.45 | 107 | 7.72
2009-06-21 | 11:46 | 8.88 | 20 | 31 | 12.42 | 107.1 | 7.72
2009-06-21 | 12:01 | 9.08 | 20 | 31 | 12.37 | 107.3 | 7.72
2009-06-21 | 12:16 | 9.3 | 20 | 32 | 12.33 | 107.4 | 7.72
2009-06-21 | 12:31 | 9.49 | 21 | 32 | 12.3 | 107.6 | 7.72
2009-06-21 | 12:46 | 9.63 | 21 | 32 | 12.29 | 107.9 | 7.73
2009-06-21 | 13:00 | 9.59 | 21 | 32 | 12.25 | 107.5 | 7.73
2009-06-21 | 13:15 | 9.8 | 21 | 32 | 12.18 | 107.4 | 7.72
2009-06-21 | 13:30 | 10.23 | 21 | 32 | 12.05 | 107.4 | 7.73
2009-06-21 | 13:45 | 10.51 | 21 | 32 | 11.94 | 107 | 7.74
2009-06-21 | 14:00 | 10.88 | 21 | 32 | 11.86 | 107.2 | 7.74
2009-06-21 | 14:15 | 11.24 | 21 | 32 | 11.75 | 107.1 | 7.75
2009-06-21 | 14:30 | 11.42 | 21 | 32 | 11.63 | 106.5 | 7.75
2009-06-21 | 14:45 | 11.71 | 21 | 32 | 11.53 | 106.3 | 7.76
2009-06-21 | 15:00 | 12.05 | 21 | 32 | 11.37 | 105.6 | 7.75
2009-06-21 | 15:15 | 12.27 | 21 | 32 | 11.32 | 105.7 | 7.76
2009-06-21 | 15:30 | 12.33 | 21 | 32 | 11.27 | 105.4 | 7.76
2009-06-21 | 15:45 | 12.6 | 21 | 32 | 11.19 | 105.2 | 7.76
2009-06-21 | 16:00 | 12.74 | 21 | 32 | 11.13 | 105 | 7.76
2009-06-21 | 16:15 | 12.87 | 21 | 32 | 11.01 | 104.2 | 7.76
2009-06-21 | 16:30 | 12.91 | 21 | 33 | 10.94 | 103.7 | 7.77
2009-06-21 | 16:45 | 12.97 | 21 | 33 | 10.95 | 103.8 | 7.76
2009-06-21 | 17:00 | 13.03 | 21 | 33 | 10.88 | 103.3 | 7.77
2009-06-21 | 17:15 | 13.04 | 21 | 33 | 10.85 | 103.1 | 7.77
2009-06-21 | 17:30 | 12.99 | 21 | 33 | 10.89 | 103.3 | 7.77
2009-06-21 | 17:45 | 12.92 | 22 | 33 | 10.89 | 103.2 | 7.76
2009-06-21 | 18:00 | 12.88 | 22 | 33 | 10.85 | 102.7 | 7.76
2009-06-21 | 18:15 | 12.84 | 22 | 34 | 10.83 | 102.4 | 7.76
2009-06-21 | 18:30 | 12.8 | 22 | 34 | 10.87 | 102.7 | 7.76
2009-06-21 | 18:45 | 12.74 | 22 | 34 | 10.85 | 102.4 | 7.76
2009-06-21 | 19:00 | 12.7 | 22 | 34 | 10.87 | 102.5 | 7.76
2009-06-21 | 19:15 | 12.67 | 22 | 34 | 10.84 | 102.1 | 7.76
2009-06-21 | 19:30 | 12.59 | 22 | 34 | 10.84 | 101.9 | 7.76
2009-06-21 | 19:45 | 12.49 | 22 | 35 | 10.91 | 102.3 | 7.76
2009-06-21 | 20:00 | 12.38 | 23 | 35 | 10.95 | 102.5 | 7.76
2009-06-21 | 20:15 | 12.25 | 23 | 35 | 10.93 | 102 | 7.76
2009-06-21 | 20:30 | 12.14 | 23 | 35 | 10.98 | 102.3 | 7.76
2009-06-21 | 20:45 | 12.04 | 23 | 35 | 10.97 | 101.8 | 7.75
2009-06-21 | 21:00 | 11.94 | 23 | 35 | 11.02 | 102.1 | 7.75
2009-06-21 | 21:15 | 11.84 | 23 | 35 | 11.05 | 102.2 | 7.74
2009-06-21 | 21:30 | 11.74 | 23 | 35 | 11.08 | 102.2 | 7.74
2009-06-21 | 21:45 | 11.64 | 23 | 35 | 11.05 | 101.7 | 7.74
2009-06-21 | 22:00 | 11.53 | 23 | 35 | 11.06 | 101.5 | 7.74
2009-06-21 | 22:15 | 11.44 | 23 | 35 | 11.07 | 101.5 | 7.74
2009-06-21 | 22:30 | 11.35 | 23 | 36 | 11.09 | 101.4 | 7.74
2009-06-21 | 22:45 | 11.25 | 23 | 36 | 11.16 | 101.8 | 7.73
2009-06-21 | 23:00 | 11.16 | 23 | 36 | 11.16 | 101.6 | 7.73
2009-06-21 | 23:15 | 11.08 | 23 | 36 | 11.18 | 101.6 | 7.73
2009-06-21 | 23:30 | 10.99 | 23 | 36 | 11.19 | 101.5 | 7.73
2009-06-21 | 23:45 | 10.89 | 23 | 36 | 11.22 | 101.4 | 7.73
2009-06-22 | 0:00 | 10.78 | 23 | 36 | 11.25 | 101.5 | 7.73
2009-06-22 | 0:15 | 10.68 | 23 | 36 | 11.28 | 101.5 | 7.73
2009-06-22 | 0:30 | 10.58 | 23 | 36 | 11.3 | 101.5 | 7.73
2009-06-22 | 0:45 | 10.47 | 23 | 36 | 11.34 | 101.5 | 7.73
2009-06-22 | 1:01 | 10.36 | 23 | 36 | 11.42 | 102 | 7.73
2009-06-22 | 1:16 | 10.26 | 24 | 36 | 11.42 | 101.8 | 7.74
2009-06-22 | 1:31 | 10.16 | 24 | 36 | 11.44 | 101.8 | 7.73
2009-06-22 | 1:46 | 10.04 | 24 | 36 | 11.53 | 102.3 | 7.73
2009-06-22 | 2:01 | 9.91 | 24 | 36 | 11.58 | 102.4 | 7.74
2009-06-22 | 2:16 | 9.78 | 24 | 36 | 11.6 | 102.2 | 7.73
2009-06-22 | 2:31 | 9.64 | 24 | 36 | 11.67 | 102.5 | 7.73
2009-06-22 | 2:46 | 9.51 | 24 | 36 | 11.73 | 102.7 | 7.72
2009-06-22 | 3:01 | 9.38 | 24 | 36 | 11.73 | 102.4 | 7.72
2009-06-22 | 3:16 | 9.25 | 24 | 36 | 11.81 | 102.7 | 7.73
2009-06-22 | 3:31 | 9.12 | 23 | 36 | 11.87 | 103 | 7.72
2009-06-22 | 3:46 | 9.01 | 23 | 36 | 11.86 | 102.6 | 7.72
2009-06-22 | 4:01 | 8.9 | 23 | 36 | 11.88 | 102.5 | 7.72
2009-06-22 | 4:16 | 8.78 | 23 | 36 | 11.96 | 102.9 | 7.71
2009-06-22 | 4:31 | 8.67 | 23 | 36 | 12.02 | 103.1 | 7.71
2009-06-22 | 4:46 | 8.56 | 23 | 36 | 12.06 | 103.2 | 7.71
2009-06-22 | 5:01 | 8.47 | 23 | 35 | 12.03 | 102.8 | 7.7
2009-06-22 | 5:16 | 8.38 | 23 | 35 | 12.06 | 102.7 | 7.7
2009-06-22 | 5:31 | 8.29 | 23 | 35 | 12.08 | 102.7 | 7.7
2009-06-22 | 5:46 | 8.2 | 23 | 35 | 12.11 | 102.7 | 7.7
2009-06-22 | 6:01 | 8.11 | 23 | 35 | 12.13 | 102.7 | 7.7
2009-06-22 | 6:16 | 8.01 | 23 | 35 | 12.24 | 103.4 | 7.71
2009-06-22 | 6:31 | 7.93 | 23 | 35 | 12.31 | 103.8 | 7.71
2009-06-22 | 6:46 | 7.86 | 23 | 35 | 12.3 | 103.5 | 7.71
2009-06-22 | 7:01 | 7.81 | 23 | 35 | 12.38 | 104.1 | 7.71
2009-06-22 | 7:16 | 7.78 | 23 | 35 | 12.44 | 104.5 | 7.71
2009-06-22 | 7:31 | 7.74 | 23 | 35 | 12.43 | 104.3 | 7.71
2009-06-22 | 7:46 | 7.7 | 23 | 35 | 12.44 | 104.3 | 7.72
2009-06-22 | 8:01 | 7.66 | 23 | 35 | 12.53 | 104.9 | 7.72
2009-06-22 | 8:16 | 7.65 | 23 | 35 | 12.59 | 105.4 | 7.73
2009-06-22 | 8:31 | 7.63 | 23 | 35 | 12.65 | 105.9 | 7.73
2009-06-22 | 8:46 | 7.62 | 23 | 35 | 12.75 | 106.6 | 7.74
2009-06-22 | 9:01 | 7.63 | 23 | 35 | 12.84 | 107.4 | 7.74
2009-06-22 | 9:16 | 7.67 | 23 | 35 | 12.9 | 108 | 7.73
2009-06-22 | 9:31 | 7.74 | 23 | 35 | 12.84 | 107.7 | 7.74
2009-06-22 | 9:46 | 7.85 | 23 | 35 | 12.8 | 107.7 | 7.75
2009-06-22 | 10:01 | 7.99 | 23 | 35 | 12.75 | 107.6 | 7.75
2009-06-22 | 10:16 | 8.17 | 23 | 35 | 12.62 | 107 | 7.75
2009-06-22 | 10:31 | 8.37 | 23 | 35 | 12.58 | 107.2 | 7.75
2009-06-22 | 10:46 | 8.6 | 23 | 35 | 12.51 | 107.2 | 7.76
2009-06-22 | 11:01 | 8.85 | 23 | 35 | 12.47 | 107.5 | 7.76
2009-06-22 | 11:16 | 9.12 | 23 | 35 | 12.47 | 108.2 | 7.77
2009-06-22 | 11:31 | 9.4 | 23 | 35 | 12.36 | 108 | 7.77
2009-06-22 | 11:46 | 9.7 | 23 | 36 | 12.33 | 108.5 | 7.78
2009-06-22 | 12:01 | 10.02 | 23 | 36 | 12.2 | 108.2 | 7.78
2009-06-22 | 12:16 | 10.33 | 23 | 36 | 12.16 | 108.5 | 7.79
2009-06-22 | 12:31 | 10.66 | 23 | 36 | 12.07 | 108.6 | 7.79
2009-06-22 | 12:46 | 11.01 | 23 | 36 | 11.92 | 108.1 | 7.8
2009-06-22 | 13:00 | 11.35 | 23 | 36 | 11.84 | 108.3 | 7.8
2009-06-22 | 13:15 | 11.72 | 23 | 36 | 11.69 | 107.8 | 7.8
2009-06-22 | 13:30 | 12.08 | 23 | 36 | 11.61 | 107.9 | 7.81
2009-06-22 | 13:45 | 12.45 | 23 | 36 | 11.5 | 107.8 | 7.82
2009-06-22 | 14:00 | 12.8 | 23 | 36 | 11.33 | 107 | 7.82
2009-06-22 | 14:15 | 13.14 | 23 | 36 | 11.23 | 106.9 | 7.82
2009-06-22 | 14:30 | 13.45 | 24 | 36 | 11.07 | 106.1 | 7.83
2009-06-22 | 14:45 | 13.74 | 24 | 36 | 10.99 | 106 | 7.83
2009-06-22 | 15:00 | 14 | 24 | 36 | 10.85 | 105.3 | 7.83
2009-06-22 | 15:15 | 14.21 | 24 | 37 | 10.75 | 104.8 | 7.84
2009-06-22 | 15:30 | 14.4 | 24 | 37 | 10.67 | 104.4 | 7.84
2009-06-22 | 15:45 | 14.54 | 24 | 37 | 10.65 | 104.5 | 7.85
2009-06-22 | 16:00 | 14.66 | 24 | 37 | 10.56 | 104 | 7.85
2009-06-22 | 16:15 | 14.75 | 24 | 37 | 10.5 | 103.5 | 7.84
2009-06-22 | 16:30 | 14.81 | 24 | 37 | 10.49 | 103.6 | 7.84
2009-06-22 | 16:45 | 14.85 | 24 | 37 | 10.48 | 103.6 | 7.84
2009-06-22 | 17:00 | 14.86 | 24 | 37 | 10.45 | 103.3 | 7.84
2009-06-22 | 17:15 | 14.85 | 24 | 37 | 10.43 | 103.2 | 7.83
2009-06-22 | 17:30 | 14.84 | 24 | 38 | 10.42 | 103 | 7.84
2009-06-22 | 17:45 | 14.82 | 25 | 38 | 10.41 | 102.9 | 7.84
2009-06-22 | 18:00 | 14.8 | 25 | 38 | 10.4 | 102.7 | 7.84
2009-06-22 | 18:15 | 14.78 | 25 | 39 | 10.34 | 102.1 | 7.84
2009-06-22 | 18:30 | 14.75 | 25 | 39 | 10.37 | 102.3 | 7.84
2009-06-22 | 18:45 | 14.74 | 25 | 39 | 10.37 | 102.2 | 7.84
2009-06-22 | 19:00 | 14.72 | 26 | 39 | 10.37 | 102.2 | 7.83
2009-06-22 | 19:15 | 14.7 | 26 | 40 | 10.31 | 101.6 | 7.84
2009-06-22 | 19:30 | 14.65 | 26 | 40 | 10.34 | 101.8 | 7.83
2009-06-22 | 19:45 | 14.6 | 26 | 40 | 10.36 | 101.9 | 7.84
2009-06-22 | 20:00 | 14.52 | 26 | 40 | 10.33 | 101.3 | 7.83
2009-06-22 | 20:15 | 14.42 | 26 | 40 | 10.37 | 101.6 | 7.82
2009-06-22 | 20:30 | 14.31 | 26 | 40 | 10.36 | 101.2 | 7.82
2009-06-22 | 20:45 | 14.19 | 26 | 40 | 10.41 | 101.4 | 7.82
2009-06-22 | 21:00 | 14.04 | 26 | 40 | 10.4 | 101 | 7.82
2009-06-22 | 21:15 | 13.91 | 26 | 40 | 10.41 | 100.8 | 7.81
2009-06-22 | 21:30 | 13.78 | 26 | 41 | 10.48 | 101.2 | 7.81
2009-06-22 | 21:45 | 13.66 | 26 | 41 | 10.52 | 101.3 | 7.81
2009-06-22 | 22:00 | 13.54 | 26 | 41 | 10.51 | 100.9 | 7.81
2009-06-22 | 22:15 | 13.43 | 27 | 41 | 10.56 | 101.2 | 7.81
2009-06-22 | 22:30 | 13.34 | 27 | 41 | 10.56 | 101 | 7.81
2009-06-22 | 22:45 | 13.24 | 27 | 41 | 10.57 | 100.9 | 7.82
2009-06-22 | 23:00 | 13.15 | 27 | 41 | 10.62 | 101.2 | 7.81
2009-06-22 | 23:15 | 13.05 | 27 | 41 | 10.62 | 100.9 | 7.81
2009-06-22 | 23:30 | 12.95 | 27 | 41 | 10.68 | 101.3 | 7.81
2009-06-22 | 23:45 | 12.84 | 27 | 41 | 10.73 | 101.4 | 7.81
2009-06-23 | 0:00 | 12.73 | 27 | 41 | 10.76 | 101.5 | 7.81
2009-06-23 | 0:15 | 12.62 | 27 | 41 | 10.75 | 101.2 | 7.81
2009-06-23 | 0:30 | 12.5 | 27 | 42 | 10.76 | 101 | 7.82
2009-06-23 | 0:45 | 12.39 | 27 | 42 | 10.78 | 100.9 | 7.81
2009-06-23 | 1:01 | 12.26 | 27 | 42 | 10.81 | 100.9 | 7.81
2009-06-23 | 1:16 | 12.12 | 27 | 42 | 10.84 | 100.9 | 7.8
2009-06-23 | 1:31 | 11.98 | 27 | 42 | 10.88 | 100.9 | 7.81
2009-06-23 | 1:46 | 11.84 | 27 | 42 | 10.93 | 101 | 7.81
2009-06-23 | 2:01 | 11.69 | 27 | 42 | 11.01 | 101.5 | 7.81
2009-06-23 | 2:16 | 11.55 | 27 | 42 | 11.03 | 101.3 | 7.8
2009-06-23 | 2:31 | 11.41 | 27 | 42 | 11.06 | 101.2 | 7.8
2009-06-23 | 2:46 | 11.27 | 28 | 42 | 11.14 | 101.7 | 7.79
2009-06-23 | 3:01 | 11.13 | 28 | 42 | 11.2 | 101.9 | 7.79
2009-06-23 | 3:16 | 11 | 28 | 43 | 11.21 | 101.6 | 7.79
2009-06-23 | 3:31 | 10.86 | 28 | 43 | 11.23 | 101.5 | 7.79
2009-06-23 | 3:46 | 10.72 | 28 | 43 | 11.27 | 101.5 | 7.79
2009-06-23 | 4:01 | 10.58 | 28 | 42 | 11.36 | 102 | 7.79
2009-06-23 | 4:16 | 10.43 | 28 | 42 | 11.42 | 102.2 | 7.78
2009-06-23 | 4:31 | 10.28 | 27 | 42 | 11.43 | 101.9 | 7.78
2009-06-23 | 4:46 | 10.13 | 27 | 42 | 11.46 | 101.8 | 7.78
2009-06-23 | 5:01 | 9.98 | 27 | 42 | 11.55 | 102.3 | 7.77
2009-06-23 | 5:16 | 9.84 | 27 | 42 | 11.56 | 102 | 7.77
2009-06-23 | 5:31 | 9.69 | 27 | 42 | 11.65 | 102.5 | 7.77
2009-06-23 | 5:46 | 9.55 | 27 | 42 | 11.67 | 102.3 | 7.76
2009-06-23 | 6:01 | 9.43 | 27 | 41 | 11.7 | 102.3 | 7.77
2009-06-23 | 6:16 | 9.31 | 27 | 41 | 11.8 | 102.9 | 7.76
2009-06-23 | 6:31 | 9.2 | 27 | 41 | 11.88 | 103.3 | 7.76
2009-06-23 | 6:46 | 9.11 | 27 | 41 | 11.88 | 103.1 | 7.76
2009-06-23 | 7:01 | 9.04 | 27 | 41 | 11.91 | 103.1 | 7.76
2009-06-23 | 7:16 | 9 | 26 | 41 | 11.99 | 103.7 | 7.77
2009-06-23 | 7:31 | 8.96 | 26 | 41 | 12.04 | 104.1 | 7.77
2009-06-23 | 7:46 | 8.93 | 26 | 40 | 12.03 | 103.9 | 7.77
2009-06-23 | 8:01 | 8.93 | 26 | 40 | 12.03 | 103.9 | 7.78
2009-06-23 | 8:16 | 8.93 | 26 | 40 | 12.12 | 104.7 | 7.78
2009-06-23 | 8:31 | 8.92 | 26 | 40 | 12.14 | 104.8 | 7.79
2009-06-23 | 8:46 | 8.93 | 26 | 40 | 12.2 | 105.3 | 7.79
2009-06-23 | 9:01 | 8.95 | 26 | 40 | 12.33 | 106.5 | 7.8
2009-06-23 | 9:16 | 9 | 26 | 40 | 12.4 | 107.2 | 7.8
2009-06-23 | 9:31 | 9.08 | 26 | 40 | 12.39 | 107.4 | 7.81
2009-06-23 | 9:46 | 9.19 | 26 | 40 | 12.28 | 106.7 | 7.81
2009-06-23 | 10:01 | 9.36 | 26 | 40 | 12.24 | 106.8 | 7.81
2009-06-23 | 10:16 | 9.57 | 26 | 40 | 12.17 | 106.8 | 7.82
2009-06-23 | 10:31 | 9.83 | 26 | 40 | 12.08 | 106.6 | 7.83
2009-06-23 | 10:46 | 10.13 | 26 | 40 | 11.99 | 106.6 | 7.82
2009-06-23 | 11:01 | 10.44 | 26 | 40 | 11.89 | 106.4 | 7.84
2009-06-23 | 11:16 | 10.76 | 26 | 40 | 11.92 | 107.5 | 7.83
2009-06-23 | 11:31 | 11.11 | 26 | 40 | 11.84 | 107.6 | 7.84
2009-06-23 | 11:46 | 11.47 | 26 | 40 | 11.8 | 108.2 | 7.84
2009-06-23 | 12:01 | 11.85 | 26 | 40 | 11.73 | 108.5 | 7.85
2009-06-23 | 12:16 | 12.25 | 26 | 40 | 11.57 | 108 | 7.85
2009-06-23 | 12:31 | 12.66 | 26 | 40 | 11.44 | 107.8 | 7.86
2009-06-23 | 12:46 | 13.07 | 26 | 40 | 11.31 | 107.5 | 7.87
2009-06-23 | 13:00 | 13.48 | 26 | 40 | 11.24 | 107.8 | 7.88
2009-06-23 | 13:15 | 13.89 | 26 | 40 | 11.14 | 107.8 | 7.88
2009-06-23 | 13:30 | 14.3 | 26 | 40 | 10.98 | 107.2 | 7.89
2009-06-23 | 13:45 | 14.69 | 26 | 40 | 10.84 | 106.7 | 7.89
2009-06-23 | 14:00 | 15.05 | 26 | 40 | 10.71 | 106.3 | 7.9
2009-06-23 | 14:15 | 15.4 | 26 | 40 | 10.57 | 105.7 | 7.9
2009-06-23 | 14:30 | 15.71 | 26 | 40 | 10.5 | 105.7 | 7.91
2009-06-23 | 14:45 | 15.98 | 26 | 40 | 10.41 | 105.5 | 7.9
2009-06-23 | 15:00 | 16.22 | 26 | 41 | 10.29 | 104.7 | 7.91
2009-06-23 | 15:15 | 16.42 | 26 | 41 | 10.23 | 104.6 | 7.91
2009-06-23 | 15:30 | 16.59 | 26 | 41 | 10.17 | 104.4 | 7.92
2009-06-23 | 15:45 | 16.71 | 26 | 41 | 10.12 | 104.1 | 7.91
2009-06-23 | 16:00 | 16.8 | 26 | 41 | 10.08 | 103.8 | 7.91
2009-06-23 | 16:15 | 16.87 | 26 | 41 | 10.03 | 103.5 | 7.92
2009-06-23 | 16:30 | 16.91 | 26 | 41 | 10 | 103.3 | 7.91
2009-06-23 | 16:45 | 16.94 | 27 | 41 | 9.97 | 103 | 7.91
2009-06-23 | 17:00 | 16.96 | 27 | 41 | 9.96 | 103 | 7.91
2009-06-23 | 17:15 | 16.94 | 27 | 41 | 9.94 | 102.7 | 7.91
2009-06-23 | 17:30 | 16.93 | 27 | 41 | 9.92 | 102.5 | 7.89
2009-06-23 | 17:45 | 16.93 | 27 | 41 | 9.87 | 101.9 | 7.9
2009-06-23 | 18:00 | 16.94 | 27 | 41 | 9.89 | 102.2 | 7.9
2009-06-23 | 18:15 | 16.92 | 27 | 41 | 9.83 | 101.6 | 7.89
2009-06-23 | 18:30 | 16.91 | 27 | 41 | 9.8 | 101.2 | 7.88
2009-06-23 | 18:45 | 16.89 | 27 | 41 | 9.78 | 101 | 7.88
2009-06-23 | 19:00 | 16.88 | 27 | 41 | 9.81 | 101.2 | 7.88
2009-06-23 | 19:15 | 16.83 | 27 | 41 | 9.82 | 101.2 | 7.87
2009-06-23 | 19:30 | 16.77 | 27 | 41 | 9.77 | 100.6 | 7.86
2009-06-23 | 19:45 | 16.73 | 27 | 41 | 9.75 | 100.3 | 7.86
2009-06-23 | 20:00 | 16.65 | 27 | 41 | 9.79 | 100.6 | 7.85
2009-06-23 | 20:15 | 16.56 | 27 | 41 | 9.82 | 100.7 | 7.85
2009-06-23 | 20:30 | 16.44 | 27 | 42 | 9.84 | 100.6 | 7.84
2009-06-23 | 20:45 | 16.32 | 27 | 42 | 9.81 | 100 | 7.84
2009-06-23 | 21:00 | 16.18 | 27 | 42 | 9.81 | 99.8 | 7.83
2009-06-23 | 21:15 | 16.04 | 27 | 42 | 9.82 | 99.6 | 7.82
2009-06-23 | 21:30 | 15.89 | 27 | 42 | 9.84 | 99.5 | 7.82
2009-06-23 | 21:45 | 15.74 | 27 | 42 | 9.87 | 99.5 | 7.82
2009-06-23 | 22:00 | 15.6 | 27 | 42 | 9.9 | 99.4 | 7.81
2009-06-23 | 22:15 | 15.47 | 27 | 42 | 9.97 | 99.8 | 7.81
2009-06-23 | 22:30 | 15.35 | 27 | 42 | 9.97 | 99.6 | 7.81
2009-06-23 | 22:45 | 15.23 | 27 | 42 | 10.03 | 99.9 | 7.81
2009-06-23 | 23:00 | 15.11 | 28 | 42 | 10.03 | 99.7 | 7.8
2009-06-23 | 23:15 | 14.99 | 28 | 42 | 10.09 | 100 | 7.81
2009-06-23 | 23:30 | 14.87 | 28 | 43 | 10.1 | 99.8 | 7.81
2009-06-23 | 23:45 | 14.74 | 28 | 43 | 10.15 | 100.1 | 7.81
2009-06-24 | 0:00 | 14.61 | 28 | 43 | 10.2 | 100.3 | 7.81
2009-06-24 | 0:15 | 14.48 | 28 | 43 | 10.19 | 99.9 | 7.81
2009-06-24 | 0:30 | 14.34 | 28 | 43 | 10.26 | 100.3 | 7.81
2009-06-24 | 0:45 | 14.21 | 28 | 43 | 10.25 | 99.9 | 7.81
2009-06-24 | 1:01 | 14.06 | 28 | 43 | 10.28 | 99.9 | 7.81
2009-06-24 | 1:16 | 13.92 | 28 | 43 | 10.32 | 99.9 | 7.81
2009-06-24 | 1:31 | 13.77 | 28 | 43 | 10.35 | 99.9 | 7.81
2009-06-24 | 1:46 | 13.63 | 28 | 43 | 10.39 | 100 | 7.81
2009-06-24 | 2:01 | 13.49 | 28 | 43 | 10.43 | 100 | 7.81
2009-06-24 | 2:16 | 13.35 | 28 | 43 | 10.51 | 100.6 | 7.81
2009-06-24 | 2:31 | 13.2 | 28 | 43 | 10.57 | 100.8 | 7.8
2009-06-24 | 2:46 | 13.06 | 28 | 43 | 10.57 | 100.4 | 7.8
2009-06-24 | 3:01 | 12.92 | 28 | 44 | 10.6 | 100.4 | 7.8
2009-06-24 | 3:16 | 12.78 | 28 | 44 | 10.68 | 100.8 | 7.8
2009-06-24 | 3:31 | 12.64 | 28 | 44 | 10.73 | 101 | 7.8
2009-06-24 | 3:46 | 12.5 | 28 | 43 | 10.73 | 100.7 | 7.79
2009-06-24 | 4:01 | 12.36 | 28 | 43 | 10.8 | 101.1 | 7.79
2009-06-24 | 4:16 | 12.22 | 28 | 43 | 10.81 | 100.8 | 7.78
2009-06-24 | 4:31 | 12.07 | 28 | 43 | 10.84 | 100.8 | 7.78
2009-06-24 | 4:46 | 11.92 | 27 | 42 | 10.88 | 100.8 | 7.77
2009-06-24 | 5:01 | 11.76 | 27 | 42 | 10.97 | 101.2 | 7.77
2009-06-24 | 5:16 | 11.6 | 27 | 42 | 10.99 | 101 | 7.76
2009-06-24 | 5:31 | 11.46 | 27 | 41 | 11.02 | 101 | 7.76
2009-06-24 | 5:46 | 11.33 | 27 | 41 | 11.06 | 101.1 | 7.76
2009-06-24 | 6:01 | 11.21 | 27 | 41 | 11.1 | 101.2 | 7.76
2009-06-24 | 6:16 | 11.1 | 26 | 41 | 11.14 | 101.3 | 7.76
2009-06-24 | 6:31 | 11.01 | 26 | 40 | 11.24 | 102 | 7.75
2009-06-24 | 6:46 | 10.92 | 26 | 40 | 11.26 | 101.9 | 7.75
2009-06-24 | 7:01 | 10.83 | 26 | 40 | 11.34 | 102.4 | 7.75
2009-06-24 | 7:16 | 10.77 | 26 | 40 | 11.39 | 102.8 | 7.75
2009-06-24 | 7:31 | 10.72 | 26 | 40 | 11.39 | 102.6 | 7.76
2009-06-24 | 7:46 | 10.68 | 26 | 40 | 11.45 | 103 | 7.76
2009-06-24 | 8:01 | 10.65 | 26 | 39 | 11.45 | 103 | 7.76
2009-06-24 | 8:16 | 10.64 | 26 | 39 | 11.53 | 103.7 | 7.76
2009-06-24 | 8:31 | 10.62 | 25 | 39 | 11.59 | 104.2 | 7.77
2009-06-24 | 8:46 | 10.61 | 25 | 39 | 11.66 | 104.8 | 7.77
2009-06-24 | 9:01 | 10.62 | 25 | 39 | 11.7 | 105.2 | 7.78
2009-06-24 | 9:16 | 10.65 | 25 | 39 | 11.79 | 106 | 7.78
2009-06-24 | 9:31 | 10.72 | 25 | 39 | 11.79 | 106.2 | 7.79
2009-06-24 | 9:46 | 10.83 | 25 | 39 | 11.69 | 105.6 | 7.79
2009-06-24 | 10:01 | 11 | 25 | 39 | 11.61 | 105.2 | 7.8
2009-06-24 | 10:16 | 11.2 | 25 | 39 | 11.57 | 105.5 | 7.8
2009-06-24 | 10:31 | 11.46 | 25 | 39 | 11.46 | 105 | 7.81
2009-06-24 | 10:46 | 11.74 | 25 | 39 | 11.35 | 104.7 | 7.81
2009-06-24 | 11:01 | 12.04 | 25 | 39 | 11.37 | 105.6 | 7.82
2009-06-24 | 11:16 | 12.36 | 25 | 39 | 11.33 | 106 | 7.83
2009-06-24 | 11:31 | 12.69 | 25 | 39 | 11.34 | 106.9 | 7.84
2009-06-24 | 11:46 | 13.05 | 25 | 39 | 11.29 | 107.3 | 7.84
2009-06-24 | 12:01 | 13.42 | 25 | 39 | 11.21 | 107.4 | 7.85
2009-06-24 | 12:16 | 13.8 | 25 | 39 | 11.07 | 107 | 7.86
2009-06-24 | 12:31 | 14.2 | 25 | 39 | 11 | 107.2 | 7.86
2009-06-24 | 12:46 | 14.6 | 25 | 39 | 10.86 | 106.7 | 7.87
2009-06-24 | 13:00 | 15 | 25 | 39 | 10.78 | 106.9 | 7.88
2009-06-24 | 13:15 | 15.4 | 25 | 39 | 10.63 | 106.3 | 7.88
2009-06-24 | 13:30 | 15.81 | 25 | 39 | 10.55 | 106.4 | 7.89
2009-06-24 | 13:45 | 16.19 | 25 | 39 | 10.45 | 106.3 | 7.9
2009-06-24 | 14:00 | 16.55 | 25 | 39 | 10.34 | 106 | 7.9
2009-06-24 | 14:15 | 16.89 | 25 | 39 | 10.18 | 105.1 | 7.9
2009-06-24 | 14:30 | 17.19 | 25 | 39 | 10.06 | 104.5 | 7.91
2009-06-24 | 14:45 | 17.47 | 25 | 39 | 10 | 104.4 | 7.9
2009-06-24 | 15:00 | 17.7 | 26 | 39 | 9.92 | 104.2 | 7.92
2009-06-24 | 15:15 | 17.91 | 25 | 39 | 9.81 | 103.4 | 7.91
2009-06-24 | 15:30 | 18.07 | 26 | 39 | 9.77 | 103.4 | 7.91
2009-06-24 | 15:45 | 18.18 | 26 | 39 | 9.68 | 102.6 | 7.91
2009-06-24 | 16:00 | 18.27 | 26 | 39 | 9.67 | 102.7 | 7.91
2009-06-24 | 16:15 | 18.34 | 26 | 39 | 9.64 | 102.5 | 7.91
2009-06-24 | 16:30 | 18.37 | 26 | 39 | 9.56 | 101.8 | 7.91
2009-06-24 | 16:45 | 18.39 | 26 | 39 | 9.56 | 101.8 | 7.9
2009-06-24 | 17:00 | 18.39 | 26 | 40 | 9.55 | 101.7 | 7.9
2009-06-24 | 17:15 | 18.36 | 26 | 40 | 9.54 | 101.5 | 7.89
2009-06-24 | 17:30 | 18.32 | 26 | 40 | 9.53 | 101.3 | 7.89
2009-06-24 | 17:45 | 18.28 | 26 | 40 | 9.47 | 100.6 | 7.88
2009-06-24 | 18:00 | 18.24 | 26 | 40 | 9.5 | 100.8 | 7.87
2009-06-24 | 18:15 | 18.16 | 26 | 40 | 9.51 | 100.7 | 7.87
2009-06-24 | 18:30 | 18.09 | 26 | 40 | 9.51 | 100.7 | 7.86
2009-06-24 | 18:45 | 18.03 | 26 | 40 | 9.47 | 100.1 | 7.86
2009-06-24 | 19:00 | 17.97 | 26 | 40 | 9.45 | 99.8 | 7.86
2009-06-24 | 19:15 | 17.89 | 26 | 40 | 9.5 | 100.1 | 7.84
2009-06-24 | 19:30 | 17.82 | 26 | 40 | 9.47 | 99.7 | 7.85
2009-06-24 | 19:45 | 17.74 | 26 | 40 | 9.51 | 99.9 | 7.85
2009-06-24 | 20:00 | 17.64 | 26 | 40 | 9.54 | 100 | 7.84
2009-06-24 | 20:15 | 17.52 | 26 | 40 | 9.56 | 100 | 7.84
2009-06-24 | 20:30 | 17.38 | 26 | 40 | 9.55 | 99.6 | 7.84
2009-06-24 | 20:45 | 17.23 | 26 | 40 | 9.59 | 99.7 | 7.83
2009-06-24 | 21:00 | 17.06 | 26 | 41 | 9.64 | 99.9 | 7.83
2009-06-24 | 21:15 | 16.9 | 26 | 41 | 9.63 | 99.4 | 7.83
2009-06-24 | 21:30 | 16.73 | 26 | 41 | 9.69 | 99.7 | 7.83
2009-06-24 | 21:45 | 16.56 | 26 | 41 | 9.7 | 99.5 | 7.83
2009-06-24 | 22:00 | 16.42 | 27 | 41 | 9.76 | 99.8 | 7.82
2009-06-24 | 22:15 | 16.29 | 27 | 41 | 9.8 | 99.9 | 7.82
2009-06-24 | 22:30 | 16.16 | 27 | 41 | 9.84 | 100.1 | 7.82
2009-06-24 | 22:45 | 16.05 | 27 | 41 | 9.83 | 99.7 | 7.83
2009-06-24 | 23:00 | 15.93 | 27 | 41 | 9.84 | 99.6 | 7.83
2009-06-24 | 23:15 | 15.81 | 27 | 41 | 9.9 | 99.9 | 7.83
2009-06-24 | 23:30 | 15.71 | 27 | 41 | 9.89 | 99.6 | 7.83
2009-06-24 | 23:45 | 15.59 | 27 | 41 | 9.92 | 99.6 | 7.82
2009-06-25 | 0:00 | 15.48 | 27 | 42 | 10 | 100.2 | 7.83
2009-06-25 | 0:15 | 15.36 | 27 | 42 | 10.04 | 100.4 | 7.82
2009-06-25 | 0:30 | 15.24 | 27 | 42 | 10.08 | 100.5 | 7.82
2009-06-25 | 0:45 | 15.11 | 27 | 42 | 10.08 | 100.2 | 7.82
2009-06-25 | 1:01 | 14.97 | 27 | 42 | 10.14 | 100.5 | 7.81
2009-06-25 | 1:16 | 14.84 | 27 | 42 | 10.15 | 100.3 | 7.82
2009-06-25 | 1:31 | 14.68 | 27 | 42 | 10.23 | 100.7 | 7.81
2009-06-25 | 1:46 | 14.54 | 27 | 42 | 10.27 | 100.8 | 7.81
2009-06-25 | 2:01 | 14.38 | 27 | 41 | 10.28 | 100.6 | 7.8
2009-06-25 | 2:16 | 14.22 | 27 | 41 | 10.37 | 101.1 | 7.8
2009-06-25 | 2:31 | 14.05 | 26 | 40 | 10.36 | 100.7 | 7.79
2009-06-25 | 2:46 | 13.89 | 26 | 40 | 10.45 | 101.1 | 7.78
2009-06-25 | 3:01 | 13.72 | 25 | 39 | 10.46 | 100.9 | 7.77
2009-06-25 | 3:16 | 13.53 | 25 | 38 | 10.55 | 101.3 | 7.76
2009-06-25 | 3:31 | 13.35 | 24 | 37 | 10.56 | 101 | 7.76
2009-06-25 | 3:46 | 13.17 | 24 | 37 | 10.65 | 101.4 | 7.75
2009-06-25 | 4:01 | 13.01 | 23 | 36 | 10.67 | 101.3 | 7.74
2009-06-25 | 4:16 | 12.87 | 23 | 36 | 10.69 | 101.1 | 7.73
2009-06-25 | 4:31 | 12.73 | 23 | 35 | 10.77 | 101.6 | 7.72
2009-06-25 | 4:46 | 12.58 | 23 | 35 | 10.84 | 102 | 7.71
2009-06-25 | 5:01 | 12.44 | 22 | 34 | 10.83 | 101.5 | 7.71
2009-06-25 | 5:16 | 12.31 | 22 | 34 | 10.91 | 101.9 | 7.7
2009-06-25 | 5:31 | 12.17 | 22 | 34 | 10.92 | 101.7 | 7.7
2009-06-25 | 5:46 | 12.03 | 22 | 34 | 11 | 102.1 | 7.7
2009-06-25 | 6:01 | 11.89 | 22 | 33 | 11.06 | 102.4 | 7.69
2009-06-25 | 6:16 | 11.76 | 22 | 33 | 11.07 | 102.2 | 7.69
2009-06-25 | 6:31 | 11.62 | 21 | 33 | 11.16 | 102.7 | 7.69
2009-06-25 | 6:46 | 11.48 | 21 | 33 | 11.23 | 103 | 7.69
2009-06-25 | 7:01 | 11.36 | 21 | 33 | 11.26 | 102.9 | 7.69
2009-06-25 | 7:16 | 11.26 | 21 | 32 | 11.33 | 103.4 | 7.69
2009-06-25 | 7:31 | 11.16 | 21 | 32 | 11.4 | 103.7 | 7.69
2009-06-25 | 7:46 | 11.07 | 21 | 32 | 11.4 | 103.5 | 7.69
2009-06-25 | 8:01 | 10.97 | 21 | 32 | 11.44 | 103.6 | 7.69
2009-06-25 | 8:16 | 10.89 | 21 | 32 | 11.48 | 103.8 | 7.7
2009-06-25 | 8:31 | 10.81 | 21 | 32 | 11.59 | 104.6 | 7.7
2009-06-25 | 8:46 | 10.74 | 21 | 32 | 11.68 | 105.3 | 7.7
2009-06-25 | 9:01 | 10.7 | 21 | 32 | 11.72 | 105.5 | 7.7
2009-06-25 | 9:16 | 10.7 | 21 | 32 | 11.82 | 106.4 | 7.7
2009-06-25 | 9:31 | 10.73 | 21 | 32 | 11.8 | 106.3 | 7.71
2009-06-25 | 9:46 | 10.78 | 21 | 32 | 11.81 | 106.5 | 7.71
2009-06-25 | 10:01 | 10.88 | 21 | 32 | 11.79 | 106.6 | 7.72
2009-06-25 | 10:16 | 11.01 | 21 | 32 | 11.7 | 106.1 | 7.73
2009-06-25 | 10:31 | 11.17 | 21 | 32 | 11.65 | 106 | 7.73
2009-06-25 | 10:46 | 11.35 | 21 | 32 | 11.61 | 106.2 | 7.74
2009-06-25 | 11:01 | 11.55 | 21 | 32 | 11.63 | 106.8 | 7.74
2009-06-25 | 11:16 | 11.76 | 21 | 32 | 11.61 | 107.1 | 7.75
2009-06-25 | 11:31 | 11.99 | 21 | 32 | 11.56 | 107.3 | 7.76
2009-06-25 | 11:46 | 12.24 | 21 | 32 | 11.59 | 108.1 | 7.76
2009-06-25 | 12:01 | 12.51 | 21 | 32 | 11.56 | 108.5 | 7.76
2009-06-25 | 12:16 | 12.79 | 21 | 32 | 11.49 | 108.5 | 7.77
2009-06-25 | 12:31 | 13.08 | 21 | 32 | 11.42 | 108.6 | 7.78
2009-06-25 | 12:46 | 13.39 | 21 | 32 | 11.34 | 108.6 | 7.78
2009-06-25 | 13:00 | 13.71 | 21 | 33 | 11.2 | 108 | 7.79
2009-06-25 | 13:15 | 14.03 | 21 | 33 | 11.14 | 108.1 | 7.79
2009-06-25 | 13:30 | 14.35 | 21 | 33 | 11.05 | 108.1 | 7.8
2009-06-25 | 13:45 | 14.66 | 21 | 33 | 10.96 | 107.9 | 7.8
2009-06-25 | 14:00 | 14.96 | 21 | 33 | 10.86 | 107.6 | 7.81
2009-06-25 | 14:15 | 15.23 | 21 | 33 | 10.76 | 107.2 | 7.81
2009-06-25 | 14:30 | 15.49 | 22 | 33 | 10.66 | 106.8 | 7.81
2009-06-25 | 14:45 | 15.72 | 22 | 33 | 10.52 | 106 | 7.81
2009-06-25 | 15:00 | 15.93 | 22 | 33 | 10.48 | 106 | 7.82
2009-06-25 | 15:15 | 16.12 | 22 | 33 | 10.36 | 105.3 | 7.82
2009-06-25 | 15:30 | 16.27 | 22 | 34 | 10.32 | 105.2 | 7.83
2009-06-25 | 15:45 | 16.4 | 22 | 34 | 10.28 | 105 | 7.83
2009-06-25 | 16:00 | 16.5 | 22 | 34 | 10.23 | 104.8 | 7.82
2009-06-25 | 16:15 | 16.58 | 22 | 34 | 10.19 | 104.6 | 7.82
2009-06-25 | 16:30 | 16.64 | 22 | 34 | 10.11 | 103.8 | 7.82
2009-06-25 | 16:45 | 16.68 | 22 | 34 | 10.06 | 103.4 | 7.82
2009-06-25 | 17:00 | 16.69 | 22 | 34 | 10.02 | 103 | 7.83
2009-06-25 | 17:15 | 16.69 | 22 | 34 | 10.04 | 103.2 | 7.82
2009-06-25 | 17:30 | 16.67 | 22 | 34 | 10 | 102.7 | 7.82
2009-06-25 | 17:45 | 16.64 | 22 | 35 | 10.01 | 102.8 | 7.82
2009-06-25 | 18:00 | 16.61 | 23 | 35 | 10.02 | 102.8 | 7.81
2009-06-25 | 18:15 | 16.57 | 23 | 35 | 9.97 | 102.2 | 7.81
2009-06-25 | 18:30 | 16.53 | 23 | 35 | 9.95 | 101.9 | 7.81
2009-06-25 | 18:45 | 16.48 | 23 | 35 | 9.98 | 102.2 | 7.8
2009-06-25 | 19:00 | 16.43 | 23 | 35 | 10 | 102.2 | 7.8
2009-06-25 | 19:15 | 16.39 | 23 | 35 | 10 | 102.2 | 7.79
2009-06-25 | 19:30 | 16.33 | 23 | 35 | 9.97 | 101.7 | 7.79
2009-06-25 | 19:45 | 16.25 | 23 | 36 | 10.01 | 101.9 | 7.79
2009-06-25 | 20:00 | 16.15 | 23 | 36 | 10.03 | 101.9 | 7.79
2009-06-25 | 20:15 | 16.03 | 23 | 36 | 10.05 | 101.9 | 7.79
2009-06-25 | 20:30 | 15.89 | 23 | 36 | 10.03 | 101.3 | 7.79
2009-06-25 | 20:45 | 15.73 | 23 | 36 | 10.08 | 101.5 | 7.78
2009-06-25 | 21:00 | 15.58 | 23 | 36 | 10.07 | 101.1 | 7.77
2009-06-25 | 21:15 | 15.45 | 24 | 36 | 10.08 | 100.9 | 7.78
2009-06-25 | 21:30 | 15.35 | 24 | 36 | 10.13 | 101.2 | 7.77
2009-06-25 | 21:45 | 15.24 | 24 | 36 | 10.16 | 101.3 | 7.77
2009-06-25 | 22:00 | 15.12 | 24 | 37 | 10.2 | 101.4 | 7.77
2009-06-25 | 22:15 | 15 | 24 | 37 | 10.24 | 101.6 | 7.77
2009-06-25 | 22:30 | 14.88 | 24 | 37 | 10.22 | 101.1 | 7.78
2009-06-25 | 22:45 | 14.75 | 24 | 37 | 10.28 | 101.4 | 7.77
2009-06-25 | 23:00 | 14.64 | 24 | 37 | 10.32 | 101.5 | 7.77
2009-06-25 | 23:15 | 14.52 | 24 | 37 | 10.35 | 101.6 | 7.78
2009-06-25 | 23:30 | 14.38 | 24 | 37 | 10.4 | 101.7 | 7.77
2009-06-25 | 23:45 | 14.23 | 24 | 37 | 10.39 | 101.3 | 7.77
2009-06-26 | 0:00 | 14.08 | 24 | 37 | 10.46 | 101.7 | 7.77
2009-06-26 | 0:15 | 13.95 | 24 | 37 | 10.51 | 101.8 | 7.77
2009-06-26 | 0:30 | 13.84 | 24 | 37 | 10.54 | 101.9 | 7.77
2009-06-26 | 0:45 | 13.75 | 24 | 37 | 10.52 | 101.5 | 7.76
2009-06-26 | 1:01 | 13.66 | 24 | 38 | 10.52 | 101.3 | 7.76
2009-06-26 | 1:16 | 13.57 | 24 | 38 | 10.59 | 101.8 | 7.76
2009-06-26 | 1:31 | 13.45 | 24 | 38 | 10.59 | 101.5 | 7.76
2009-06-26 | 1:46 | 13.32 | 24 | 38 | 10.61 | 101.4 | 7.76
2009-06-26 | 2:01 | 13.15 | 24 | 37 | 10.66 | 101.5 | 7.75
2009-06-26 | 2:16 | 12.97 | 24 | 37 | 10.71 | 101.5 | 7.75
2009-06-26 | 2:31 | 12.77 | 24 | 37 | 10.8 | 102 | 7.74
2009-06-26 | 2:46 | 12.59 | 24 | 37 | 10.84 | 101.9 | 7.73
2009-06-26 | 3:01 | 12.41 | 24 | 37 | 10.87 | 101.8 | 7.73
2009-06-26 | 3:16 | 12.24 | 24 | 36 | 10.92 | 101.9 | 7.72
2009-06-26 | 3:31 | 12.07 | 23 | 36 | 10.96 | 101.9 | 7.72
2009-06-26 | 3:46 | 11.91 | 23 | 36 | 11.06 | 102.4 | 7.71
2009-06-26 | 4:01 | 11.75 | 23 | 35 | 11.08 | 102.2 | 7.7
2009-06-26 | 4:16 | 11.6 | 23 | 35 | 11.17 | 102.7 | 7.7
2009-06-26 | 4:31 | 11.46 | 23 | 35 | 11.17 | 102.4 | 7.69
2009-06-26 | 4:46 | 11.32 | 22 | 35 | 11.25 | 102.8 | 7.68
2009-06-26 | 5:01 | 11.19 | 22 | 34 | 11.26 | 102.5 | 7.68
2009-06-26 | 5:16 | 11.06 | 22 | 34 | 11.34 | 103 | 7.67
2009-06-26 | 5:31 | 10.93 | 22 | 34 | 11.34 | 102.7 | 7.67
2009-06-26 | 5:46 | 10.79 | 22 | 34 | 11.43 | 103.2 | 7.67
2009-06-26 | 6:01 | 10.65 | 22 | 34 | 11.52 | 103.7 | 7.67
2009-06-26 | 6:16 | 10.51 | 22 | 33 | 11.58 | 103.8 | 7.67
2009-06-26 | 6:31 | 10.38 | 22 | 33 | 11.64 | 104 | 7.66
2009-06-26 | 6:46 | 10.27 | 22 | 33 | 11.64 | 103.7 | 7.67
2009-06-26 | 7:01 | 10.17 | 22 | 33 | 11.67 | 103.8 | 7.67
2009-06-26 | 7:16 | 10.09 | 21 | 33 | 11.71 | 104 | 7.67
2009-06-26 | 7:31 | 10.01 | 21 | 33 | 11.81 | 104.7 | 7.68
2009-06-26 | 7:46 | 9.95 | 21 | 33 | 11.88 | 105.1 | 7.69
2009-06-26 | 8:01 | 9.89 | 21 | 33 | 11.94 | 105.5 | 7.69
2009-06-26 | 8:16 | 9.84 | 21 | 33 | 12 | 105.9 | 7.69
2009-06-26 | 8:31 | 9.79 | 21 | 33 | 11.99 | 105.7 | 7.7
2009-06-26 | 8:46 | 9.75 | 21 | 33 | 12.04 | 106.1 | 7.71
2009-06-26 | 9:01 | 9.74 | 21 | 33 | 12.11 | 106.6 | 7.71
2009-06-26 | 9:16 | 9.76 | 21 | 33 | 12.21 | 107.5 | 7.71
2009-06-26 | 9:31 | 9.81 | 21 | 33 | 12.24 | 107.9 | 7.72
2009-06-26 | 9:46 | 9.9 | 21 | 33 | 12.21 | 107.9 | 7.72
2009-06-26 | 10:01 | 10.03 | 21 | 33 | 12.1 | 107.2 | 7.72
2009-06-26 | 10:16 | 10.19 | 21 | 33 | 12.03 | 107 | 7.73
2009-06-26 | 10:31 | 10.39 | 21 | 33 | 11.94 | 106.8 | 7.73
2009-06-26 | 10:46 | 10.6 | 21 | 33 | 11.87 | 106.6 | 7.74
2009-06-26 | 11:01 | 10.84 | 21 | 33 | 11.84 | 107 | 7.75
2009-06-26 | 11:16 | 11.1 | 21 | 33 | 11.9 | 108.2 | 7.75
2009-06-26 | 11:31 | 11.38 | 21 | 33 | 11.85 | 108.4 | 7.75
2009-06-26 | 11:46 | 11.68 | 22 | 33 | 11.78 | 108.5 | 7.76
2009-06-26 | 12:01 | 12 | 22 | 33 | 11.76 | 109.1 | 7.77
2009-06-26 | 12:16 | 12.31 | 22 | 33 | 11.7 | 109.3 | 7.77
2009-06-26 | 12:31 | 12.64 | 22 | 33 | 11.56 | 108.9 | 7.77
2009-06-26 | 12:46 | 12.99 | 22 | 33 | 11.5 | 109.1 | 7.78
2009-06-26 | 13:00 | 13.34 | 22 | 34 | 11.35 | 108.6 | 7.79
2009-06-26 | 13:15 | 13.69 | 22 | 34 | 11.23 | 108.2 | 7.8
2009-06-26 | 13:30 | 14.05 | 22 | 34 | 11.17 | 108.5 | 7.81
2009-06-26 | 13:45 | 14.38 | 22 | 34 | 11.09 | 108.5 | 7.81
2009-06-26 | 14:00 | 14.69 | 22 | 34 | 10.99 | 108.2 | 7.81
2009-06-26 | 14:15 | 14.98 | 22 | 34 | 10.81 | 107.2 | 7.82
2009-06-26 | 14:30 | 15.24 | 22 | 34 | 10.73 | 107 | 7.82
2009-06-26 | 14:45 | 15.49 | 22 | 34 | 10.65 | 106.7 | 7.83
2009-06-26 | 15:00 | 15.69 | 22 | 34 | 10.58 | 106.4 | 7.83
2009-06-26 | 15:15 | 15.87 | 22 | 34 | 10.5 | 106.1 | 7.83
2009-06-26 | 15:30 | 16.02 | 22 | 35 | 10.38 | 105.2 | 7.84
2009-06-26 | 15:45 | 16.12 | 23 | 35 | 10.36 | 105.2 | 7.83
2009-06-26 | 16:00 | 16.2 | 23 | 35 | 10.28 | 104.6 | 7.83
2009-06-26 | 16:15 | 16.25 | 23 | 35 | 10.22 | 104.1 | 7.83
2009-06-26 | 16:30 | 16.3 | 23 | 35 | 10.18 | 103.8 | 7.84
2009-06-26 | 16:45 | 16.3 | 23 | 35 | 10.19 | 103.9 | 7.83
2009-06-26 | 17:00 | 16.3 | 23 | 35 | 10.19 | 103.8 | 7.83
2009-06-26 | 17:15 | 16.25 | 23 | 35 | 10.18 | 103.7 | 7.83
2009-06-26 | 17:30 | 16.2 | 23 | 35 | 10.18 | 103.6 | 7.82
2009-06-26 | 17:45 | 16.15 | 23 | 35 | 10.12 | 102.9 | 7.81
2009-06-26 | 18:00 | 16.09 | 23 | 36 | 10.11 | 102.7 | 7.81
2009-06-26 | 18:15 | 16.04 | 23 | 36 | 10.15 | 102.9 | 7.81
2009-06-26 | 18:30 | 15.98 | 23 | 36 | 10.1 | 102.2 | 7.8
2009-06-26 | 18:45 | 15.93 | 23 | 36 | 10.13 | 102.5 | 7.8
2009-06-26 | 19:00 | 15.88 | 23 | 36 | 10.15 | 102.5 | 7.79
2009-06-26 | 19:15 | 15.82 | 23 | 36 | 10.12 | 102.1 | 7.79
2009-06-26 | 19:30 | 15.76 | 24 | 36 | 10.09 | 101.7 | 7.79
2009-06-26 | 19:45 | 15.67 | 24 | 36 | 10.16 | 102.2 | 7.79
2009-06-26 | 20:00 | 15.56 | 24 | 36 | 10.14 | 101.8 | 7.78
2009-06-26 | 20:15 | 15.43 | 24 | 37 | 10.17 | 101.7 | 7.78
2009-06-26 | 20:30 | 15.3 | 24 | 37 | 10.21 | 102 | 7.78
2009-06-26 | 20:45 | 15.14 | 24 | 37 | 10.25 | 102 | 7.77
2009-06-26 | 21:00 | 14.96 | 24 | 37 | 10.26 | 101.7 | 7.77
2009-06-26 | 21:15 | 14.79 | 24 | 37 | 10.27 | 101.4 | 7.76
2009-06-26 | 21:30 | 14.63 | 24 | 37 | 10.29 | 101.3 | 7.76
2009-06-26 | 21:45 | 14.48 | 24 | 37 | 10.38 | 101.8 | 7.75
2009-06-26 | 22:00 | 14.34 | 24 | 37 | 10.38 | 101.5 | 7.75
2009-06-26 | 22:15 | 14.2 | 24 | 37 | 10.41 | 101.4 | 7.75
2009-06-26 | 22:30 | 14.06 | 24 | 38 | 10.49 | 101.9 | 7.75
2009-06-26 | 22:45 | 13.92 | 24 | 38 | 10.52 | 101.9 | 7.75
2009-06-26 | 23:00 | 13.78 | 25 | 38 | 10.58 | 102.2 | 7.75
2009-06-26 | 23:15 | 13.62 | 25 | 38 | 10.58 | 101.8 | 7.75
2009-06-26 | 23:30 | 13.48 | 25 | 38 | 10.61 | 101.7 | 7.75
2009-06-26 | 23:45 | 13.33 | 25 | 38 | 10.69 | 102.2 | 7.75
2009-06-27 | 0:00 | 13.19 | 25 | 38 | 10.69 | 101.9 | 7.75
2009-06-27 | 0:15 | 13.05 | 25 | 38 | 10.77 | 102.3 | 7.75
2009-06-27 | 0:30 | 12.89 | 25 | 38 | 10.78 | 102.1 | 7.75
2009-06-27 | 0:45 | 12.74 | 25 | 39 | 10.86 | 102.5 | 7.74
2009-06-27 | 1:01 | 12.59 | 25 | 39 | 10.87 | 102.2 | 7.74
2009-06-27 | 1:16 | 12.44 | 25 | 39 | 10.91 | 102.2 | 7.74
2009-06-27 | 1:31 | 12.28 | 25 | 39 | 10.99 | 102.7 | 7.74
2009-06-27 | 1:46 | 12.13 | 25 | 39 | 11.01 | 102.5 | 7.74
2009-06-27 | 2:01 | 11.97 | 25 | 39 | 11.1 | 102.9 | 7.74
2009-06-27 | 2:16 | 11.81 | 25 | 39 | 11.16 | 103.1 | 7.74
2009-06-27 | 2:31 | 11.65 | 25 | 39 | 11.16 | 102.8 | 7.73
2009-06-27 | 2:46 | 11.5 | 25 | 39 | 11.25 | 103.2 | 7.73
2009-06-27 | 3:01 | 11.34 | 25 | 39 | 11.26 | 102.9 | 7.73
2009-06-27 | 3:16 | 11.19 | 25 | 39 | 11.29 | 102.8 | 7.73
2009-06-27 | 3:31 | 11.04 | 25 | 38 | 11.38 | 103.3 | 7.73
2009-06-27 | 3:46 | 10.89 | 25 | 38 | 11.46 | 103.7 | 7.72
2009-06-27 | 4:01 | 10.74 | 25 | 38 | 11.46 | 103.3 | 7.72
2009-06-27 | 4:16 | 10.59 | 25 | 38 | 11.49 | 103.2 | 7.72
2009-06-27 | 4:31 | 10.44 | 24 | 38 | 11.58 | 103.7 | 7.71
2009-06-27 | 4:46 | 10.3 | 24 | 37 | 11.65 | 103.9 | 7.7
2009-06-27 | 5:01 | 10.16 | 24 | 37 | 11.71 | 104.1 | 7.69
2009-06-27 | 5:16 | 10.03 | 24 | 37 | 11.7 | 103.8 | 7.7
2009-06-27 | 5:31 | 9.91 | 24 | 37 | 11.78 | 104.1 | 7.69
2009-06-27 | 5:46 | 9.78 | 24 | 36 | 11.84 | 104.4 | 7.69
2009-06-27 | 6:01 | 9.67 | 24 | 36 | 11.9 | 104.6 | 7.69
2009-06-27 | 6:16 | 9.56 | 23 | 36 | 11.96 | 104.8 | 7.69
2009-06-27 | 6:31 | 9.46 | 23 | 36 | 12.01 | 105 | 7.68
2009-06-27 | 6:46 | 9.37 | 23 | 36 | 12.01 | 104.8 | 7.68
2009-06-27 | 7:01 | 9.29 | 23 | 36 | 12.09 | 105.3 | 7.68
2009-06-27 | 7:16 | 9.22 | 23 | 36 | 12.1 | 105.2 | 7.68
2009-06-27 | 7:31 | 9.18 | 23 | 36 | 12.17 | 105.7 | 7.69
2009-06-27 | 7:46 | 9.14 | 23 | 35 | 12.24 | 106.2 | 7.69
2009-06-27 | 8:01 | 9.12 | 23 | 35 | 12.21 | 106 | 7.7
2009-06-27 | 8:16 | 9.1 | 23 | 35 | 12.24 | 106.1 | 7.7
2009-06-27 | 8:31 | 9.08 | 23 | 35 | 12.26 | 106.3 | 7.71
2009-06-27 | 8:46 | 9.08 | 23 | 35 | 12.32 | 106.8 | 7.71
2009-06-27 | 9:01 | 9.09 | 23 | 35 | 12.45 | 108 | 7.72
2009-06-27 | 9:16 | 9.13 | 23 | 35 | 12.52 | 108.6 | 7.72
2009-06-27 | 9:31 | 9.2 | 23 | 35 | 12.47 | 108.3 | 7.73
2009-06-27 | 9:46 | 9.31 | 23 | 35 | 12.47 | 108.7 | 7.73
2009-06-27 | 10:01 | 9.48 | 23 | 35 | 12.37 | 108.2 | 7.74
2009-06-27 | 10:16 | 9.68 | 23 | 35 | 12.33 | 108.4 | 7.75
2009-06-27 | 10:31 | 9.93 | 23 | 35 | 12.26 | 108.5 | 7.75
2009-06-27 | 10:46 | 10.2 | 23 | 35 | 12.19 | 108.5 | 7.75
2009-06-27 | 11:01 | 10.49 | 23 | 35 | 12.15 | 108.9 | 7.76
2009-06-27 | 11:16 | 10.8 | 23 | 35 | 12.12 | 109.4 | 7.77
2009-06-27 | 11:31 | 11.13 | 23 | 35 | 12.08 | 109.8 | 7.77
2009-06-27 | 11:46 | 11.48 | 23 | 35 | 12.02 | 110.2 | 7.77
2009-06-27 | 12:01 | 11.84 | 23 | 35 | 11.84 | 109.5 | 7.78
2009-06-27 | 12:16 | 12.22 | 23 | 36 | 11.74 | 109.5 | 7.79
2009-06-27 | 12:31 | 12.6 | 23 | 36 | 11.66 | 109.7 | 7.8
2009-06-27 | 12:46 | 12.99 | 23 | 36 | 11.57 | 109.8 | 7.81
2009-06-27 | 13:00 | 13.38 | 23 | 36 | 11.46 | 109.7 | 7.81
2009-06-27 | 13:15 | 13.78 | 23 | 36 | 11.35 | 109.6 | 7.81
2009-06-27 | 13:30 | 14.19 | 23 | 36 | 11.19 | 109 | 7.82
2009-06-27 | 13:45 | 14.58 | 23 | 36 | 11.05 | 108.6 | 7.83
2009-06-27 | 14:00 | 14.94 | 23 | 36 | 10.96 | 108.5 | 7.84
2009-06-27 | 14:15 | 15.29 | 24 | 36 | 10.82 | 107.9 | 7.83
2009-06-27 | 14:30 | 15.59 | 24 | 36 | 10.69 | 107.3 | 7.83
2009-06-27 | 14:45 | 15.88 | 24 | 36 | 10.58 | 107 | 7.84
2009-06-27 | 15:00 | 16.13 | 24 | 37 | 10.5 | 106.7 | 7.85
2009-06-27 | 15:15 | 16.34 | 24 | 37 | 10.37 | 105.8 | 7.85
2009-06-27 | 15:30 | 16.53 | 24 | 37 | 10.29 | 105.5 | 7.85
2009-06-27 | 15:45 | 16.68 | 24 | 37 | 10.22 | 105 | 7.85
2009-06-27 | 16:00 | 16.8 | 24 | 37 | 10.2 | 105.1 | 7.85
2009-06-27 | 16:15 | 16.9 | 24 | 37 | 10.17 | 105 | 7.84
2009-06-27 | 16:30 | 16.97 | 24 | 37 | 10.13 | 104.7 | 7.85
2009-06-27 | 16:45 | 17.02 | 24 | 37 | 10.05 | 104 | 7.83
2009-06-27 | 17:00 | 17.05 | 24 | 37 | 10.04 | 104 | 7.83
2009-06-27 | 17:15 | 17.06 | 24 | 37 | 9.97 | 103.3 | 7.84
2009-06-27 | 17:30 | 17.06 | 24 | 37 | 9.93 | 102.9 | 7.84
2009-06-27 | 17:45 | 17.07 | 24 | 37 | 9.9 | 102.6 | 7.83
2009-06-27 | 18:00 | 17.06 | 24 | 37 | 9.93 | 102.8 | 7.82
2009-06-27 | 18:15 | 17.04 | 24 | 38 | 9.93 | 102.8 | 7.82
2009-06-27 | 18:30 | 17.02 | 24 | 38 | 9.92 | 102.7 | 7.78
2009-06-27 | 18:45 | 17.01 | 25 | 38 | 9.87 | 102.1 | 7.79
2009-06-27 | 19:00 | 16.99 | 25 | 38 | 9.88 | 102.3 | 7.79
2009-06-27 | 19:15 | 16.97 | 25 | 38 | 9.89 | 102.2 | 7.78
2009-06-27 | 19:30 | 16.94 | 25 | 38 | 9.84 | 101.6 | 7.78
2009-06-27 | 19:45 | 16.88 | 25 | 38 | 9.82 | 101.4 | 7.78
2009-06-27 | 20:00 | 16.8 | 25 | 38 | 9.82 | 101.2 | 7.77
2009-06-27 | 20:15 | 16.69 | 25 | 38 | 9.82 | 101 | 7.76
2009-06-27 | 20:30 | 16.58 | 25 | 38 | 9.83 | 100.8 | 7.77
2009-06-27 | 20:45 | 16.45 | 25 | 38 | 9.85 | 100.8 | 7.76
2009-06-27 | 21:00 | 16.3 | 25 | 39 | 9.9 | 101 | 7.76
2009-06-27 | 21:15 | 16.17 | 25 | 39 | 9.91 | 100.7 | 7.76
2009-06-27 | 21:30 | 16.05 | 25 | 39 | 9.92 | 100.6 | 7.76
2009-06-27 | 21:45 | 15.94 | 25 | 39 | 9.94 | 100.6 | 7.76
2009-06-27 | 22:00 | 15.83 | 25 | 39 | 10.01 | 101 | 7.75
2009-06-27 | 22:15 | 15.73 | 25 | 39 | 10 | 100.8 | 7.75
2009-06-27 | 22:30 | 15.62 | 26 | 39 | 10.01 | 100.6 | 7.75
2009-06-27 | 22:45 | 15.52 | 26 | 39 | 10.08 | 101.1 | 7.75
2009-06-27 | 23:00 | 15.43 | 26 | 40 | 10.11 | 101.2 | 7.75
2009-06-27 | 23:15 | 15.33 | 26 | 40 | 10.14 | 101.3 | 7.76
2009-06-27 | 23:30 | 15.22 | 26 | 40 | 10.13 | 101 | 7.77
2009-06-27 | 23:45 | 15.12 | 26 | 40 | 10.15 | 100.9 | 7.76
2009-06-28 | 0:00 | 15.01 | 26 | 40 | 10.21 | 101.2 | 7.76
2009-06-28 | 0:15 | 14.89 | 26 | 40 | 10.21 | 101 | 7.77
2009-06-28 | 0:30 | 14.77 | 26 | 40 | 10.28 | 101.4 | 7.77
2009-06-28 | 0:45 | 14.65 | 26 | 40 | 10.28 | 101.2 | 7.77
2009-06-28 | 1:01 | 14.52 | 26 | 40 | 10.35 | 101.6 | 7.78
2009-06-28 | 1:16 | 14.38 | 26 | 41 | 10.4 | 101.8 | 7.78
2009-06-28 | 1:31 | 14.24 | 26 | 41 | 10.4 | 101.4 | 7.78
2009-06-28 | 1:46 | 14.1 | 26 | 41 | 10.47 | 101.8 | 7.78
2009-06-28 | 2:01 | 13.96 | 27 | 41 | 10.52 | 102 | 7.78
2009-06-28 | 2:16 | 13.82 | 27 | 41 | 10.56 | 102 | 7.78
2009-06-28 | 2:31 | 13.67 | 27 | 41 | 10.57 | 101.8 | 7.77
2009-06-28 | 2:46 | 13.51 | 26 | 41 | 10.64 | 102.2 | 7.78
2009-06-28 | 3:01 | 13.35 | 26 | 41 | 10.65 | 101.9 | 7.77
2009-06-28 | 3:16 | 13.19 | 26 | 41 | 10.73 | 102.3 | 7.76
2009-06-28 | 3:31 | 13.03 | 26 | 40 | 10.74 | 102 | 7.76
2009-06-28 | 3:46 | 12.86 | 26 | 40 | 10.85 | 102.6 | 7.75
2009-06-28 | 4:01 | 12.69 | 26 | 40 | 10.86 | 102.3 | 7.75
2009-06-28 | 4:16 | 12.52 | 26 | 40 | 10.93 | 102.6 | 7.75
2009-06-28 | 4:31 | 12.34 | 25 | 39 | 10.96 | 102.5 | 7.75
2009-06-28 | 4:46 | 12.17 | 25 | 39 | 11.04 | 102.8 | 7.73
2009-06-28 | 5:01 | 12.01 | 25 | 38 | 11.05 | 102.5 | 7.72
2009-06-28 | 5:16 | 11.85 | 25 | 38 | 11.09 | 102.6 | 7.72
2009-06-28 | 5:31 | 11.7 | 25 | 38 | 11.13 | 102.6 | 7.72
2009-06-28 | 5:46 | 11.57 | 24 | 38 | 11.17 | 102.7 | 7.71
2009-06-28 | 6:01 | 11.44 | 24 | 37 | 11.27 | 103.3 | 7.71
2009-06-28 | 6:16 | 11.32 | 24 | 37 | 11.29 | 103.2 | 7.71
2009-06-28 | 6:31 | 11.21 | 24 | 37 | 11.38 | 103.7 | 7.7
2009-06-28 | 6:46 | 11.12 | 24 | 37 | 11.4 | 103.6 | 7.71
2009-06-28 | 7:01 | 11.03 | 24 | 37 | 11.42 | 103.7 | 7.7
2009-06-28 | 7:16 | 10.97 | 24 | 36 | 11.52 | 104.4 | 7.71
2009-06-28 | 7:31 | 10.91 | 24 | 36 | 11.53 | 104.4 | 7.71
2009-06-28 | 7:46 | 10.86 | 24 | 36 | 11.55 | 104.4 | 7.71
2009-06-28 | 8:01 | 10.82 | 23 | 36 | 11.64 | 105.1 | 7.72
2009-06-28 | 8:16 | 10.79 | 23 | 36 | 11.64 | 105 | 7.72
2009-06-28 | 8:31 | 10.75 | 23 | 36 | 11.73 | 105.7 | 7.73
2009-06-28 | 8:46 | 10.73 | 23 | 36 | 11.78 | 106.2 | 7.73
2009-06-28 | 9:01 | 10.73 | 23 | 36 | 11.91 | 107.3 | 7.74
2009-06-28 | 9:16 | 10.76 | 23 | 36 | 12 | 108.2 | 7.75
2009-06-28 | 9:31 | 10.82 | 23 | 36 | 11.94 | 107.9 | 7.75
2009-06-28 | 9:46 | 10.93 | 23 | 36 | 11.92 | 107.9 | 7.76
2009-06-28 | 10:01 | 11.09 | 23 | 36 | 11.8 | 107.2 | 7.76
2009-06-28 | 10:16 | 11.27 | 23 | 36 | 11.77 | 107.4 | 7.77
2009-06-28 | 10:31 | 11.5 | 23 | 36 | 11.72 | 107.5 | 7.76
2009-06-28 | 10:46 | 11.76 | 23 | 36 | 11.63 | 107.4 | 7.77
2009-06-28 | 11:01 | 12.03 | 23 | 36 | 11.57 | 107.4 | 7.78
2009-06-28 | 11:16 | 12.32 | 23 | 36 | 11.61 | 108.5 | 7.79
2009-06-28 | 11:31 | 12.63 | 23 | 36 | 11.59 | 109.1 | 7.79
2009-06-28 | 11:46 | 12.96 | 23 | 36 | 11.53 | 109.4 | 7.78
2009-06-28 | 12:01 | 13.3 | 23 | 36 | 11.41 | 109 | 7.8
2009-06-28 | 12:16 | 13.66 | 23 | 36 | 11.31 | 108.9 | 7.8
2009-06-28 | 12:31 | 14.02 | 23 | 36 | 11.24 | 109.1 | 7.81
2009-06-28 | 12:46 | 14.39 | 23 | 36 | 11.15 | 109.2 | 7.81
2009-06-28 | 13:00 | 14.76 | 23 | 36 | 11.06 | 109.1 | 7.81
2009-06-28 | 13:15 | 15.14 | 24 | 36 | 10.94 | 108.8 | 7.8
2009-06-28 | 13:30 | 15.51 | 24 | 36 | 10.79 | 108.2 | 7.83
2009-06-28 | 13:45 | 15.88 | 24 | 36 | 10.65 | 107.6 | 7.83
2009-06-28 | 14:00 | 16.23 | 24 | 37 | 10.59 | 107.8 | 7.83
2009-06-28 | 14:15 | 16.55 | 24 | 37 | 10.47 | 107.3 | 7.83
2009-06-28 | 14:30 | 16.84 | 24 | 37 | 10.31 | 106.4 | 7.83
2009-06-28 | 14:45 | 17.11 | 24 | 37 | 10.25 | 106.3 | 7.83
2009-06-28 | 15:00 | 17.34 | 24 | 37 | 10.18 | 106.1 | 7.84
2009-06-28 | 15:15 | 17.55 | 24 | 37 | 10.05 | 105.2 | 7.84
2009-06-28 | 15:30 | 17.72 | 24 | 37 | 10.01 | 105.1 | 7.84
2009-06-28 | 15:45 | 17.84 | 24 | 37 | 9.96 | 104.9 | 7.83
2009-06-28 | 16:00 | 17.94 | 24 | 37 | 9.87 | 104.1 | 7.85
2009-06-28 | 16:15 | 18.02 | 24 | 37 | 9.86 | 104.1 | 7.85
2009-06-28 | 16:30 | 18.06 | 24 | 37 | 9.82 | 103.9 | 7.85
2009-06-28 | 16:45 | 18.08 | 24 | 37 | 9.79 | 103.6 | 7.84
2009-06-28 | 17:00 | 18.09 | 24 | 38 | 9.77 | 103.4 | 7.85
2009-06-28 | 17:15 | 18.04 | 24 | 38 | 9.75 | 103.1 | 7.83
2009-06-28 | 17:30 | 17.98 | 24 | 38 | 9.74 | 102.8 | 7.83
2009-06-28 | 17:45 | 17.93 | 24 | 38 | 9.73 | 102.7 | 7.84
2009-06-28 | 18:00 | 17.88 | 25 | 38 | 9.73 | 102.5 | 7.82
2009-06-28 | 18:15 | 17.85 | 25 | 38 | 9.71 | 102.2 | 7.81
2009-06-28 | 18:30 | 17.82 | 25 | 38 | 9.67 | 101.7 | 7.8
2009-06-28 | 18:45 | 17.76 | 25 | 38 | 9.64 | 101.3 | 7.82
2009-06-28 | 19:00 | 17.74 | 25 | 38 | 9.68 | 101.7 | 7.81
2009-06-28 | 19:15 | 17.68 | 25 | 38 | 9.7 | 101.8 | 7.81
2009-06-28 | 19:30 | 17.62 | 25 | 38 | 9.7 | 101.7 | 7.79
2009-06-28 | 19:45 | 17.54 | 25 | 38 | 9.7 | 101.5 | 7.77
2009-06-28 | 20:00 | 17.44 | 25 | 38 | 9.68 | 101.1 | 7.79
2009-06-28 | 20:15 | 17.31 | 25 | 39 | 9.73 | 101.3 | 7.78
2009-06-28 | 20:30 | 17.19 | 25 | 39 | 9.76 | 101.4 | 7.75
2009-06-28 | 20:45 | 17.04 | 25 | 39 | 9.74 | 100.9 | 7.75
2009-06-28 | 21:00 | 16.87 | 25 | 39 | 9.78 | 100.9 | 7.77
2009-06-28 | 21:15 | 16.71 | 25 | 39 | 9.83 | 101.1 | 7.78
2009-06-28 | 21:30 | 16.56 | 25 | 39 | 9.83 | 100.8 | 7.79
2009-06-28 | 21:45 | 16.41 | 25 | 39 | 9.88 | 101 | 7.78
2009-06-28 | 22:00 | 16.27 | 25 | 39 | 9.88 | 100.7 | 7.78
2009-06-28 | 22:15 | 16.14 | 26 | 39 | 9.9 | 100.6 | 7.78
2009-06-28 | 22:30 | 16.01 | 26 | 39 | 9.97 | 101 | 7.78
2009-06-28 | 22:45 | 15.88 | 26 | 39 | 9.96 | 100.7 | 7.78
2009-06-28 | 23:00 | 15.77 | 26 | 40 | 9.97 | 100.5 | 7.78
2009-06-28 | 23:15 | 15.66 | 26 | 40 | 9.99 | 100.5 | 7.78
2009-06-28 | 23:30 | 15.54 | 26 | 40 | 10.02 | 100.6 | 7.78
2009-06-28 | 23:45 | 15.41 | 26 | 40 | 10.1 | 101.1 | 7.78
2009-06-29 | 0:00 | 15.27 | 26 | 40 | 10.11 | 100.9 | 7.78
2009-06-29 | 0:15 | 15.14 | 26 | 40 | 10.14 | 100.8 | 7.78
2009-06-29 | 0:30 | 15 | 26 | 40 | 10.21 | 101.3 | 7.78
2009-06-29 | 0:45 | 14.85 | 26 | 40 | 10.28 | 101.6 | 7.78
2009-06-29 | 1:01 | 14.69 | 26 | 40 | 10.33 | 101.7 | 7.78
2009-06-29 | 1:16 | 14.54 | 26 | 40 | 10.38 | 101.9 | 7.77
2009-06-29 | 1:31 | 14.39 | 26 | 40 | 10.42 | 102 | 7.77
2009-06-29 | 1:46 | 14.22 | 26 | 40 | 10.47 | 102.1 | 7.76
2009-06-29 | 2:01 | 14.06 | 26 | 40 | 10.47 | 101.7 | 7.77
2009-06-29 | 2:16 | 13.9 | 26 | 40 | 10.54 | 102.1 | 7.76
2009-06-29 | 2:31 | 13.73 | 26 | 40 | 10.55 | 101.7 | 7.76
2009-06-29 | 2:46 | 13.56 | 26 | 39 | 10.59 | 101.7 | 7.75
2009-06-29 | 3:01 | 13.38 | 25 | 39 | 10.68 | 102.2 | 7.74
2009-06-29 | 3:16 | 13.2 | 25 | 39 | 10.75 | 102.4 | 7.75
2009-06-29 | 3:31 | 13.01 | 25 | 38 | 10.81 | 102.6 | 7.73
2009-06-29 | 3:46 | 12.82 | 25 | 38 | 10.82 | 102.3 | 7.73
2009-06-29 | 4:01 | 12.64 | 24 | 37 | 10.85 | 102.2 | 7.72
2009-06-29 | 4:16 | 12.46 | 24 | 37 | 10.95 | 102.7 | 7.72
2009-06-29 | 4:31 | 12.29 | 24 | 37 | 11.02 | 102.9 | 7.71
2009-06-29 | 4:46 | 12.13 | 24 | 36 | 11.02 | 102.6 | 7.71
2009-06-29 | 5:01 | 11.97 | 23 | 36 | 11.09 | 102.8 | 7.7
2009-06-29 | 5:16 | 11.82 | 23 | 36 | 11.16 | 103.2 | 7.69
2009-06-29 | 5:31 | 11.68 | 23 | 35 | 11.22 | 103.3 | 7.69
2009-06-29 | 5:46 | 11.55 | 23 | 35 | 11.26 | 103.4 | 7.68
2009-06-29 | 6:01 | 11.42 | 23 | 35 | 11.27 | 103.2 | 7.69
2009-06-29 | 6:16 | 11.29 | 23 | 35 | 11.35 | 103.6 | 7.68
2009-06-29 | 6:31 | 11.16 | 23 | 35 | 11.43 | 104 | 7.68
2009-06-29 | 6:46 | 11.05 | 22 | 35 | 11.42 | 103.7 | 7.68
2009-06-29 | 7:01 | 10.95 | 22 | 34 | 11.51 | 104.3 | 7.67
2009-06-29 | 7:16 | 10.86 | 22 | 34 | 11.52 | 104.1 | 7.68
2009-06-29 | 7:31 | 10.78 | 22 | 34 | 11.61 | 104.7 | 7.67
2009-06-29 | 7:46 | 10.71 | 22 | 34 | 11.61 | 104.6 | 7.68
2009-06-29 | 8:01 | 10.64 | 22 | 34 | 11.64 | 104.6 | 7.68
2009-06-29 | 8:16 | 10.59 | 22 | 34 | 11.68 | 104.9 | 7.68
2009-06-29 | 8:31 | 10.53 | 22 | 34 | 11.79 | 105.8 | 7.69
2009-06-29 | 8:46 | 10.49 | 22 | 34 | 11.9 | 106.7 | 7.69
2009-06-29 | 9:01 | 10.47 | 22 | 34 | 11.96 | 107.1 | 7.7
2009-06-29 | 9:16 | 10.48 | 22 | 34 | 12.07 | 108.2 | 7.71
2009-06-29 | 9:31 | 10.53 | 22 | 34 | 12.08 | 108.4 | 7.72
2009-06-29 | 9:46 | 10.63 | 22 | 34 | 12.04 | 108.3 | 7.72
2009-06-29 | 10:01 | 10.76 | 22 | 34 | 11.98 | 108 | 7.72
2009-06-29 | 10:16 | 10.94 | 22 | 34 | 11.91 | 107.8 | 7.72
2009-06-29 | 10:31 | 11.15 | 22 | 34 | 11.84 | 107.7 | 7.74
2009-06-29 | 10:46 | 11.39 | 22 | 34 | 11.76 | 107.6 | 7.74
2009-06-29 | 11:01 | 11.64 | 22 | 34 | 11.75 | 108.1 | 7.74
2009-06-29 | 11:16 | 11.91 | 22 | 34 | 11.77 | 109 | 7.75
2009-06-29 | 11:31 | 12.2 | 22 | 34 | 11.69 | 109 | 7.75
2009-06-29 | 11:46 | 12.51 | 23 | 35 | 11.66 | 109.4 | 7.77
2009-06-29 | 12:01 | 12.83 | 23 | 36 | 11.61 | 109.8 | 7.77