## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AB & YSI 556 MPS & Sonde AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-18 |  | 2009-04-18 |  | A
Time | Computer set to atomic time |  | 13:30 |  | 13:30 |  | A
Temp | In Situ |  | 9.37 |  | 9.36 |  | A
DO | Winkler titration of surface water |  | 11.4 |  | 12.71 |  | Fail
SpC | In Situ |  | 65 |  | 66 |  | A
Ph | In Situ |  | 6.2 |  | 7.98 |  | Fail
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-04-18 13:30 | 9.37 | 6.6 | 17.38 | 151.7 | 65 | 45
Sonde | 2009-04-18 13:30 | 9.36 | 7.98 | 12.71 | 110.9 | 66 | 43
Grade |  | A |  |  |  | A
Audit 2 | 2009-04-21 15:45 | 17.81 | 7.56 | 13.33 | 140.3 | 58 | 41
Sonde | 2009-04-21 15:45 | 17.84 | 8.07 | 10.17 | 107.1 | 59 | 38
Grade |  | A |  |  |  | A
Audit 3 | 2009-04-25 12:00 | 8.08 | 7.03 | 17.38 | 147.7 | 67 | 47
Sonde | 2009-04-25 12:00 | 8.04 | 8.04 | 13.8 | 116.5 | 69 | 43
Grade |  | A |  |  |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AB & YSI 556 MPS & Sonde AA
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-25 |  | 2009-04-25 |  | A
Time | Computer set to atomic time |  | 13:24 |  | 13:27 |  | A
Temp | VWR NIST traceable surface water |  | 12.1 |  | 11.93 |  | A
DO | Winkler titration of surface water |  | 10.8 |  | 13.58 |  | 0
SpC | In Situ |  | 69 |  | 69 |  | A
Ph | In Situ |  | 8.78 |  | 8.8 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 Was not working for pH and DO. Sonde data is disqualified as a precaution |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-18 |  | 2009-04-18 |  | A
Time | Computer set to atomic time |  | 11:59 |  | 11:49 |  | A
Temp | VWR NIST traceable surface water |  | 23.5 |  | 23.41 |  | A
DO | Winkler titration of surface water |  | 8.5 |  | 7.45 |  | Fail
SpC | 148 µS/mL |  | 153.6 |  | 144 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-04-18 13:30 | 9.37 | 6.6 | 17.38 | 151.7 | 65 | 45
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-18 |  | 2009-04-18 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | VWR NIST traceable surface water |  | 12.67 |  | 12.71 |  | A
DO | Winkler titration of surface water |  | 11.5 |  | 18.63 |  | 0
SpC | 148 µS/mL |  | 153.6 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.002 |  | 7.13 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-21 |  | 2009-04-21 |  | A
Time | Computer set to atomic time |  | 8:56 |  | 8:56 |  | A
Temp | VWR NIST traceable surface water |  | 11 |  | 10.84 |  | A
DO | Winkler titration of surface water |  | 10.3 |  | 11.95 |  | Fail
SpC | 148 µS/mL |  | 153.6 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.13 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-04-21 15:45 | 17.81 | 7.56 | 13.33 | 140.3 | 58 | 41
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-21 |  | 2009-04-21 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 19.6 |  | 19.61 |  | A
DO | Winkler titration of surface water |  | 9.5 |  | 13.14 |  | 0
SpC | 148 µS/mL |  | 153.6 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-21 |  | 2009-04-21 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 19.6 |  | 19.61 |  | A
DO | Winkler titration of surface water |  | 9.5 |  | 13.14 |  | 0
SpC | 148 µS/mL |  | 153.6 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-04-25 12:00 | 8.08 | 7.03 | 17.38 | 147.7 | 67 | 47
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-25 |  | 2009-04-25 |  | A
Time | Computer set to atomic time |  | 13:24 |  | 13:24 |  | A
Temp | VWR NIST traceable surface water |  | 12.1 |  | 11.96 |  | A
DO | Winkler titration of surface water |  | 10.8 |  | 15.62 |  | 0
SpC | 148 µS/mL |  | 69 |  | 69 |  | A
Ph | pH 7 standard solution |  | 8.8 |  | 8.29 |  | Fail
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | TDS | SpCond | DO Conc | DO% | pH
M/D/Y |  | °C | mg/L | uS/cm | mg/L | %
2009-04-18 | 13:30 | 9.36 | 43 | 66
2009-04-18 | 13:45 | 9.77 | 43 | 66
2009-04-18 | 14:00 | 10.2 | 43 | 66
2009-04-18 | 14:15 | 10.66 | 43 | 65
2009-04-18 | 14:30 | 11.13 | 42 | 65
2009-04-18 | 14:45 | 11.56 | 42 | 65
2009-04-18 | 15:00 | 11.97 | 42 | 65
2009-04-18 | 15:15 | 12.29 | 42 | 65
2009-04-18 | 15:30 | 12.59 | 42 | 65
2009-04-18 | 15:45 | 12.8 | 42 | 65
2009-04-18 | 16:00 | 12.95 | 42 | 65
2009-04-18 | 16:15 | 12.99 | 42 | 65
2009-04-18 | 16:30 | 13.02 | 42 | 65
2009-04-18 | 16:45 | 13.02 | 42 | 65
2009-04-18 | 17:00 | 12.97 | 42 | 65
2009-04-18 | 17:15 | 12.91 | 42 | 65
2009-04-18 | 17:30 | 12.86 | 42 | 65
2009-04-18 | 17:45 | 12.79 | 42 | 65
2009-04-18 | 18:00 | 12.73 | 42 | 65
2009-04-18 | 18:15 | 12.65 | 42 | 65
2009-04-18 | 18:30 | 12.6 | 42 | 65
2009-04-18 | 18:45 | 12.56 | 42 | 65
2009-04-18 | 19:00 | 12.56 | 42 | 65
2009-04-18 | 19:15 | 12.55 | 42 | 65
2009-04-18 | 19:30 | 12.52 | 42 | 65
2009-04-18 | 19:45 | 12.48 | 42 | 65
2009-04-18 | 20:00 | 12.38 | 42 | 65
2009-04-18 | 20:15 | 12.25 | 42 | 65
2009-04-18 | 20:30 | 12.12 | 42 | 65
2009-04-18 | 20:45 | 11.95 | 42 | 65
2009-04-18 | 21:00 | 11.76 | 42 | 65
2009-04-18 | 21:15 | 11.56 | 42 | 65
2009-04-18 | 21:30 | 11.36 | 42 | 65
2009-04-18 | 21:45 | 11.17 | 42 | 65
2009-04-18 | 22:00 | 10.98 | 42 | 65
2009-04-18 | 22:15 | 10.81 | 42 | 65
2009-04-18 | 22:30 | 10.63 | 42 | 65
2009-04-18 | 22:45 | 10.46 | 42 | 65
2009-04-18 | 23:00 | 10.28 | 42 | 65
2009-04-18 | 23:15 | 10.11 | 42 | 65
2009-04-18 | 23:30 | 9.95 | 42 | 65
2009-04-18 | 23:45 | 9.8 | 43 | 66
2009-04-19 | 0:00 | 9.64 | 43 | 66
2009-04-19 | 0:15 | 9.5 | 43 | 66
2009-04-19 | 0:30 | 9.36 | 43 | 66
2009-04-19 | 0:45 | 9.22 | 43 | 66
2009-04-19 | 1:01 | 9.08 | 43 | 66
2009-04-19 | 1:16 | 8.95 | 43 | 66
2009-04-19 | 1:31 | 8.83 | 43 | 66
2009-04-19 | 1:46 | 8.7 | 43 | 66
2009-04-19 | 2:01 | 8.58 | 43 | 66
2009-04-19 | 2:16 | 8.46 | 43 | 66
2009-04-19 | 2:31 | 8.33 | 43 | 66
2009-04-19 | 2:46 | 8.21 | 43 | 66
2009-04-19 | 3:01 | 8.08 | 43 | 66
2009-04-19 | 3:16 | 7.94 | 43 | 66
2009-04-19 | 3:31 | 7.81 | 43 | 66
2009-04-19 | 3:46 | 7.67 | 43 | 66
2009-04-19 | 4:01 | 7.53 | 43 | 66
2009-04-19 | 4:16 | 7.38 | 43 | 66
2009-04-19 | 4:31 | 7.24 | 43 | 66
2009-04-19 | 4:46 | 7.1 | 43 | 66
2009-04-19 | 5:01 | 6.96 | 43 | 66
2009-04-19 | 5:16 | 6.82 | 43 | 66
2009-04-19 | 5:31 | 6.69 | 43 | 66
2009-04-19 | 5:46 | 6.56 | 43 | 67
2009-04-19 | 6:01 | 6.43 | 43 | 67
2009-04-19 | 6:16 | 6.3 | 43 | 66
2009-04-19 | 6:31 | 6.19 | 43 | 67
2009-04-19 | 6:46 | 6.08 | 43 | 67
2009-04-19 | 7:01 | 5.98 | 43 | 67
2009-04-19 | 7:16 | 5.88 | 43 | 67
2009-04-19 | 7:31 | 5.8 | 43 | 67
2009-04-19 | 7:46 | 5.72 | 43 | 67
2009-04-19 | 8:01 | 5.66 | 43 | 67
2009-04-19 | 8:16 | 5.62 | 43 | 67
2009-04-19 | 8:31 | 5.61 | 43 | 66
2009-04-19 | 8:46 | 5.62 | 43 | 66
2009-04-19 | 9:01 | 5.68 | 43 | 66
2009-04-19 | 9:16 | 5.71 | 43 | 66
2009-04-19 | 9:31 | 5.77 | 43 | 66
2009-04-19 | 9:46 | 5.85 | 43 | 66
2009-04-19 | 10:01 | 5.95 | 43 | 66
2009-04-19 | 10:16 | 6.05 | 43 | 66
2009-04-19 | 10:31 | 6.21 | 43 | 66
2009-04-19 | 10:46 | 6.38 | 43 | 66
2009-04-19 | 11:01 | 6.62 | 43 | 66
2009-04-19 | 11:16 | 6.95 | 43 | 66
2009-04-19 | 11:31 | 7.39 | 43 | 66
2009-04-19 | 11:46 | 7.92 | 43 | 66
2009-04-19 | 12:01 | 8.5 | 43 | 65
2009-04-19 | 12:16 | 9.11 | 43 | 65
2009-04-19 | 12:31 | 9.72 | 43 | 65
2009-04-19 | 12:46 | 10.32 | 42 | 65
2009-04-19 | 13:00 | 10.89 | 42 | 65
2009-04-19 | 13:15 | 11.45 | 42 | 65
2009-04-19 | 13:30 | 12.01 | 42 | 65
2009-04-19 | 13:45 | 12.58 | 42 | 65
2009-04-19 | 14:00 | 13.12 | 42 | 65
2009-04-19 | 14:15 | 13.64 | 42 | 65
2009-04-19 | 14:30 | 14.09 | 42 | 65
2009-04-19 | 14:45 | 14.5 | 42 | 65
2009-04-19 | 15:00 | 14.85 | 42 | 65
2009-04-19 | 15:15 | 15.15 | 42 | 65
2009-04-19 | 15:30 | 15.41 | 42 | 65
2009-04-19 | 15:45 | 15.59 | 42 | 65
2009-04-19 | 16:00 | 15.68 | 42 | 65
2009-04-19 | 16:15 | 15.7 | 42 | 65
2009-04-19 | 16:30 | 15.7 | 42 | 65
2009-04-19 | 16:45 | 15.66 | 42 | 65
2009-04-19 | 17:00 | 15.57 | 42 | 65
2009-04-19 | 17:15 | 15.45 | 42 | 65
2009-04-19 | 17:30 | 15.34 | 42 | 65
2009-04-19 | 17:45 | 15.23 | 42 | 65
2009-04-19 | 18:00 | 15.12 | 42 | 65
2009-04-19 | 18:15 | 15 | 42 | 65
2009-04-19 | 18:30 | 14.94 | 42 | 65
2009-04-19 | 18:45 | 14.9 | 42 | 65
2009-04-19 | 19:00 | 14.9 | 42 | 65
2009-04-19 | 19:15 | 14.88 | 42 | 65
2009-04-19 | 19:30 | 14.84 | 42 | 65
2009-04-19 | 19:45 | 14.77 | 42 | 65
2009-04-19 | 20:00 | 14.67 | 42 | 65
2009-04-19 | 20:15 | 14.55 | 42 | 65
2009-04-19 | 20:30 | 14.39 | 42 | 65
2009-04-19 | 20:45 | 14.22 | 42 | 65
2009-04-19 | 21:00 | 14.02 | 42 | 65
2009-04-19 | 21:15 | 13.82 | 42 | 65
2009-04-19 | 21:30 | 13.61 | 42 | 65
2009-04-19 | 21:45 | 13.4 | 42 | 65
2009-04-19 | 22:00 | 13.2 | 42 | 65
2009-04-19 | 22:15 | 13.01 | 42 | 65
2009-04-19 | 22:30 | 12.82 | 42 | 65
2009-04-19 | 22:45 | 12.64 | 42 | 65
2009-04-19 | 23:00 | 12.46 | 42 | 65
2009-04-19 | 23:15 | 12.3 | 42 | 65
2009-04-19 | 23:30 | 12.14 | 42 | 65
2009-04-19 | 23:45 | 12 | 42 | 65
2009-04-20 | 0:00 | 11.86 | 43 | 65
2009-04-20 | 0:15 | 11.72 | 42 | 65
2009-04-20 | 0:30 | 11.6 | 43 | 65
2009-04-20 | 0:45 | 11.48 | 43 | 65
2009-04-20 | 1:01 | 11.36 | 43 | 65
2009-04-20 | 1:16 | 11.23 | 43 | 65
2009-04-20 | 1:31 | 11.1 | 43 | 66
2009-04-20 | 1:46 | 10.97 | 43 | 66
2009-04-20 | 2:01 | 10.84 | 43 | 66
2009-04-20 | 2:16 | 10.71 | 43 | 66
2009-04-20 | 2:31 | 10.58 | 43 | 66
2009-04-20 | 2:46 | 10.44 | 43 | 66
2009-04-20 | 3:01 | 10.31 | 43 | 66
2009-04-20 | 3:16 | 10.17 | 43 | 66
2009-04-20 | 3:31 | 10.04 | 43 | 66
2009-04-20 | 3:46 | 9.9 | 43 | 66
2009-04-20 | 4:01 | 9.75 | 43 | 66
2009-04-20 | 4:16 | 9.62 | 43 | 66
2009-04-20 | 4:31 | 9.47 | 43 | 66
2009-04-20 | 4:46 | 9.32 | 43 | 66
2009-04-20 | 5:01 | 9.18 | 43 | 66
2009-04-20 | 5:16 | 9.03 | 43 | 66
2009-04-20 | 5:31 | 8.89 | 43 | 66
2009-04-20 | 5:46 | 8.75 | 43 | 66
2009-04-20 | 6:01 | 8.62 | 43 | 66
2009-04-20 | 6:16 | 8.49 | 43 | 66
2009-04-20 | 6:31 | 8.36 | 43 | 66
2009-04-20 | 6:46 | 8.25 | 43 | 66
2009-04-20 | 7:01 | 8.14 | 43 | 66
2009-04-20 | 7:16 | 8.04 | 43 | 66
2009-04-20 | 7:31 | 7.96 | 43 | 66
2009-04-20 | 7:46 | 7.88 | 43 | 66
2009-04-20 | 8:01 | 7.82 | 43 | 65
2009-04-20 | 8:16 | 7.78 | 42 | 65
2009-04-20 | 8:31 | 7.76 | 42 | 65
2009-04-20 | 8:46 | 7.76 | 42 | 65
2009-04-20 | 9:01 | 7.77 | 42 | 65
2009-04-20 | 9:16 | 7.77 | 42 | 64
2009-04-20 | 9:31 | 7.8 | 42 | 64
2009-04-20 | 9:46 | 7.85 | 41 | 64
2009-04-20 | 10:01 | 7.92 | 41 | 64
2009-04-20 | 10:16 | 8.01 | 41 | 63
2009-04-20 | 10:31 | 8.14 | 41 | 63
2009-04-20 | 10:46 | 8.33 | 41 | 63
2009-04-20 | 11:01 | 8.58 | 41 | 62
2009-04-20 | 11:16 | 8.95 | 40 | 62
2009-04-20 | 11:31 | 9.42 | 40 | 62
2009-04-20 | 11:46 | 9.95 | 40 | 62
2009-04-20 | 12:01 | 10.48 | 40 | 61
2009-04-20 | 12:16 | 11.03 | 40 | 61
2009-04-20 | 12:31 | 11.57 | 40 | 61
2009-04-20 | 12:46 | 12.09 | 40 | 61
2009-04-20 | 13:00 | 12.6 | 40 | 61
2009-04-20 | 13:15 | 13.12 | 39 | 61
2009-04-20 | 13:30 | 13.64 | 39 | 61
2009-04-20 | 13:45 | 14.17 | 39 | 60
2009-04-20 | 14:00 | 14.68 | 39 | 60
2009-04-20 | 14:15 | 15.15 | 39 | 60
2009-04-20 | 14:30 | 15.57 | 39 | 60
2009-04-20 | 14:45 | 15.92 | 39 | 60
2009-04-20 | 15:00 | 16.23 | 39 | 60
2009-04-20 | 15:15 | 16.5 | 39 | 60
2009-04-20 | 15:30 | 16.71 | 39 | 60
2009-04-20 | 15:45 | 16.86 | 39 | 60
2009-04-20 | 16:00 | 16.94 | 39 | 60
2009-04-20 | 16:15 | 16.99 | 39 | 60
2009-04-20 | 16:30 | 16.99 | 39 | 60
2009-04-20 | 16:45 | 16.98 | 39 | 60
2009-04-20 | 17:00 | 16.89 | 39 | 60
2009-04-20 | 17:15 | 16.79 | 39 | 60
2009-04-20 | 17:30 | 16.69 | 39 | 60
2009-04-20 | 17:45 | 16.58 | 39 | 60
2009-04-20 | 18:00 | 16.47 | 39 | 60
2009-04-20 | 18:15 | 16.38 | 39 | 60
2009-04-20 | 18:30 | 16.32 | 39 | 60
2009-04-20 | 18:45 | 16.27 | 39 | 60
2009-04-20 | 19:00 | 16.26 | 39 | 60
2009-04-20 | 19:15 | 16.26 | 39 | 60
2009-04-20 | 19:30 | 16.23 | 39 | 60
2009-04-20 | 19:45 | 16.17 | 39 | 60
2009-04-20 | 20:00 | 16.08 | 39 | 60
2009-04-20 | 20:15 | 15.96 | 39 | 60
2009-04-20 | 20:30 | 15.82 | 39 | 60
2009-04-20 | 20:45 | 15.65 | 39 | 61
2009-04-20 | 21:00 | 15.46 | 39 | 61
2009-04-20 | 21:15 | 15.24 | 39 | 61
2009-04-20 | 21:30 | 15.03 | 40 | 61
2009-04-20 | 21:45 | 14.82 | 40 | 61
2009-04-20 | 22:00 | 14.62 | 40 | 61
2009-04-20 | 22:15 | 14.43 | 40 | 61
2009-04-20 | 22:30 | 14.24 | 40 | 62
2009-04-20 | 22:45 | 14.05 | 40 | 62
2009-04-20 | 23:00 | 13.87 | 41 | 63
2009-04-20 | 23:15 | 13.69 | 41 | 63
2009-04-20 | 23:30 | 13.51 | 41 | 63
2009-04-20 | 23:45 | 13.35 | 41 | 64
2009-04-21 | 0:00 | 13.19 | 41 | 64
2009-04-21 | 0:15 | 13.04 | 42 | 64
2009-04-21 | 0:30 | 12.9 | 42 | 64
2009-04-21 | 0:45 | 12.76 | 42 | 64
2009-04-21 | 1:01 | 12.62 | 42 | 65
2009-04-21 | 1:16 | 12.48 | 42 | 65
2009-04-21 | 1:31 | 12.34 | 42 | 65
2009-04-21 | 1:46 | 12.19 | 42 | 65
2009-04-21 | 2:01 | 12.06 | 42 | 65
2009-04-21 | 2:16 | 11.93 | 42 | 65
2009-04-21 | 2:31 | 11.8 | 42 | 65
2009-04-21 | 2:46 | 11.67 | 43 | 66
2009-04-21 | 3:01 | 11.54 | 43 | 66
2009-04-21 | 3:16 | 11.4 | 43 | 66
2009-04-21 | 3:31 | 11.27 | 43 | 66
2009-04-21 | 3:46 | 11.14 | 43 | 66
2009-04-21 | 4:01 | 11 | 43 | 66
2009-04-21 | 4:16 | 10.86 | 43 | 66
2009-04-21 | 4:31 | 10.72 | 43 | 66
2009-04-21 | 4:46 | 10.57 | 43 | 66
2009-04-21 | 5:01 | 10.42 | 43 | 67
2009-04-21 | 5:16 | 10.27 | 43 | 67
2009-04-21 | 5:31 | 10.12 | 43 | 67
2009-04-21 | 5:46 | 9.97 | 43 | 67
2009-04-21 | 6:01 | 9.82 | 43 | 67
2009-04-21 | 6:16 | 9.67 | 43 | 67
2009-04-21 | 6:31 | 9.53 | 44 | 67
2009-04-21 | 6:46 | 9.4 | 43 | 67
2009-04-21 | 7:01 | 9.29 | 43 | 67
2009-04-21 | 7:16 | 9.18 | 43 | 67
2009-04-21 | 7:31 | 9.08 | 43 | 67
2009-04-21 | 7:46 | 9 | 43 | 67
2009-04-21 | 8:01 | 8.93 | 43 | 66
2009-04-21 | 8:16 | 8.88 | 43 | 66
2009-04-21 | 8:31 | 8.84 | 43 | 66
2009-04-21 | 8:46 | 8.83 | 42 | 65
2009-04-21 | 9:01 | 8.83 | 42 | 65
2009-04-21 | 9:16 | 8.83 | 42 | 64
2009-04-21 | 9:31 | 8.84 | 42 | 64
2009-04-21 | 9:46 | 8.87 | 41 | 63
2009-04-21 | 10:01 | 8.93 | 41 | 63
2009-04-21 | 10:16 | 9.02 | 41 | 62
2009-04-21 | 10:31 | 9.16 | 40 | 62
2009-04-21 | 10:46 | 9.36 | 40 | 62
2009-04-21 | 11:01 | 9.64 | 40 | 61
2009-04-21 | 11:16 | 10.03 | 40 | 61
2009-04-21 | 11:31 | 10.52 | 39 | 61
2009-04-21 | 11:46 | 11.05 | 39 | 60
2009-04-21 | 12:01 | 11.6 | 39 | 60
2009-04-21 | 12:16 | 12.14 | 39 | 60
2009-04-21 | 12:31 | 12.67 | 39 | 60
2009-04-21 | 12:46 | 13.19 | 39 | 60
2009-04-21 | 13:00 | 13.69 | 39 | 60
2009-04-21 | 13:15 | 14.21 | 39 | 60
2009-04-21 | 13:30 | 14.73 | 39 | 59
2009-04-21 | 13:45 | 15.24 | 39 | 59
2009-04-21 | 14:00 | 15.75 | 39 | 59
2009-04-21 | 14:15 | 16.2 | 38 | 59
2009-04-21 | 14:30 | 16.6 | 38 | 59
2009-04-21 | 14:45 | 16.94 | 38 | 59
2009-04-21 | 15:00 | 17.25 | 38 | 59
2009-04-21 | 15:15 | 17.53 | 38 | 59
2009-04-21 | 15:30 | 17.71 | 38 | 59
2009-04-21 | 15:45 | 17.84 | 38 | 59
2009-04-21 | 16:00 | 17.92 | 38 | 59
2009-04-21 | 16:15 | 17.94 | 38 | 59
2009-04-21 | 16:30 | 17.9 | 38 | 59
2009-04-21 | 16:45 | 17.8 | 38 | 59
2009-04-21 | 17:00 | 17.65 | 38 | 59
2009-04-21 | 17:15 | 17.47 | 38 | 59
2009-04-21 | 17:30 | 17.31 | 38 | 59
2009-04-21 | 17:45 | 17.15 | 38 | 59
2009-04-21 | 18:00 | 17 | 38 | 59
2009-04-21 | 18:15 | 16.87 | 38 | 59
2009-04-21 | 18:30 | 16.8 | 38 | 59
2009-04-21 | 18:45 | 16.72 | 38 | 59
2009-04-21 | 19:00 | 16.65 | 38 | 59
2009-04-21 | 19:15 | 16.63 | 38 | 59
2009-04-21 | 19:30 | 16.59 | 38 | 59
2009-04-21 | 19:45 | 16.5 | 39 | 59
2009-04-21 | 20:00 | 16.39 | 39 | 59
2009-04-21 | 20:15 | 16.25 | 39 | 59
2009-04-21 | 20:30 | 16.08 | 39 | 60
2009-04-21 | 20:45 | 15.88 | 39 | 60
2009-04-21 | 21:00 | 15.65 | 39 | 60
2009-04-21 | 21:15 | 15.41 | 39 | 60
2009-04-21 | 21:30 | 15.17 | 39 | 60
2009-04-21 | 21:45 | 14.93 | 39 | 60
2009-04-21 | 22:00 | 14.69 | 39 | 61
2009-04-21 | 22:15 | 14.46 | 40 | 61
2009-04-21 | 22:30 | 14.24 | 40 | 61
2009-04-21 | 22:45 | 14.02 | 40 | 62
2009-04-21 | 23:00 | 13.8 | 40 | 62
2009-04-21 | 23:15 | 13.59 | 40 | 62
2009-04-21 | 23:30 | 13.39 | 40 | 62
2009-04-21 | 23:45 | 13.19 | 40 | 62
2009-04-22 | 0:00 | 13 | 41 | 62
2009-04-22 | 0:15 | 12.8 | 41 | 63
2009-04-22 | 0:30 | 12.62 | 41 | 63
2009-04-22 | 0:45 | 12.45 | 41 | 63
2009-04-22 | 1:01 | 12.27 | 41 | 63
2009-04-22 | 1:16 | 12.1 | 41 | 63
2009-04-22 | 1:31 | 11.93 | 41 | 63
2009-04-22 | 1:46 | 11.76 | 41 | 64
2009-04-22 | 2:01 | 11.6 | 41 | 64
2009-04-22 | 2:16 | 11.43 | 42 | 64
2009-04-22 | 2:31 | 11.27 | 42 | 64
2009-04-22 | 2:46 | 11.1 | 42 | 64
2009-04-22 | 3:01 | 10.94 | 42 | 64
2009-04-22 | 3:16 | 10.78 | 42 | 65
2009-04-22 | 3:31 | 10.61 | 42 | 65
2009-04-22 | 3:46 | 10.44 | 42 | 65
2009-04-22 | 4:01 | 10.28 | 42 | 65
2009-04-22 | 4:16 | 10.13 | 42 | 65
2009-04-22 | 4:31 | 9.97 | 42 | 65
2009-04-22 | 4:46 | 9.82 | 43 | 65
2009-04-22 | 5:01 | 9.67 | 43 | 66
2009-04-22 | 5:16 | 9.51 | 43 | 66
2009-04-22 | 5:31 | 9.35 | 43 | 66
2009-04-22 | 5:46 | 9.17 | 43 | 66
2009-04-22 | 6:01 | 9 | 43 | 66
2009-04-22 | 6:16 | 8.83 | 43 | 67
2009-04-22 | 6:31 | 8.68 | 43 | 67
2009-04-22 | 6:46 | 8.54 | 43 | 67
2009-04-22 | 7:01 | 8.42 | 43 | 67
2009-04-22 | 7:16 | 8.31 | 44 | 67
2009-04-22 | 7:31 | 8.22 | 43 | 67
2009-04-22 | 7:46 | 8.14 | 43 | 67
2009-04-22 | 8:01 | 8.07 | 43 | 67
2009-04-22 | 8:16 | 8.02 | 43 | 66
2009-04-22 | 8:31 | 8 | 43 | 66
2009-04-22 | 8:46 | 8.01 | 43 | 66
2009-04-22 | 9:01 | 8.04 | 43 | 65
2009-04-22 | 9:16 | 8.07 | 42 | 65
2009-04-22 | 9:31 | 8.11 | 42 | 65
2009-04-22 | 9:46 | 8.16 | 42 | 64
2009-04-22 | 10:01 | 8.24 | 41 | 64
2009-04-22 | 10:16 | 8.35 | 41 | 63
2009-04-22 | 10:31 | 8.39 | 41 | 63
2009-04-22 | 10:46 | 8.54 | 40 | 62
2009-04-22 | 11:01 | 8.75 | 40 | 62
2009-04-22 | 11:16 | 9.01 | 40 | 61
2009-04-22 | 11:31 | 9.3 | 40 | 61
2009-04-22 | 11:46 | 9.59 | 39 | 61
2009-04-22 | 12:01 | 9.93 | 39 | 60
2009-04-22 | 12:16 | 10.4 | 39 | 60
2009-04-22 | 12:31 | 10.71 | 39 | 60
2009-04-22 | 12:46 | 10.97 | 39 | 60
2009-04-22 | 13:00 | 11.23 | 39 | 60
2009-04-22 | 13:15 | 11.54 | 39 | 59
2009-04-22 | 13:30 | 11.79 | 39 | 59
2009-04-22 | 13:45 | 11.93 | 39 | 59
2009-04-22 | 14:00 | 12.05 | 38 | 59
2009-04-22 | 14:15 | 12.15 | 38 | 59
2009-04-22 | 14:30 | 12.28 | 38 | 59
2009-04-22 | 14:45 | 12.41 | 38 | 59
2009-04-22 | 15:00 | 12.97 | 38 | 59
2009-04-22 | 15:15 | 13.59 | 38 | 59
2009-04-22 | 15:30 | 14.11 | 38 | 59
2009-04-22 | 15:45 | 14.51 | 38 | 59
2009-04-22 | 16:00 | 14.75 | 38 | 59
2009-04-22 | 16:15 | 14.87 | 38 | 59
2009-04-22 | 16:30 | 14.87 | 38 | 59
2009-04-22 | 16:45 | 14.77 | 38 | 59
2009-04-22 | 17:00 | 14.69 | 38 | 59
2009-04-22 | 17:15 | 14.59 | 38 | 59
2009-04-22 | 17:30 | 14.5 | 38 | 59
2009-04-22 | 17:45 | 14.43 | 39 | 59
2009-04-22 | 18:00 | 14.37 | 39 | 59
2009-04-22 | 18:15 | 14.26 | 39 | 59
2009-04-22 | 18:30 | 14.2 | 39 | 59
2009-04-22 | 18:45 | 14.14 | 39 | 59
2009-04-22 | 19:00 | 14.08 | 39 | 59
2009-04-22 | 19:15 | 14.03 | 39 | 59
2009-04-22 | 19:30 | 13.99 | 39 | 60
2009-04-22 | 19:45 | 13.91 | 39 | 60
2009-04-22 | 20:00 | 13.83 | 39 | 60
2009-04-22 | 20:15 | 13.75 | 39 | 60
2009-04-22 | 20:30 | 13.66 | 39 | 60
2009-04-22 | 20:45 | 13.54 | 39 | 60
2009-04-22 | 21:00 | 13.42 | 39 | 60
2009-04-22 | 21:15 | 13.27 | 39 | 60
2009-04-22 | 21:30 | 13.11 | 39 | 60
2009-04-22 | 21:45 | 12.94 | 39 | 61
2009-04-22 | 22:00 | 12.8 | 40 | 61
2009-04-22 | 22:15 | 12.64 | 40 | 61
2009-04-22 | 22:30 | 12.45 | 40 | 62
2009-04-22 | 22:45 | 12.29 | 40 | 62
2009-04-22 | 23:00 | 12.15 | 40 | 62
2009-04-22 | 23:15 | 12 | 40 | 62
2009-04-22 | 23:30 | 11.82 | 41 | 62
2009-04-22 | 23:45 | 11.65 | 41 | 63
2009-04-23 | 0:00 | 11.46 | 41 | 63
2009-04-23 | 0:15 | 11.28 | 41 | 63
2009-04-23 | 0:30 | 11.11 | 41 | 63
2009-04-23 | 0:45 | 10.94 | 41 | 63
2009-04-23 | 1:01 | 10.78 | 41 | 63
2009-04-23 | 1:16 | 10.63 | 41 | 63
2009-04-23 | 1:31 | 10.51 | 41 | 63
2009-04-23 | 1:46 | 10.33 | 41 | 64
2009-04-23 | 2:01 | 10.18 | 41 | 64
2009-04-23 | 2:16 | 10.02 | 41 | 64
2009-04-23 | 2:31 | 9.9 | 42 | 64
2009-04-23 | 2:46 | 9.75 | 42 | 64
2009-04-23 | 3:01 | 9.58 | 42 | 64
2009-04-23 | 3:16 | 9.43 | 42 | 64
2009-04-23 | 3:31 | 9.29 | 42 | 64
2009-04-23 | 3:46 | 9.14 | 42 | 65
2009-04-23 | 4:01 | 8.98 | 42 | 65
2009-04-23 | 4:16 | 8.81 | 42 | 65
2009-04-23 | 4:31 | 8.63 | 42 | 65
2009-04-23 | 4:46 | 8.44 | 42 | 65
2009-04-23 | 5:01 | 8.26 | 42 | 65
2009-04-23 | 5:16 | 8.05 | 42 | 65
2009-04-23 | 5:31 | 7.84 | 42 | 65
2009-04-23 | 5:46 | 7.64 | 42 | 65
2009-04-23 | 6:01 | 7.46 | 43 | 65
2009-04-23 | 6:16 | 7.29 | 43 | 65
2009-04-23 | 6:31 | 7.12 | 43 | 66
2009-04-23 | 6:46 | 6.97 | 43 | 66
2009-04-23 | 7:01 | 6.84 | 43 | 66
2009-04-23 | 7:16 | 6.73 | 43 | 66
2009-04-23 | 7:31 | 6.63 | 43 | 66
2009-04-23 | 7:46 | 6.52 | 43 | 66
2009-04-23 | 8:01 | 6.42 | 43 | 66
2009-04-23 | 8:16 | 6.34 | 43 | 66
2009-04-23 | 8:31 | 6.27 | 43 | 66
2009-04-23 | 8:46 | 6.24 | 43 | 65
2009-04-23 | 9:01 | 6.22 | 42 | 65
2009-04-23 | 9:16 | 6.2 | 42 | 65
2009-04-23 | 9:31 | 6.19 | 42 | 65
2009-04-23 | 9:46 | 6.2 | 42 | 65
2009-04-23 | 10:01 | 6.23 | 42 | 65
2009-04-23 | 10:16 | 6.3 | 42 | 64
2009-04-23 | 10:31 | 6.4 | 42 | 64
2009-04-23 | 10:46 | 6.54 | 42 | 64
2009-04-23 | 11:01 | 6.59 | 41 | 64
2009-04-23 | 11:16 | 6.8 | 41 | 64
2009-04-23 | 11:31 | 7.35 | 41 | 63
2009-04-23 | 11:46 | 7.87 | 41 | 63
2009-04-23 | 12:01 | 8.19 | 41 | 63
2009-04-23 | 12:16 | 8.38 | 41 | 63
2009-04-23 | 12:31 | 8.83 | 41 | 63
2009-04-23 | 12:46 | 9.41 | 41 | 63
2009-04-23 | 13:00 | 10.04 | 41 | 63
2009-04-23 | 13:15 | 10.26 | 41 | 62
2009-04-23 | 13:30 | 10.78 | 40 | 62
2009-04-23 | 13:45 | 11.26 | 40 | 62
2009-04-23 | 14:00 | 11.53 | 40 | 62
2009-04-23 | 14:15 | 11.97 | 40 | 62
2009-04-23 | 14:30 | 12.48 | 40 | 62
2009-04-23 | 14:45 | 12.75 | 40 | 62
2009-04-23 | 15:00 | 12.98 | 40 | 62
2009-04-23 | 15:15 | 13.31 | 40 | 62
2009-04-23 | 15:30 | 13.51 | 40 | 62
2009-04-23 | 15:45 | 13.66 | 40 | 62
2009-04-23 | 16:00 | 13.67 | 40 | 62
2009-04-23 | 16:15 | 13.69 | 40 | 62
2009-04-23 | 16:30 | 13.72 | 40 | 62
2009-04-23 | 16:45 | 13.66 | 40 | 62
2009-04-23 | 17:00 | 13.41 | 40 | 62
2009-04-23 | 17:15 | 13.15 | 40 | 62
2009-04-23 | 17:30 | 12.88 | 40 | 62
2009-04-23 | 17:45 | 12.61 | 40 | 62
2009-04-23 | 18:00 | 12.41 | 40 | 62
2009-04-23 | 18:15 | 12.19 | 40 | 62
2009-04-23 | 18:30 | 12.07 | 40 | 62
2009-04-23 | 18:45 | 11.98 | 40 | 62
2009-04-23 | 19:00 | 11.87 | 40 | 62
2009-04-23 | 19:15 | 11.8 | 40 | 62
2009-04-23 | 19:30 | 11.68 | 40 | 62
2009-04-23 | 19:45 | 11.56 | 40 | 62
2009-04-23 | 20:00 | 11.43 | 40 | 62
2009-04-23 | 20:15 | 11.25 | 40 | 62
2009-04-23 | 20:30 | 11.07 | 40 | 62
2009-04-23 | 20:45 | 10.86 | 40 | 62
2009-04-23 | 21:00 | 10.63 | 40 | 62
2009-04-23 | 21:15 | 10.41 | 40 | 62
2009-04-23 | 21:30 | 10.18 | 40 | 62
2009-04-23 | 21:45 | 9.94 | 40 | 62
2009-04-23 | 22:00 | 9.72 | 40 | 62
2009-04-23 | 22:15 | 9.5 | 40 | 62
2009-04-23 | 22:30 | 9.27 | 41 | 62
2009-04-23 | 22:45 | 9.04 | 41 | 63
2009-04-23 | 23:00 | 8.83 | 41 | 63
2009-04-23 | 23:15 | 8.63 | 41 | 63
2009-04-23 | 23:30 | 8.44 | 41 | 63
2009-04-23 | 23:45 | 8.26 | 41 | 63
2009-04-24 | 0:00 | 8.07 | 41 | 63
2009-04-24 | 0:15 | 7.89 | 41 | 63
2009-04-24 | 0:30 | 7.72 | 41 | 63
2009-04-24 | 0:45 | 7.55 | 41 | 63
2009-04-24 | 1:01 | 7.39 | 41 | 63
2009-04-24 | 1:16 | 7.21 | 41 | 63
2009-04-24 | 1:31 | 7.05 | 41 | 64
2009-04-24 | 1:46 | 6.9 | 41 | 64
2009-04-24 | 2:01 | 6.74 | 41 | 64
2009-04-24 | 2:16 | 6.59 | 42 | 64
2009-04-24 | 2:31 | 6.44 | 42 | 64
2009-04-24 | 2:46 | 6.3 | 42 | 64
2009-04-24 | 3:01 | 6.15 | 42 | 64
2009-04-24 | 3:16 | 6 | 42 | 64
2009-04-24 | 3:31 | 5.85 | 42 | 65
2009-04-24 | 3:46 | 5.7 | 42 | 65
2009-04-24 | 4:01 | 5.56 | 42 | 65
2009-04-24 | 4:16 | 5.42 | 42 | 65
2009-04-24 | 4:31 | 5.28 | 42 | 65
2009-04-24 | 4:46 | 5.14 | 42 | 65
2009-04-24 | 5:01 | 5 | 42 | 65
2009-04-24 | 5:16 | 4.87 | 42 | 65
2009-04-24 | 5:31 | 4.74 | 42 | 65
2009-04-24 | 5:46 | 4.6 | 42 | 65
2009-04-24 | 6:01 | 4.48 | 42 | 65
2009-04-24 | 6:16 | 4.34 | 43 | 66
2009-04-24 | 6:31 | 4.21 | 43 | 66
2009-04-24 | 6:46 | 4.1 | 43 | 66
2009-04-24 | 7:01 | 3.99 | 43 | 66
2009-04-24 | 7:16 | 3.9 | 43 | 66
2009-04-24 | 7:31 | 3.83 | 43 | 66
2009-04-24 | 7:46 | 3.77 | 43 | 66
2009-04-24 | 8:01 | 3.73 | 43 | 66
2009-04-24 | 8:16 | 3.71 | 43 | 66
2009-04-24 | 8:31 | 3.72 | 43 | 66
2009-04-24 | 8:46 | 3.76 | 43 | 66
2009-04-24 | 9:01 | 3.84 | 43 | 66
2009-04-24 | 9:16 | 3.92 | 43 | 66
2009-04-24 | 9:31 | 4.03 | 43 | 66
2009-04-24 | 9:46 | 4.15 | 43 | 66
2009-04-24 | 10:01 | 4.29 | 43 | 66
2009-04-24 | 10:16 | 4.45 | 43 | 66
2009-04-24 | 10:31 | 4.6 | 43 | 66
2009-04-24 | 10:46 | 4.8 | 43 | 66
2009-04-24 | 11:01 | 5.04 | 43 | 66
2009-04-24 | 11:16 | 5.32 | 43 | 66
2009-04-24 | 11:31 | 5.71 | 43 | 66
2009-04-24 | 11:46 | 6.17 | 43 | 66
2009-04-24 | 12:01 | 6.7 | 43 | 66
2009-04-24 | 12:16 | 7.26 | 43 | 66
2009-04-24 | 12:31 | 7.84 | 43 | 66
2009-04-24 | 12:46 | 8.43 | 43 | 66
2009-04-24 | 13:00 | 8.99 | 43 | 66
2009-04-24 | 13:15 | 9.55 | 43 | 66
2009-04-24 | 13:30 | 10.09 | 43 | 66
2009-04-24 | 13:45 | 10.64 | 43 | 66
2009-04-24 | 14:00 | 11.2 | 43 | 66
2009-04-24 | 14:15 | 11.71 | 43 | 66
2009-04-24 | 14:30 | 12.13 | 43 | 66
2009-04-24 | 14:45 | 12.5 | 43 | 66
2009-04-24 | 15:00 | 12.82 | 43 | 66
2009-04-24 | 15:15 | 13.07 | 43 | 66
2009-04-24 | 15:30 | 13.28 | 43 | 66
2009-04-24 | 15:45 | 13.39 | 43 | 66
2009-04-24 | 16:00 | 13.41 | 43 | 66
2009-04-24 | 16:15 | 13.36 | 43 | 66
2009-04-24 | 16:30 | 13.29 | 43 | 66
2009-04-24 | 16:45 | 13.17 | 43 | 66
2009-04-24 | 17:00 | 13 | 43 | 66
2009-04-24 | 17:15 | 12.82 | 43 | 66
2009-04-24 | 17:30 | 12.64 | 43 | 66
2009-04-24 | 17:45 | 12.48 | 43 | 66
2009-04-24 | 18:00 | 12.31 | 43 | 66
2009-04-24 | 18:15 | 12.16 | 43 | 66
2009-04-24 | 18:30 | 12.03 | 43 | 66
2009-04-24 | 18:45 | 11.93 | 43 | 66
2009-04-24 | 19:00 | 11.89 | 43 | 66
2009-04-24 | 19:15 | 11.85 | 43 | 66
2009-04-24 | 19:30 | 11.8 | 43 | 66
2009-04-24 | 19:45 | 11.75 | 43 | 66
2009-04-24 | 20:00 | 11.67 | 43 | 66
2009-04-24 | 20:15 | 11.54 | 43 | 66
2009-04-24 | 20:30 | 11.39 | 43 | 66
2009-04-24 | 20:45 | 11.22 | 43 | 66
2009-04-24 | 21:00 | 11.02 | 43 | 66
2009-04-24 | 21:15 | 10.8 | 43 | 66
2009-04-24 | 21:30 | 10.57 | 43 | 66
2009-04-24 | 21:45 | 10.34 | 43 | 66
2009-04-24 | 22:00 | 10.11 | 43 | 66
2009-04-24 | 22:15 | 9.88 | 43 | 66
2009-04-24 | 22:30 | 9.67 | 43 | 66
2009-04-24 | 22:45 | 9.47 | 43 | 66
2009-04-24 | 23:00 | 9.28 | 43 | 66
2009-04-24 | 23:15 | 9.1 | 43 | 67
2009-04-24 | 23:30 | 8.92 | 43 | 67
2009-04-24 | 23:45 | 8.76 | 43 | 67
2009-04-25 | 0:00 | 8.59 | 43 | 67
2009-04-25 | 0:15 | 8.43 | 43 | 67
2009-04-25 | 0:30 | 8.27 | 43 | 67
2009-04-25 | 0:45 | 8.12 | 44 | 67
2009-04-25 | 1:01 | 7.98 | 44 | 67
2009-04-25 | 1:16 | 7.84 | 44 | 67
2009-04-25 | 1:31 | 7.71 | 44 | 67
2009-04-25 | 1:46 | 7.58 | 44 | 67
2009-04-25 | 2:01 | 7.47 | 44 | 67
2009-04-25 | 2:16 | 7.37 | 44 | 67
2009-04-25 | 2:31 | 7.28 | 44 | 67
2009-04-25 | 2:46 | 7.19 | 44 | 68
2009-04-25 | 3:01 | 7.09 | 44 | 68
2009-04-25 | 3:16 | 7 | 44 | 68
2009-04-25 | 3:31 | 6.91 | 44 | 68
2009-04-25 | 3:46 | 6.8 | 44 | 68
2009-04-25 | 4:01 | 6.69 | 44 | 68
2009-04-25 | 4:16 | 6.57 | 44 | 68
2009-04-25 | 4:31 | 6.45 | 44 | 68
2009-04-25 | 4:46 | 6.32 | 44 | 68
2009-04-25 | 5:01 | 6.19 | 44 | 68
2009-04-25 | 5:16 | 6.06 | 44 | 68
2009-04-25 | 5:31 | 5.95 | 44 | 68
2009-04-25 | 5:46 | 5.85 | 44 | 68
2009-04-25 | 6:01 | 5.75 | 44 | 68
2009-04-25 | 6:16 | 5.67 | 44 | 68
2009-04-25 | 6:31 | 5.59 | 45 | 68
2009-04-25 | 6:46 | 5.51 | 45 | 69
2009-04-25 | 7:01 | 5.47 | 45 | 69
2009-04-25 | 7:16 | 5.45 | 45 | 69
2009-04-25 | 7:31 | 5.44 | 45 | 69
2009-04-25 | 7:46 | 5.43 | 45 | 69
2009-04-25 | 8:01 | 5.41 | 45 | 69
2009-04-25 | 8:16 | 5.44 | 45 | 69
2009-04-25 | 8:31 | 5.53 | 45 | 69
2009-04-25 | 8:46 | 5.57 | 45 | 69
2009-04-25 | 9:01 | 5.6 | 45 | 69
2009-04-25 | 9:16 | 5.69 | 45 | 69
2009-04-25 | 9:31 | 5.84 | 45 | 69
2009-04-25 | 9:46 | 5.99 | 45 | 69
2009-04-25 | 10:01 | 6.16 | 45 | 69
2009-04-25 | 10:16 | 6.31 | 45 | 69
2009-04-25 | 10:31 | 6.41 | 45 | 69
2009-04-25 | 10:46 | 6.54 | 45 | 69
2009-04-25 | 11:01 | 6.72 | 45 | 69
2009-04-25 | 11:16 | 6.92 | 45 | 69
2009-04-25 | 11:31 | 7.25 | 45 | 69
2009-04-25 | 11:46 | 7.67 | 45 | 69
2009-04-25 | 12:01 | 8.04 | 45 | 69