## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AB Audited against Sonde AA and Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-05 |  | 2009-06-05 |  | A
Time | Computer set to atomic time |  | 13:05 |  | 13:05 |  | A
Temp | VWR NIST traceable surface water |  | 9.3 |  | 9.3 |  | A
DO | In Situ |  | 11.16 |  | 11.5 |  | B
SpC | In SItu |  | 29.9 |  | 29 |  | A
Ph | In Situ |  | 7.44 |  | 7.4 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-06-05 13:16 | 9.4 | 7.48 | 11.21 | 98 | 29.9 | 19.5
Sonde | 2009-06-05 13:15 | 9.43 | 7.53 | 11.95 | 104.5 | 29 | 19
Grade |  | A | A | B |  | A
Audit 2 | 2009-06-15 13:00 | 11.1 | 7.74 | 10.93 | 99.3 | 30.7 | 20.1
Sonde | 2009-06-15 13:00 | 11.17 | 7.67 | 12.46 | 113.4 | 27 | 17
Grade |  | A | A | Fail |  | B
Audit 3
Sonde
Grade |  | A | A | A |  | #DIV/0!
Post-sampling Field Equipment Check
Field equipment ID | Sonde AB Audited against Sonde AA and Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-15 |  | 2009-06-15 |  | A
Time | Computer set to atomic time |  | 14:46 |  | 14:46 |  | A
Temp | VWR NIST traceable surface water |  | 11.3 |  | 11.43 |  | A
DO | In Situ |  | 10.05 |  | 11.83 |  | Fail
SpC | In Situ |  | 23.5 |  | 22 |  | A
Ph | In Situ |  | 7.41 |  | 7.54 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
DO Data Trimmed due to Multiple Failing Audits |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-05-31 |  | 2009-05-31 |  | A
Time | Computer set to atomic time |  | 15:15 |  | 15:15 |  | A
Temp | VWR NIST traceable surface water |  | 19.4 |  | 19.3 |  | A
DO | Winkler titration of surface water |  | 7.9 |  | 8.03 |  | A
SpC | 148 µS/mL |  | 153.6 |  | 177.1 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-06-05 13:16 | 9.4 | 7.48 | 11.21 | 98 | 29.9 | 19.5
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-05 |  | 2009-06-05 |  | A
Time | Computer set to atomic time |  | 14:15 |  | 14:15 |  | A
Temp | VWR NIST traceable surface water |  | 10.8 |  | 10.4 |  | A
DO | Winkler titration of surface water |  | 10.6 |  | 10.62 |  | A
SpC | 148 µS/mL |  | 153.6 |  | 178.9 |  | Fail
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
SpC on Pro-Series can be trusted; funky audit solution is likely cause of failing grades |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 006.00
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, Rd 6360
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-09 |  | 2009-06-09 |  | A
Time | Computer set to atomic time |  | 13:39 |  | 13:39 |  | A
Temp | VWR NIST traceable surface water |  | 11.4 |  | 11.2 |  | A
DO | Winkler titration of surface water |  | 10.3 |  | 10.76 |  | B
SpC | 148 µS/mL |  | 156.7 |  | 174.7 |  | B
Ph | pH 7 standard solution |  | 7.02 |  | 7.03 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-06-15 13:00 | 11.1 | 7.74 | 10.93 | 99.3 | 30.7 | 20.1
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-06-15 |  | 2009-06-15 |  | A
Time | Computer set to atomic time |  | 14:48 |  | 14:48 |  | A
Temp | VWR NIST traceable surface water |  | 11.3 |  | 11.5 |  | A
DO | Winkler titration of surface water |  | 10.2 |  | 10.14 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 154.3 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.02 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | TDS | SpCond | DO Conc | DO% | pH
M/D/Y |  | °C | mg/L | uS/cm | mg/L | %
2009-06-05 | 13:15 | 9.43 | 19 | 29 |  |  | 7.53
2009-06-05 | 13:30 | 9.7 | 19 | 29 |  |  | 7.54
2009-06-05 | 13:45 | 10.2 | 19 | 29 |  |  | 7.57
2009-06-05 | 14:00 | 10 | 19 | 29 |  |  | 7.56
2009-06-05 | 14:15 | 9.67 | 19 | 29 |  |  | 7.56
2009-06-05 | 14:30 | 9.62 | 19 | 29 |  |  | 7.56
2009-06-05 | 14:45 | 9.73 | 19 | 29 |  |  | 7.57
2009-06-05 | 15:00 | 10.07 | 19 | 29 |  |  | 7.57
2009-06-05 | 15:15 | 10.58 | 19 | 29 |  |  | 7.58
2009-06-05 | 15:30 | 10.96 | 19 | 29 |  |  | 7.58
2009-06-05 | 15:45 | 11.49 | 19 | 29 |  |  | 7.59
2009-06-05 | 16:00 | 11.97 | 19 | 29 |  |  | 7.59
2009-06-05 | 16:15 | 12.3 | 19 | 29 |  |  | 7.6
2009-06-05 | 16:30 | 12.48 | 19 | 29 |  |  | 7.6
2009-06-05 | 16:45 | 12.6 | 19 | 29 |  |  | 7.6
2009-06-05 | 17:00 | 12.67 | 19 | 29 |  |  | 7.6
2009-06-05 | 17:15 | 12.7 | 19 | 29 |  |  | 7.6
2009-06-05 | 17:30 | 12.82 | 19 | 30 |  |  | 7.61
2009-06-05 | 17:45 | 12.99 | 20 | 30 |  |  | 7.62
2009-06-05 | 18:00 | 13.19 | 20 | 30 |  |  | 7.62
2009-06-05 | 18:15 | 13.28 | 20 | 30 |  |  | 7.62
2009-06-05 | 18:30 | 13.33 | 20 | 30 |  |  | 7.63
2009-06-05 | 18:45 | 13.33 | 20 | 30 |  |  | 7.63
2009-06-05 | 19:00 | 13.31 | 20 | 30 |  |  | 7.63
2009-06-05 | 19:15 | 13.29 | 20 | 30 |  |  | 7.63
2009-06-05 | 19:30 | 13.26 | 20 | 30 |  |  | 7.63
2009-06-05 | 19:45 | 13.18 | 20 | 30 |  |  | 7.63
2009-06-05 | 20:00 | 13.1 | 20 | 31 |  |  | 7.62
2009-06-05 | 20:15 | 13 | 20 | 31 |  |  | 7.62
2009-06-05 | 20:30 | 12.92 | 20 | 31 |  |  | 7.61
2009-06-05 | 20:45 | 12.86 | 20 | 31 |  |  | 7.62
2009-06-05 | 21:00 | 12.81 | 20 | 31 |  |  | 7.62
2009-06-05 | 21:15 | 12.76 | 20 | 31 |  |  | 7.61
2009-06-05 | 21:30 | 12.7 | 20 | 31 |  |  | 7.61
2009-06-05 | 21:45 | 12.62 | 20 | 31 |  |  | 7.61
2009-06-05 | 22:00 | 12.54 | 20 | 31 |  |  | 7.61
2009-06-05 | 22:15 | 12.45 | 20 | 31 |  |  | 7.6
2009-06-05 | 22:30 | 12.32 | 20 | 30 |  |  | 7.6
2009-06-05 | 22:45 | 12.17 | 20 | 30 |  |  | 7.59
2009-06-05 | 23:00 | 12.04 | 20 | 30 |  |  | 7.59
2009-06-05 | 23:15 | 11.95 | 20 | 30 |  |  | 7.59
2009-06-05 | 23:30 | 11.86 | 19 | 30 |  |  | 7.58
2009-06-05 | 23:45 | 11.77 | 19 | 30 |  |  | 7.58
2009-06-06 | 0:00 | 11.66 | 19 | 30 |  |  | 7.58
2009-06-06 | 0:15 | 11.55 | 19 | 29 |  |  | 7.57
2009-06-06 | 0:30 | 11.43 | 19 | 29 |  |  | 7.57
2009-06-06 | 0:45 | 11.31 | 19 | 29 |  |  | 7.57
2009-06-06 | 1:01 | 11.2 | 19 | 29 |  |  | 7.56
2009-06-06 | 1:16 | 11.08 | 19 | 29 |  |  | 7.55
2009-06-06 | 1:31 | 10.96 | 19 | 29 |  |  | 7.56
2009-06-06 | 1:46 | 10.83 | 19 | 29 |  |  | 7.55
2009-06-06 | 2:01 | 10.7 | 18 | 28 |  |  | 7.55
2009-06-06 | 2:16 | 10.56 | 18 | 28 |  |  | 7.54
2009-06-06 | 2:31 | 10.42 | 18 | 28 |  |  | 7.54
2009-06-06 | 2:46 | 10.28 | 18 | 28 |  |  | 7.54
2009-06-06 | 3:01 | 10.14 | 18 | 28 |  |  | 7.52
2009-06-06 | 3:16 | 9.99 | 18 | 28 |  |  | 7.53
2009-06-06 | 3:31 | 9.83 | 18 | 28 |  |  | 7.53
2009-06-06 | 3:46 | 9.67 | 18 | 28 |  |  | 7.52
2009-06-06 | 4:01 | 9.52 | 18 | 28 |  |  | 7.51
2009-06-06 | 4:16 | 9.36 | 18 | 28 |  |  | 7.51
2009-06-06 | 4:31 | 9.22 | 18 | 28 |  |  | 7.51
2009-06-06 | 4:46 | 9.07 | 18 | 28 |  |  | 7.52
2009-06-06 | 5:01 | 8.93 | 18 | 27 |  |  | 7.53
2009-06-06 | 5:16 | 8.79 | 18 | 27 |  |  | 7.5
2009-06-06 | 5:31 | 8.66 | 18 | 27 |  |  | 7.49
2009-06-06 | 5:46 | 8.53 | 18 | 27 |  |  | 7.52
2009-06-06 | 6:01 | 8.41 | 18 | 27 |  |  | 7.52
2009-06-06 | 6:16 | 8.3 | 18 | 27 |  |  | 7.51
2009-06-06 | 6:31 | 8.2 | 18 | 27 |  |  | 7.51
2009-06-06 | 6:46 | 8.1 | 18 | 27 |  |  | 7.51
2009-06-06 | 7:01 | 8.02 | 18 | 27 |  |  | 7.51
2009-06-06 | 7:16 | 7.94 | 18 | 27 |  |  | 7.5
2009-06-06 | 7:31 | 7.87 | 18 | 27 |  |  | 7.52
2009-06-06 | 7:46 | 7.81 | 18 | 28 |  |  | 7.54
2009-06-06 | 8:01 | 7.76 | 18 | 28 |  |  | 7.53
2009-06-06 | 8:16 | 7.73 | 18 | 28 |  |  | 7.54
2009-06-06 | 8:31 | 7.71 | 18 | 28 |  |  | 7.53
2009-06-06 | 8:46 | 7.7 | 18 | 28 |  |  | 7.53
2009-06-06 | 9:01 | 7.72 | 18 | 28 |  |  | 7.55
2009-06-06 | 9:16 | 7.75 | 18 | 28 |  |  | 7.53
2009-06-06 | 9:31 | 7.81 | 18 | 28 |  |  | 7.55
2009-06-06 | 9:46 | 7.9 | 18 | 28 |  |  | 7.54
2009-06-06 | 10:01 | 8.01 | 18 | 28 |  |  | 7.56
2009-06-06 | 10:16 | 8.14 | 18 | 28 |  |  | 7.55
2009-06-06 | 10:31 | 8.29 | 18 | 28 |  |  | 7.56
2009-06-06 | 10:46 | 8.47 | 18 | 28 |  |  | 7.57
2009-06-06 | 11:01 | 8.66 | 18 | 28 |  |  | 7.57
2009-06-06 | 11:16 | 8.87 | 18 | 28 |  |  | 7.58
2009-06-06 | 11:31 | 9.09 | 18 | 28 |  |  | 7.59
2009-06-06 | 11:46 | 9.33 | 18 | 28 |  |  | 7.59
2009-06-06 | 12:01 | 9.58 | 18 | 28 |  |  | 7.59
2009-06-06 | 12:16 | 9.83 | 18 | 28 |  |  | 7.6
2009-06-06 | 12:31 | 10.1 | 18 | 28 |  |  | 7.6
2009-06-06 | 12:46 | 10.37 | 18 | 28 |  |  | 7.61
2009-06-06 | 13:00 | 10.64 | 18 | 28 |  |  | 7.62
2009-06-06 | 13:15 | 10.9 | 18 | 28 |  |  | 7.61
2009-06-06 | 13:30 | 11.16 | 18 | 28 |  |  | 7.62
2009-06-06 | 13:45 | 11.41 | 18 | 28 |  |  | 7.63
2009-06-06 | 14:00 | 11.64 | 18 | 28 |  |  | 7.62
2009-06-06 | 14:15 | 11.84 | 18 | 28 |  |  | 7.63
2009-06-06 | 14:30 | 12.03 | 18 | 28 |  |  | 7.63
2009-06-06 | 14:45 | 12.22 | 18 | 28 |  |  | 7.63
2009-06-06 | 15:00 | 12.39 | 18 | 28 |  |  | 7.63
2009-06-06 | 15:15 | 12.53 | 19 | 29 |  |  | 7.64
2009-06-06 | 15:30 | 12.66 | 19 | 29 |  |  | 7.64
2009-06-06 | 15:45 | 12.76 | 18 | 28 |  |  | 7.64
2009-06-06 | 16:00 | 12.84 | 18 | 28 |  |  | 7.65
2009-06-06 | 16:15 | 12.91 | 18 | 28 |  |  | 7.64
2009-06-06 | 16:30 | 12.96 | 18 | 28 |  |  | 7.64
2009-06-06 | 16:45 | 13 | 18 | 28 |  |  | 7.64
2009-06-06 | 17:00 | 13.02 | 18 | 28 |  |  | 7.64
2009-06-06 | 17:15 | 13.02 | 18 | 28 |  |  | 7.65
2009-06-06 | 17:30 | 13.04 | 19 | 29 |  |  | 7.64
2009-06-06 | 17:45 | 13.05 | 18 | 28 |  |  | 7.65
2009-06-06 | 18:00 | 13.05 | 18 | 28 |  |  | 7.64
2009-06-06 | 18:15 | 13.04 | 18 | 28 |  |  | 7.65
2009-06-06 | 18:30 | 13.02 | 18 | 28 |  |  | 7.64
2009-06-06 | 18:45 | 12.99 | 18 | 28 |  |  | 7.65
2009-06-06 | 19:00 | 12.94 | 18 | 28 |  |  | 7.64
2009-06-06 | 19:15 | 12.9 | 18 | 28 |  |  | 7.64
2009-06-06 | 19:30 | 12.85 | 18 | 28 |  |  | 7.64
2009-06-06 | 19:45 | 12.79 | 18 | 28 |  |  | 7.64
2009-06-06 | 20:00 | 12.73 | 18 | 28 |  |  | 7.63
2009-06-06 | 20:15 | 12.66 | 18 | 28 |  |  | 7.64
2009-06-06 | 20:30 | 12.59 | 18 | 28 |  |  | 7.63
2009-06-06 | 20:45 | 12.52 | 18 | 28 |  |  | 7.64
2009-06-06 | 21:00 | 12.43 | 18 | 28 |  |  | 7.63
2009-06-06 | 21:15 | 12.36 | 18 | 28 |  |  | 7.62
2009-06-06 | 21:30 | 12.28 | 18 | 28 |  |  | 7.63
2009-06-06 | 21:45 | 12.18 | 18 | 28 |  |  | 7.62
2009-06-06 | 22:00 | 12.06 | 18 | 28 |  |  | 7.62
2009-06-06 | 22:15 | 11.94 | 18 | 28 |  |  | 7.61
2009-06-06 | 22:30 | 11.82 | 18 | 28 |  |  | 7.61
2009-06-06 | 22:45 | 11.69 | 18 | 28 |  |  | 7.63
2009-06-06 | 23:00 | 11.56 | 18 | 27 |  |  | 7.6
2009-06-06 | 23:15 | 11.43 | 18 | 27 |  |  | 7.61
2009-06-06 | 23:30 | 11.31 | 18 | 27 |  |  | 7.59
2009-06-06 | 23:45 | 11.18 | 18 | 27 |  |  | 7.59
2009-06-07 | 0:00 | 11.07 | 17 | 27 |  |  | 7.59
2009-06-07 | 0:15 | 10.96 | 17 | 27 |  |  | 7.58
2009-06-07 | 0:30 | 10.85 | 17 | 26 |  |  | 7.58
2009-06-07 | 0:45 | 10.73 | 17 | 26 |  |  | 7.55
2009-06-07 | 1:01 | 10.62 | 17 | 26 |  |  | 7.56
2009-06-07 | 1:16 | 10.5 | 17 | 26 |  |  | 7.57
2009-06-07 | 1:31 | 10.37 | 17 | 26 |  |  | 7.55
2009-06-07 | 1:46 | 10.24 | 17 | 26 |  |  | 7.54
2009-06-07 | 2:01 | 10.11 | 17 | 26 |  |  | 7.54
2009-06-07 | 2:16 | 9.97 | 17 | 25 |  |  | 7.54
2009-06-07 | 2:31 | 9.82 | 17 | 25 |  |  | 7.54
2009-06-07 | 2:46 | 9.66 | 16 | 25 |  |  | 7.53
2009-06-07 | 3:01 | 9.49 | 16 | 25 |  |  | 7.53
2009-06-07 | 3:16 | 9.31 | 16 | 25 |  |  | 7.52
2009-06-07 | 3:31 | 9.14 | 16 | 25 |  |  | 7.53
2009-06-07 | 3:46 | 8.97 | 16 | 25 |  |  | 7.52
2009-06-07 | 4:01 | 8.81 | 16 | 25 |  |  | 7.52
2009-06-07 | 4:16 | 8.65 | 16 | 25 |  |  | 7.52
2009-06-07 | 4:31 | 8.5 | 16 | 24 |  |  | 7.52
2009-06-07 | 4:46 | 8.36 | 16 | 25 |  |  | 7.51
2009-06-07 | 5:01 | 8.24 | 16 | 25 |  |  | 7.51
2009-06-07 | 5:16 | 8.12 | 16 | 25 |  |  | 7.5
2009-06-07 | 5:31 | 8.01 | 16 | 25 |  |  | 7.51
2009-06-07 | 5:46 | 7.89 | 16 | 25 |  |  | 7.51
2009-06-07 | 6:01 | 7.79 | 16 | 25 |  |  | 7.5
2009-06-07 | 6:16 | 7.69 | 16 | 25 |  |  | 7.49
2009-06-07 | 6:31 | 7.58 | 16 | 25 |  |  | 7.49
2009-06-07 | 6:46 | 7.49 | 16 | 25 |  |  | 7.5
2009-06-07 | 7:01 | 7.4 | 16 | 25 |  |  | 7.51
2009-06-07 | 7:16 | 7.32 | 16 | 25 |  |  | 7.51
2009-06-07 | 7:31 | 7.25 | 16 | 25 |  |  | 7.51
2009-06-07 | 7:46 | 7.19 | 16 | 25 |  |  | 7.52
2009-06-07 | 8:01 | 7.14 | 16 | 25 |  |  | 7.53
2009-06-07 | 8:16 | 7.11 | 16 | 25 |  |  | 7.53
2009-06-07 | 8:31 | 7.1 | 16 | 25 |  |  | 7.54
2009-06-07 | 8:46 | 7.09 | 16 | 25 |  |  | 7.54
2009-06-07 | 9:01 | 7.07 | 16 | 25 |  |  | 7.56
2009-06-07 | 9:16 | 7.08 | 16 | 25 |  |  | 7.55
2009-06-07 | 9:31 | 7.09 | 16 | 25 |  |  | 7.55
2009-06-07 | 9:46 | 7.2 | 16 | 25 |  |  | 7.55
2009-06-07 | 10:01 | 7.26 | 16 | 25 |  |  | 7.56
2009-06-07 | 10:16 | 7.36 | 17 | 25 |  |  | 7.57
2009-06-07 | 10:31 | 7.51 | 16 | 25 |  |  | 7.58
2009-06-07 | 10:46 | 7.66 | 16 | 25 |  |  | 7.56
2009-06-07 | 11:01 | 7.71 | 17 | 25 |  |  | 7.56
2009-06-07 | 11:16 | 7.62 | 17 | 26 |  |  | 7.57
2009-06-07 | 11:31 | 7.89 | 17 | 26 |  |  | 7.57
2009-06-07 | 11:46 | 8.19 | 17 | 26 |  |  | 7.58
2009-06-07 | 12:01 | 8.22 | 17 | 26 |  |  | 7.58
2009-06-07 | 12:16 | 8.45 | 17 | 26 |  |  | 7.58
2009-06-07 | 12:31 | 8.74 | 17 | 26 |  |  | 7.59
2009-06-07 | 12:46 | 9.09 | 17 | 26 |  |  | 7.6
2009-06-07 | 13:00 | 9.24 | 17 | 26 |  |  | 7.6
2009-06-07 | 13:15 | 9.45 | 17 | 26 |  |  | 7.6
2009-06-07 | 13:30 | 9.87 | 17 | 26 |  |  | 7.61
2009-06-07 | 13:45 | 10.2 | 17 | 26 |  |  | 7.62
2009-06-07 | 14:00 | 10.39 | 17 | 26 |  |  | 7.62
2009-06-07 | 14:15 | 10.49 | 17 | 26 |  |  | 7.62
2009-06-07 | 14:30 | 10.79 | 17 | 26 |  |  | 7.62
2009-06-07 | 14:45 | 11.09 | 17 | 26 |  |  | 7.63
2009-06-07 | 15:00 | 11.49 | 17 | 27 |  |  | 7.63
2009-06-07 | 15:15 | 11.63 | 17 | 27 |  |  | 7.64
2009-06-07 | 15:30 | 11.66 | 17 | 27 |  |  | 7.64
2009-06-07 | 15:45 | 11.83 | 17 | 27 |  |  | 7.65
2009-06-07 | 16:00 | 12 | 17 | 27 |  |  | 7.64
2009-06-07 | 16:15 | 12.13 | 17 | 27 |  |  | 7.64
2009-06-07 | 16:30 | 12.2 | 17 | 27 |  |  | 7.65
2009-06-07 | 16:45 | 12.23 | 17 | 27 |  |  | 7.65
2009-06-07 | 17:00 | 12.23 | 17 | 27 |  |  | 7.66
2009-06-07 | 17:15 | 12.27 | 17 | 27 |  |  | 7.65
2009-06-07 | 17:30 | 12.4 | 18 | 27 |  |  | 7.65
2009-06-07 | 17:45 | 12.48 | 18 | 27 |  |  | 7.65
2009-06-07 | 18:00 | 12.47 | 18 | 27 |  |  | 7.65
2009-06-07 | 18:15 | 12.4 | 18 | 27 |  |  | 7.65
2009-06-07 | 18:30 | 12.33 | 18 | 27 |  |  | 7.65
2009-06-07 | 18:45 | 12.33 | 18 | 27 |  |  | 7.66
2009-06-07 | 19:00 | 12.32 | 18 | 28 |  |  | 7.66
2009-06-07 | 19:15 | 12.31 | 18 | 27 |  |  | 7.65
2009-06-07 | 19:30 | 12.26 | 18 | 27 |  |  | 7.64
2009-06-07 | 19:45 | 12.21 | 18 | 27 |  |  | 7.65
2009-06-07 | 20:00 | 12.19 | 18 | 27 |  |  | 7.64
2009-06-07 | 20:15 | 12.16 | 18 | 27 |  |  | 7.65
2009-06-07 | 20:30 | 12.14 | 18 | 27 |  |  | 7.64
2009-06-07 | 20:45 | 12.11 | 18 | 27 |  |  | 7.64
2009-06-07 | 21:00 | 12.05 | 18 | 27 |  |  | 7.64
2009-06-07 | 21:15 | 11.99 | 18 | 28 |  |  | 7.63
2009-06-07 | 21:30 | 11.92 | 18 | 28 |  |  | 7.63
2009-06-07 | 21:45 | 11.82 | 18 | 28 |  |  | 7.62
2009-06-07 | 22:00 | 11.73 | 18 | 28 |  |  | 7.62
2009-06-07 | 22:15 | 11.68 | 18 | 28 |  |  | 7.62
2009-06-07 | 22:30 | 11.63 | 18 | 28 |  |  | 7.62
2009-06-07 | 22:45 | 11.57 | 18 | 28 |  |  | 7.62
2009-06-07 | 23:00 | 11.48 | 18 | 28 |  |  | 7.62
2009-06-07 | 23:15 | 11.38 | 18 | 28 |  |  | 7.61
2009-06-07 | 23:30 | 11.27 | 18 | 28 |  |  | 7.61
2009-06-07 | 23:45 | 11.16 | 18 | 28 |  |  | 7.61
2009-06-08 | 0:00 | 11.04 | 18 | 28 |  |  | 7.6
2009-06-08 | 0:15 | 10.92 | 18 | 27 |  |  | 7.6
2009-06-08 | 0:30 | 10.81 | 18 | 27 |  |  | 7.6
2009-06-08 | 0:45 | 10.69 | 17 | 27 |  |  | 7.59
2009-06-08 | 1:01 | 10.58 | 17 | 27 |  |  | 7.59
2009-06-08 | 1:16 | 10.46 | 17 | 26 |  |  | 7.59
2009-06-08 | 1:31 | 10.35 | 17 | 26 |  |  | 7.58
2009-06-08 | 1:46 | 10.23 | 16 | 25 |  |  | 7.59
2009-06-08 | 2:01 | 10.11 | 16 | 24 |  |  | 7.58
2009-06-08 | 2:16 | 9.99 | 16 | 24 |  |  | 7.57
2009-06-08 | 2:31 | 9.86 | 16 | 24 |  |  | 7.58
2009-06-08 | 2:46 | 9.73 | 15 | 24 |  |  | 7.57
2009-06-08 | 3:01 | 9.6 | 15 | 24 |  |  | 7.56
2009-06-08 | 3:16 | 9.47 | 15 | 23 |  |  | 7.56
2009-06-08 | 3:31 | 9.33 | 15 | 23 |  |  | 7.55
2009-06-08 | 3:46 | 9.18 | 15 | 23 |  |  | 7.55
2009-06-08 | 4:01 | 9.03 | 15 | 23 |  |  | 7.55
2009-06-08 | 4:16 | 8.89 | 15 | 23 |  |  | 7.55
2009-06-08 | 4:31 | 8.76 | 15 | 23 |  |  | 7.54
2009-06-08 | 4:46 | 8.64 | 15 | 23 |  |  | 7.54
2009-06-08 | 5:01 | 8.53 | 15 | 23 |  |  | 7.54
2009-06-08 | 5:16 | 8.42 | 15 | 23 |  |  | 7.54
2009-06-08 | 5:31 | 8.33 | 15 | 23 |  |  | 7.54
2009-06-08 | 5:46 | 8.23 | 15 | 23 |  |  | 7.54
2009-06-08 | 6:01 | 8.13 | 15 | 23 |  |  | 7.54
2009-06-08 | 6:16 | 8.02 | 15 | 23 |  |  | 7.54
2009-06-08 | 6:31 | 7.91 | 15 | 22 |  |  | 7.54
2009-06-08 | 6:46 | 7.79 | 15 | 22 |  |  | 7.55
2009-06-08 | 7:01 | 7.68 | 15 | 22 |  |  | 7.54
2009-06-08 | 7:16 | 7.58 | 15 | 22 |  |  | 7.54
2009-06-08 | 7:31 | 7.49 | 15 | 22 |  |  | 7.56
2009-06-08 | 7:46 | 7.42 | 15 | 22 |  |  | 7.55
2009-06-08 | 8:01 | 7.36 | 15 | 22 |  |  | 7.56
2009-06-08 | 8:16 | 7.31 | 15 | 22 |  |  | 7.56
2009-06-08 | 8:31 | 7.27 | 15 | 22 |  |  | 7.57
2009-06-08 | 8:46 | 7.26 | 15 | 22 |  |  | 7.58
2009-06-08 | 9:01 | 7.29 | 15 | 22 |  |  | 7.58
2009-06-08 | 9:16 | 7.35 | 15 | 22 |  |  | 7.59
2009-06-08 | 9:31 | 7.42 | 15 | 22 |  |  | 7.59
2009-06-08 | 9:46 | 7.48 | 15 | 23 |  |  | 7.59
2009-06-08 | 10:01 | 7.62 | 15 | 23 |  |  | 7.61
2009-06-08 | 10:16 | 7.8 | 15 | 23 |  |  | 7.6
2009-06-08 | 10:31 | 7.83 | 16 | 24 |  |  | 7.61
2009-06-08 | 10:46 | 8.07 | 17 | 26 |  |  | 7.61
2009-06-08 | 11:01 | 8.22 | 17 | 26 |  |  | 7.61
2009-06-08 | 11:16 | 8.21 | 17 | 26 |  |  | 7.62
2009-06-08 | 11:31 | 8.42 | 17 | 26 |  |  | 7.62
2009-06-08 | 11:46 | 8.64 | 17 | 26 |  |  | 7.63
2009-06-08 | 12:01 | 8.89 | 17 | 26 |  |  | 7.62
2009-06-08 | 12:16 | 9.19 | 17 | 26 |  |  | 7.63
2009-06-08 | 12:31 | 9.44 | 17 | 26 |  |  | 7.63
2009-06-08 | 12:46 | 9.81 | 17 | 26 |  |  | 7.64
2009-06-08 | 13:00 | 10.21 | 17 | 26 |  |  | 7.64
2009-06-08 | 13:15 | 10.53 | 17 | 26 |  |  | 7.65
2009-06-08 | 13:30 | 10.78 | 17 | 26 |  |  | 7.66
2009-06-08 | 13:45 | 11.01 | 17 | 26 |  |  | 7.66
2009-06-08 | 14:00 | 11.28 | 16 | 25 |  |  | 7.67
2009-06-08 | 14:15 | 11.54 | 16 | 25 |  |  | 7.68
2009-06-08 | 14:30 | 11.78 | 16 | 25 |  |  | 7.67
2009-06-08 | 14:45 | 12.01 | 16 | 25 |  |  | 7.67
2009-06-08 | 15:00 | 12.26 | 16 | 25 |  |  | 7.68
2009-06-08 | 15:15 | 12.48 | 16 | 25 |  |  | 7.68
2009-06-08 | 15:30 | 12.68 | 16 | 25 |  |  | 7.68
2009-06-08 | 15:45 | 12.84 | 16 | 25 |  |  | 7.69
2009-06-08 | 16:00 | 12.98 | 17 | 26 |  |  | 7.69
2009-06-08 | 16:15 | 13.09 | 17 | 26 |  |  | 7.69
2009-06-08 | 16:30 | 13.19 | 17 | 26 |  |  | 7.69
2009-06-08 | 16:45 | 13.3 | 17 | 26 |  |  | 7.69
2009-06-08 | 17:00 | 13.33 | 17 | 26 |  |  | 7.69
2009-06-08 | 17:15 | 13.32 | 17 | 26 |  |  | 7.69
2009-06-08 | 17:30 | 13.34 | 17 | 26 |  |  | 7.69
2009-06-08 | 17:45 | 13.3 | 17 | 26 |  |  | 7.68
2009-06-08 | 18:00 | 13.26 | 17 | 26 |  |  | 7.68
2009-06-08 | 18:15 | 13.14 | 17 | 26 |  |  | 7.68
2009-06-08 | 18:30 | 13.09 | 17 | 26 |  |  | 7.67
2009-06-08 | 18:45 | 13.06 | 17 | 26 |  |  | 7.67
2009-06-08 | 19:00 | 12.97 | 17 | 26 |  |  | 7.66
2009-06-08 | 19:15 | 12.96 | 17 | 26 |  |  | 7.66
2009-06-08 | 19:30 | 12.92 | 17 | 26 |  |  | 7.66
2009-06-08 | 19:45 | 12.87 | 17 | 26 |  |  | 7.66
2009-06-08 | 20:00 | 12.83 | 17 | 26 |  |  | 7.66
2009-06-08 | 20:15 | 12.8 | 17 | 26 |  |  | 7.65
2009-06-08 | 20:30 | 12.77 | 17 | 27 |  |  | 7.65
2009-06-08 | 20:45 | 12.73 | 17 | 27 |  |  | 7.65
2009-06-08 | 21:00 | 12.66 | 17 | 27 |  |  | 7.64
2009-06-08 | 21:15 | 12.58 | 17 | 27 |  |  | 7.63
2009-06-08 | 21:30 | 12.53 | 17 | 27 |  |  | 7.63
2009-06-08 | 21:45 | 12.48 | 17 | 27 |  |  | 7.63
2009-06-08 | 22:00 | 12.46 | 17 | 26 |  |  | 7.63
2009-06-08 | 22:15 | 12.45 | 17 | 25 |  |  | 7.63
2009-06-08 | 22:30 | 12.43 | 16 | 25 |  |  | 7.63
2009-06-08 | 22:45 | 12.38 | 16 | 25 |  |  | 7.62
2009-06-08 | 23:00 | 12.32 | 16 | 25 |  |  | 7.62
2009-06-08 | 23:15 | 12.28 | 16 | 25 |  |  | 7.62
2009-06-08 | 23:30 | 12.25 | 16 | 25 |  |  | 7.62
2009-06-08 | 23:45 | 12.21 | 16 | 25 |  |  | 7.62
2009-06-09 | 0:00 | 12.12 | 16 | 24 |  |  | 7.6
2009-06-09 | 0:15 | 12 | 16 | 24 |  |  | 7.6
2009-06-09 | 0:30 | 11.87 | 16 | 24 |  |  | 7.6
2009-06-09 | 0:45 | 11.75 | 15 | 24 |  |  | 7.6
2009-06-09 | 1:01 | 11.64 | 15 | 24 |  |  | 7.59
2009-06-09 | 1:16 | 11.52 | 15 | 23 |  |  | 7.59
2009-06-09 | 1:31 | 11.41 | 15 | 23 |  |  | 7.58
2009-06-09 | 1:46 | 11.29 | 15 | 23 |  |  | 7.59
2009-06-09 | 2:01 | 11.18 | 15 | 23 |  |  | 7.58
2009-06-09 | 2:16 | 11.08 | 15 | 23 |  |  | 7.57
2009-06-09 | 2:31 | 10.98 | 15 | 23 |  |  | 7.58
2009-06-09 | 2:46 | 10.87 | 15 | 23 |  |  | 7.57
2009-06-09 | 3:01 | 10.75 | 15 | 23 |  |  | 7.57
2009-06-09 | 3:16 | 10.62 | 15 | 23 |  |  | 7.57
2009-06-09 | 3:31 | 10.49 | 15 | 23 |  |  | 7.56
2009-06-09 | 3:46 | 10.37 | 15 | 23 |  |  | 7.57
2009-06-09 | 4:01 | 10.25 | 15 | 22 |  |  | 7.56
2009-06-09 | 4:16 | 10.14 | 15 | 22 |  |  | 7.56
2009-06-09 | 4:31 | 10.03 | 15 | 22 |  |  | 7.56
2009-06-09 | 4:46 | 9.92 | 14 | 22 |  |  | 7.56
2009-06-09 | 5:01 | 9.8 | 15 | 22 |  |  | 7.56
2009-06-09 | 5:16 | 9.68 | 14 | 22 |  |  | 7.55
2009-06-09 | 5:31 | 9.54 | 14 | 22 |  |  | 7.56
2009-06-09 | 5:46 | 9.42 | 14 | 22 |  |  | 7.55
2009-06-09 | 6:01 | 9.29 | 14 | 22 |  |  | 7.55
2009-06-09 | 6:16 | 9.17 | 14 | 22 |  |  | 7.55
2009-06-09 | 6:31 | 9.05 | 14 | 22 |  |  | 7.54
2009-06-09 | 6:46 | 8.94 | 14 | 22 |  |  | 7.55
2009-06-09 | 7:01 | 8.83 | 14 | 22 |  |  | 7.55
2009-06-09 | 7:16 | 8.73 | 14 | 22 |  |  | 7.56
2009-06-09 | 7:31 | 8.64 | 14 | 22 |  |  | 7.55
2009-06-09 | 7:46 | 8.56 | 14 | 22 |  |  | 7.56
2009-06-09 | 8:01 | 8.49 | 14 | 22 |  |  | 7.56
2009-06-09 | 8:16 | 8.43 | 14 | 22 |  |  | 7.56
2009-06-09 | 8:31 | 8.38 | 14 | 22 |  |  | 7.58
2009-06-09 | 8:46 | 8.35 | 14 | 22 |  |  | 7.58
2009-06-09 | 9:01 | 8.32 | 14 | 22 |  |  | 7.58
2009-06-09 | 9:16 | 8.34 | 14 | 22 |  |  | 7.59
2009-06-09 | 9:31 | 8.4 | 14 | 22 |  |  | 7.6
2009-06-09 | 9:46 | 8.51 | 14 | 22 |  |  | 7.61
2009-06-09 | 10:01 | 8.59 | 14 | 22 |  |  | 7.61
2009-06-09 | 10:16 | 8.69 | 14 | 22 |  |  | 7.62
2009-06-09 | 10:31 | 8.83 | 14 | 22 |  |  | 7.62
2009-06-09 | 10:46 | 9 | 14 | 22 |  |  | 7.63
2009-06-09 | 11:01 | 9.17 | 14 | 22 |  |  | 7.63
2009-06-09 | 11:16 | 9.36 | 14 | 22 |  |  | 7.64
2009-06-09 | 11:31 | 9.57 | 14 | 22 |  |  | 7.63
2009-06-09 | 11:46 | 9.78 | 14 | 22 |  |  | 7.64
2009-06-09 | 12:01 | 9.97 | 14 | 22 |  |  | 7.64
2009-06-09 | 12:16 | 10.16 | 14 | 22 |  |  | 7.65
2009-06-09 | 12:31 | 10.37 | 14 | 22 |  |  | 7.65
2009-06-09 | 12:46 | 10.6 | 14 | 22 |  |  | 7.66
2009-06-09 | 13:00 | 10.79 | 14 | 22 |  |  | 7.66
2009-06-09 | 13:15 | 10.95 | 14 | 22 |  |  | 7.66
2009-06-09 | 13:30 | 11.13 | 15 | 22 |  |  | 7.67
2009-06-09 | 13:45 | 11.3 | 15 | 22 |  |  | 7.67
2009-06-09 | 14:00 | 11.49 | 15 | 22 |  |  | 7.68
2009-06-09 | 14:15 | 11.65 | 15 | 22 |  |  | 7.67
2009-06-09 | 14:30 | 11.81 | 15 | 23 |  |  | 7.67
2009-06-09 | 14:45 | 11.83 | 15 | 23 |  |  | 7.67
2009-06-09 | 15:00 | 11.84 | 15 | 23 |  |  | 7.67
2009-06-09 | 15:15 | 11.81 | 15 | 23 |  |  | 7.67
2009-06-09 | 15:30 | 11.79 | 15 | 23 |  |  | 7.67
2009-06-09 | 15:45 | 11.75 | 15 | 23 |  |  | 7.67
2009-06-09 | 16:00 | 11.68 | 15 | 23 |  |  | 7.66
2009-06-09 | 16:15 | 11.58 | 15 | 23 |  |  | 7.67
2009-06-09 | 16:30 | 11.49 | 15 | 23 |  |  | 7.67
2009-06-09 | 16:45 | 11.41 | 15 | 23 |  |  | 7.67
2009-06-09 | 17:00 | 11.36 | 15 | 23 |  |  | 7.66
2009-06-09 | 17:15 | 11.32 | 15 | 23 |  |  | 7.67
2009-06-09 | 17:30 | 11.26 | 15 | 23 |  |  | 7.67
2009-06-09 | 17:45 | 11.22 | 15 | 24 |  |  | 7.66
2009-06-09 | 18:00 | 11.18 | 15 | 24 |  |  | 7.66
2009-06-09 | 18:15 | 11.16 | 15 | 24 |  |  | 7.66
2009-06-09 | 18:30 | 11.16 | 15 | 24 |  |  | 7.66
2009-06-09 | 18:45 | 11.15 | 15 | 24 |  |  | 7.67
2009-06-09 | 19:00 | 11.16 | 15 | 24 |  |  | 7.66
2009-06-09 | 19:15 | 11.16 | 15 | 24 |  |  | 7.67
2009-06-09 | 19:30 | 11.16 | 15 | 24 |  |  | 7.66
2009-06-09 | 19:45 | 11.15 | 15 | 24 |  |  | 7.66
2009-06-09 | 20:00 | 11.12 | 15 | 24 |  |  | 7.66
2009-06-09 | 20:15 | 11.08 | 15 | 24 |  |  | 7.66
2009-06-09 | 20:30 | 11.03 | 15 | 24 |  |  | 7.66
2009-06-09 | 20:45 | 10.97 | 16 | 24 |  |  | 7.65
2009-06-09 | 21:00 | 10.92 | 16 | 24 |  |  | 7.65
2009-06-09 | 21:15 | 10.86 | 16 | 24 |  |  | 7.65
2009-06-09 | 21:30 | 10.8 | 16 | 24 |  |  | 7.65
2009-06-09 | 21:45 | 10.72 | 16 | 24 |  |  | 7.64
2009-06-09 | 22:00 | 10.65 | 16 | 24 |  |  | 7.64
2009-06-09 | 22:15 | 10.58 | 16 | 24 |  |  | 7.64
2009-06-09 | 22:30 | 10.51 | 16 | 24 |  |  | 7.64
2009-06-09 | 22:45 | 10.43 | 16 | 24 |  |  | 7.64
2009-06-09 | 23:00 | 10.34 | 16 | 24 |  |  | 7.63
2009-06-09 | 23:15 | 10.24 | 16 | 24 |  |  | 7.63
2009-06-09 | 23:30 | 10.14 | 16 | 24 |  |  | 7.63
2009-06-09 | 23:45 | 10.05 | 16 | 24 |  |  | 7.63
2009-06-10 | 0:00 | 9.97 | 16 | 24 |  |  | 7.62
2009-06-10 | 0:15 | 9.9 | 16 | 24 |  |  | 7.62
2009-06-10 | 0:30 | 9.84 | 16 | 24 |  |  | 7.63
2009-06-10 | 0:45 | 9.78 | 16 | 24 |  |  | 7.62
2009-06-10 | 1:01 | 9.72 | 16 | 24 |  |  | 7.62
2009-06-10 | 1:16 | 9.67 | 16 | 24 |  |  | 7.62
2009-06-10 | 1:31 | 9.62 | 16 | 24 |  |  | 7.62
2009-06-10 | 1:46 | 9.57 | 16 | 24 |  |  | 7.62
2009-06-10 | 2:01 | 9.52 | 16 | 24 |  |  | 7.61
2009-06-10 | 2:16 | 9.47 | 16 | 24 |  |  | 7.61
2009-06-10 | 2:31 | 9.42 | 16 | 24 |  |  | 7.61
2009-06-10 | 2:46 | 9.38 | 16 | 24 |  |  | 7.61
2009-06-10 | 3:01 | 9.33 | 16 | 24 |  |  | 7.61
2009-06-10 | 3:16 | 9.29 | 15 | 24 |  |  | 7.61
2009-06-10 | 3:31 | 9.24 | 15 | 24 |  |  | 7.6
2009-06-10 | 3:46 | 9.19 | 15 | 24 |  |  | 7.6
2009-06-10 | 4:01 | 9.14 | 15 | 24 |  |  | 7.6
2009-06-10 | 4:16 | 9.1 | 15 | 24 |  |  | 7.6
2009-06-10 | 4:31 | 9.05 | 15 | 24 |  |  | 7.6
2009-06-10 | 4:46 | 9.01 | 15 | 24 |  |  | 7.59
2009-06-10 | 5:01 | 8.97 | 15 | 24 |  |  | 7.6
2009-06-10 | 5:16 | 8.92 | 15 | 24 |  |  | 7.59
2009-06-10 | 5:31 | 8.88 | 15 | 24 |  |  | 7.59
2009-06-10 | 5:46 | 8.83 | 15 | 24 |  |  | 7.6
2009-06-10 | 6:01 | 8.78 | 15 | 23 |  |  | 7.59
2009-06-10 | 6:16 | 8.73 | 15 | 23 |  |  | 7.59
2009-06-10 | 6:31 | 8.67 | 15 | 23 |  |  | 7.59
2009-06-10 | 6:46 | 8.61 | 15 | 23 |  |  | 7.58
2009-06-10 | 7:01 | 8.56 | 15 | 23 |  |  | 7.59
2009-06-10 | 7:16 | 8.51 | 15 | 23 |  |  | 7.6
2009-06-10 | 7:31 | 8.47 | 15 | 23 |  |  | 7.6
2009-06-10 | 7:46 | 8.44 | 15 | 23 |  |  | 7.6
2009-06-10 | 8:01 | 8.4 | 15 | 23 |  |  | 7.61
2009-06-10 | 8:16 | 8.37 | 15 | 23 |  |  | 7.61
2009-06-10 | 8:31 | 8.34 | 15 | 23 |  |  | 7.6
2009-06-10 | 8:46 | 8.31 | 15 | 23 |  |  | 7.61
2009-06-10 | 9:01 | 8.28 | 15 | 23 |  |  | 7.62
2009-06-10 | 9:16 | 8.26 | 15 | 23 |  |  | 7.62
2009-06-10 | 9:31 | 8.23 | 15 | 23 |  |  | 7.62
2009-06-10 | 9:46 | 8.21 | 15 | 23 |  |  | 7.62
2009-06-10 | 10:01 | 8.19 | 15 | 23 |  |  | 7.62
2009-06-10 | 10:16 | 8.19 | 15 | 23 |  |  | 7.63
2009-06-10 | 10:31 | 8.19 | 15 | 23 |  |  | 7.63
2009-06-10 | 10:46 | 8.2 | 15 | 23 |  |  | 7.63
2009-06-10 | 11:01 | 8.21 | 15 | 23 |  |  | 7.64
2009-06-10 | 11:16 | 8.24 | 15 | 23 |  |  | 7.64
2009-06-10 | 11:31 | 8.29 | 15 | 23 |  |  | 7.64
2009-06-10 | 11:46 | 8.35 | 15 | 23 |  |  | 7.65
2009-06-10 | 12:01 | 8.44 | 15 | 23 |  |  | 7.65
2009-06-10 | 12:16 | 8.5 | 15 | 23 |  |  | 7.65
2009-06-10 | 12:31 | 8.57 | 15 | 23 |  |  | 7.66
2009-06-10 | 12:46 | 8.66 | 15 | 23 |  |  | 7.67
2009-06-10 | 13:00 | 8.75 | 15 | 23 |  |  | 7.67
2009-06-10 | 13:15 | 8.89 | 15 | 23 |  |  | 7.67
2009-06-10 | 13:30 | 9.01 | 15 | 23 |  |  | 7.68
2009-06-10 | 13:45 | 9.19 | 15 | 23 |  |  | 7.69
2009-06-10 | 14:00 | 9.41 | 15 | 23 |  |  | 7.69
2009-06-10 | 14:15 | 9.6 | 15 | 23 |  |  | 7.69
2009-06-10 | 14:30 | 9.75 | 15 | 23 |  |  | 7.69
2009-06-10 | 14:45 | 10.02 | 15 | 23 |  |  | 7.69
2009-06-10 | 15:00 | 10.31 | 15 | 23 |  |  | 7.7
2009-06-10 | 15:15 | 10.6 | 15 | 23 |  |  | 7.7
2009-06-10 | 15:30 | 10.74 | 15 | 23 |  |  | 7.7
2009-06-10 | 15:45 | 10.84 | 15 | 23 |  |  | 7.7
2009-06-10 | 16:00 | 10.94 | 15 | 23 |  |  | 7.7
2009-06-10 | 16:15 | 11.16 | 15 | 24 |  |  | 7.7
2009-06-10 | 16:30 | 11.34 | 15 | 24 |  |  | 7.7
2009-06-10 | 16:45 | 11.43 | 15 | 24 |  |  | 7.7
2009-06-10 | 17:00 | 11.5 | 15 | 24 |  |  | 7.7
2009-06-10 | 17:15 | 11.52 | 15 | 24 |  |  | 7.7
2009-06-10 | 17:30 | 11.66 | 15 | 24 |  |  | 7.71
2009-06-10 | 17:45 | 11.94 | 15 | 24 |  |  | 7.71
2009-06-10 | 18:00 | 12.18 | 15 | 24 |  |  | 7.71
2009-06-10 | 18:15 | 12.43 | 15 | 24 |  |  | 7.72
2009-06-10 | 18:30 | 12.58 | 15 | 24 |  |  | 7.72
2009-06-10 | 18:45 | 12.68 | 15 | 24 |  |  | 7.71
2009-06-10 | 19:00 | 12.61 | 15 | 24 |  |  | 7.72
2009-06-10 | 19:15 | 12.54 | 15 | 24 |  |  | 7.71
2009-06-10 | 19:30 | 12.4 | 15 | 24 |  |  | 7.72
2009-06-10 | 19:45 | 12.22 | 15 | 24 |  |  | 7.72
2009-06-10 | 20:00 | 12.01 | 15 | 24 |  |  | 7.71
2009-06-10 | 20:15 | 11.81 | 15 | 24 |  |  | 7.71
2009-06-10 | 20:30 | 11.59 | 15 | 24 |  |  | 7.71
2009-06-10 | 20:45 | 11.41 | 16 | 24 |  |  | 7.7
2009-06-10 | 21:00 | 11.3 | 16 | 24 |  |  | 7.7
2009-06-10 | 21:15 | 11.23 | 16 | 24 |  |  | 7.69
2009-06-10 | 21:30 | 11.2 | 16 | 24 |  |  | 7.69
2009-06-10 | 21:45 | 11.21 | 16 | 24 |  |  | 7.69
2009-06-10 | 22:00 | 11.24 | 16 | 24 |  |  | 7.68
2009-06-10 | 22:15 | 11.25 | 16 | 24 |  |  | 7.68
2009-06-10 | 22:30 | 11.21 | 16 | 24 |  |  | 7.67
2009-06-10 | 22:45 | 11.14 | 16 | 24 |  |  | 7.68
2009-06-10 | 23:00 | 11.07 | 16 | 24 |  |  | 7.67
2009-06-10 | 23:15 | 11.01 | 16 | 24 |  |  | 7.66
2009-06-10 | 23:30 | 10.95 | 16 | 24 |  |  | 7.66
2009-06-10 | 23:45 | 10.89 | 16 | 24 |  |  | 7.66
2009-06-11 | 0:00 | 10.84 | 16 | 24 |  |  | 7.66
2009-06-11 | 0:15 | 10.81 | 16 | 24 |  |  | 7.66
2009-06-11 | 0:30 | 10.77 | 16 | 24 |  |  | 7.66
2009-06-11 | 0:45 | 10.73 | 16 | 24 |  |  | 7.66
2009-06-11 | 1:01 | 10.68 | 16 | 24 |  |  | 7.66
2009-06-11 | 1:16 | 10.62 | 16 | 24 |  |  | 7.66
2009-06-11 | 1:31 | 10.56 | 16 | 24 |  |  | 7.66
2009-06-11 | 1:46 | 10.49 | 16 | 24 |  |  | 7.66
2009-06-11 | 2:01 | 10.44 | 16 | 24 |  |  | 7.66
2009-06-11 | 2:16 | 10.38 | 16 | 24 |  |  | 7.65
2009-06-11 | 2:31 | 10.31 | 16 | 24 |  |  | 7.64
2009-06-11 | 2:46 | 10.25 | 16 | 24 |  |  | 7.65
2009-06-11 | 3:01 | 10.19 | 16 | 24 |  |  | 7.64
2009-06-11 | 3:16 | 10.12 | 16 | 24 |  |  | 7.64
2009-06-11 | 3:31 | 10.04 | 16 | 24 |  |  | 7.64
2009-06-11 | 3:46 | 9.95 | 15 | 24 |  |  | 7.64
2009-06-11 | 4:01 | 9.85 | 16 | 24 |  |  | 7.64
2009-06-11 | 4:16 | 9.74 | 15 | 24 |  |  | 7.63
2009-06-11 | 4:31 | 9.62 | 15 | 24 |  |  | 7.63
2009-06-11 | 4:46 | 9.52 | 15 | 23 |  |  | 7.62
2009-06-11 | 5:01 | 9.44 | 15 | 23 |  |  | 7.62
2009-06-11 | 5:16 | 9.38 | 15 | 23 |  |  | 7.62
2009-06-11 | 5:31 | 9.32 | 15 | 23 |  |  | 7.61
2009-06-11 | 5:46 | 9.25 | 15 | 23 |  |  | 7.61
2009-06-11 | 6:01 | 9.19 | 15 | 23 |  |  | 7.61
2009-06-11 | 6:16 | 9.14 | 15 | 23 |  |  | 7.62
2009-06-11 | 6:31 | 9.08 | 15 | 23 |  |  | 7.62
2009-06-11 | 6:46 | 9.03 | 15 | 23 |  |  | 7.62
2009-06-11 | 7:01 | 8.99 | 15 | 23 |  |  | 7.63
2009-06-11 | 7:16 | 8.95 | 15 | 23 |  |  | 7.63
2009-06-11 | 7:31 | 8.91 | 15 | 23 |  |  | 7.64
2009-06-11 | 7:46 | 8.89 | 15 | 23 |  |  | 7.63
2009-06-11 | 8:01 | 8.87 | 15 | 23 |  |  | 7.64
2009-06-11 | 8:16 | 8.85 | 15 | 23 |  |  | 7.65
2009-06-11 | 8:31 | 8.87 | 15 | 23 |  |  | 7.65
2009-06-11 | 8:46 | 8.92 | 15 | 23 |  |  | 7.65
2009-06-11 | 9:01 | 8.98 | 15 | 22 |  |  | 7.65
2009-06-11 | 9:16 | 9.05 | 15 | 22 |  |  | 7.66
2009-06-11 | 9:31 | 9.1 | 15 | 22 |  |  | 7.66
2009-06-11 | 9:46 | 9.14 | 15 | 22 |  |  | 7.66
2009-06-11 | 10:01 | 9.23 | 15 | 22 |  |  | 7.66
2009-06-11 | 10:16 | 9.32 | 15 | 22 |  |  | 7.67
2009-06-11 | 10:31 | 9.38 | 15 | 22 |  |  | 7.67
2009-06-11 | 10:46 | 9.48 | 15 | 22 |  |  | 7.67
2009-06-11 | 11:01 | 9.59 | 15 | 22 |  |  | 7.68
2009-06-11 | 11:16 | 9.75 | 15 | 22 |  |  | 7.67
2009-06-11 | 11:31 | 9.79 | 15 | 23 |  |  | 7.67
2009-06-11 | 11:46 | 9.99 | 15 | 23 |  |  | 7.68
2009-06-11 | 12:01 | 10.13 | 15 | 23 |  |  | 7.68
2009-06-11 | 12:16 | 10.34 | 15 | 23 |  |  | 7.68
2009-06-11 | 12:31 | 10.62 | 15 | 23 |  |  | 7.69
2009-06-11 | 12:46 | 10.9 | 15 | 23 |  |  | 7.68
2009-06-11 | 13:00 | 11.09 | 15 | 23 |  |  | 7.7
2009-06-11 | 13:15 | 11.37 | 15 | 23 |  |  | 7.7
2009-06-11 | 13:30 | 11.56 | 15 | 23 |  |  | 7.7
2009-06-11 | 13:45 | 11.71 | 15 | 23 |  |  | 7.7
2009-06-11 | 14:00 | 11.94 | 15 | 23 |  |  | 7.7
2009-06-11 | 14:15 | 12.25 | 15 | 23 |  |  | 7.71
2009-06-11 | 14:30 | 12.44 | 15 | 23 |  |  | 7.71
2009-06-11 | 14:45 | 12.52 | 15 | 23 |  |  | 7.71
2009-06-11 | 15:00 | 12.55 | 15 | 23 |  |  | 7.7
2009-06-11 | 15:15 | 12.62 | 15 | 23 |  |  | 7.71
2009-06-11 | 15:30 | 12.79 | 15 | 23 |  |  | 7.71
2009-06-11 | 15:45 | 12.81 | 15 | 23 |  |  | 7.71
2009-06-11 | 16:00 | 12.76 | 15 | 23 |  |  | 7.71
2009-06-11 | 16:15 | 12.67 | 15 | 23 |  |  | 7.7
2009-06-11 | 16:30 | 12.39 | 15 | 23 |  |  | 7.7
2009-06-11 | 16:45 | 12.04 | 15 | 23 |  |  | 7.7
2009-06-11 | 17:00 | 11.82 | 15 | 23 |  |  | 7.69
2009-06-11 | 17:15 | 11.83 | 15 | 23 |  |  | 7.7
2009-06-11 | 17:30 | 11.99 | 15 | 23 |  |  | 7.7
2009-06-11 | 17:45 | 12.06 | 15 | 23 |  |  | 7.69
2009-06-11 | 18:00 | 12.02 | 15 | 23 |  |  | 7.69
2009-06-11 | 18:15 | 11.95 | 15 | 23 |  |  | 7.68
2009-06-11 | 18:30 | 11.89 | 15 | 23 |  |  | 7.69
2009-06-11 | 18:45 | 11.61 | 14 | 22 |  |  | 7.66
2009-06-11 | 19:00 | 11.28 | 14 | 22 |  |  | 7.65
2009-06-11 | 19:15 | 11.33 | 14 | 22 |  |  | 7.66
2009-06-11 | 19:30 | 11.51 | 15 | 22 |  |  | 7.66
2009-06-11 | 19:45 | 11.62 | 15 | 23 |  |  | 7.67
2009-06-11 | 20:00 | 11.66 | 15 | 22 |  |  | 7.67
2009-06-11 | 20:15 | 11.66 | 15 | 23 |  |  | 7.67
2009-06-11 | 20:30 | 11.61 | 15 | 23 |  |  | 7.67
2009-06-11 | 20:45 | 11.52 | 15 | 23 |  |  | 7.66
2009-06-11 | 21:00 | 11.42 | 15 | 23 |  |  | 7.66
2009-06-11 | 21:15 | 11.34 | 15 | 23 |  |  | 7.66
2009-06-11 | 21:30 | 11.25 | 15 | 23 |  |  | 7.66
2009-06-11 | 21:45 | 11.14 | 15 | 23 |  |  | 7.65
2009-06-11 | 22:00 | 11.04 | 15 | 23 |  |  | 7.65
2009-06-11 | 22:15 | 10.94 | 15 | 23 |  |  | 7.65
2009-06-11 | 22:30 | 10.87 | 15 | 23 |  |  | 7.65
2009-06-11 | 22:45 | 10.81 | 15 | 23 |  |  | 7.64
2009-06-11 | 23:00 | 10.76 | 15 | 23 |  |  | 7.64
2009-06-11 | 23:15 | 10.68 | 15 | 23 |  |  | 7.64
2009-06-11 | 23:30 | 10.56 | 15 | 23 |  |  | 7.64
2009-06-11 | 23:45 | 10.42 | 15 | 23 |  |  | 7.64
2009-06-12 | 0:00 | 10.27 | 15 | 23 |  |  | 7.63
2009-06-12 | 0:15 | 10.13 | 15 | 23 |  |  | 7.64
2009-06-12 | 0:30 | 9.99 | 15 | 23 |  |  | 7.63
2009-06-12 | 0:45 | 9.86 | 15 | 23 |  |  | 7.63
2009-06-12 | 1:01 | 9.75 | 15 | 23 |  |  | 7.62
2009-06-12 | 1:16 | 9.64 | 15 | 23 |  |  | 7.62
2009-06-12 | 1:31 | 9.54 | 15 | 23 |  |  | 7.62
2009-06-12 | 1:46 | 9.44 | 15 | 23 |  |  | 7.62
2009-06-12 | 2:01 | 9.34 | 15 | 23 |  |  | 7.61
2009-06-12 | 2:16 | 9.23 | 15 | 23 |  |  | 7.6
2009-06-12 | 2:31 | 9.13 | 15 | 23 |  |  | 7.6
2009-06-12 | 2:46 | 9.04 | 15 | 23 |  |  | 7.6
2009-06-12 | 3:01 | 8.96 | 15 | 23 |  |  | 7.59
2009-06-12 | 3:16 | 8.9 | 15 | 22 |  |  | 7.59
2009-06-12 | 3:31 | 8.85 | 15 | 22 |  |  | 7.59
2009-06-12 | 3:46 | 8.81 | 15 | 22 |  |  | 7.58
2009-06-12 | 4:01 | 8.79 | 15 | 22 |  |  | 7.58
2009-06-12 | 4:16 | 8.77 | 15 | 22 |  |  | 7.58
2009-06-12 | 4:31 | 8.75 | 15 | 22 |  |  | 7.58
2009-06-12 | 4:46 | 8.72 | 15 | 22 |  |  | 7.57
2009-06-12 | 5:01 | 8.68 | 15 | 22 |  |  | 7.57
2009-06-12 | 5:16 | 8.64 | 14 | 22 |  |  | 7.57
2009-06-12 | 5:31 | 8.58 | 14 | 22 |  |  | 7.57
2009-06-12 | 5:46 | 8.52 | 14 | 22 |  |  | 7.57
2009-06-12 | 6:01 | 8.46 | 14 | 22 |  |  | 7.56
2009-06-12 | 6:16 | 8.4 | 14 | 22 |  |  | 7.56
2009-06-12 | 6:31 | 8.35 | 14 | 22 |  |  | 7.57
2009-06-12 | 6:46 | 8.3 | 14 | 22 |  |  | 7.57
2009-06-12 | 7:01 | 8.26 | 14 | 22 |  |  | 7.57
2009-06-12 | 7:16 | 8.22 | 14 | 22 |  |  | 7.57
2009-06-12 | 7:31 | 8.18 | 14 | 22 |  |  | 7.58
2009-06-12 | 7:46 | 8.15 | 14 | 22 |  |  | 7.58
2009-06-12 | 8:01 | 8.14 | 14 | 22 |  |  | 7.59
2009-06-12 | 8:16 | 8.13 | 14 | 22 |  |  | 7.6
2009-06-12 | 8:31 | 8.14 | 14 | 22 |  |  | 7.61
2009-06-12 | 8:46 | 8.17 | 14 | 22 |  |  | 7.61
2009-06-12 | 9:01 | 8.2 | 14 | 22 |  |  | 7.62
2009-06-12 | 9:16 | 8.25 | 14 | 22 |  |  | 7.62
2009-06-12 | 9:31 | 8.31 | 15 | 22 |  |  | 7.62
2009-06-12 | 9:46 | 8.37 | 15 | 22 |  |  | 7.63
2009-06-12 | 10:01 | 8.44 | 14 | 22 |  |  | 7.64
2009-06-12 | 10:16 | 8.58 | 14 | 22 |  |  | 7.65
2009-06-12 | 10:31 | 8.67 | 14 | 22 |  |  | 7.65
2009-06-12 | 10:46 | 8.79 | 14 | 22 |  |  | 7.66
2009-06-12 | 11:01 | 8.92 | 14 | 22 |  |  | 7.66
2009-06-12 | 11:16 | 9.13 | 14 | 22 |  |  | 7.66
2009-06-12 | 11:31 | 9.27 | 14 | 22 |  |  | 7.67
2009-06-12 | 11:46 | 9.57 | 14 | 22 |  |  | 7.68
2009-06-12 | 12:01 | 9.94 | 14 | 22 |  |  | 7.68
2009-06-12 | 12:16 | 10.32 | 15 | 22 |  |  | 7.69
2009-06-12 | 12:31 | 10.56 | 15 | 22 |  |  | 7.69
2009-06-12 | 12:46 | 10.81 | 15 | 22 |  |  | 7.7
2009-06-12 | 13:00 | 11.19 | 15 | 22 |  |  | 7.7
2009-06-12 | 13:15 | 11.33 | 15 | 22 |  |  | 7.7
2009-06-12 | 13:30 | 11.78 | 15 | 22 |  |  | 7.71
2009-06-12 | 13:45 | 11.93 | 15 | 23 |  |  | 7.71
2009-06-12 | 14:00 | 11.71 | 15 | 23 |  |  | 7.71
2009-06-12 | 14:15 | 11.87 | 15 | 23 |  |  | 7.71
2009-06-12 | 14:30 | 12.14 | 15 | 23 |  |  | 7.72
2009-06-12 | 14:45 | 11.97 | 15 | 23 |  |  | 7.71
2009-06-12 | 15:00 | 11.83 | 15 | 23 |  |  | 7.71
2009-06-12 | 15:15 | 11.71 | 15 | 23 |  |  | 7.71
2009-06-12 | 15:30 | 11.85 | 15 | 23 |  |  | 7.72
2009-06-12 | 15:45 | 11.95 | 15 | 23 |  |  | 7.72
2009-06-12 | 16:00 | 11.98 | 15 | 23 |  |  | 7.71
2009-06-12 | 16:15 | 12.16 | 15 | 22 |  |  | 7.72
2009-06-12 | 16:30 | 12.37 | 15 | 22 |  |  | 7.72
2009-06-12 | 16:45 | 12.3 | 15 | 23 |  |  | 7.72
2009-06-12 | 17:00 | 12.31 | 15 | 23 |  |  | 7.71
2009-06-12 | 17:15 | 12.29 | 15 | 23 |  |  | 7.71
2009-06-12 | 17:30 | 12.16 | 14 | 22 |  |  | 7.71
2009-06-12 | 17:45 | 12.09 | 15 | 23 |  |  | 7.71
2009-06-12 | 18:00 | 12 | 15 | 23 |  |  | 7.7
2009-06-12 | 18:15 | 11.94 | 15 | 23 |  |  | 7.7
2009-06-12 | 18:30 | 11.9 | 15 | 23 |  |  | 7.7
2009-06-12 | 18:45 | 11.88 | 15 | 23 |  |  | 7.7
2009-06-12 | 19:00 | 11.89 | 15 | 23 |  |  | 7.7
2009-06-12 | 19:15 | 11.9 | 15 | 22 |  |  | 7.7
2009-06-12 | 19:30 | 11.84 | 15 | 22 |  |  | 7.7
2009-06-12 | 19:45 | 11.73 | 15 | 23 |  |  | 7.7
2009-06-12 | 20:00 | 11.63 | 15 | 22 |  |  | 7.7
2009-06-12 | 20:15 | 11.58 | 15 | 23 |  |  | 7.7
2009-06-12 | 20:30 | 11.55 | 15 | 23 |  |  | 7.69
2009-06-12 | 20:45 | 11.51 | 15 | 23 |  |  | 7.68
2009-06-12 | 21:00 | 11.49 | 15 | 23 |  |  | 7.67
2009-06-12 | 21:15 | 11.48 | 15 | 23 |  |  | 7.67
2009-06-12 | 21:30 | 11.45 | 15 | 23 |  |  | 7.67
2009-06-12 | 21:45 | 11.39 | 15 | 23 |  |  | 7.67
2009-06-12 | 22:00 | 11.32 | 15 | 23 |  |  | 7.67
2009-06-12 | 22:15 | 11.25 | 15 | 23 |  |  | 7.65
2009-06-12 | 22:30 | 11.2 | 15 | 23 |  |  | 7.65
2009-06-12 | 22:45 | 11.15 | 15 | 23 |  |  | 7.64
2009-06-12 | 23:00 | 11.1 | 15 | 23 |  |  | 7.64
2009-06-12 | 23:15 | 11.01 | 15 | 23 |  |  | 7.64
2009-06-12 | 23:30 | 10.91 | 15 | 23 |  |  | 7.63
2009-06-12 | 23:45 | 10.81 | 15 | 23 |  |  | 7.63
2009-06-13 | 0:00 | 10.71 | 15 | 23 |  |  | 7.63
2009-06-13 | 0:15 | 10.6 | 15 | 23 |  |  | 7.63
2009-06-13 | 0:30 | 10.48 | 15 | 23 |  |  | 7.62
2009-06-13 | 0:45 | 10.34 | 15 | 23 |  |  | 7.62
2009-06-13 | 1:01 | 10.18 | 15 | 23 |  |  | 7.61
2009-06-13 | 1:16 | 9.99 | 15 | 23 |  |  | 7.6
2009-06-13 | 1:31 | 9.78 | 15 | 23 |  |  | 7.6
2009-06-13 | 1:46 | 9.58 | 15 | 23 |  |  | 7.6
2009-06-13 | 2:01 | 9.41 | 15 | 23 |  |  | 7.59
2009-06-13 | 2:16 | 9.26 | 16 | 24 |  |  | 7.59
2009-06-13 | 2:31 | 9.13 | 16 | 24 |  |  | 7.59
2009-06-13 | 2:46 | 9.03 | 16 | 24 |  |  | 7.58
2009-06-13 | 3:01 | 8.95 | 16 | 24 |  |  | 7.58
2009-06-13 | 3:16 | 8.88 | 16 | 24 |  |  | 7.57
2009-06-13 | 3:31 | 8.81 | 16 | 24 |  |  | 7.58
2009-06-13 | 3:46 | 8.74 | 16 | 24 |  |  | 7.57
2009-06-13 | 4:01 | 8.66 | 16 | 24 |  |  | 7.57
2009-06-13 | 4:16 | 8.56 | 16 | 24 |  |  | 7.57
2009-06-13 | 4:31 | 8.46 | 16 | 24 |  |  | 7.57
2009-06-13 | 4:46 | 8.35 | 16 | 24 |  |  | 7.57
2009-06-13 | 5:01 | 8.25 | 16 | 24 |  |  | 7.56
2009-06-13 | 5:16 | 8.16 | 16 | 24 |  |  | 7.56
2009-06-13 | 5:31 | 8.07 | 16 | 24 |  |  | 7.56
2009-06-13 | 5:46 | 8 | 16 | 24 |  |  | 7.56
2009-06-13 | 6:01 | 7.94 | 16 | 24 |  |  | 7.57
2009-06-13 | 6:16 | 7.89 | 16 | 24 |  |  | 7.56
2009-06-13 | 6:31 | 7.84 | 16 | 24 |  |  | 7.56
2009-06-13 | 6:46 | 7.78 | 16 | 24 |  |  | 7.57
2009-06-13 | 7:01 | 7.74 | 16 | 24 |  |  | 7.58
2009-06-13 | 7:16 | 7.69 | 16 | 24 |  |  | 7.57
2009-06-13 | 7:31 | 7.65 | 16 | 24 |  |  | 7.58
2009-06-13 | 7:46 | 7.62 | 16 | 24 |  |  | 7.59
2009-06-13 | 8:01 | 7.59 | 16 | 24 |  |  | 7.59
2009-06-13 | 8:16 | 7.58 | 16 | 24 |  |  | 7.6
2009-06-13 | 8:31 | 7.56 | 16 | 24 |  |  | 7.61
2009-06-13 | 8:46 | 7.56 | 16 | 24 |  |  | 7.61
2009-06-13 | 9:01 | 7.61 | 16 | 24 |  |  | 7.62
2009-06-13 | 9:16 | 7.7 | 16 | 24 |  |  | 7.63
2009-06-13 | 9:31 | 7.81 | 16 | 24 |  |  | 7.63
2009-06-13 | 9:46 | 7.92 | 16 | 24 |  |  | 7.65
2009-06-13 | 10:01 | 8.04 | 16 | 24 |  |  | 7.64
2009-06-13 | 10:16 | 8.14 | 16 | 24 |  |  | 7.65
2009-06-13 | 10:31 | 8.28 | 16 | 24 |  |  | 7.66
2009-06-13 | 10:46 | 8.53 | 16 | 24 |  |  | 7.67
2009-06-13 | 11:01 | 8.8 | 16 | 24 |  |  | 7.67
2009-06-13 | 11:16 | 9.06 | 16 | 24 |  |  | 7.68
2009-06-13 | 11:31 | 9.28 | 16 | 24 |  |  | 7.68
2009-06-13 | 11:46 | 9.47 | 16 | 24 |  |  | 7.69
2009-06-13 | 12:01 | 9.74 | 16 | 24 |  |  | 7.69
2009-06-13 | 12:16 | 10.03 | 16 | 24 |  |  | 7.7
2009-06-13 | 12:31 | 10.38 | 16 | 24 |  |  | 7.7
2009-06-13 | 12:46 | 10.7 | 16 | 25 |  |  | 7.71
2009-06-13 | 13:00 | 10.98 | 16 | 25 |  |  | 7.72
2009-06-13 | 13:15 | 11.33 | 16 | 25 |  |  | 7.72
2009-06-13 | 13:30 | 11.63 | 16 | 25 |  |  | 7.73
2009-06-13 | 13:45 | 11.92 | 16 | 25 |  |  | 7.73
2009-06-13 | 14:00 | 12.13 | 16 | 25 |  |  | 7.73
2009-06-13 | 14:15 | 11.93 | 16 | 25 |  |  | 7.73
2009-06-13 | 14:30 | 12.15 | 16 | 25 |  |  | 7.73
2009-06-13 | 14:45 | 12.34 | 16 | 25 |  |  | 7.73
2009-06-13 | 15:00 | 12.61 | 16 | 25 |  |  | 7.74
2009-06-13 | 15:15 | 13.03 | 16 | 25 |  |  | 7.74
2009-06-13 | 15:30 | 13.45 | 16 | 25 |  |  | 7.75
2009-06-13 | 15:45 | 13.56 | 16 | 25 |  |  | 7.76
2009-06-13 | 16:00 | 13.42 | 16 | 25 |  |  | 7.74
2009-06-13 | 16:15 | 13.24 | 16 | 25 |  |  | 7.74
2009-06-13 | 16:30 | 13.29 | 16 | 25 |  |  | 7.74
2009-06-13 | 16:45 | 13.19 | 16 | 25 |  |  | 7.74
2009-06-13 | 17:00 | 13.25 | 16 | 25 |  |  | 7.74
2009-06-13 | 17:15 | 13.31 | 16 | 25 |  |  | 7.74
2009-06-13 | 17:30 | 13.12 | 16 | 24 |  |  | 7.73
2009-06-13 | 17:45 | 12.92 | 16 | 24 |  |  | 7.72
2009-06-13 | 18:00 | 12.76 | 16 | 25 |  |  | 7.73
2009-06-13 | 18:15 | 12.64 | 16 | 24 |  |  | 7.73
2009-06-13 | 18:30 | 12.64 | 16 | 25 |  |  | 7.73
2009-06-13 | 18:45 | 12.73 | 16 | 25 |  |  | 7.74
2009-06-13 | 19:00 | 12.9 | 16 | 25 |  |  | 7.74
2009-06-13 | 19:15 | 13.02 | 16 | 24 |  |  | 7.74
2009-06-13 | 19:30 | 13.12 | 16 | 25 |  |  | 7.74
2009-06-13 | 19:45 | 13.25 | 16 | 25 |  |  | 7.74
2009-06-13 | 20:00 | 13.29 | 16 | 25 |  |  | 7.74
2009-06-13 | 20:15 | 13.3 | 16 | 25 |  |  | 7.73
2009-06-13 | 20:30 | 13.32 | 16 | 25 |  |  | 7.73
2009-06-13 | 20:45 | 13.28 | 16 | 25 |  |  | 7.72
2009-06-13 | 21:00 | 13.16 | 16 | 25 |  |  | 7.71
2009-06-13 | 21:15 | 13.03 | 16 | 24 |  |  | 7.71
2009-06-13 | 21:30 | 12.95 | 16 | 25 |  |  | 7.71
2009-06-13 | 21:45 | 12.9 | 16 | 24 |  |  | 7.7
2009-06-13 | 22:00 | 12.84 | 16 | 24 |  |  | 7.7
2009-06-13 | 22:15 | 12.78 | 16 | 24 |  |  | 7.7
2009-06-13 | 22:30 | 12.72 | 16 | 24 |  |  | 7.69
2009-06-13 | 22:45 | 12.64 | 16 | 24 |  |  | 7.68
2009-06-13 | 23:00 | 12.58 | 16 | 24 |  |  | 7.68
2009-06-13 | 23:15 | 12.5 | 16 | 24 |  |  | 7.68
2009-06-13 | 23:30 | 12.39 | 16 | 24 |  |  | 7.68
2009-06-13 | 23:45 | 12.25 | 16 | 25 |  |  | 7.67
2009-06-14 | 0:00 | 12.1 | 16 | 25 |  |  | 7.67
2009-06-14 | 0:15 | 11.93 | 16 | 25 |  |  | 7.67
2009-06-14 | 0:30 | 11.8 | 16 | 25 |  |  | 7.67
2009-06-14 | 0:45 | 11.69 | 16 | 25 |  |  | 7.66
2009-06-14 | 1:01 | 11.59 | 16 | 25 |  |  | 7.66
2009-06-14 | 1:16 | 11.48 | 16 | 25 |  |  | 7.66
2009-06-14 | 1:31 | 11.34 | 16 | 24 |  |  | 7.66
2009-06-14 | 1:46 | 11.2 | 16 | 24 |  |  | 7.66
2009-06-14 | 2:01 | 11.08 | 16 | 24 |  |  | 7.65
2009-06-14 | 2:16 | 10.97 | 16 | 24 |  |  | 7.65
2009-06-14 | 2:31 | 10.87 | 16 | 24 |  |  | 7.64
2009-06-14 | 2:46 | 10.78 | 16 | 24 |  |  | 7.64
2009-06-14 | 3:01 | 10.7 | 16 | 24 |  |  | 7.63
2009-06-14 | 3:16 | 10.61 | 16 | 24 |  |  | 7.63
2009-06-14 | 3:31 | 10.53 | 15 | 24 |  |  | 7.62
2009-06-14 | 3:46 | 10.43 | 15 | 24 |  |  | 7.62
2009-06-14 | 4:01 | 10.33 | 15 | 24 |  |  | 7.62
2009-06-14 | 4:16 | 10.22 | 15 | 24 |  |  | 7.62
2009-06-14 | 4:31 | 10.11 | 15 | 24 |  |  | 7.62
2009-06-14 | 4:46 | 10.01 | 15 | 24 |  |  | 7.61
2009-06-14 | 5:01 | 9.92 | 15 | 23 |  |  | 7.61
2009-06-14 | 5:16 | 9.84 | 15 | 23 |  |  | 7.6
2009-06-14 | 5:31 | 9.78 | 15 | 23 |  |  | 7.61
2009-06-14 | 5:46 | 9.72 | 15 | 23 |  |  | 7.6
2009-06-14 | 6:01 | 9.67 | 15 | 23 |  |  | 7.6
2009-06-14 | 6:16 | 9.62 | 15 | 23 |  |  | 7.6
2009-06-14 | 6:31 | 9.57 | 15 | 23 |  |  | 7.61
2009-06-14 | 6:46 | 9.53 | 15 | 23 |  |  | 7.6
2009-06-14 | 7:01 | 9.5 | 15 | 23 |  |  | 7.61
2009-06-14 | 7:16 | 9.46 | 15 | 23 |  |  | 7.61
2009-06-14 | 7:31 | 9.44 | 15 | 23 |  |  | 7.61
2009-06-14 | 7:46 | 9.41 | 15 | 23 |  |  | 7.61
2009-06-14 | 8:01 | 9.41 | 15 | 23 |  |  | 7.63
2009-06-14 | 8:16 | 9.39 | 15 | 23 |  |  | 7.63
2009-06-14 | 8:31 | 9.38 | 15 | 23 |  |  | 7.64
2009-06-14 | 8:46 | 9.41 | 15 | 23 |  |  | 7.64
2009-06-14 | 9:01 | 9.42 | 15 | 23 |  |  | 7.64
2009-06-14 | 9:16 | 9.47 | 15 | 23 |  |  | 7.66
2009-06-14 | 9:31 | 9.57 | 15 | 23 |  |  | 7.67
2009-06-14 | 9:46 | 9.67 | 15 | 23 |  |  | 7.68
2009-06-14 | 10:01 | 9.74 | 15 | 23 |  |  | 7.67
2009-06-14 | 10:16 | 9.81 | 15 | 23 |  |  | 7.68
2009-06-14 | 10:31 | 9.93 | 15 | 23 |  |  | 7.67
2009-06-14 | 10:46 | 10.01 | 15 | 23 |  |  | 7.68
2009-06-14 | 11:01 | 10.14 | 15 | 23 |  |  | 7.69
2009-06-14 | 11:16 | 10.23 | 15 | 23 |  |  | 7.69
2009-06-14 | 11:31 | 10.48 | 15 | 23 |  |  | 7.69
2009-06-14 | 11:46 | 10.59 | 15 | 23 |  |  | 7.69
2009-06-14 | 12:01 | 10.8 | 15 | 23 |  |  | 7.69
2009-06-14 | 12:16 | 10.79 | 15 | 23 |  |  | 7.69
2009-06-14 | 12:31 | 10.77 | 15 | 23 |  |  | 7.7
2009-06-14 | 12:46 | 10.9 | 15 | 23 |  |  | 7.7
2009-06-14 | 13:00 | 11 | 15 | 23 |  |  | 7.7
2009-06-14 | 13:15 | 11.2 | 15 | 23 |  |  | 7.7
2009-06-14 | 13:30 | 11.36 | 15 | 23 |  |  | 7.71
2009-06-14 | 13:45 | 11.37 | 15 | 23 |  |  | 7.71
2009-06-14 | 14:00 | 11.43 | 15 | 23 |  |  | 7.71
2009-06-14 | 14:15 | 11.7 | 15 | 23 |  |  | 7.72
2009-06-14 | 14:30 | 11.77 | 15 | 23 |  |  | 7.71
2009-06-14 | 14:45 | 11.97 | 15 | 23 |  |  | 7.72
2009-06-14 | 15:00 | 12.28 | 15 | 23 |  |  | 7.73
2009-06-14 | 15:15 | 12.62 | 15 | 23 |  |  | 7.73
2009-06-14 | 15:30 | 13.23 | 15 | 23 |  |  | 7.74
2009-06-14 | 15:45 | 13.45 | 15 | 23 |  |  | 7.74
2009-06-14 | 16:00 | 13.88 | 15 | 23 |  |  | 7.75
2009-06-14 | 16:15 | 13.99 | 15 | 24 |  |  | 7.74
2009-06-14 | 16:30 | 14.09 | 15 | 23 |  |  | 7.75
2009-06-14 | 16:45 | 14.38 | 15 | 24 |  |  | 7.75
2009-06-14 | 17:00 | 14.38 | 15 | 24 |  |  | 7.76
2009-06-14 | 17:15 | 14.53 | 15 | 24 |  |  | 7.75
2009-06-14 | 17:30 | 14.61 | 15 | 24 |  |  | 7.75
2009-06-14 | 17:45 | 14.5 | 15 | 24 |  |  | 7.75
2009-06-14 | 18:00 | 14.54 | 15 | 24 |  |  | 7.74
2009-06-14 | 18:15 | 14.65 | 15 | 24 |  |  | 7.74
2009-06-14 | 18:30 | 14.72 | 16 | 24 |  |  | 7.74
2009-06-14 | 18:45 | 14.81 | 16 | 24 |  |  | 7.74
2009-06-14 | 19:00 | 14.89 | 16 | 24 |  |  | 7.74
2009-06-14 | 19:15 | 14.87 | 16 | 24 |  |  | 7.74
2009-06-14 | 19:30 | 14.83 | 15 | 24 |  |  | 7.74
2009-06-14 | 19:45 | 14.74 | 15 | 24 |  |  | 7.73
2009-06-14 | 20:00 | 14.64 | 15 | 24 |  |  | 7.73
2009-06-14 | 20:15 | 14.52 | 16 | 24 |  |  | 7.73
2009-06-14 | 20:30 | 14.41 | 16 | 24 |  |  | 7.72
2009-06-14 | 20:45 | 14.28 | 16 | 24 |  |  | 7.72
2009-06-14 | 21:00 | 14.15 | 16 | 24 |  |  | 7.71
2009-06-14 | 21:15 | 14.04 | 16 | 24 |  |  | 7.71
2009-06-14 | 21:30 | 13.97 | 16 | 24 |  |  | 7.71
2009-06-14 | 21:45 | 13.88 | 16 | 24 |  |  | 7.71
2009-06-14 | 22:00 | 13.74 | 16 | 24 |  |  | 7.7
2009-06-14 | 22:15 | 13.54 | 16 | 24 |  |  | 7.7
2009-06-14 | 22:30 | 13.34 | 16 | 24 |  |  | 7.7
2009-06-14 | 22:45 | 13.12 | 16 | 24 |  |  | 7.69
2009-06-14 | 23:00 | 12.9 | 16 | 24 |  |  | 7.69
2009-06-14 | 23:15 | 12.73 | 16 | 24 |  |  | 7.68
2009-06-14 | 23:30 | 12.61 | 16 | 24 |  |  | 7.68
2009-06-14 | 23:45 | 12.53 | 16 | 24 |  |  | 7.67
2009-06-15 | 0:00 | 12.44 | 16 | 24 |  |  | 7.67
2009-06-15 | 0:15 | 12.33 | 16 | 24 |  |  | 7.66
2009-06-15 | 0:30 | 12.2 | 15 | 24 |  |  | 7.66
2009-06-15 | 0:45 | 12.06 | 15 | 24 |  |  | 7.66
2009-06-15 | 1:01 | 11.89 | 15 | 24 |  |  | 7.65
2009-06-15 | 1:16 | 11.73 | 15 | 23 |  |  | 7.64
2009-06-15 | 1:31 | 11.58 | 15 | 23 |  |  | 7.64
2009-06-15 | 1:46 | 11.46 | 15 | 23 |  |  | 7.63
2009-06-15 | 2:01 | 11.38 | 15 | 23 |  |  | 7.62
2009-06-15 | 2:16 | 11.32 | 15 | 23 |  |  | 7.62
2009-06-15 | 2:31 | 11.28 | 15 | 23 |  |  | 7.62
2009-06-15 | 2:46 | 11.23 | 15 | 23 |  |  | 7.62
2009-06-15 | 3:01 | 11.17 | 15 | 23 |  |  | 7.6
2009-06-15 | 3:16 | 11.1 | 15 | 23 |  |  | 7.6
2009-06-15 | 3:31 | 11.03 | 15 | 23 |  |  | 7.6
2009-06-15 | 3:46 | 10.93 | 14 | 21 |  |  | 7.6
2009-06-15 | 4:01 | 10.82 | 14 | 21 |  |  | 7.59
2009-06-15 | 4:16 | 10.7 | 14 | 21 |  |  | 7.59
2009-06-15 | 4:31 | 10.58 | 14 | 21 |  |  | 7.58
2009-06-15 | 4:46 | 10.44 | 14 | 21 |  |  | 7.59
2009-06-15 | 5:01 | 10.31 | 14 | 21 |  |  | 7.57
2009-06-15 | 5:16 | 10.17 | 13 | 21 |  |  | 7.57
2009-06-15 | 5:31 | 10.04 | 13 | 21 |  |  | 7.57
2009-06-15 | 5:46 | 9.9 | 13 | 21 |  |  | 7.56
2009-06-15 | 6:01 | 9.78 | 13 | 21 |  |  | 7.57
2009-06-15 | 6:16 | 9.66 | 13 | 20 |  |  | 7.56
2009-06-15 | 6:31 | 9.53 | 13 | 20 |  |  | 7.57
2009-06-15 | 6:46 | 9.41 | 13 | 20 |  |  | 7.57
2009-06-15 | 7:01 | 9.29 | 13 | 20 |  |  | 7.57
2009-06-15 | 7:16 | 9.16 | 13 | 20 |  |  | 7.57
2009-06-15 | 7:31 | 9.04 | 13 | 20 |  |  | 7.57
2009-06-15 | 7:46 | 8.93 | 13 | 21 |  |  | 7.58
2009-06-15 | 8:01 | 8.84 | 13 | 21 |  |  | 7.57
2009-06-15 | 8:16 | 8.77 | 13 | 21 |  |  | 7.58
2009-06-15 | 8:31 | 8.71 | 13 | 21 |  |  | 7.59
2009-06-15 | 8:46 | 8.68 | 13 | 21 |  |  | 7.6
2009-06-15 | 9:01 | 8.65 | 13 | 21 |  |  | 7.6
2009-06-15 | 9:16 | 8.71 | 14 | 21 |  |  | 7.6
2009-06-15 | 9:31 | 8.68 | 14 | 21 |  |  | 7.6
2009-06-15 | 9:46 | 8.92 | 14 | 21 |  |  | 7.62
2009-06-15 | 10:01 | 9.07 | 14 | 21 |  |  | 7.62
2009-06-15 | 10:16 | 9.32 | 14 | 21 |  |  | 7.63
2009-06-15 | 10:31 | 9.37 | 14 | 21 |  |  | 7.63
2009-06-15 | 10:46 | 9.43 | 14 | 21 |  |  | 7.63
2009-06-15 | 11:01 | 9.56 | 14 | 21 |  |  | 7.64
2009-06-15 | 11:16 | 9.76 | 14 | 21 |  |  | 7.64
2009-06-15 | 11:31 | 10.14 | 14 | 21 |  |  | 7.65
2009-06-15 | 11:46 | 10.15 | 14 | 21 |  |  | 7.64
2009-06-15 | 12:01 | 10.27 | 14 | 21 |  |  | 7.64
2009-06-15 | 12:16 | 10.8 | 14 | 21 |  |  | 7.66
2009-06-15 | 12:31 | 10.93 | 14 | 21 |  |  | 7.65
2009-06-15 | 12:46 | 10.93 | 14 | 21 |  |  | 7.65
2009-06-15 | 13:00 | 11.17 | 17 | 27 |  |  | 7.67
2009-06-15 | 13:15 | 11.46 | 19 | 30 |  |  | 7.68