## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AA, YSI 556 MPS & Sonde AB
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-18 |  | 2009-04-18 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | In Situ |  | 12.71 |  | 12.67 |  | A
DO | Winkler titration of surface water |  | 11.5 |  | 12.88 |  | Fail
SpC | In SItu |  | 72 |  | 74 |  | A
Ph | In Situ |  | 8.12 |  | 8.77 |  | Fail
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-04-18 15:00 | 12.71 | 8.12 | 18.63 | 145.7 | 72 | 51
Sonde | 2009-04-18 15:00 | 12.67 | 8.77 | 12.88 | 121.3 | 74 | 48
Grade |  | A |  |  |  | A
Audit 2 | 2009-04-21 17:00 | 19.61 | 8.23 | 13.14 | 143.4 | 75 | 53
Sonde | 2009-04-21 17:00 | 19.6 | 9.05 | 10.57 | 115.4 | 76 | 49
Grade |  | A |  |  |  | A
Audit 3 | 2009-04-30 13:00 | 11.92 | 7.99 | 16.02 | 148.4 | 66 | 46
Sonde | 2009-04-25 13:00 | 11.87 | 8.78 | 12.84 | 118.8 | 67 | 44
Grade |  | A |  |  |  | A
Post-sampling Field Equipment Check
Field equipment ID | Sonde AA, YSI 556 MPS & Sonde AB
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-25 |  | 2009-04-25 |  | A
Time | Computer set to atomic time |  | 13:16 |  | 13:21 |  | A
Temp | VWR NIST traceable surface water |  | 12.1 |  | 11.88 |  | A
DO | Winkler titration of surface water |  | 10.8 |  | 12.66 |  | Fail
SpC | In Situ |  | 69 |  | 70 |  | A
Ph | In SItu |  | 8.8 |  | 8.78 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
556 Was not working for pH and DO. Sonde data will be disqualified as a precaution |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-18 |  | 2009-04-18 |  | A
Time | Computer set to atomic time |  | 11:59 |  | 11:49 |  | A
Temp | VWR NIST traceable surface water |  | 23.5 |  | 23.41 |  | A
DO | Winkler titration of surface water |  | 8.5 |  | 7.45 |  | Fail
SpC | 148 µS/mL |  | 153.6 |  | 144 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-04-18 15:00 | 12.71 | 8.12 | 18.63 | 145.7 | 72 | 51
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-18 |  | 2009-04-18 |  | A
Time | Computer set to atomic time |  | 15:00 |  | 15:00 |  | A
Temp | VWR NIST traceable surface water |  | 12.67 |  | 12.71 |  | A
DO | Winkler titration of surface water |  | 11.5 |  | 18.63 |  | 0
SpC | 148 µS/mL |  | 153.6 |  | 147 |  | A
Ph | pH 7 standard solution |  | 7.002 |  | 7.13 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-21 |  | 2009-04-21 |  | A
Time | Computer set to atomic time |  | 8:56 |  | 8:56 |  | A
Temp | VWR NIST traceable surface water |  | 11 |  | 10.84 |  | A
DO | Winkler titration of surface water |  | 10.3 |  | 11.95 |  | Fail
SpC | 148 µS/mL |  | 153.6 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.13 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-04-21 17:00 | 19.61 | 8.23 | 13.14 | 143.4 | 75 | 53
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-21 |  | 2009-04-21 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 19.6 |  | 19.61 |  | A
DO | Winkler titration of surface water |  | 9.5 |  | 13.14 |  | 0
SpC | 148 µS/mL |  | 153.6 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 3
Organization | Upper Deschutes Watershed Council |  | Station ID |  | DR 133.50
Program Name | Water Quality Monitoring Program |  | Description |  | Deschutes River, Lower Bridge
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | YSI 556 MPS
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-21 |  | 2009-04-21 |  | A
Time | Computer set to atomic time |  | 17:00 |  | 17:00 |  | A
Temp | VWR NIST traceable surface water |  | 19.6 |  | 19.61 |  | A
DO | Winkler titration of surface water |  | 9.5 |  | 13.14 |  | 0
SpC | 148 µS/mL |  | 153.6 |  | 148 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-04-30 13:00 | 11.92 | 7.99 | 16.02 | 148.4 | 66 | 46
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-04-25 |  | 2009-04-25 |  | A
Time | Computer set to atomic time |  | 13:24 |  | 13:24 |  | A
Temp | VWR NIST traceable surface water |  | 12.1 |  | 11.96 |  | A
DO | Winkler titration of surface water |  | 10.8 |  | 15.62 |  | 0
SpC | 148 µS/mL |  | 69 |  | 69 |  | A
Ph | pH 7 standard solution |  | 8.8 |  | 8.29 |  | Fail
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | TDS | SpCond | DO Conc | DO% | pH
M/D/Y |  | °C | mg/L | uS/cm | mg/L | %
2009-04-18 | 15:00 | 12.67 | 48 | 74
2009-04-18 | 15:15 | 12.81 | 48 | 74
2009-04-18 | 15:30 | 12.91 | 49 | 75
2009-04-18 | 15:45 | 13 | 49 | 75
2009-04-18 | 16:00 | 13.11 | 49 | 75
2009-04-18 | 16:15 | 13.18 | 49 | 75
2009-04-18 | 16:30 | 13.24 | 49 | 75
2009-04-18 | 16:45 | 13.29 | 49 | 75
2009-04-18 | 17:00 | 13.3 | 49 | 75
2009-04-18 | 17:15 | 13.29 | 49 | 75
2009-04-18 | 17:30 | 13.25 | 49 | 75
2009-04-18 | 17:45 | 13.21 | 49 | 75
2009-04-18 | 18:00 | 13.15 | 49 | 75
2009-04-18 | 18:15 | 13.06 | 49 | 76
2009-04-18 | 18:30 | 12.96 | 49 | 75
2009-04-18 | 18:45 | 12.82 | 49 | 75
2009-04-18 | 19:00 | 12.66 | 49 | 75
2009-04-18 | 19:15 | 12.53 | 49 | 75
2009-04-18 | 19:30 | 12.37 | 49 | 75
2009-04-18 | 19:45 | 12.22 | 49 | 75
2009-04-18 | 20:00 | 12.06 | 49 | 75
2009-04-18 | 20:15 | 11.9 | 49 | 75
2009-04-18 | 20:30 | 11.73 | 49 | 75
2009-04-18 | 20:45 | 11.57 | 49 | 75
2009-04-18 | 21:00 | 11.43 | 49 | 75
2009-04-18 | 21:15 | 11.27 | 49 | 75
2009-04-18 | 21:30 | 11.13 | 49 | 75
2009-04-18 | 21:45 | 10.99 | 49 | 75
2009-04-18 | 22:00 | 10.85 | 49 | 75
2009-04-18 | 22:15 | 10.71 | 49 | 75
2009-04-18 | 22:30 | 10.59 | 49 | 75
2009-04-18 | 22:45 | 10.47 | 49 | 75
2009-04-18 | 23:00 | 10.37 | 49 | 75
2009-04-18 | 23:15 | 10.28 | 49 | 75
2009-04-18 | 23:30 | 10.19 | 49 | 75
2009-04-18 | 23:45 | 10.12 | 48 | 75
2009-04-19 | 0:00 | 10.06 | 48 | 74
2009-04-19 | 0:15 | 10.01 | 48 | 74
2009-04-19 | 0:30 | 9.97 | 48 | 74
2009-04-19 | 0:45 | 9.93 | 48 | 74
2009-04-19 | 1:01 | 9.89 | 48 | 74
2009-04-19 | 1:16 | 9.87 | 48 | 74
2009-04-19 | 1:31 | 9.86 | 48 | 74
2009-04-19 | 1:46 | 9.84 | 48 | 74
2009-04-19 | 2:01 | 9.83 | 48 | 74
2009-04-19 | 2:16 | 9.82 | 48 | 74
2009-04-19 | 2:31 | 9.81 | 48 | 74
2009-04-19 | 2:46 | 9.81 | 48 | 73
2009-04-19 | 3:01 | 9.81 | 48 | 73
2009-04-19 | 3:16 | 9.79 | 48 | 73
2009-04-19 | 3:31 | 9.79 | 48 | 73
2009-04-19 | 3:46 | 9.78 | 48 | 73
2009-04-19 | 4:01 | 9.77 | 48 | 74
2009-04-19 | 4:16 | 9.77 | 48 | 74
2009-04-19 | 4:31 | 9.74 | 48 | 74
2009-04-19 | 4:46 | 9.73 | 48 | 74
2009-04-19 | 5:01 | 9.71 | 48 | 74
2009-04-19 | 5:16 | 9.69 | 48 | 74
2009-04-19 | 5:31 | 9.68 | 48 | 74
2009-04-19 | 5:46 | 9.65 | 48 | 74
2009-04-19 | 6:01 | 9.63 | 48 | 74
2009-04-19 | 6:16 | 9.6 | 48 | 74
2009-04-19 | 6:31 | 9.59 | 48 | 74
2009-04-19 | 6:46 | 9.56 | 48 | 74
2009-04-19 | 7:01 | 9.55 | 48 | 74
2009-04-19 | 7:16 | 9.56 | 48 | 74
2009-04-19 | 7:31 | 9.58 | 48 | 74
2009-04-19 | 7:46 | 9.61 | 48 | 74
2009-04-19 | 8:01 | 9.67 | 48 | 75
2009-04-19 | 8:16 | 9.74 | 48 | 74
2009-04-19 | 8:31 | 9.83 | 48 | 74
2009-04-19 | 8:46 | 9.93 | 48 | 75
2009-04-19 | 9:01 | 10.05 | 49 | 75
2009-04-19 | 9:16 | 10.2 | 49 | 75
2009-04-19 | 9:31 | 10.37 | 49 | 75
2009-04-19 | 9:46 | 10.56 | 49 | 75
2009-04-19 | 10:01 | 10.78 | 49 | 75
2009-04-19 | 10:16 | 10.98 | 49 | 75
2009-04-19 | 10:31 | 11.2 | 49 | 75
2009-04-19 | 10:46 | 11.44 | 49 | 75
2009-04-19 | 11:01 | 11.69 | 49 | 75
2009-04-19 | 11:16 | 11.94 | 49 | 75
2009-04-19 | 11:31 | 12.21 | 49 | 75
2009-04-19 | 11:46 | 12.47 | 49 | 75
2009-04-19 | 12:01 | 12.74 | 49 | 75
2009-04-19 | 12:16 | 13 | 49 | 75
2009-04-19 | 12:31 | 13.27 | 49 | 76
2009-04-19 | 12:46 | 13.54 | 49 | 76
2009-04-19 | 13:00 | 13.79 | 49 | 76
2009-04-19 | 13:15 | 14.05 | 49 | 76
2009-04-19 | 13:30 | 14.3 | 49 | 76
2009-04-19 | 13:45 | 14.51 | 49 | 76
2009-04-19 | 14:00 | 14.74 | 49 | 76
2009-04-19 | 14:15 | 14.94 | 49 | 76
2009-04-19 | 14:30 | 15.14 | 50 | 76
2009-04-19 | 14:45 | 15.32 | 49 | 76
2009-04-19 | 15:00 | 15.47 | 49 | 76
2009-04-19 | 15:15 | 15.61 | 49 | 76
2009-04-19 | 15:30 | 15.73 | 49 | 76
2009-04-19 | 15:45 | 15.83 | 50 | 76
2009-04-19 | 16:00 | 15.9 | 49 | 76
2009-04-19 | 16:15 | 15.96 | 50 | 76
2009-04-19 | 16:30 | 16 | 50 | 76
2009-04-19 | 16:45 | 16.02 | 50 | 76
2009-04-19 | 17:00 | 16.03 | 50 | 76
2009-04-19 | 17:15 | 16.02 | 50 | 76
2009-04-19 | 17:30 | 16 | 50 | 77
2009-04-19 | 17:45 | 15.95 | 50 | 76
2009-04-19 | 18:00 | 15.88 | 50 | 77
2009-04-19 | 18:15 | 15.81 | 50 | 77
2009-04-19 | 18:30 | 15.71 | 50 | 77
2009-04-19 | 18:45 | 15.59 | 50 | 77
2009-04-19 | 19:00 | 15.47 | 50 | 77
2009-04-19 | 19:15 | 15.33 | 50 | 77
2009-04-19 | 19:30 | 15.2 | 50 | 77
2009-04-19 | 19:45 | 15.09 | 50 | 77
2009-04-19 | 20:00 | 14.96 | 50 | 77
2009-04-19 | 20:15 | 14.82 | 50 | 77
2009-04-19 | 20:30 | 14.66 | 50 | 78
2009-04-19 | 20:45 | 14.52 | 51 | 78
2009-04-19 | 21:00 | 14.37 | 51 | 78
2009-04-19 | 21:15 | 14.22 | 51 | 78
2009-04-19 | 21:30 | 14.08 | 51 | 78
2009-04-19 | 21:45 | 13.93 | 51 | 78
2009-04-19 | 22:00 | 13.77 | 51 | 79
2009-04-19 | 22:15 | 13.61 | 51 | 78
2009-04-19 | 22:30 | 13.45 | 51 | 79
2009-04-19 | 22:45 | 13.29 | 51 | 79
2009-04-19 | 23:00 | 13.15 | 51 | 79
2009-04-19 | 23:15 | 13 | 51 | 79
2009-04-19 | 23:30 | 12.86 | 51 | 79
2009-04-19 | 23:45 | 12.71 | 51 | 79
2009-04-20 | 0:00 | 12.57 | 51 | 79
2009-04-20 | 0:15 | 12.43 | 51 | 79
2009-04-20 | 0:30 | 12.29 | 51 | 79
2009-04-20 | 0:45 | 12.16 | 51 | 79
2009-04-20 | 1:01 | 12.04 | 52 | 79
2009-04-20 | 1:16 | 11.93 | 51 | 79
2009-04-20 | 1:31 | 11.82 | 51 | 79
2009-04-20 | 1:46 | 11.71 | 51 | 79
2009-04-20 | 2:01 | 11.61 | 51 | 79
2009-04-20 | 2:16 | 11.5 | 51 | 78
2009-04-20 | 2:31 | 11.41 | 51 | 78
2009-04-20 | 2:46 | 11.31 | 51 | 78
2009-04-20 | 3:01 | 11.22 | 51 | 78
2009-04-20 | 3:16 | 11.13 | 51 | 78
2009-04-20 | 3:31 | 11.05 | 51 | 78
2009-04-20 | 3:46 | 10.97 | 51 | 78
2009-04-20 | 4:01 | 10.9 | 51 | 78
2009-04-20 | 4:16 | 10.84 | 51 | 78
2009-04-20 | 4:31 | 10.78 | 50 | 78
2009-04-20 | 4:46 | 10.74 | 50 | 78
2009-04-20 | 5:01 | 10.7 | 50 | 78
2009-04-20 | 5:16 | 10.68 | 51 | 78
2009-04-20 | 5:31 | 10.66 | 51 | 78
2009-04-20 | 5:46 | 10.66 | 50 | 77
2009-04-20 | 6:01 | 10.65 | 50 | 78
2009-04-20 | 6:16 | 10.64 | 50 | 77
2009-04-20 | 6:31 | 10.67 | 50 | 77
2009-04-20 | 6:46 | 10.68 | 50 | 77
2009-04-20 | 7:01 | 10.72 | 50 | 77
2009-04-20 | 7:16 | 10.79 | 50 | 77
2009-04-20 | 7:31 | 10.87 | 50 | 77
2009-04-20 | 7:46 | 10.98 | 50 | 77
2009-04-20 | 8:01 | 11.1 | 50 | 77
2009-04-20 | 8:16 | 11.24 | 50 | 77
2009-04-20 | 8:31 | 11.4 | 50 | 77
2009-04-20 | 8:46 | 11.57 | 50 | 77
2009-04-20 | 9:01 | 11.75 | 50 | 76
2009-04-20 | 9:16 | 11.95 | 50 | 76
2009-04-20 | 9:31 | 12.19 | 50 | 76
2009-04-20 | 9:46 | 12.44 | 49 | 76
2009-04-20 | 10:01 | 12.71 | 49 | 76
2009-04-20 | 10:16 | 12.98 | 49 | 76
2009-04-20 | 10:31 | 13.25 | 49 | 76
2009-04-20 | 10:46 | 13.54 | 49 | 76
2009-04-20 | 11:01 | 13.83 | 49 | 76
2009-04-20 | 11:16 | 14.13 | 49 | 76
2009-04-20 | 11:31 | 14.42 | 49 | 76
2009-04-20 | 11:46 | 14.72 | 49 | 76
2009-04-20 | 12:01 | 15.02 | 49 | 76
2009-04-20 | 12:16 | 15.31 | 49 | 76
2009-04-20 | 12:31 | 15.58 | 49 | 76
2009-04-20 | 12:46 | 15.86 | 49 | 76
2009-04-20 | 13:00 | 16.13 | 49 | 76
2009-04-20 | 13:15 | 16.37 | 49 | 76
2009-04-20 | 13:30 | 16.61 | 49 | 76
2009-04-20 | 13:45 | 16.84 | 49 | 76
2009-04-20 | 14:00 | 17.03 | 49 | 76
2009-04-20 | 14:15 | 17.23 | 49 | 76
2009-04-20 | 14:30 | 17.41 | 49 | 76
2009-04-20 | 14:45 | 17.55 | 49 | 76
2009-04-20 | 15:00 | 17.69 | 49 | 76
2009-04-20 | 15:15 | 17.81 | 49 | 76
2009-04-20 | 15:30 | 17.9 | 49 | 76
2009-04-20 | 15:45 | 17.98 | 49 | 76
2009-04-20 | 16:00 | 18.04 | 49 | 76
2009-04-20 | 16:15 | 18.09 | 49 | 76
2009-04-20 | 16:30 | 18.11 | 49 | 76
2009-04-20 | 16:45 | 18.12 | 49 | 76
2009-04-20 | 17:00 | 18.13 | 49 | 76
2009-04-20 | 17:15 | 18.11 | 49 | 76
2009-04-20 | 17:30 | 18.07 | 49 | 76
2009-04-20 | 17:45 | 18.02 | 50 | 76
2009-04-20 | 18:00 | 17.95 | 49 | 76
2009-04-20 | 18:15 | 17.86 | 50 | 76
2009-04-20 | 18:30 | 17.76 | 50 | 76
2009-04-20 | 18:45 | 17.65 | 50 | 76
2009-04-20 | 19:00 | 17.52 | 50 | 76
2009-04-20 | 19:15 | 17.38 | 50 | 76
2009-04-20 | 19:30 | 17.25 | 50 | 77
2009-04-20 | 19:45 | 17.12 | 50 | 77
2009-04-20 | 20:00 | 16.99 | 50 | 77
2009-04-20 | 20:15 | 16.84 | 50 | 77
2009-04-20 | 20:30 | 16.69 | 50 | 77
2009-04-20 | 20:45 | 16.54 | 50 | 77
2009-04-20 | 21:00 | 16.39 | 50 | 77
2009-04-20 | 21:15 | 16.24 | 50 | 77
2009-04-20 | 21:30 | 16.1 | 50 | 78
2009-04-20 | 21:45 | 15.95 | 51 | 78
2009-04-20 | 22:00 | 15.81 | 50 | 78
2009-04-20 | 22:15 | 15.65 | 50 | 78
2009-04-20 | 22:30 | 15.5 | 51 | 78
2009-04-20 | 22:45 | 15.36 | 51 | 78
2009-04-20 | 23:00 | 15.21 | 51 | 78
2009-04-20 | 23:15 | 15.06 | 51 | 78
2009-04-20 | 23:30 | 14.92 | 51 | 78
2009-04-20 | 23:45 | 14.79 | 51 | 78
2009-04-21 | 0:00 | 14.66 | 50 | 78
2009-04-21 | 0:15 | 14.53 | 50 | 78
2009-04-21 | 0:30 | 14.42 | 50 | 77
2009-04-21 | 0:45 | 14.3 | 50 | 77
2009-04-21 | 1:01 | 14.19 | 50 | 77
2009-04-21 | 1:16 | 14.08 | 50 | 77
2009-04-21 | 1:31 | 13.98 | 50 | 77
2009-04-21 | 1:46 | 13.89 | 50 | 77
2009-04-21 | 2:01 | 13.79 | 50 | 77
2009-04-21 | 2:16 | 13.7 | 50 | 76
2009-04-21 | 2:31 | 13.62 | 49 | 76
2009-04-21 | 2:46 | 13.55 | 50 | 77
2009-04-21 | 3:01 | 13.48 | 49 | 76
2009-04-21 | 3:16 | 13.42 | 49 | 76
2009-04-21 | 3:31 | 13.37 | 49 | 76
2009-04-21 | 3:46 | 13.32 | 49 | 76
2009-04-21 | 4:01 | 13.28 | 49 | 76
2009-04-21 | 4:16 | 13.24 | 49 | 76
2009-04-21 | 4:31 | 13.2 | 49 | 76
2009-04-21 | 4:46 | 13.17 | 49 | 76
2009-04-21 | 5:01 | 13.15 | 49 | 76
2009-04-21 | 5:16 | 13.13 | 49 | 75
2009-04-21 | 5:31 | 13.11 | 49 | 75
2009-04-21 | 5:46 | 13.09 | 49 | 75
2009-04-21 | 6:01 | 13.08 | 49 | 75
2009-04-21 | 6:16 | 13.07 | 49 | 75
2009-04-21 | 6:31 | 13.07 | 49 | 75
2009-04-21 | 6:46 | 13.06 | 49 | 75
2009-04-21 | 7:01 | 13.1 | 49 | 75
2009-04-21 | 7:16 | 13.13 | 49 | 75
2009-04-21 | 7:31 | 13.18 | 49 | 75
2009-04-21 | 7:46 | 13.26 | 49 | 75
2009-04-21 | 8:01 | 13.35 | 49 | 75
2009-04-21 | 8:16 | 13.46 | 49 | 75
2009-04-21 | 8:31 | 13.58 | 49 | 75
2009-04-21 | 8:46 | 13.72 | 49 | 75
2009-04-21 | 9:01 | 13.89 | 49 | 75
2009-04-21 | 9:16 | 14.06 | 49 | 75
2009-04-21 | 9:31 | 14.26 | 49 | 75
2009-04-21 | 9:46 | 14.48 | 49 | 75
2009-04-21 | 10:01 | 14.7 | 49 | 75
2009-04-21 | 10:16 | 14.94 | 49 | 75
2009-04-21 | 10:31 | 15.2 | 48 | 75
2009-04-21 | 10:46 | 15.45 | 49 | 75
2009-04-21 | 11:01 | 15.72 | 49 | 75
2009-04-21 | 11:16 | 15.99 | 49 | 75
2009-04-21 | 11:31 | 16.27 | 48 | 74
2009-04-21 | 11:46 | 16.54 | 48 | 75
2009-04-21 | 12:01 | 16.82 | 49 | 75
2009-04-21 | 12:16 | 17.1 | 49 | 75
2009-04-21 | 12:31 | 17.36 | 49 | 75
2009-04-21 | 12:46 | 17.62 | 49 | 75
2009-04-21 | 13:00 | 17.87 | 49 | 75
2009-04-21 | 13:15 | 18.12 | 49 | 75
2009-04-21 | 13:30 | 18.35 | 49 | 75
2009-04-21 | 13:45 | 18.55 | 49 | 75
2009-04-21 | 14:00 | 18.75 | 49 | 75
2009-04-21 | 14:15 | 18.92 | 49 | 75
2009-04-21 | 14:30 | 19.08 | 49 | 75
2009-04-21 | 14:45 | 19.21 | 49 | 75
2009-04-21 | 15:00 | 19.31 | 49 | 75
2009-04-21 | 15:15 | 19.45 | 49 | 75
2009-04-21 | 15:30 | 19.52 | 49 | 75
2009-04-21 | 15:45 | 19.59 | 49 | 75
2009-04-21 | 16:00 | 19.62 | 49 | 75
2009-04-21 | 16:15 | 19.61 | 49 | 76
2009-04-21 | 16:30 | 19.61 | 49 | 75
2009-04-21 | 16:45 | 19.63 | 49 | 76
2009-04-21 | 17:00 | 19.6 | 49 | 76
2009-04-21 | 17:15 | 19.56 | 49 | 76
2009-04-21 | 17:30 | 19.5 | 49 | 76
2009-04-21 | 17:45 | 19.43 | 49 | 76
2009-04-21 | 18:00 | 19.31 | 49 | 76
2009-04-21 | 18:15 | 19.19 | 49 | 76
2009-04-21 | 18:30 | 19.04 | 49 | 76
2009-04-21 | 18:45 | 18.89 | 49 | 76
2009-04-21 | 19:00 | 18.73 | 49 | 76
2009-04-21 | 19:15 | 18.61 | 50 | 76
2009-04-21 | 19:30 | 18.47 | 50 | 76
2009-04-21 | 19:45 | 18.3 | 50 | 76
2009-04-21 | 20:00 | 18.15 | 50 | 76
2009-04-21 | 20:15 | 18.02 | 50 | 77
2009-04-21 | 20:30 | 17.86 | 50 | 77
2009-04-21 | 20:45 | 17.69 | 50 | 77
2009-04-21 | 21:00 | 17.51 | 50 | 77
2009-04-21 | 21:15 | 17.35 | 50 | 77
2009-04-21 | 21:30 | 17.19 | 50 | 77
2009-04-21 | 21:45 | 17.02 | 50 | 77
2009-04-21 | 22:00 | 16.86 | 50 | 77
2009-04-21 | 22:15 | 16.69 | 50 | 77
2009-04-21 | 22:30 | 16.51 | 50 | 77
2009-04-21 | 22:45 | 16.34 | 50 | 77
2009-04-21 | 23:00 | 16.16 | 50 | 77
2009-04-21 | 23:15 | 15.99 | 50 | 77
2009-04-21 | 23:30 | 15.83 | 50 | 77
2009-04-21 | 23:45 | 15.67 | 50 | 77
2009-04-22 | 0:00 | 15.51 | 50 | 77
2009-04-22 | 0:15 | 15.35 | 50 | 77
2009-04-22 | 0:30 | 15.21 | 50 | 77
2009-04-22 | 0:45 | 15.07 | 50 | 77
2009-04-22 | 1:01 | 14.94 | 50 | 76
2009-04-22 | 1:16 | 14.81 | 50 | 76
2009-04-22 | 1:31 | 14.68 | 49 | 76
2009-04-22 | 1:46 | 14.57 | 49 | 76
2009-04-22 | 2:01 | 14.45 | 49 | 76
2009-04-22 | 2:16 | 14.35 | 50 | 76
2009-04-22 | 2:31 | 14.25 | 49 | 75
2009-04-22 | 2:46 | 14.16 | 49 | 75
2009-04-22 | 3:01 | 14.08 | 49 | 75
2009-04-22 | 3:16 | 14.02 | 49 | 75
2009-04-22 | 3:31 | 13.95 | 49 | 75
2009-04-22 | 3:46 | 13.88 | 49 | 75
2009-04-22 | 4:01 | 13.83 | 49 | 75
2009-04-22 | 4:16 | 13.78 | 49 | 75
2009-04-22 | 4:31 | 13.75 | 48 | 75
2009-04-22 | 4:46 | 13.7 | 48 | 75
2009-04-22 | 5:01 | 13.66 | 48 | 74
2009-04-22 | 5:16 | 13.63 | 48 | 74
2009-04-22 | 5:31 | 13.59 | 48 | 74
2009-04-22 | 5:46 | 13.57 | 48 | 74
2009-04-22 | 6:01 | 13.56 | 48 | 74
2009-04-22 | 6:16 | 13.55 | 48 | 74
2009-04-22 | 6:31 | 13.53 | 48 | 74
2009-04-22 | 6:46 | 13.52 | 48 | 74
2009-04-22 | 7:01 | 13.53 | 48 | 74
2009-04-22 | 7:16 | 13.56 | 48 | 74
2009-04-22 | 7:31 | 13.61 | 48 | 74
2009-04-22 | 7:46 | 13.67 | 48 | 74
2009-04-22 | 8:01 | 13.75 | 48 | 74
2009-04-22 | 8:16 | 13.85 | 48 | 73
2009-04-22 | 8:31 | 13.96 | 48 | 73
2009-04-22 | 8:46 | 14.09 | 48 | 73
2009-04-22 | 9:01 | 14.24 | 47 | 73
2009-04-22 | 9:16 | 14.39 | 47 | 73
2009-04-22 | 9:31 | 14.57 | 47 | 73
2009-04-22 | 9:46 | 14.73 | 47 | 73
2009-04-22 | 10:01 | 14.96 | 47 | 73
2009-04-22 | 10:16 | 15.2 | 47 | 73
2009-04-22 | 10:31 | 15.46 | 47 | 73
2009-04-22 | 10:46 | 15.75 | 47 | 73
2009-04-22 | 11:01 | 15.94 | 47 | 73
2009-04-22 | 11:16 | 16.18 | 47 | 72
2009-04-22 | 11:31 | 16.36 | 47 | 72
2009-04-22 | 11:46 | 16.48 | 47 | 73
2009-04-22 | 12:01 | 16.62 | 47 | 73
2009-04-22 | 12:16 | 16.8 | 47 | 73
2009-04-22 | 12:31 | 16.87 | 47 | 73
2009-04-22 | 12:46 | 16.9 | 47 | 73
2009-04-22 | 13:00 | 16.87 | 47 | 73
2009-04-22 | 13:15 | 16.88 | 47 | 73
2009-04-22 | 13:30 | 16.98 | 48 | 73
2009-04-22 | 13:45 | 17.09 | 47 | 73
2009-04-22 | 14:00 | 17.07 | 47 | 73
2009-04-22 | 14:15 | 17.07 | 47 | 73
2009-04-22 | 14:30 | 17.03 | 47 | 73
2009-04-22 | 14:45 | 17.01 | 47 | 73
2009-04-22 | 15:00 | 16.93 | 47 | 73
2009-04-22 | 15:15 | 16.8 | 47 | 73
2009-04-22 | 15:30 | 16.7 | 47 | 73
2009-04-22 | 15:45 | 16.8 | 47 | 73
2009-04-22 | 16:00 | 16.95 | 47 | 73
2009-04-22 | 16:15 | 17.03 | 47 | 73
2009-04-22 | 16:30 | 16.92 | 48 | 73
2009-04-22 | 16:45 | 16.87 | 48 | 74
2009-04-22 | 17:00 | 16.9 | 48 | 73
2009-04-22 | 17:15 | 16.89 | 48 | 73
2009-04-22 | 17:30 | 16.83 | 48 | 73
2009-04-22 | 17:45 | 16.75 | 48 | 73
2009-04-22 | 18:00 | 16.67 | 48 | 73
2009-04-22 | 18:15 | 16.59 | 48 | 73
2009-04-22 | 18:30 | 16.5 | 47 | 73
2009-04-22 | 18:45 | 16.38 | 47 | 73
2009-04-22 | 19:00 | 16.27 | 47 | 73
2009-04-22 | 19:15 | 16.12 | 47 | 73
2009-04-22 | 19:30 | 15.97 | 47 | 73
2009-04-22 | 19:45 | 15.83 | 47 | 73
2009-04-22 | 20:00 | 15.71 | 47 | 73
2009-04-22 | 20:15 | 15.56 | 47 | 73
2009-04-22 | 20:30 | 15.39 | 47 | 73
2009-04-22 | 20:45 | 15.24 | 48 | 73
2009-04-22 | 21:00 | 15.09 | 48 | 73
2009-04-22 | 21:15 | 14.96 | 48 | 73
2009-04-22 | 21:30 | 14.86 | 48 | 73
2009-04-22 | 21:45 | 14.74 | 48 | 73
2009-04-22 | 22:00 | 14.63 | 48 | 73
2009-04-22 | 22:15 | 14.51 | 48 | 73
2009-04-22 | 22:30 | 14.4 | 48 | 74
2009-04-22 | 22:45 | 14.27 | 48 | 74
2009-04-22 | 23:00 | 14.18 | 48 | 74
2009-04-22 | 23:15 | 14.1 | 48 | 74
2009-04-22 | 23:30 | 14.02 | 48 | 74
2009-04-22 | 23:45 | 13.97 | 48 | 74
2009-04-23 | 0:00 | 13.94 | 48 | 73
2009-04-23 | 0:15 | 13.86 | 48 | 73
2009-04-23 | 0:30 | 13.81 | 48 | 73
2009-04-23 | 0:45 | 13.76 | 48 | 73
2009-04-23 | 1:01 | 13.75 | 48 | 73
2009-04-23 | 1:16 | 13.7 | 47 | 73
2009-04-23 | 1:31 | 13.63 | 47 | 73
2009-04-23 | 1:46 | 13.54 | 47 | 73
2009-04-23 | 2:01 | 13.49 | 47 | 73
2009-04-23 | 2:16 | 13.45 | 47 | 73
2009-04-23 | 2:31 | 13.42 | 47 | 73
2009-04-23 | 2:46 | 13.37 | 47 | 73
2009-04-23 | 3:01 | 13.32 | 47 | 73
2009-04-23 | 3:16 | 13.26 | 47 | 73
2009-04-23 | 3:31 | 13.19 | 47 | 73
2009-04-23 | 3:46 | 13.16 | 47 | 73
2009-04-23 | 4:01 | 13.13 | 47 | 73
2009-04-23 | 4:16 | 13.09 | 47 | 73
2009-04-23 | 4:31 | 13.03 | 47 | 73
2009-04-23 | 4:46 | 12.97 | 47 | 73
2009-04-23 | 5:01 | 12.91 | 47 | 73
2009-04-23 | 5:16 | 12.84 | 47 | 73
2009-04-23 | 5:31 | 12.78 | 47 | 73
2009-04-23 | 5:46 | 12.72 | 47 | 73
2009-04-23 | 6:01 | 12.66 | 47 | 73
2009-04-23 | 6:16 | 12.58 | 47 | 73
2009-04-23 | 6:31 | 12.51 | 47 | 73
2009-04-23 | 6:46 | 12.45 | 47 | 73
2009-04-23 | 7:01 | 12.41 | 47 | 73
2009-04-23 | 7:16 | 12.39 | 47 | 73
2009-04-23 | 7:31 | 12.36 | 47 | 73
2009-04-23 | 7:46 | 12.34 | 47 | 73
2009-04-23 | 8:01 | 12.35 | 47 | 73
2009-04-23 | 8:16 | 12.39 | 47 | 73
2009-04-23 | 8:31 | 12.45 | 47 | 73
2009-04-23 | 8:46 | 12.52 | 47 | 73
2009-04-23 | 9:01 | 12.61 | 47 | 73
2009-04-23 | 9:16 | 12.71 | 47 | 73
2009-04-23 | 9:31 | 12.82 | 47 | 73
2009-04-23 | 9:46 | 12.95 | 47 | 73
2009-04-23 | 10:01 | 13.09 | 47 | 73
2009-04-23 | 10:16 | 13.24 | 47 | 73
2009-04-23 | 10:31 | 13.4 | 47 | 73
2009-04-23 | 10:46 | 13.59 | 47 | 73
2009-04-23 | 11:01 | 13.76 | 47 | 73
2009-04-23 | 11:16 | 13.94 | 47 | 72
2009-04-23 | 11:31 | 14.14 | 47 | 72
2009-04-23 | 11:46 | 14.33 | 47 | 72
2009-04-23 | 12:01 | 14.52 | 47 | 72
2009-04-23 | 12:16 | 14.73 | 47 | 73
2009-04-23 | 12:31 | 14.83 | 47 | 72
2009-04-23 | 12:46 | 14.8 | 47 | 73
2009-04-23 | 13:00 | 14.94 | 47 | 73
2009-04-23 | 13:15 | 15.06 | 47 | 72
2009-04-23 | 13:30 | 15.27 | 47 | 72
2009-04-23 | 13:45 | 15.55 | 47 | 72
2009-04-23 | 14:00 | 15.73 | 47 | 72
2009-04-23 | 14:15 | 15.86 | 47 | 72
2009-04-23 | 14:30 | 15.94 | 47 | 72
2009-04-23 | 14:45 | 16 | 47 | 73
2009-04-23 | 15:00 | 16.11 | 47 | 73
2009-04-23 | 15:15 | 16.24 | 47 | 72
2009-04-23 | 15:30 | 16.32 | 47 | 72
2009-04-23 | 15:45 | 16.4 | 47 | 72
2009-04-23 | 16:00 | 16.51 | 47 | 72
2009-04-23 | 16:15 | 16.59 | 47 | 72
2009-04-23 | 16:30 | 16.42 | 47 | 72
2009-04-23 | 16:45 | 16.23 | 47 | 72
2009-04-23 | 17:00 | 16.16 | 47 | 72
2009-04-23 | 17:15 | 16.16 | 47 | 72
2009-04-23 | 17:30 | 16.16 | 47 | 72
2009-04-23 | 17:45 | 16.02 | 47 | 72
2009-04-23 | 18:00 | 15.8 | 47 | 72
2009-04-23 | 18:15 | 15.61 | 47 | 72
2009-04-23 | 18:30 | 15.46 | 47 | 72
2009-04-23 | 18:45 | 15.31 | 47 | 72
2009-04-23 | 19:00 | 15.11 | 47 | 72
2009-04-23 | 19:15 | 14.88 | 47 | 72
2009-04-23 | 19:30 | 14.66 | 47 | 72
2009-04-23 | 19:45 | 14.45 | 47 | 72
2009-04-23 | 20:00 | 14.25 | 47 | 72
2009-04-23 | 20:15 | 14.03 | 46 | 71
2009-04-23 | 20:30 | 13.83 | 46 | 71
2009-04-23 | 20:45 | 13.66 | 46 | 71
2009-04-23 | 21:00 | 13.52 | 46 | 71
2009-04-23 | 21:15 | 13.37 | 46 | 71
2009-04-23 | 21:30 | 13.22 | 46 | 71
2009-04-23 | 21:45 | 13.09 | 46 | 71
2009-04-23 | 22:00 | 12.94 | 46 | 71
2009-04-23 | 22:15 | 12.81 | 46 | 71
2009-04-23 | 22:30 | 12.67 | 46 | 71
2009-04-23 | 22:45 | 12.51 | 46 | 71
2009-04-23 | 23:00 | 12.37 | 46 | 70
2009-04-23 | 23:15 | 12.2 | 46 | 70
2009-04-23 | 23:30 | 12.04 | 46 | 70
2009-04-23 | 23:45 | 11.89 | 46 | 70
2009-04-24 | 0:00 | 11.75 | 45 | 70
2009-04-24 | 0:15 | 11.63 | 45 | 70
2009-04-24 | 0:30 | 11.53 | 45 | 70
2009-04-24 | 0:45 | 11.43 | 45 | 70
2009-04-24 | 1:01 | 11.34 | 45 | 69
2009-04-24 | 1:16 | 11.27 | 45 | 69
2009-04-24 | 1:31 | 11.2 | 45 | 69
2009-04-24 | 1:46 | 11.12 | 45 | 69
2009-04-24 | 2:01 | 11.06 | 45 | 69
2009-04-24 | 2:16 | 11 | 45 | 69
2009-04-24 | 2:31 | 10.95 | 45 | 69
2009-04-24 | 2:46 | 10.89 | 45 | 69
2009-04-24 | 3:01 | 10.85 | 45 | 69
2009-04-24 | 3:16 | 10.79 | 45 | 69
2009-04-24 | 3:31 | 10.77 | 45 | 69
2009-04-24 | 3:46 | 10.73 | 45 | 69
2009-04-24 | 4:01 | 10.7 | 45 | 69
2009-04-24 | 4:16 | 10.69 | 45 | 69
2009-04-24 | 4:31 | 10.66 | 45 | 69
2009-04-24 | 4:46 | 10.62 | 45 | 69
2009-04-24 | 5:01 | 10.6 | 45 | 69
2009-04-24 | 5:16 | 10.58 | 45 | 69
2009-04-24 | 5:31 | 10.54 | 45 | 69
2009-04-24 | 5:46 | 10.52 | 45 | 69
2009-04-24 | 6:01 | 10.49 | 45 | 69
2009-04-24 | 6:16 | 10.45 | 45 | 69
2009-04-24 | 6:31 | 10.43 | 45 | 68
2009-04-24 | 6:46 | 10.4 | 44 | 68
2009-04-24 | 7:01 | 10.41 | 44 | 68
2009-04-24 | 7:16 | 10.41 | 44 | 68
2009-04-24 | 7:31 | 10.43 | 44 | 68
2009-04-24 | 7:46 | 10.46 | 44 | 68
2009-04-24 | 8:01 | 10.5 | 44 | 68
2009-04-24 | 8:16 | 10.57 | 44 | 68
2009-04-24 | 8:31 | 10.65 | 44 | 68
2009-04-24 | 8:46 | 10.74 | 44 | 68
2009-04-24 | 9:01 | 10.85 | 44 | 68
2009-04-24 | 9:16 | 10.98 | 44 | 68
2009-04-24 | 9:31 | 11.11 | 44 | 68
2009-04-24 | 9:46 | 11.27 | 44 | 68
2009-04-24 | 10:01 | 11.43 | 44 | 68
2009-04-24 | 10:16 | 11.58 | 44 | 68
2009-04-24 | 10:31 | 11.75 | 44 | 68
2009-04-24 | 10:46 | 11.94 | 44 | 68
2009-04-24 | 11:01 | 12.13 | 45 | 68
2009-04-24 | 11:16 | 12.32 | 45 | 69
2009-04-24 | 11:31 | 12.52 | 45 | 69
2009-04-24 | 11:46 | 12.72 | 45 | 69
2009-04-24 | 12:01 | 12.93 | 45 | 69
2009-04-24 | 12:16 | 13.12 | 45 | 69
2009-04-24 | 12:31 | 13.3 | 45 | 69
2009-04-24 | 12:46 | 13.5 | 45 | 69
2009-04-24 | 13:00 | 13.68 | 45 | 69
2009-04-24 | 13:15 | 13.86 | 45 | 69
2009-04-24 | 13:30 | 14.05 | 45 | 69
2009-04-24 | 13:45 | 14.2 | 45 | 69
2009-04-24 | 14:00 | 14.37 | 45 | 69
2009-04-24 | 14:15 | 14.52 | 45 | 69
2009-04-24 | 14:30 | 14.65 | 45 | 69
2009-04-24 | 14:45 | 14.76 | 45 | 69
2009-04-24 | 15:00 | 14.86 | 45 | 70
2009-04-24 | 15:15 | 14.96 | 45 | 69
2009-04-24 | 15:30 | 15.03 | 45 | 69
2009-04-24 | 15:45 | 15.09 | 45 | 70
2009-04-24 | 16:00 | 15.13 | 45 | 70
2009-04-24 | 16:15 | 15.16 | 45 | 70
2009-04-24 | 16:30 | 15.18 | 45 | 70
2009-04-24 | 16:45 | 15.18 | 45 | 70
2009-04-24 | 17:00 | 15.16 | 45 | 70
2009-04-24 | 17:15 | 15.12 | 46 | 70
2009-04-24 | 17:30 | 15.06 | 46 | 70
2009-04-24 | 17:45 | 14.99 | 46 | 70
2009-04-24 | 18:00 | 14.9 | 46 | 70
2009-04-24 | 18:15 | 14.78 | 46 | 70
2009-04-24 | 18:30 | 14.67 | 46 | 70
2009-04-24 | 18:45 | 14.54 | 46 | 70
2009-04-24 | 19:00 | 14.36 | 46 | 70
2009-04-24 | 19:15 | 14.2 | 46 | 71
2009-04-24 | 19:30 | 14.05 | 46 | 70
2009-04-24 | 19:45 | 13.89 | 46 | 71
2009-04-24 | 20:00 | 13.74 | 46 | 71
2009-04-24 | 20:15 | 13.57 | 46 | 71
2009-04-24 | 20:30 | 13.39 | 46 | 71
2009-04-24 | 20:45 | 13.22 | 46 | 71
2009-04-24 | 21:00 | 13.05 | 46 | 71
2009-04-24 | 21:15 | 12.89 | 46 | 71
2009-04-24 | 21:30 | 12.72 | 46 | 71
2009-04-24 | 21:45 | 12.56 | 46 | 71
2009-04-24 | 22:00 | 12.39 | 46 | 71
2009-04-24 | 22:15 | 12.23 | 46 | 71
2009-04-24 | 22:30 | 12.06 | 46 | 71
2009-04-24 | 22:45 | 11.9 | 46 | 71
2009-04-24 | 23:00 | 11.73 | 46 | 71
2009-04-24 | 23:15 | 11.58 | 46 | 71
2009-04-24 | 23:30 | 11.45 | 46 | 71
2009-04-24 | 23:45 | 11.31 | 46 | 71
2009-04-25 | 0:00 | 11.17 | 46 | 71
2009-04-25 | 0:15 | 11.04 | 46 | 71
2009-04-25 | 0:30 | 10.92 | 46 | 71
2009-04-25 | 0:45 | 10.81 | 46 | 71
2009-04-25 | 1:01 | 10.71 | 46 | 70
2009-04-25 | 1:16 | 10.61 | 46 | 70
2009-04-25 | 1:31 | 10.5 | 46 | 70
2009-04-25 | 1:46 | 10.42 | 45 | 70
2009-04-25 | 2:01 | 10.35 | 45 | 70
2009-04-25 | 2:16 | 10.29 | 45 | 70
2009-04-25 | 2:31 | 10.23 | 45 | 70
2009-04-25 | 2:46 | 10.18 | 45 | 70
2009-04-25 | 3:01 | 10.14 | 45 | 69
2009-04-25 | 3:16 | 10.1 | 45 | 69
2009-04-25 | 3:31 | 10.09 | 45 | 69
2009-04-25 | 3:46 | 10.07 | 45 | 69
2009-04-25 | 4:01 | 10.06 | 45 | 69
2009-04-25 | 4:16 | 10.04 | 45 | 69
2009-04-25 | 4:31 | 10.04 | 45 | 69
2009-04-25 | 4:46 | 10.01 | 45 | 69
2009-04-25 | 5:01 | 10 | 45 | 69
2009-04-25 | 5:16 | 9.98 | 45 | 69
2009-04-25 | 5:31 | 10 | 45 | 69
2009-04-25 | 5:46 | 10.01 | 45 | 69
2009-04-25 | 6:01 | 9.99 | 45 | 69
2009-04-25 | 6:16 | 10.01 | 45 | 69
2009-04-25 | 6:31 | 10.02 | 45 | 68
2009-04-25 | 6:46 | 10.04 | 44 | 68
2009-04-25 | 7:01 | 10.07 | 44 | 68
2009-04-25 | 7:16 | 10.12 | 44 | 68
2009-04-25 | 7:31 | 10.17 | 44 | 68
2009-04-25 | 7:46 | 10.21 | 44 | 68
2009-04-25 | 8:01 | 10.27 | 44 | 68
2009-04-25 | 8:16 | 10.34 | 44 | 68
2009-04-25 | 8:31 | 10.45 | 44 | 68
2009-04-25 | 8:46 | 10.53 | 44 | 68
2009-04-25 | 9:01 | 10.55 | 44 | 68
2009-04-25 | 9:16 | 10.64 | 44 | 68
2009-04-25 | 9:31 | 10.81 | 44 | 67
2009-04-25 | 9:46 | 10.92 | 44 | 67
2009-04-25 | 10:01 | 10.96 | 44 | 67
2009-04-25 | 10:16 | 11.01 | 44 | 67
2009-04-25 | 10:31 | 11.09 | 44 | 67
2009-04-25 | 10:46 | 11.15 | 44 | 67
2009-04-25 | 11:01 | 11.27 | 44 | 67
2009-04-25 | 11:16 | 11.37 | 44 | 67
2009-04-25 | 11:31 | 11.4 | 44 | 67
2009-04-25 | 11:46 | 11.53 | 44 | 67
2009-04-25 | 12:01 | 11.79 | 44 | 67
2009-04-25 | 12:16 | 12.06 | 44 | 67
2009-04-25 | 12:31 | 12.09 | 44 | 67
2009-04-25 | 12:46 | 11.92 | 44 | 67
2009-04-25 | 13:00 | 11.87 | 44 | 67