## cont multiparameter audit
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, City Park Gauge
Contact Name | Mike Logan |  | Latitude
Number | 541-322-3135 |  | Longitude
Pre-sampling Field Equipment Check
Field equipment ID | Sonde AA (Audited against Pro-Series (loner), NIST, & Sonde AB)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-07-08 |  | 2009-07-08 |  | A
Time | Computer set to atomic time |  | 14:30 |  | 14:30 |  | A
Temp | VWR NIST traceable surface water |  | 9.2 |  | 9.39 |  | A
DO | In Situ |  | 10.91 |  | 11.04 |  | A
SpC | In Situ |  | 30.6 |  | 30 |  | A
Ph | In Situ |  | 7.44 |  | 7.68 |  | A
Field Equipment Audit Results
Audit # | Date/time | Temp | pH | DO | %SAT | SpC | TDS
Audit 1 | 2009-07-08 17:45 | 11.2 | 7.43 | 10.17 | 92.8 | 25.5 | 16.9
Sonde | 2009-07-08 17:45 | 11.25 | 7.59 | 10.39 | 94.8 | 25 | 16
Grade |  | A | A | A |  | A
Audit 2 | 2009-07-15 12:51 | 13.1 | 7.49 | 9.95 | 94.6 | 25.1 | 16.3
Sonde | 2009-07-15 12:45 | 12.89 | 7.65 | 11.16 | 105.6 | 24 | 16
Grade |  | A | A | Fail |  | A
Audit 3
Sonde
Grade
Post-sampling Field Equipment Check
Field equipment ID | Sonde AA (Audited against Pro-Series (loner), NIST, & Titration)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-07-15 |  | 2009-07-15 |  | A
Time | Computer set to atomic time |  | 16:03 |  | 16:03 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17.44 |  | A
DO | Winkler titration of surface water |  | 8.8 |  | 9.4 |  | B
SpC | In Situ |  | 26.3 |  | 25 |  | A
Ph | In Situ |  | 7.48 |  | 7.7 |  | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
Considering DO data doesn't indicate any standards were violated and it passed all but one audit, data will be kept |  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 1
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, City Park Gauge
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-07-07 |  | 2009-07-07 |  | A
Time | Computer set to atomic time |  | 11:58 |  | 11:58 |  | A
Temp | VWR NIST traceable surface water |  | 14.7 |  | 14.9 |  | A
DO | Winkler titration of surface water |  | 9.3 |  | 9.14 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 177.1 |  | B
Ph | pH 7 standard solution |  | 7.02 |  | 7.06 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-07-08 17:45 | 11.2 | 7.43 | 10.17 | 92.8 | 25.5 | 16.9
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-07-09 |  | 2009-07-09 |  | A
Time | Computer set to atomic time |  | 11:57 |  | 11:57 |  | A
Temp | VWR NIST traceable surface water |  | 14.9 |  | 14.7 |  | A
DO | Winkler titration of surface water |  | 9.4 |  | 8.49 |  | B
SpC | 148 µS/mL |  | 156.7 |  | 165.2 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## audit 2
Organization | Upper Deschutes Watershed Council |  | Station ID |  | WC 024.25
Program Name | Water Quality Monitoring Program |  | Description |  | Whychus Creek, City Park Gauge
Contact Name | Mike Logan |  | Latitude |  | 0
Number | 541-322-3135 |  | Longitude |  | 0
Pre-sampling Field Equipment Check
Field equipment ID | Pro-Series (Loner)
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-07-09 |  | 2009-07-09 |  | A
Time | Computer set to atomic time |  | 11:57 |  | 11:57 |  | A
Temp | VWR NIST traceable surface water |  | 14.9 |  | 14.7 |  | A
DO | Winkler titration of surface water |  | 9.4 |  | 8.49 |  | B
SpC | 148 µS/mL |  | 156.7 |  | 165.2 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.08 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Field Equipment Grab Sampling Results
| Date/time | Temp | pH | DO | %SAT | SpC | TDS | Turb
| 2009-07-15 12:51 | 13.1 | 7.49 | 9.95 | 94.6 | 25.1 | 16.3
Post-sampling Field Equipment Check
Parameter | Audit reference |  | Audit value |  | Field equip. value |  | Grade
Date | Computer set to atomic time |  | 2009-07-15 |  | 2009-07-15 |  | A
Time | Computer set to atomic time |  | 16:04 |  | 16:04 |  | A
Temp | VWR NIST traceable surface water |  | 17.2 |  | 17.4 |  | A
DO | Winkler titration of surface water |  | 8.8 |  | 8.69 |  | A
SpC | 148 µS/mL |  | 156.7 |  | 156 |  | A
Ph | pH 7 standard solution |  | 7.02 |  | 7.07 |  | A
Turb | Gelex secondary 0-5 NTU | [5 empty] | A
Turb | Gelex secondary 0-500 NTU | [5 empty] | A
Turb | Gelex secondary 0-1000 NTU | [5 empty] | A
Comments: |  |  | Grade | Temp | DO | SpC | pH | Turb
|  |  |  | C | mg/L | µS/mL | SU | NTU
|  |  | A | <1.50 | <=0.3 | <=10% | <=0.30 | <=1.0
|  |  | B | 1.51 to 2.00 | 0.31 to 1.0 | 10.1% to 15% | 0.31 to 0.50 | 1.1 to 2.0
|  |  | Fail | >2.01 | >1.1 | >15.1% | >0.51 | >2.1

## Data
Date | Time | Temp | SpCond | TDS | DO Conc | DO% | pH
M/D/Y |  | °C | uS/cm | mg/L | mg/L | %
2009-07-08 | 17:45 | 11.25 | 25 | 16 | 10.39 | 94.8 | 7.59
2009-07-08 | 18:00 | 11 | 25 | 16 | 10.63 | 96.4 | 7.64
2009-07-08 | 18:15 | 10.89 | 25 | 16 | 10.68 | 96.6 | 7.64
2009-07-08 | 18:30 | 10.84 | 25 | 16 | 10.66 | 96.3 | 7.64
2009-07-08 | 18:45 | 10.84 | 25 | 16 | 10.69 | 96.6 | 7.64
2009-07-08 | 19:00 | 10.85 | 24 | 16 | 10.67 | 96.5 | 7.65
2009-07-08 | 19:15 | 10.84 | 24 | 16 | 10.68 | 96.5 | 7.65
2009-07-08 | 19:30 | 10.8 | 24 | 16 | 10.68 | 96.4 | 7.66
2009-07-08 | 19:45 | 10.76 | 24 | 16 | 10.67 | 96.3 | 7.65
2009-07-08 | 20:00 | 10.76 | 24 | 16 | 10.66 | 96.1 | 7.65
2009-07-08 | 20:15 | 10.77 | 24 | 16 | 10.64 | 96 | 7.65
2009-07-08 | 20:30 | 10.78 | 24 | 16 | 10.63 | 95.9 | 7.64
2009-07-08 | 20:45 | 10.76 | 24 | 16 | 10.66 | 96.1 | 7.65
2009-07-08 | 21:00 | 10.69 | 24 | 16 | 10.65 | 95.9 | 7.64
2009-07-08 | 21:15 | 10.61 | 24 | 16 | 10.71 | 96.2 | 7.64
2009-07-08 | 21:30 | 10.53 | 24 | 16 | 10.72 | 96.1 | 7.64
2009-07-08 | 21:45 | 10.47 | 24 | 16 | 10.75 | 96.2 | 7.62
2009-07-08 | 22:00 | 10.41 | 24 | 16 | 10.71 | 95.8 | 7.63
2009-07-08 | 22:15 | 10.37 | 24 | 16 | 10.71 | 95.7 | 7.62
2009-07-08 | 22:30 | 10.32 | 24 | 16 | 10.7 | 95.5 | 7.62
2009-07-08 | 22:45 | 10.25 | 24 | 16 | 10.77 | 96 | 7.61
2009-07-08 | 23:00 | 10.16 | 24 | 16 | 10.81 | 96.2 | 7.6
2009-07-08 | 23:15 | 10.04 | 24 | 16 | 10.81 | 95.9 | 7.59
2009-07-08 | 23:30 | 9.9 | 24 | 16 | 10.88 | 96.2 | 7.61
2009-07-08 | 23:45 | 9.76 | 24 | 16 | 10.95 | 96.5 | 7.6
2009-07-09 | 0:00 | 9.62 | 24 | 16 | 10.99 | 96.5 | 7.59
2009-07-09 | 0:15 | 9.49 | 24 | 16 | 11.03 | 96.6 | 7.58
2009-07-09 | 0:30 | 9.37 | 24 | 16 | 11.02 | 96.2 | 7.59
2009-07-09 | 0:45 | 9.22 | 24 | 16 | 11.09 | 96.4 | 7.58
2009-07-09 | 1:01 | 9.07 | 24 | 16 | 11.19 | 96.9 | 7.57
2009-07-09 | 1:16 | 8.91 | 24 | 16 | 11.2 | 96.6 | 7.56
2009-07-09 | 1:31 | 8.75 | 24 | 16 | 11.2 | 96.3 | 7.56
2009-07-09 | 1:46 | 8.58 | 24 | 16 | 11.25 | 96.4 | 7.54
2009-07-09 | 2:01 | 8.41 | 24 | 16 | 11.31 | 96.4 | 7.53
2009-07-09 | 2:16 | 8.25 | 24 | 16 | 11.43 | 97.1 | 7.54
2009-07-09 | 2:31 | 8.09 | 24 | 16 | 11.48 | 97.2 | 7.52
2009-07-09 | 2:46 | 7.95 | 24 | 16 | 11.57 | 97.6 | 7.52
2009-07-09 | 3:01 | 7.81 | 24 | 16 | 11.55 | 97.1 | 7.52
2009-07-09 | 3:16 | 7.67 | 24 | 16 | 11.64 | 97.5 | 7.52
2009-07-09 | 3:31 | 7.54 | 24 | 16 | 11.69 | 97.6 | 7.51
2009-07-09 | 3:46 | 7.43 | 24 | 16 | 11.73 | 97.6 | 7.51
2009-07-09 | 4:01 | 7.31 | 24 | 16 | 11.79 | 97.8 | 7.51
2009-07-09 | 4:16 | 7.2 | 24 | 16 | 11.76 | 97.4 | 7.51
2009-07-09 | 4:31 | 7.09 | 24 | 16 | 11.83 | 97.7 | 7.51
2009-07-09 | 4:46 | 6.99 | 24 | 16 | 11.88 | 97.9 | 7.5
2009-07-09 | 5:01 | 6.89 | 24 | 16 | 11.92 | 97.9 | 7.5
2009-07-09 | 5:16 | 6.79 | 24 | 16 | 11.96 | 98 | 7.5
2009-07-09 | 5:31 | 6.69 | 24 | 16 | 11.93 | 97.5 | 7.5
2009-07-09 | 5:46 | 6.6 | 24 | 16 | 12.01 | 97.9 | 7.49
2009-07-09 | 6:01 | 6.5 | 24 | 16 | 12.02 | 97.8 | 7.49
2009-07-09 | 6:16 | 6.41 | 24 | 16 | 12.05 | 97.8 | 7.49
2009-07-09 | 6:31 | 6.33 | 24 | 16 | 12.16 | 98.5 | 7.49
2009-07-09 | 6:46 | 6.25 | 24 | 16 | 12.21 | 98.7 | 7.5
2009-07-09 | 7:01 | 6.19 | 24 | 16 | 12.22 | 98.6 | 7.51
2009-07-09 | 7:16 | 6.14 | 24 | 16 | 12.33 | 99.4 | 7.5
2009-07-09 | 7:31 | 6.12 | 24 | 16 | 12.39 | 99.8 | 7.52
2009-07-09 | 7:46 | 6.1 | 24 | 16 | 12.42 | 100.1 | 7.53
2009-07-09 | 8:01 | 6.1 | 24 | 16 | 12.42 | 100 | 7.54
2009-07-09 | 8:16 | 6.13 | 24 | 16 | 12.41 | 100 | 7.55
2009-07-09 | 8:31 | 6.17 | 24 | 16 | 12.48 | 100.7 | 7.56
2009-07-09 | 8:46 | 6.26 | 24 | 16 | 12.54 | 101.4 | 7.57
2009-07-09 | 9:01 | 6.35 | 24 | 16 | 12.54 | 101.6 | 7.57
2009-07-09 | 9:16 | 6.47 | 24 | 16 | 12.52 | 101.7 | 7.6
2009-07-09 | 9:31 | 6.62 | 24 | 16 | 12.51 | 102.1 | 7.61
2009-07-09 | 9:46 | 6.8 | 24 | 16 | 12.48 | 102.3 | 7.61
2009-07-09 | 10:01 | 6.99 | 25 | 16 | 12.37 | 101.9 | 7.63
2009-07-09 | 10:16 | 7.22 | 25 | 16 | 12.28 | 101.7 | 7.63
2009-07-09 | 10:31 | 7.49 | 25 | 16 | 12.27 | 102.3 | 7.62
2009-07-09 | 10:46 | 7.8 | 25 | 16 | 12.18 | 102.3 | 7.63
2009-07-09 | 11:01 | 8.12 | 25 | 16 | 12.07 | 102.2 | 7.64
2009-07-09 | 11:16 | 8.46 | 25 | 16 | 12.04 | 102.8 | 7.64
2009-07-09 | 11:31 | 8.82 | 25 | 16 | 11.92 | 102.7 | 7.64
2009-07-09 | 11:46 | 9.21 | 25 | 16 | 11.8 | 102.6 | 7.66
2009-07-09 | 12:01 | 9.61 | 25 | 16 | 11.73 | 103 | 7.66
2009-07-09 | 12:16 | 10.06 | 25 | 16 | 11.67 | 103.5 | 7.67
2009-07-09 | 12:31 | 10.35 | 25 | 16 | 11.62 | 103.8 | 7.67
2009-07-09 | 12:46 | 10.6 | 25 | 16 | 11.44 | 102.8 | 7.68
2009-07-09 | 13:00 | 11.03 | 25 | 16 | 11.45 | 103.9 | 7.68
2009-07-09 | 13:15 | 11.38 | 25 | 16 | 11.52 | 105.4 | 7.68
2009-07-09 | 13:30 | 11.59 | 25 | 16 | 11.38 | 104.6 | 7.7
2009-07-09 | 13:45 | 11.78 | 25 | 16 | 11.2 | 103.4 | 7.69
2009-07-09 | 14:00 | 11.98 | 25 | 16 | 11.17 | 103.6 | 7.7
2009-07-09 | 14:15 | 12.15 | 25 | 16 | 11.15 | 103.8 | 7.7
2009-07-09 | 14:30 | 12.25 | 25 | 16 | 11.14 | 103.9 | 7.7
2009-07-09 | 14:45 | 12.46 | 25 | 16 | 11.05 | 103.6 | 7.71
2009-07-09 | 15:00 | 12.59 | 25 | 16 | 10.86 | 102.1 | 7.69
2009-07-09 | 15:15 | 12.93 | 25 | 16 | 10.77 | 102.1 | 7.7
2009-07-09 | 15:30 | 12.92 | 25 | 16 | 10.7 | 101.4 | 7.71
2009-07-09 | 15:45 | 12.97 | 25 | 16 | 10.69 | 101.4 | 7.71
2009-07-09 | 16:00 | 12.96 | 25 | 16 | 10.69 | 101.3 | 7.71
2009-07-09 | 16:15 | 12.92 | 25 | 16 | 10.64 | 100.8 | 7.7
2009-07-09 | 16:30 | 12.9 | 25 | 16 | 10.59 | 100.2 | 7.7
2009-07-09 | 16:45 | 12.92 | 25 | 16 | 10.49 | 99.3 | 7.7
2009-07-09 | 17:00 | 12.96 | 25 | 16 | 10.44 | 99 | 7.7
2009-07-09 | 17:15 | 12.96 | 25 | 16 | 10.45 | 99.1 | 7.7
2009-07-09 | 17:30 | 12.91 | 25 | 16 | 10.45 | 99 | 7.7
2009-07-09 | 17:45 | 12.79 | 25 | 16 | 10.48 | 99 | 7.69
2009-07-09 | 18:00 | 12.62 | 25 | 16 | 10.45 | 98.3 | 7.68
2009-07-09 | 18:15 | 12.42 | 25 | 16 | 10.51 | 98.5 | 7.67
2009-07-09 | 18:30 | 12.23 | 25 | 16 | 10.5 | 98 | 7.66
2009-07-09 | 18:45 | 12.01 | 25 | 16 | 10.57 | 98.1 | 7.65
2009-07-09 | 19:00 | 11.82 | 25 | 16 | 10.62 | 98.2 | 7.65
2009-07-09 | 19:15 | 11.65 | 25 | 16 | 10.66 | 98.1 | 7.63
2009-07-09 | 19:30 | 11.5 | 25 | 16 | 10.69 | 98 | 7.64
2009-07-09 | 19:45 | 11.38 | 25 | 16 | 10.66 | 97.6 | 7.64
2009-07-09 | 20:00 | 11.27 | 25 | 16 | 10.72 | 97.8 | 7.62
2009-07-09 | 20:15 | 11.17 | 25 | 16 | 10.7 | 97.4 | 7.61
2009-07-09 | 20:30 | 11.08 | 25 | 16 | 10.74 | 97.6 | 7.61
2009-07-09 | 20:45 | 11.04 | 25 | 16 | 10.75 | 97.5 | 7.61
2009-07-09 | 21:00 | 11.02 | 25 | 16 | 10.74 | 97.5 | 7.6
2009-07-09 | 21:15 | 11 | 25 | 16 | 10.7 | 97 | 7.61
2009-07-09 | 21:30 | 10.98 | 25 | 16 | 10.72 | 97.2 | 7.6
2009-07-09 | 21:45 | 10.98 | 25 | 16 | 10.68 | 96.8 | 7.61
2009-07-09 | 22:00 | 10.96 | 25 | 16 | 10.71 | 97 | 7.59
2009-07-09 | 22:15 | 10.94 | 25 | 16 | 10.68 | 96.7 | 7.6
2009-07-09 | 22:30 | 10.91 | 25 | 16 | 10.72 | 97 | 7.61
2009-07-09 | 22:45 | 10.87 | 25 | 16 | 10.74 | 97.1 | 7.61
2009-07-09 | 23:00 | 10.83 | 25 | 16 | 10.72 | 96.8 | 7.62
2009-07-09 | 23:15 | 10.77 | 25 | 16 | 10.72 | 96.7 | 7.61
2009-07-09 | 23:30 | 10.67 | 25 | 16 | 10.73 | 96.6 | 7.61
2009-07-09 | 23:45 | 10.54 | 25 | 16 | 10.77 | 96.6 | 7.6
2009-07-10 | 0:00 | 10.39 | 25 | 16 | 10.82 | 96.7 | 7.59
2009-07-10 | 0:15 | 10.25 | 25 | 16 | 10.85 | 96.7 | 7.58
2009-07-10 | 0:30 | 10.11 | 25 | 16 | 10.93 | 97.1 | 7.59
2009-07-10 | 0:45 | 9.98 | 25 | 16 | 10.97 | 97.1 | 7.58
2009-07-10 | 1:01 | 9.85 | 25 | 16 | 11.01 | 97.2 | 7.58
2009-07-10 | 1:16 | 9.73 | 25 | 16 | 11.09 | 97.6 | 7.57
2009-07-10 | 1:31 | 9.62 | 25 | 16 | 11.07 | 97.2 | 7.57
2009-07-10 | 1:46 | 9.52 | 25 | 16 | 11.14 | 97.5 | 7.55
2009-07-10 | 2:01 | 9.42 | 25 | 16 | 11.13 | 97.3 | 7.55
2009-07-10 | 2:16 | 9.32 | 25 | 16 | 11.14 | 97.1 | 7.54
2009-07-10 | 2:31 | 9.23 | 25 | 16 | 11.16 | 97.1 | 7.54
2009-07-10 | 2:46 | 9.14 | 25 | 16 | 11.18 | 97.1 | 7.54
2009-07-10 | 3:01 | 9.06 | 25 | 16 | 11.21 | 97.1 | 7.54
2009-07-10 | 3:16 | 8.97 | 25 | 16 | 11.28 | 97.6 | 7.52
2009-07-10 | 3:31 | 8.89 | 25 | 16 | 11.29 | 97.4 | 7.53
2009-07-10 | 3:46 | 8.81 | 24 | 16 | 11.3 | 97.3 | 7.52
2009-07-10 | 4:01 | 8.73 | 24 | 16 | 11.32 | 97.3 | 7.52
2009-07-10 | 4:16 | 8.66 | 24 | 16 | 11.35 | 97.4 | 7.52
2009-07-10 | 4:31 | 8.59 | 24 | 16 | 11.38 | 97.5 | 7.52
2009-07-10 | 4:46 | 8.53 | 24 | 16 | 11.39 | 97.4 | 7.52
2009-07-10 | 5:01 | 8.47 | 24 | 16 | 11.45 | 97.8 | 7.52
2009-07-10 | 5:16 | 8.41 | 24 | 16 | 11.45 | 97.7 | 7.52
2009-07-10 | 5:31 | 8.36 | 24 | 16 | 11.5 | 97.9 | 7.52
2009-07-10 | 5:46 | 8.3 | 24 | 16 | 11.53 | 98.1 | 7.52
2009-07-10 | 6:01 | 8.25 | 24 | 16 | 11.51 | 97.8 | 7.54
2009-07-10 | 6:16 | 8.21 | 24 | 16 | 11.58 | 98.3 | 7.55
2009-07-10 | 6:31 | 8.15 | 24 | 16 | 11.57 | 98.1 | 7.54
2009-07-10 | 6:46 | 8.1 | 24 | 16 | 11.59 | 98.1 | 7.56
2009-07-10 | 7:01 | 8.07 | 24 | 16 | 11.67 | 98.7 | 7.55
2009-07-10 | 7:16 | 8.05 | 24 | 16 | 11.73 | 99.2 | 7.57
2009-07-10 | 7:31 | 8.04 | 24 | 16 | 11.79 | 99.6 | 7.57
2009-07-10 | 7:46 | 8.04 | 24 | 16 | 11.76 | 99.4 | 7.57
2009-07-10 | 8:01 | 8.04 | 24 | 16 | 11.77 | 99.5 | 7.58
2009-07-10 | 8:16 | 8.05 | 24 | 16 | 11.78 | 99.6 | 7.59
2009-07-10 | 8:31 | 8.07 | 24 | 16 | 11.85 | 100.2 | 7.59
2009-07-10 | 8:46 | 8.11 | 24 | 16 | 11.91 | 100.8 | 7.61
2009-07-10 | 9:01 | 8.18 | 24 | 16 | 11.97 | 101.5 | 7.62
2009-07-10 | 9:16 | 8.24 | 24 | 16 | 11.9 | 101.1 | 7.62
2009-07-10 | 9:31 | 8.35 | 25 | 16 | 11.92 | 101.5 | 7.63
2009-07-10 | 9:46 | 8.47 | 25 | 16 | 11.91 | 101.7 | 7.63
2009-07-10 | 10:01 | 8.64 | 25 | 16 | 11.88 | 101.9 | 7.64
2009-07-10 | 10:16 | 8.83 | 25 | 16 | 11.8 | 101.7 | 7.64
2009-07-10 | 10:31 | 9.04 | 25 | 16 | 11.77 | 101.9 | 7.64
2009-07-10 | 10:46 | 9.25 | 25 | 16 | 11.7 | 101.8 | 7.65
2009-07-10 | 11:01 | 9.58 | 25 | 16 | 11.66 | 102.2 | 7.65
2009-07-10 | 11:16 | 9.86 | 25 | 16 | 11.58 | 102.3 | 7.66
2009-07-10 | 11:31 | 10.2 | 25 | 16 | 11.44 | 101.8 | 7.67
2009-07-10 | 11:46 | 10.5 | 25 | 16 | 11.32 | 101.5 | 7.67
2009-07-10 | 12:01 | 10.91 | 25 | 16 | 11.23 | 101.6 | 7.67
2009-07-10 | 12:16 | 11.35 | 25 | 16 | 11.18 | 102.2 | 7.69
2009-07-10 | 12:31 | 11.8 | 25 | 16 | 11.23 | 103.7 | 7.69
2009-07-10 | 12:46 | 12.23 | 25 | 16 | 11.18 | 104.3 | 7.68
2009-07-10 | 13:00 | 12.74 | 25 | 16 | 11.1 | 104.7 | 7.69
2009-07-10 | 13:15 | 13.23 | 25 | 16 | 11 | 104.9 | 7.7
2009-07-10 | 13:30 | 13.72 | 25 | 16 | 10.92 | 105.3 | 7.71
2009-07-10 | 13:45 | 14.28 | 25 | 16 | 10.76 | 105.1 | 7.71
2009-07-10 | 14:00 | 14.8 | 25 | 16 | 10.53 | 103.9 | 7.73
2009-07-10 | 14:15 | 15.26 | 25 | 16 | 10.33 | 103 | 7.72
2009-07-10 | 14:30 | 15.66 | 25 | 16 | 10.2 | 102.6 | 7.74
2009-07-10 | 14:45 | 16 | 25 | 16 | 10.05 | 101.8 | 7.74
2009-07-10 | 15:00 | 16.31 | 25 | 17 | 9.86 | 100.6 | 7.74
2009-07-10 | 15:15 | 16.55 | 25 | 17 | 9.74 | 99.8 | 7.75
2009-07-10 | 15:30 | 16.65 | 25 | 17 | 9.64 | 99 | 7.74
2009-07-10 | 15:45 | 16.73 | 26 | 17 | 9.58 | 98.6 | 7.73
2009-07-10 | 16:00 | 16.78 | 26 | 17 | 9.54 | 98.3 | 7.75
2009-07-10 | 16:15 | 16.73 | 25 | 17 | 9.53 | 98.1 | 7.74
2009-07-10 | 16:30 | 16.58 | 26 | 17 | 9.56 | 98.1 | 7.74
2009-07-10 | 16:45 | 16.34 | 25 | 17 | 9.59 | 97.8 | 7.73
2009-07-10 | 17:00 | 16.12 | 25 | 17 | 9.61 | 97.6 | 7.72
2009-07-10 | 17:15 | 15.84 | 25 | 17 | 9.59 | 96.8 | 7.71
2009-07-10 | 17:30 | 15.56 | 25 | 17 | 9.61 | 96.5 | 7.7
2009-07-10 | 17:45 | 15.28 | 25 | 16 | 9.7 | 96.8 | 7.69
2009-07-10 | 18:00 | 14.98 | 25 | 16 | 9.77 | 96.8 | 7.69
2009-07-10 | 18:15 | 14.69 | 25 | 16 | 9.84 | 96.9 | 7.68
2009-07-10 | 18:30 | 14.44 | 25 | 16 | 9.92 | 97.2 | 7.67
2009-07-10 | 18:45 | 14.27 | 25 | 16 | 9.93 | 96.9 | 7.68
2009-07-10 | 19:00 | 14.06 | 25 | 16 | 9.98 | 97 | 7.66
2009-07-10 | 19:15 | 13.89 | 25 | 16 | 10.02 | 97 | 7.66
2009-07-10 | 19:30 | 13.71 | 25 | 16 | 10.07 | 97.1 | 7.65
2009-07-10 | 19:45 | 13.54 | 25 | 16 | 10.11 | 97.1 | 7.65
2009-07-10 | 20:00 | 13.38 | 25 | 16 | 10.15 | 97.1 | 7.64
2009-07-10 | 20:15 | 13.23 | 25 | 16 | 10.17 | 97.1 | 7.64
2009-07-10 | 20:30 | 13.12 | 25 | 16 | 10.15 | 96.6 | 7.64
2009-07-10 | 20:45 | 13.04 | 25 | 16 | 10.18 | 96.7 | 7.63
2009-07-10 | 21:00 | 12.97 | 25 | 16 | 10.15 | 96.3 | 7.63
2009-07-10 | 21:15 | 12.9 | 25 | 16 | 10.19 | 96.5 | 7.63
2009-07-10 | 21:30 | 12.81 | 25 | 16 | 10.22 | 96.6 | 7.62
2009-07-10 | 21:45 | 12.71 | 25 | 16 | 10.25 | 96.6 | 7.61
2009-07-10 | 22:00 | 12.57 | 25 | 16 | 10.24 | 96.3 | 7.6
2009-07-10 | 22:15 | 12.39 | 25 | 16 | 10.27 | 96.1 | 7.59
2009-07-10 | 22:30 | 12.18 | 25 | 16 | 10.32 | 96.2 | 7.58
2009-07-10 | 22:45 | 11.97 | 25 | 16 | 10.38 | 96.2 | 7.57
2009-07-10 | 23:00 | 11.77 | 25 | 16 | 10.43 | 96.3 | 7.57
2009-07-10 | 23:15 | 11.55 | 24 | 16 | 10.49 | 96.3 | 7.57
2009-07-10 | 23:30 | 11.35 | 24 | 16 | 10.55 | 96.4 | 7.56
2009-07-10 | 23:45 | 11.14 | 24 | 16 | 10.66 | 97 | 7.56
2009-07-11 | 0:00 | 10.95 | 24 | 16 | 10.68 | 96.8 | 7.55
2009-07-11 | 0:15 | 10.77 | 24 | 16 | 10.72 | 96.7 | 7.53
2009-07-11 | 0:30 | 10.6 | 24 | 16 | 10.82 | 97.2 | 7.53
2009-07-11 | 0:45 | 10.44 | 24 | 16 | 10.83 | 96.9 | 7.53
2009-07-11 | 1:01 | 10.3 | 24 | 16 | 10.87 | 96.9 | 7.52
2009-07-11 | 1:16 | 10.16 | 24 | 16 | 10.9 | 96.9 | 7.52
2009-07-11 | 1:31 | 10.04 | 24 | 16 | 10.94 | 97 | 7.53
2009-07-11 | 1:46 | 9.91 | 24 | 16 | 11.02 | 97.4 | 7.53
2009-07-11 | 2:01 | 9.8 | 24 | 15 | 11.06 | 97.6 | 7.52
2009-07-11 | 2:16 | 9.68 | 24 | 15 | 11.09 | 97.5 | 7.52
2009-07-11 | 2:31 | 9.56 | 24 | 15 | 11.14 | 97.7 | 7.51
2009-07-11 | 2:46 | 9.44 | 24 | 15 | 11.2 | 97.9 | 7.52
2009-07-11 | 3:01 | 9.32 | 24 | 15 | 11.25 | 98 | 7.52
2009-07-11 | 3:16 | 9.2 | 24 | 15 | 11.24 | 97.7 | 7.51
2009-07-11 | 3:31 | 9.09 | 24 | 15 | 11.25 | 97.5 | 7.51
2009-07-11 | 3:46 | 8.97 | 24 | 15 | 11.29 | 97.6 | 7.5
2009-07-11 | 4:01 | 8.86 | 24 | 15 | 11.3 | 97.5 | 7.51
2009-07-11 | 4:16 | 8.75 | 24 | 15 | 11.4 | 98 | 7.5
2009-07-11 | 4:31 | 8.64 | 24 | 15 | 11.45 | 98.2 | 7.5
2009-07-11 | 4:46 | 8.54 | 24 | 15 | 11.44 | 97.9 | 7.5
2009-07-11 | 5:01 | 8.43 | 24 | 15 | 11.52 | 98.3 | 7.49
2009-07-11 | 5:16 | 8.33 | 24 | 15 | 11.57 | 98.5 | 7.49
2009-07-11 | 5:31 | 8.23 | 24 | 15 | 11.55 | 98.1 | 7.5
2009-07-11 | 5:46 | 8.13 | 24 | 15 | 11.63 | 98.5 | 7.5
2009-07-11 | 6:01 | 8.04 | 24 | 15 | 11.68 | 98.7 | 7.49
2009-07-11 | 6:16 | 7.95 | 24 | 15 | 11.68 | 98.5 | 7.5
2009-07-11 | 6:31 | 7.87 | 24 | 15 | 11.7 | 98.5 | 7.5
2009-07-11 | 6:46 | 7.8 | 24 | 15 | 11.74 | 98.7 | 7.5
2009-07-11 | 7:01 | 7.74 | 24 | 15 | 11.84 | 99.3 | 7.5
2009-07-11 | 7:16 | 7.69 | 24 | 15 | 11.9 | 99.7 | 7.5
2009-07-11 | 7:31 | 7.65 | 24 | 15 | 11.95 | 100 | 7.51
2009-07-11 | 7:46 | 7.63 | 24 | 16 | 12 | 100.4 | 7.52
2009-07-11 | 8:01 | 7.62 | 24 | 16 | 11.98 | 100.2 | 7.53
2009-07-11 | 8:16 | 7.64 | 24 | 16 | 12.04 | 100.7 | 7.53
2009-07-11 | 8:31 | 7.67 | 24 | 16 | 12.01 | 100.6 | 7.55
2009-07-11 | 8:46 | 7.74 | 24 | 16 | 12.1 | 101.5 | 7.56
2009-07-11 | 9:01 | 7.82 | 24 | 16 | 12.18 | 102.4 | 7.57
2009-07-11 | 9:16 | 7.93 | 24 | 16 | 12.12 | 102.1 | 7.58
2009-07-11 | 9:31 | 8.08 | 24 | 16 | 12.09 | 102.3 | 7.59
2009-07-11 | 9:46 | 8.23 | 24 | 16 | 12.06 | 102.4 | 7.59
2009-07-11 | 10:01 | 8.43 | 24 | 16 | 11.95 | 101.9 | 7.61
2009-07-11 | 10:16 | 8.64 | 24 | 16 | 11.92 | 102.2 | 7.61
2009-07-11 | 10:31 | 8.88 | 24 | 16 | 11.81 | 101.8 | 7.61
2009-07-11 | 10:46 | 9.15 | 24 | 16 | 11.75 | 102 | 7.63
2009-07-11 | 11:01 | 9.47 | 24 | 16 | 11.73 | 102.6 | 7.64
2009-07-11 | 11:16 | 9.79 | 24 | 16 | 11.65 | 102.7 | 7.64
2009-07-11 | 11:31 | 10.17 | 24 | 16 | 11.53 | 102.6 | 7.63
2009-07-11 | 11:46 | 10.56 | 24 | 16 | 11.36 | 101.9 | 7.65
2009-07-11 | 12:01 | 10.98 | 24 | 16 | 11.29 | 102.3 | 7.65
2009-07-11 | 12:16 | 11.42 | 24 | 16 | 11.24 | 102.9 | 7.67
2009-07-11 | 12:31 | 11.95 | 24 | 16 | 11.22 | 104 | 7.67
2009-07-11 | 12:46 | 12.48 | 24 | 16 | 11.17 | 104.7 | 7.68
2009-07-11 | 13:00 | 13.03 | 25 | 16 | 11.13 | 105.7 | 7.69
2009-07-11 | 13:15 | 13.55 | 25 | 16 | 10.99 | 105.6 | 7.7
2009-07-11 | 13:30 | 14.07 | 25 | 16 | 10.91 | 106 | 7.69
2009-07-11 | 13:45 | 14.56 | 25 | 16 | 10.71 | 105.2 | 7.71
2009-07-11 | 14:00 | 15.02 | 25 | 16 | 10.52 | 104.4 | 7.71
2009-07-11 | 14:15 | 15.44 | 25 | 16 | 10.32 | 103.3 | 7.71
2009-07-11 | 14:30 | 15.82 | 25 | 16 | 10.16 | 102.5 | 7.74
2009-07-11 | 14:45 | 16.14 | 25 | 16 | 9.99 | 101.5 | 7.73
2009-07-11 | 15:00 | 16.42 | 25 | 16 | 9.88 | 101 | 7.74
2009-07-11 | 15:15 | 16.63 | 25 | 16 | 9.8 | 100.6 | 7.75
2009-07-11 | 15:30 | 16.78 | 25 | 16 | 9.72 | 100.1 | 7.74
2009-07-11 | 15:45 | 16.87 | 25 | 16 | 9.67 | 99.7 | 7.73
2009-07-11 | 16:00 | 16.91 | 25 | 16 | 9.58 | 99 | 7.75
2009-07-11 | 16:15 | 16.92 | 25 | 16 | 9.57 | 98.8 | 7.73
2009-07-11 | 16:30 | 16.84 | 25 | 16 | 9.53 | 98.3 | 7.72
2009-07-11 | 16:45 | 16.74 | 25 | 16 | 9.52 | 97.9 | 7.71
2009-07-11 | 17:00 | 16.58 | 25 | 16 | 9.56 | 98.1 | 7.72
2009-07-11 | 17:15 | 16.4 | 25 | 16 | 9.59 | 98 | 7.71
2009-07-11 | 17:30 | 16.2 | 25 | 16 | 9.59 | 97.5 | 7.7
2009-07-11 | 17:45 | 16 | 25 | 16 | 9.64 | 97.6 | 7.7
2009-07-11 | 18:00 | 15.81 | 25 | 16 | 9.67 | 97.6 | 7.69
2009-07-11 | 18:15 | 15.61 | 25 | 16 | 9.72 | 97.7 | 7.7
2009-07-11 | 18:30 | 15.42 | 25 | 16 | 9.71 | 97.2 | 7.68
2009-07-11 | 18:45 | 15.25 | 25 | 16 | 9.76 | 97.3 | 7.68
2009-07-11 | 19:00 | 15.1 | 25 | 16 | 9.79 | 97.3 | 7.68
2009-07-11 | 19:15 | 14.94 | 25 | 16 | 9.82 | 97.3 | 7.67
2009-07-11 | 19:30 | 14.77 | 25 | 16 | 9.85 | 97.2 | 7.67
2009-07-11 | 19:45 | 14.63 | 25 | 16 | 9.83 | 96.7 | 7.67
2009-07-11 | 20:00 | 14.5 | 25 | 16 | 9.82 | 96.3 | 7.66
2009-07-11 | 20:15 | 14.39 | 25 | 16 | 9.87 | 96.6 | 7.65
2009-07-11 | 20:30 | 14.29 | 25 | 16 | 9.85 | 96.2 | 7.66
2009-07-11 | 20:45 | 14.2 | 25 | 16 | 9.85 | 96 | 7.64
2009-07-11 | 21:00 | 14.1 | 25 | 16 | 9.85 | 95.8 | 7.64
2009-07-11 | 21:15 | 14 | 25 | 16 | 9.91 | 96.1 | 7.64
2009-07-11 | 21:30 | 13.89 | 25 | 16 | 9.95 | 96.3 | 7.65
2009-07-11 | 21:45 | 13.75 | 25 | 16 | 9.99 | 96.4 | 7.64
2009-07-11 | 22:00 | 13.61 | 24 | 16 | 10.03 | 96.5 | 7.62
2009-07-11 | 22:15 | 13.48 | 24 | 16 | 10.04 | 96.3 | 7.63
2009-07-11 | 22:30 | 13.35 | 24 | 16 | 10.09 | 96.5 | 7.62
2009-07-11 | 22:45 | 13.21 | 24 | 16 | 10.14 | 96.7 | 7.62
2009-07-11 | 23:00 | 13.06 | 24 | 16 | 10.14 | 96.3 | 7.61
2009-07-11 | 23:15 | 12.92 | 24 | 16 | 10.21 | 96.7 | 7.6
2009-07-11 | 23:30 | 12.75 | 24 | 16 | 10.21 | 96.4 | 7.6
2009-07-11 | 23:45 | 12.61 | 24 | 16 | 10.28 | 96.7 | 7.58
2009-07-12 | 0:00 | 12.45 | 24 | 16 | 10.29 | 96.5 | 7.57
2009-07-12 | 0:15 | 12.3 | 24 | 16 | 10.32 | 96.4 | 7.56
2009-07-12 | 0:30 | 12.15 | 24 | 15 | 10.39 | 96.8 | 7.55
2009-07-12 | 0:45 | 12 | 24 | 15 | 10.46 | 97.1 | 7.54
2009-07-12 | 1:01 | 11.85 | 24 | 15 | 10.51 | 97.2 | 7.54
2009-07-12 | 1:16 | 11.69 | 24 | 15 | 10.52 | 96.9 | 7.52
2009-07-12 | 1:31 | 11.54 | 23 | 15 | 10.53 | 96.7 | 7.52
2009-07-12 | 1:46 | 11.38 | 23 | 15 | 10.62 | 97.1 | 7.52
2009-07-12 | 2:01 | 11.21 | 23 | 15 | 10.64 | 97 | 7.51
2009-07-12 | 2:16 | 11.05 | 23 | 15 | 10.72 | 97.3 | 7.5
2009-07-12 | 2:31 | 10.88 | 23 | 15 | 10.74 | 97.1 | 7.48
2009-07-12 | 2:46 | 10.73 | 23 | 15 | 10.77 | 97.1 | 7.48
2009-07-12 | 3:01 | 10.57 | 23 | 15 | 10.86 | 97.5 | 7.48
2009-07-12 | 3:16 | 10.41 | 23 | 15 | 10.92 | 97.7 | 7.47
2009-07-12 | 3:31 | 10.26 | 23 | 15 | 10.92 | 97.3 | 7.48
2009-07-12 | 3:46 | 10.09 | 23 | 15 | 10.96 | 97.3 | 7.47
2009-07-12 | 4:01 | 9.93 | 23 | 15 | 11 | 97.3 | 7.48
2009-07-12 | 4:16 | 9.78 | 23 | 15 | 11.05 | 97.4 | 7.48
2009-07-12 | 4:31 | 9.62 | 23 | 15 | 11.09 | 97.4 | 7.47
2009-07-12 | 4:46 | 9.46 | 23 | 15 | 11.2 | 97.9 | 7.46
2009-07-12 | 5:01 | 9.31 | 23 | 15 | 11.26 | 98.2 | 7.46
2009-07-12 | 5:16 | 9.16 | 23 | 15 | 11.27 | 97.9 | 7.46
2009-07-12 | 5:31 | 9.02 | 23 | 15 | 11.3 | 97.8 | 7.46
2009-07-12 | 5:46 | 8.87 | 23 | 15 | 11.35 | 97.8 | 7.46
2009-07-12 | 6:01 | 8.73 | 23 | 15 | 11.45 | 98.4 | 7.46
2009-07-12 | 6:16 | 8.6 | 23 | 15 | 11.47 | 98.3 | 7.46
2009-07-12 | 6:31 | 8.48 | 23 | 15 | 11.53 | 98.5 | 7.47
2009-07-12 | 6:46 | 8.38 | 23 | 15 | 11.63 | 99.1 | 7.46
2009-07-12 | 7:01 | 8.29 | 23 | 15 | 11.65 | 99.1 | 7.46
2009-07-12 | 7:16 | 8.22 | 23 | 15 | 11.75 | 99.7 | 7.47
2009-07-12 | 7:31 | 8.17 | 23 | 15 | 11.82 | 100.2 | 7.48
2009-07-12 | 7:46 | 8.15 | 23 | 15 | 11.87 | 100.6 | 7.49
2009-07-12 | 8:01 | 8.14 | 23 | 15 | 11.86 | 100.4 | 7.5
2009-07-12 | 8:16 | 8.16 | 23 | 15 | 11.86 | 100.5 | 7.52
2009-07-12 | 8:31 | 8.22 | 23 | 15 | 11.86 | 100.7 | 7.54
2009-07-12 | 8:46 | 8.31 | 23 | 15 | 11.96 | 101.8 | 7.56
2009-07-12 | 9:01 | 8.42 | 23 | 15 | 12.03 | 102.6 | 7.57
2009-07-12 | 9:16 | 8.56 | 23 | 15 | 11.96 | 102.3 | 7.59
2009-07-12 | 9:31 | 8.74 | 23 | 15 | 11.93 | 102.6 | 7.59
2009-07-12 | 9:46 | 8.93 | 23 | 15 | 11.88 | 102.6 | 7.6
2009-07-12 | 10:01 | 9.16 | 23 | 15 | 11.83 | 102.7 | 7.6
2009-07-12 | 10:16 | 9.41 | 23 | 15 | 11.7 | 102.2 | 7.63
2009-07-12 | 10:31 | 9.7 | 23 | 15 | 11.65 | 102.5 | 7.61
2009-07-12 | 10:46 | 10.02 | 23 | 15 | 11.6 | 102.8 | 7.63
2009-07-12 | 11:01 | 10.36 | 23 | 15 | 11.53 | 103 | 7.64
2009-07-12 | 11:16 | 10.73 | 24 | 15 | 11.42 | 102.9 | 7.64
2009-07-12 | 11:31 | 11.1 | 24 | 15 | 11.25 | 102.2 | 7.64
2009-07-12 | 11:46 | 11.49 | 24 | 15 | 11.17 | 102.5 | 7.66
2009-07-12 | 12:01 | 11.91 | 24 | 15 | 11.04 | 102.2 | 7.66
2009-07-12 | 12:16 | 12.35 | 24 | 15 | 11.05 | 103.3 | 7.67
2009-07-12 | 12:31 | 12.82 | 24 | 15 | 11.08 | 104.7 | 7.69
2009-07-12 | 12:46 | 13.3 | 24 | 16 | 10.99 | 105 | 7.68
2009-07-12 | 13:00 | 13.77 | 24 | 16 | 10.95 | 105.7 | 7.7
2009-07-12 | 13:15 | 14.23 | 24 | 16 | 10.87 | 106 | 7.69
2009-07-12 | 13:30 | 14.66 | 24 | 16 | 10.72 | 105.6 | 7.7
2009-07-12 | 13:45 | 15.08 | 24 | 16 | 10.56 | 104.9 | 7.71
2009-07-12 | 14:00 | 15.45 | 24 | 16 | 10.43 | 104.4 | 7.72
2009-07-12 | 14:15 | 15.78 | 24 | 16 | 10.28 | 103.6 | 7.72
2009-07-12 | 14:30 | 16.08 | 24 | 16 | 10.1 | 102.4 | 7.74
2009-07-12 | 14:45 | 16.34 | 24 | 16 | 10 | 102 | 7.72
2009-07-12 | 15:00 | 16.54 | 24 | 16 | 9.9 | 101.5 | 7.72
2009-07-12 | 15:15 | 16.7 | 24 | 16 | 9.77 | 100.5 | 7.73
2009-07-12 | 15:30 | 16.81 | 24 | 16 | 9.73 | 100.3 | 7.73
2009-07-12 | 15:45 | 16.84 | 24 | 16 | 9.65 | 99.6 | 7.73
2009-07-12 | 16:00 | 16.8 | 24 | 16 | 9.62 | 99.1 | 7.73
2009-07-12 | 16:15 | 16.73 | 24 | 16 | 9.66 | 99.4 | 7.72
2009-07-12 | 16:30 | 16.59 | 24 | 16 | 9.69 | 99.4 | 7.73
2009-07-12 | 16:45 | 16.41 | 24 | 16 | 9.74 | 99.5 | 7.72
2009-07-12 | 17:00 | 16.21 | 24 | 16 | 9.77 | 99.5 | 7.72
2009-07-12 | 17:15 | 15.97 | 24 | 16 | 9.81 | 99.3 | 7.71
2009-07-12 | 17:30 | 15.7 | 24 | 16 | 9.79 | 98.5 | 7.7
2009-07-12 | 17:45 | 15.4 | 24 | 16 | 9.81 | 98.2 | 7.7
2009-07-12 | 18:00 | 15.16 | 24 | 16 | 9.92 | 98.7 | 7.68
2009-07-12 | 18:15 | 14.92 | 24 | 16 | 9.95 | 98.5 | 7.68
2009-07-12 | 18:30 | 14.62 | 24 | 16 | 10.04 | 98.8 | 7.68
2009-07-12 | 18:45 | 14.33 | 24 | 16 | 10.11 | 98.8 | 7.66
2009-07-12 | 19:00 | 14.05 | 24 | 16 | 10.17 | 98.8 | 7.66
2009-07-12 | 19:15 | 13.8 | 24 | 16 | 10.21 | 98.7 | 7.65
2009-07-12 | 19:30 | 13.6 | 24 | 16 | 10.2 | 98.1 | 7.66
2009-07-12 | 19:45 | 13.44 | 24 | 16 | 10.25 | 98.3 | 7.64
2009-07-12 | 20:00 | 13.3 | 24 | 16 | 10.28 | 98.2 | 7.63
2009-07-12 | 20:15 | 13.21 | 24 | 16 | 10.29 | 98.1 | 7.65
2009-07-12 | 20:30 | 13.1 | 24 | 16 | 10.31 | 98 | 7.64
2009-07-12 | 20:45 | 12.97 | 24 | 16 | 10.28 | 97.5 | 7.63
2009-07-12 | 21:00 | 12.85 | 24 | 16 | 10.28 | 97.2 | 7.63
2009-07-12 | 21:15 | 12.75 | 24 | 16 | 10.29 | 97.1 | 7.63
2009-07-12 | 21:30 | 12.63 | 24 | 16 | 10.36 | 97.5 | 7.62
2009-07-12 | 21:45 | 12.49 | 24 | 16 | 10.41 | 97.7 | 7.61
2009-07-12 | 22:00 | 12.36 | 24 | 16 | 10.46 | 97.9 | 7.61
2009-07-12 | 22:15 | 12.23 | 24 | 16 | 10.5 | 97.9 | 7.6
2009-07-12 | 22:30 | 12.09 | 24 | 15 | 10.54 | 98 | 7.6
2009-07-12 | 22:45 | 11.95 | 24 | 15 | 10.59 | 98.1 | 7.6
2009-07-12 | 23:00 | 11.82 | 24 | 15 | 10.58 | 97.8 | 7.6
2009-07-12 | 23:15 | 11.68 | 24 | 15 | 10.65 | 98.1 | 7.59
2009-07-12 | 23:30 | 11.54 | 24 | 15 | 10.66 | 97.9 | 7.59
2009-07-12 | 23:45 | 11.41 | 23 | 15 | 10.68 | 97.8 | 7.6
2009-07-13 | 0:00 | 11.29 | 23 | 15 | 10.77 | 98.3 | 7.58
2009-07-13 | 0:15 | 11.17 | 23 | 15 | 10.77 | 98.1 | 7.59
2009-07-13 | 0:30 | 11.04 | 23 | 15 | 10.84 | 98.4 | 7.58
2009-07-13 | 0:45 | 10.91 | 23 | 15 | 10.84 | 98.1 | 7.58
2009-07-13 | 1:01 | 10.78 | 23 | 15 | 10.87 | 98.1 | 7.57
2009-07-13 | 1:16 | 10.64 | 23 | 15 | 10.95 | 98.5 | 7.58
2009-07-13 | 1:31 | 10.51 | 23 | 15 | 11.01 | 98.7 | 7.57
2009-07-13 | 1:46 | 10.36 | 23 | 15 | 11.07 | 98.9 | 7.56
2009-07-13 | 2:01 | 10.23 | 23 | 15 | 11.06 | 98.5 | 7.56
2009-07-13 | 2:16 | 10.09 | 23 | 15 | 11.09 | 98.4 | 7.55
2009-07-13 | 2:31 | 9.96 | 23 | 15 | 11.17 | 98.9 | 7.54
2009-07-13 | 2:46 | 9.84 | 23 | 15 | 11.23 | 99.1 | 7.54
2009-07-13 | 3:01 | 9.72 | 23 | 15 | 11.23 | 98.8 | 7.55
2009-07-13 | 3:16 | 9.61 | 23 | 15 | 11.24 | 98.7 | 7.55
2009-07-13 | 3:31 | 9.52 | 23 | 15 | 11.27 | 98.7 | 7.53
2009-07-13 | 3:46 | 9.41 | 22 | 15 | 11.3 | 98.7 | 7.54
2009-07-13 | 4:01 | 9.32 | 22 | 15 | 11.38 | 99.2 | 7.54
2009-07-13 | 4:16 | 9.23 | 22 | 15 | 11.38 | 99 | 7.54
2009-07-13 | 4:31 | 9.14 | 22 | 15 | 11.39 | 98.9 | 7.54
2009-07-13 | 4:46 | 9.06 | 22 | 15 | 11.42 | 98.9 | 7.54
2009-07-13 | 5:01 | 8.98 | 22 | 15 | 11.44 | 98.9 | 7.54
2009-07-13 | 5:16 | 8.89 | 22 | 15 | 11.52 | 99.4 | 7.54
2009-07-13 | 5:31 | 8.8 | 22 | 15 | 11.52 | 99.2 | 7.54
2009-07-13 | 5:46 | 8.71 | 22 | 15 | 11.59 | 99.6 | 7.54
2009-07-13 | 6:01 | 8.62 | 23 | 15 | 11.65 | 99.8 | 7.54
2009-07-13 | 6:16 | 8.54 | 22 | 15 | 11.64 | 99.6 | 7.53
2009-07-13 | 6:31 | 8.46 | 23 | 15 | 11.72 | 100.1 | 7.52
2009-07-13 | 6:46 | 8.39 | 23 | 15 | 11.72 | 99.9 | 7.53
2009-07-13 | 7:01 | 8.32 | 23 | 15 | 11.75 | 100 | 7.53
2009-07-13 | 7:16 | 8.26 | 23 | 15 | 11.84 | 100.6 | 7.54
2009-07-13 | 7:31 | 8.23 | 23 | 15 | 11.84 | 100.6 | 7.55
2009-07-13 | 7:46 | 8.22 | 23 | 15 | 11.88 | 100.8 | 7.56
2009-07-13 | 8:01 | 8.23 | 23 | 15 | 11.89 | 101 | 7.56
2009-07-13 | 8:16 | 8.27 | 23 | 15 | 11.99 | 101.9 | 7.57
2009-07-13 | 8:31 | 8.33 | 23 | 15 | 11.98 | 102 | 7.57
2009-07-13 | 8:46 | 8.44 | 23 | 15 | 12.01 | 102.5 | 7.57
2009-07-13 | 9:01 | 8.56 | 23 | 15 | 12.1 | 103.6 | 7.59
2009-07-13 | 9:16 | 8.69 | 23 | 15 | 12.08 | 103.7 | 7.57
2009-07-13 | 9:31 | 8.86 | 23 | 15 | 12.04 | 103.8 | 7.6
2009-07-13 | 9:46 | 8.91 | 23 | 15 | 11.93 | 103 | 7.59
2009-07-13 | 10:01 | 9 | 23 | 15 | 11.92 | 103.1 | 7.6
2009-07-13 | 10:16 | 9.05 | 23 | 15 | 11.91 | 103.1 | 7.61
2009-07-13 | 10:31 | 9.15 | 23 | 15 | 11.91 | 103.4 | 7.59
2009-07-13 | 10:46 | 9.23 | 23 | 15 | 11.91 | 103.6 | 7.6
2009-07-13 | 11:01 | 9.38 | 23 | 15 | 11.91 | 103.9 | 7.62
2009-07-13 | 11:16 | 9.55 | 23 | 15 | 11.86 | 104 | 7.6
2009-07-13 | 11:31 | 9.86 | 23 | 15 | 11.78 | 104 | 7.62
2009-07-13 | 11:46 | 10.19 | 23 | 15 | 11.67 | 103.9 | 7.61
2009-07-13 | 12:01 | 10.58 | 23 | 15 | 11.53 | 103.6 | 7.62
2009-07-13 | 12:16 | 11.03 | 23 | 15 | 11.44 | 103.8 | 7.63
2009-07-13 | 12:31 | 11.58 | 23 | 15 | 11.46 | 105.3 | 7.64
2009-07-13 | 12:46 | 12.2 | 23 | 15 | 11.35 | 105.8 | 7.64
2009-07-13 | 13:00 | 12.8 | 23 | 15 | 11.27 | 106.5 | 7.66
2009-07-13 | 13:15 | 13.36 | 23 | 15 | 11.17 | 106.9 | 7.67
2009-07-13 | 13:30 | 13.86 | 23 | 15 | 11.05 | 106.9 | 7.65
2009-07-13 | 13:45 | 14.31 | 23 | 15 | 10.84 | 105.9 | 7.67
2009-07-13 | 14:00 | 14.73 | 24 | 15 | 10.63 | 104.8 | 7.69
2009-07-13 | 14:15 | 15.1 | 24 | 15 | 10.5 | 104.3 | 7.69
2009-07-13 | 14:30 | 15.44 | 24 | 15 | 10.38 | 103.9 | 7.7
2009-07-13 | 14:45 | 15.71 | 24 | 15 | 10.2 | 102.7 | 7.7
2009-07-13 | 15:00 | 15.92 | 24 | 15 | 10.13 | 102.4 | 7.7
2009-07-13 | 15:15 | 16.07 | 24 | 15 | 10.03 | 101.8 | 7.71
2009-07-13 | 15:30 | 16.18 | 24 | 16 | 9.98 | 101.4 | 7.71
2009-07-13 | 15:45 | 16.24 | 24 | 16 | 9.89 | 100.7 | 7.72
2009-07-13 | 16:00 | 16.25 | 24 | 16 | 9.84 | 100.2 | 7.72
2009-07-13 | 16:15 | 16.19 | 24 | 16 | 9.81 | 99.8 | 7.71
2009-07-13 | 16:30 | 16.08 | 24 | 16 | 9.86 | 100 | 7.71
2009-07-13 | 16:45 | 15.91 | 24 | 16 | 9.89 | 100 | 7.68
2009-07-13 | 17:00 | 15.71 | 24 | 16 | 9.92 | 99.9 | 7.69
2009-07-13 | 17:15 | 15.47 | 24 | 16 | 9.96 | 99.8 | 7.69
2009-07-13 | 17:30 | 15.23 | 24 | 16 | 9.99 | 99.6 | 7.68
2009-07-13 | 17:45 | 15 | 24 | 16 | 10.04 | 99.6 | 7.68
2009-07-13 | 18:00 | 14.77 | 24 | 16 | 10.08 | 99.4 | 7.66
2009-07-13 | 18:15 | 14.54 | 24 | 16 | 10.08 | 99 | 7.66
2009-07-13 | 18:30 | 14.31 | 24 | 16 | 10.11 | 98.8 | 7.65
2009-07-13 | 18:45 | 14.11 | 24 | 16 | 10.18 | 99 | 7.65
2009-07-13 | 19:00 | 13.91 | 24 | 15 | 10.23 | 99.1 | 7.65
2009-07-13 | 19:15 | 13.71 | 24 | 15 | 10.23 | 98.7 | 7.63
2009-07-13 | 19:30 | 13.53 | 24 | 15 | 10.3 | 98.9 | 7.63
2009-07-13 | 19:45 | 13.38 | 24 | 15 | 10.29 | 98.5 | 7.62
2009-07-13 | 20:00 | 13.23 | 24 | 15 | 10.35 | 98.7 | 7.62
2009-07-13 | 20:15 | 13.11 | 24 | 15 | 10.33 | 98.3 | 7.64
2009-07-13 | 20:30 | 12.99 | 24 | 15 | 10.39 | 98.6 | 7.62
2009-07-13 | 20:45 | 12.88 | 24 | 15 | 10.42 | 98.6 | 7.6
2009-07-13 | 21:00 | 12.75 | 24 | 15 | 10.45 | 98.7 | 7.61
2009-07-13 | 21:15 | 12.63 | 24 | 15 | 10.43 | 98.2 | 7.61
2009-07-13 | 21:30 | 12.5 | 23 | 15 | 10.49 | 98.5 | 7.6
2009-07-13 | 21:45 | 12.37 | 23 | 15 | 10.54 | 98.6 | 7.6
2009-07-13 | 22:00 | 12.23 | 23 | 15 | 10.58 | 98.7 | 7.59
2009-07-13 | 22:15 | 12.09 | 23 | 15 | 10.62 | 98.7 | 7.58
2009-07-13 | 22:30 | 11.93 | 23 | 15 | 10.62 | 98.4 | 7.56
2009-07-13 | 22:45 | 11.75 | 23 | 15 | 10.66 | 98.3 | 7.56
2009-07-13 | 23:00 | 11.58 | 23 | 15 | 10.68 | 98.1 | 7.55
2009-07-13 | 23:15 | 11.39 | 23 | 15 | 10.78 | 98.6 | 7.53
2009-07-13 | 23:30 | 11.21 | 23 | 15 | 10.81 | 98.5 | 7.53
2009-07-13 | 23:45 | 11.03 | 23 | 15 | 10.9 | 98.9 | 7.52
2009-07-14 | 0:00 | 10.85 | 23 | 15 | 10.91 | 98.6 | 7.52
2009-07-14 | 0:15 | 10.68 | 23 | 15 | 11.01 | 99.1 | 7.52
2009-07-14 | 0:30 | 10.49 | 23 | 15 | 11.08 | 99.3 | 7.51
2009-07-14 | 0:45 | 10.31 | 22 | 15 | 11.14 | 99.4 | 7.52
2009-07-14 | 1:01 | 10.09 | 22 | 15 | 11.16 | 99.1 | 7.49
2009-07-14 | 1:16 | 9.88 | 22 | 15 | 11.26 | 99.5 | 7.49
2009-07-14 | 1:31 | 9.68 | 22 | 14 | 11.34 | 99.7 | 7.49
2009-07-14 | 1:46 | 9.48 | 22 | 14 | 11.36 | 99.4 | 7.48
2009-07-14 | 2:01 | 9.29 | 22 | 14 | 11.41 | 99.4 | 7.47
2009-07-14 | 2:16 | 9.11 | 22 | 14 | 11.45 | 99.3 | 7.47
2009-07-14 | 2:31 | 8.93 | 22 | 14 | 11.57 | 99.9 | 7.46
2009-07-14 | 2:46 | 8.75 | 22 | 14 | 11.59 | 99.6 | 7.47
2009-07-14 | 3:01 | 8.56 | 22 | 14 | 11.64 | 99.6 | 7.46
2009-07-14 | 3:16 | 8.38 | 22 | 14 | 11.69 | 99.6 | 7.45
2009-07-14 | 3:31 | 8.2 | 22 | 14 | 11.75 | 99.7 | 7.45
2009-07-14 | 3:46 | 8.02 | 22 | 14 | 11.87 | 100.3 | 7.45
2009-07-14 | 4:01 | 7.84 | 22 | 14 | 11.89 | 100 | 7.46
2009-07-14 | 4:16 | 7.67 | 22 | 14 | 11.94 | 100 | 7.44
2009-07-14 | 4:31 | 7.49 | 22 | 14 | 12 | 100 | 7.44
2009-07-14 | 4:46 | 7.33 | 22 | 14 | 12.05 | 100.1 | 7.44
2009-07-14 | 5:01 | 7.17 | 22 | 14 | 12.17 | 100.6 | 7.44
2009-07-14 | 5:16 | 7.01 | 22 | 14 | 12.25 | 100.9 | 7.44
2009-07-14 | 5:31 | 6.85 | 22 | 14 | 12.27 | 100.7 | 7.44
2009-07-14 | 5:46 | 6.71 | 22 | 14 | 12.35 | 101 | 7.43
2009-07-14 | 6:01 | 6.57 | 22 | 14 | 12.44 | 101.3 | 7.42
2009-07-14 | 6:16 | 6.44 | 22 | 14 | 12.45 | 101.1 | 7.42
2009-07-14 | 6:31 | 6.32 | 22 | 14 | 12.55 | 101.6 | 7.43
2009-07-14 | 6:46 | 6.23 | 22 | 14 | 12.56 | 101.5 | 7.44
2009-07-14 | 7:01 | 6.15 | 22 | 14 | 12.6 | 101.6 | 7.44
2009-07-14 | 7:16 | 6.08 | 22 | 14 | 12.62 | 101.6 | 7.45
2009-07-14 | 7:31 | 6.03 | 22 | 14 | 12.74 | 102.4 | 7.44
2009-07-14 | 7:46 | 6 | 22 | 14 | 12.75 | 102.4 | 7.48
2009-07-14 | 8:01 | 5.99 | 22 | 15 | 12.83 | 103 | 7.47
2009-07-14 | 8:16 | 6.01 | 22 | 14 | 12.82 | 103 | 7.47
2009-07-14 | 8:31 | 6.06 | 22 | 15 | 12.8 | 103 | 7.49
2009-07-14 | 8:46 | 6.15 | 22 | 15 | 12.84 | 103.5 | 7.5
2009-07-14 | 9:01 | 6.26 | 22 | 15 | 12.94 | 104.6 | 7.5
2009-07-14 | 9:16 | 6.4 | 22 | 15 | 12.86 | 104.4 | 7.52
2009-07-14 | 9:31 | 6.57 | 23 | 15 | 12.84 | 104.6 | 7.52
2009-07-14 | 9:46 | 6.76 | 23 | 15 | 12.73 | 104.3 | 7.52
2009-07-14 | 10:01 | 6.97 | 23 | 15 | 12.64 | 104 | 7.55
2009-07-14 | 10:16 | 7.22 | 23 | 15 | 12.6 | 104.3 | 7.55
2009-07-14 | 10:31 | 7.5 | 23 | 15 | 12.53 | 104.5 | 7.55
2009-07-14 | 10:46 | 7.81 | 23 | 15 | 12.46 | 104.7 | 7.58
2009-07-14 | 11:01 | 8.16 | 23 | 15 | 12.36 | 104.8 | 7.58
2009-07-14 | 11:16 | 8.53 | 23 | 15 | 12.24 | 104.7 | 7.59
2009-07-14 | 11:31 | 8.93 | 23 | 15 | 12.1 | 104.4 | 7.59
2009-07-14 | 11:46 | 9.33 | 23 | 15 | 11.96 | 104.3 | 7.58
2009-07-14 | 12:01 | 9.76 | 23 | 15 | 11.78 | 103.8 | 7.59
2009-07-14 | 12:16 | 10.23 | 23 | 15 | 11.73 | 104.5 | 7.61
2009-07-14 | 12:31 | 10.77 | 23 | 15 | 11.68 | 105.4 | 7.61
2009-07-14 | 12:46 | 11.3 | 23 | 15 | 11.69 | 106.8 | 7.64
2009-07-14 | 13:00 | 11.85 | 23 | 15 | 11.6 | 107.3 | 7.62
2009-07-14 | 13:15 | 12.38 | 23 | 15 | 11.5 | 107.6 | 7.66
2009-07-14 | 13:30 | 12.9 | 23 | 15 | 11.35 | 107.5 | 7.66
2009-07-14 | 13:45 | 13.38 | 24 | 15 | 11.19 | 107.1 | 7.66
2009-07-14 | 14:00 | 13.85 | 24 | 15 | 10.98 | 106.2 | 7.65
2009-07-14 | 14:15 | 14.28 | 24 | 15 | 10.8 | 105.4 | 7.65
2009-07-14 | 14:30 | 14.67 | 24 | 15 | 10.62 | 104.6 | 7.67
2009-07-14 | 14:45 | 15 | 24 | 15 | 10.41 | 103.2 | 7.68
2009-07-14 | 15:00 | 15.29 | 24 | 15 | 10.25 | 102.3 | 7.69
2009-07-14 | 15:15 | 15.51 | 24 | 16 | 10.13 | 101.5 | 7.7
2009-07-14 | 15:30 | 15.67 | 24 | 16 | 10.07 | 101.3 | 7.69
2009-07-14 | 15:45 | 15.78 | 24 | 16 | 9.98 | 100.6 | 7.7
2009-07-14 | 16:00 | 15.82 | 24 | 16 | 9.92 | 100.1 | 7.7
2009-07-14 | 16:15 | 15.81 | 24 | 16 | 9.89 | 99.8 | 7.69
2009-07-14 | 16:30 | 15.72 | 24 | 16 | 9.87 | 99.4 | 7.68
2009-07-14 | 16:45 | 15.6 | 24 | 16 | 9.87 | 99.1 | 7.67
2009-07-14 | 17:00 | 15.44 | 24 | 16 | 9.92 | 99.3 | 7.69
2009-07-14 | 17:15 | 15.25 | 24 | 16 | 9.92 | 98.9 | 7.67
2009-07-14 | 17:30 | 15.06 | 24 | 16 | 9.98 | 99.1 | 7.66
2009-07-14 | 17:45 | 14.86 | 24 | 16 | 9.99 | 98.7 | 7.65
2009-07-14 | 18:00 | 14.66 | 24 | 16 | 10 | 98.5 | 7.65
2009-07-14 | 18:15 | 14.46 | 24 | 16 | 10.03 | 98.3 | 7.64
2009-07-14 | 18:30 | 14.27 | 24 | 16 | 10.08 | 98.3 | 7.65
2009-07-14 | 18:45 | 14.1 | 24 | 16 | 10.13 | 98.6 | 7.64
2009-07-14 | 19:00 | 13.95 | 24 | 16 | 10.18 | 98.7 | 7.64
2009-07-14 | 19:15 | 13.81 | 24 | 16 | 10.2 | 98.6 | 7.63
2009-07-14 | 19:30 | 13.67 | 24 | 16 | 10.19 | 98.2 | 7.62
2009-07-14 | 19:45 | 13.54 | 24 | 16 | 10.24 | 98.4 | 7.6
2009-07-14 | 20:00 | 13.44 | 24 | 15 | 10.27 | 98.4 | 7.6
2009-07-14 | 20:15 | 13.33 | 24 | 15 | 10.25 | 98 | 7.6
2009-07-14 | 20:30 | 13.24 | 24 | 15 | 10.28 | 98.1 | 7.6
2009-07-14 | 20:45 | 13.16 | 24 | 16 | 10.3 | 98.1 | 7.6
2009-07-14 | 21:00 | 13.08 | 24 | 16 | 10.32 | 98.1 | 7.58
2009-07-14 | 21:15 | 12.99 | 24 | 15 | 10.28 | 97.5 | 7.57
2009-07-14 | 21:30 | 12.91 | 24 | 16 | 10.29 | 97.5 | 7.57
2009-07-14 | 21:45 | 12.83 | 24 | 15 | 10.29 | 97.3 | 7.57
2009-07-14 | 22:00 | 12.75 | 24 | 15 | 10.36 | 97.8 | 7.54
2009-07-14 | 22:15 | 12.66 | 24 | 15 | 10.35 | 97.5 | 7.54
2009-07-14 | 22:30 | 12.56 | 24 | 15 | 10.41 | 97.9 | 7.52
2009-07-14 | 22:45 | 12.45 | 24 | 15 | 10.45 | 97.9 | 7.52
2009-07-14 | 23:00 | 12.34 | 24 | 15 | 10.44 | 97.6 | 7.5
2009-07-14 | 23:15 | 12.21 | 24 | 15 | 10.46 | 97.6 | 7.51
2009-07-14 | 23:30 | 12.07 | 24 | 15 | 10.55 | 98.1 | 7.52
2009-07-14 | 23:45 | 11.93 | 24 | 15 | 10.6 | 98.2 | 7.5
2009-07-15 | 0:00 | 11.78 | 24 | 15 | 10.7 | 98.7 | 7.5
2009-07-15 | 0:15 | 11.62 | 24 | 15 | 10.7 | 98.5 | 7.49
2009-07-15 | 0:30 | 11.47 | 24 | 15 | 10.69 | 98 | 7.48
2009-07-15 | 0:45 | 11.33 | 24 | 15 | 10.72 | 98 | 7.48
2009-07-15 | 1:01 | 11.19 | 23 | 15 | 10.74 | 97.8 | 7.48
2009-07-15 | 1:16 | 11.04 | 23 | 15 | 10.84 | 98.4 | 7.47
2009-07-15 | 1:31 | 10.9 | 23 | 15 | 10.85 | 98.2 | 7.48
2009-07-15 | 1:46 | 10.78 | 23 | 15 | 10.88 | 98.1 | 7.47
2009-07-15 | 2:01 | 10.65 | 23 | 15 | 10.92 | 98.2 | 7.47
2009-07-15 | 2:16 | 10.52 | 23 | 15 | 11.01 | 98.7 | 7.45
2009-07-15 | 2:31 | 10.39 | 23 | 15 | 11.01 | 98.4 | 7.47
2009-07-15 | 2:46 | 10.27 | 23 | 15 | 11.03 | 98.4 | 7.46
2009-07-15 | 3:01 | 10.13 | 23 | 15 | 11.13 | 98.9 | 7.45
2009-07-15 | 3:16 | 9.98 | 23 | 15 | 11.19 | 99.1 | 7.44
2009-07-15 | 3:31 | 9.86 | 23 | 15 | 11.18 | 98.7 | 7.46
2009-07-15 | 3:46 | 9.73 | 23 | 15 | 11.26 | 99.1 | 7.45
2009-07-15 | 4:01 | 9.58 | 23 | 15 | 11.28 | 98.9 | 7.44
2009-07-15 | 4:16 | 9.44 | 23 | 15 | 11.3 | 98.8 | 7.46
2009-07-15 | 4:31 | 9.3 | 23 | 15 | 11.34 | 98.8 | 7.46
2009-07-15 | 4:46 | 9.15 | 23 | 15 | 11.44 | 99.3 | 7.43
2009-07-15 | 5:01 | 9.01 | 23 | 15 | 11.46 | 99.2 | 7.44
2009-07-15 | 5:16 | 8.89 | 23 | 15 | 11.49 | 99.1 | 7.44
2009-07-15 | 5:31 | 8.75 | 23 | 15 | 11.53 | 99.1 | 7.43
2009-07-15 | 5:46 | 8.6 | 23 | 15 | 11.58 | 99.2 | 7.44
2009-07-15 | 6:01 | 8.46 | 23 | 15 | 11.62 | 99.2 | 7.44
2009-07-15 | 6:16 | 8.34 | 23 | 15 | 11.68 | 99.4 | 7.44
2009-07-15 | 6:31 | 8.23 | 23 | 15 | 11.79 | 100.1 | 7.45
2009-07-15 | 6:46 | 8.13 | 23 | 15 | 11.88 | 100.6 | 7.45
2009-07-15 | 7:01 | 8.03 | 23 | 15 | 11.94 | 100.9 | 7.46
2009-07-15 | 7:16 | 7.97 | 23 | 15 | 11.93 | 100.6 | 7.47
2009-07-15 | 7:31 | 7.91 | 23 | 15 | 12.04 | 101.4 | 7.47
2009-07-15 | 7:46 | 7.89 | 23 | 15 | 12.02 | 101.2 | 7.49
2009-07-15 | 8:01 | 7.87 | 23 | 15 | 12.1 | 101.8 | 7.5
2009-07-15 | 8:16 | 7.88 | 23 | 15 | 12.07 | 101.6 | 7.51
2009-07-15 | 8:31 | 7.92 | 23 | 15 | 12.13 | 102.2 | 7.5
2009-07-15 | 8:46 | 7.98 | 23 | 15 | 12.16 | 102.6 | 7.52
2009-07-15 | 9:01 | 8.07 | 23 | 15 | 12.17 | 102.9 | 7.54
2009-07-15 | 9:16 | 8.19 | 23 | 15 | 12.19 | 103.4 | 7.54
2009-07-15 | 9:31 | 8.33 | 23 | 15 | 12.1 | 103 | 7.55
2009-07-15 | 9:46 | 8.51 | 23 | 15 | 12.09 | 103.3 | 7.57
2009-07-15 | 10:01 | 8.71 | 23 | 15 | 11.99 | 103 | 7.58
2009-07-15 | 10:16 | 8.95 | 23 | 15 | 11.89 | 102.7 | 7.57
2009-07-15 | 10:31 | 9.2 | 23 | 15 | 11.88 | 103.2 | 7.58
2009-07-15 | 10:46 | 9.49 | 23 | 15 | 11.82 | 103.4 | 7.58
2009-07-15 | 11:01 | 9.82 | 23 | 15 | 11.7 | 103.2 | 7.6
2009-07-15 | 11:16 | 10.19 | 24 | 15 | 11.55 | 102.8 | 7.61
2009-07-15 | 11:31 | 10.55 | 24 | 15 | 11.48 | 103.1 | 7.63
2009-07-15 | 11:46 | 10.97 | 24 | 15 | 11.38 | 103.1 | 7.62
2009-07-15 | 12:01 | 11.39 | 24 | 15 | 11.27 | 103.1 | 7.63
2009-07-15 | 12:16 | 11.86 | 24 | 15 | 11.12 | 102.8 | 7.62
2009-07-15 | 12:31 | 12.36 | 24 | 15 | 11.18 | 104.6 | 7.66
2009-07-15 | 12:46 | 12.89 | 24 | 16 | 11.16 | 105.6 | 7.65